## 10_Profils position chambre ref
Measurement time: | 12/08/2022 19:20:27
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Position | Coin
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-100 | 0 | 100 | 2.58 | 6387.8 | 3.81
-99.3 | 0 | 100 | 2.65 | 6554.7 | 3.95
-97.9 | 0 | 100 | 2.79 | 6909.7 | 4.17
-96.9 | 0 | 100 | 2.88 | 7114.7 | 4.29
-95.7 | 0 | 100 | 2.97 | 7333.3 | 4.42
-94.4 | 0 | 100 | 3.05 | 7531.9 | 4.54
-93.2 | 0 | 100 | 3.17 | 7833.9 | 4.73
-92 | 0 | 100 | 3.23 | 7986.9 | 4.82
-90.8 | 0 | 100 | 3.38 | 8356.4 | 5.05
-89.6 | 0 | 100 | 3.5 | 8651.5 | 5.22
-88.4 | 0 | 100 | 3.59 | 8868 | 5.34
-87.2 | 0 | 100 | 3.71 | 9179.2 | 5.54
-86 | 0 | 100 | 3.87 | 9566.5 | 5.77
-84.8 | 0 | 100 | 4.01 | 9901.9 | 5.97
-83.6 | 0 | 100 | 4.1 | 10126.5 | 6.11
-82.4 | 0 | 100 | 4.24 | 10488.8 | 6.33
-81.2 | 0 | 100 | 4.42 | 10929.9 | 6.59
-80 | 0 | 100 | 4.59 | 11337.9 | 6.84
-78.8 | 0 | 100 | 4.76 | 11758.5 | 7.11
-77.6 | 0 | 100 | 4.93 | 12186 | 7.36
-76.3 | 0 | 100 | 5.16 | 12759.8 | 7.71
-75.2 | 0 | 100 | 5.35 | 13216.4 | 7.97
-73.9 | 0 | 100 | 5.55 | 13719 | 8.29
-72.8 | 0 | 100 | 5.81 | 14375.5 | 8.67
-71.6 | 0 | 100 | 5.99 | 14800.6 | 8.93
-70.3 | 0 | 100 | 6.26 | 15486.5 | 9.35
-69.1 | 0 | 100 | 6.51 | 16105.2 | 9.73
-67.9 | 0 | 100 | 6.83 | 16878.9 | 10.19
-66.7 | 0 | 100 | 7.19 | 17773.9 | 10.75
-65.5 | 0 | 100 | 7.53 | 18608 | 11.25
-64.3 | 0 | 100 | 8.05 | 19895.9 | 12.02
-63.1 | 0 | 100 | 8.73 | 21574.5 | 13.04
-61.9 | 0 | 100 | 9.76 | 24125.2 | 14.58
-60.7 | 0 | 100 | 11.57 | 28615.6 | 17.27
-59.5 | 0 | 100 | 14.5 | 35856.5 | 21.66
-58.3 | 0 | 100 | 19.01 | 46987.8 | 28.39
-57.1 | 0 | 100 | 24.96 | 61706.7 | 37.21
-55.9 | 0 | 100 | 32.41 | 80126.8 | 48.45
-54.7 | 0 | 100 | 40.85 | 100991.3 | 61.05
-53.5 | 0 | 100 | 48.28 | 119363.6 | 72.24
-52.3 | 0 | 100 | 54.88 | 135681.9 | 81.99
-51 | 0 | 100 | 59.11 | 146148.3 | 88.33
-49.9 | 0 | 100 | 61.44 | 151895.3 | 91.87
-48.6 | 0 | 100 | 62.7 | 155013 | 93.74
-47.5 | 0 | 100 | 63.36 | 156649.1 | 94.68
-46.2 | 0 | 100 | 63.78 | 157681.8 | 95.4
-45 | 0 | 100 | 64.26 | 158885 | 96.08
-43.8 | 0 | 100 | 64.48 | 159416 | 96.52
-42.6 | 0 | 100 | 64.72 | 160018.6 | 96.62
-41.4 | 0 | 100 | 64.79 | 160190.4 | 96.87
-40.2 | 0 | 100 | 64.95 | 160572.3 | 97.19
-39 | 0 | 100 | 65.21 | 161219 | 97.52
-37.8 | 0 | 100 | 65.45 | 161806.1 | 97.69
-36.6 | 0 | 100 | 65.49 | 161910.3 | 97.97
-35.4 | 0 | 100 | 65.68 | 162387.3 | 98.22
-34.2 | 0 | 100 | 65.77 | 162602.6 | 98.51
-33 | 0 | 100 | 65.78 | 162633 | 98.47
-31.8 | 0 | 100 | 66.09 | 163393.7 | 98.71
-30.6 | 0 | 100 | 66.09 | 163395.2 | 98.79
-29.4 | 0 | 100 | 66.2 | 163666 | 98.98
-28.2 | 0 | 100 | 66.35 | 164041 | 99.43
-27 | 0 | 100 | 66.33 | 164002.1 | 99.08
-25.8 | 0 | 100 | 66.34 | 164026.6 | 99.18
-24.6 | 0 | 100 | 66.49 | 164394.1 | 99.37
-23.4 | 0 | 100 | 66.55 | 164536.7 | 99.48
-22.2 | 0 | 100 | 66.63 | 164729.5 | 99.57
-21 | 0 | 100 | 66.76 | 165051.5 | 99.7
-19.8 | 0 | 100 | 66.72 | 164945.4 | 99.78
-18.6 | 0 | 100 | 66.9 | 165409.2 | 100
-17.4 | 0 | 100 | 66.79 | 165136.3 | 99.77
-16.2 | 0 | 100 | 66.73 | 164973.7 | 99.85
-15 | 0 | 100 | 66.72 | 164945.8 | 99.62
-13.7 | 0 | 100 | 66.71 | 164919.9 | 99.66
-12.6 | 0 | 100 | 66.77 | 165083.5 | 99.91
-11.3 | 0 | 100 | 66.72 | 164955.9 | 99.7
-10.1 | 0 | 100 | 66.57 | 164593.8 | 99.71
-8.9 | 0 | 100 | 66.56 | 164570.5 | 99.5
-7.7 | 0 | 100 | 66.55 | 164536.5 | 99.67
-6.5 | 0 | 100 | 66.52 | 164460.5 | 99.42
-5.3 | 0 | 100 | 66.43 | 164233.8 | 99.36
-4.1 | 0 | 100 | 66.33 | 163986.3 | 99.15
-2.9 | 0 | 100 | 66.4 | 164162.3 | 99.27
-1.7 | 0 | 100 | 66.41 | 164196.4 | 99.41
-0.5 | 0 | 100 | 66.42 | 164217.2 | 99.38
0.7 | 0 | 100 | 66.43 | 164237 | 99.36
1.9 | 0 | 100 | 66.27 | 163838.2 | 99.24
3.1 | 0 | 100 | 66.24 | 163780.8 | 99.22
4.3 | 0 | 100 | 66.33 | 163996.4 | 99.25
5.5 | 0 | 100 | 66.24 | 163774.7 | 99.18
6.7 | 0 | 100 | 66.31 | 163951.3 | 99.27
7.9 | 0 | 100 | 66.25 | 163782.4 | 99.2
9.1 | 0 | 100 | 66.31 | 163943 | 99.13
10.3 | 0 | 100 | 66.22 | 163708.7 | 99.07
11.5 | 0 | 100 | 66.24 | 163773.5 | 99.15
12.7 | 0 | 100 | 66.16 | 163564.4 | 99.09
13.9 | 0 | 100 | 66.16 | 163561.2 | 99.08
15.2 | 0 | 100 | 66.16 | 163566.1 | 98.84
16.3 | 0 | 100 | 66.16 | 163571.2 | 98.89
17.6 | 0 | 100 | 66.23 | 163756.1 | 98.99
18.7 | 0 | 100 | 66.12 | 163482.8 | 99.06
20 | 0 | 100 | 66.12 | 163482.4 | 98.77
21.2 | 0 | 100 | 66.02 | 163229.3 | 98.79
22.4 | 0 | 100 | 65.94 | 163038.3 | 98.67
23.6 | 0 | 100 | 65.99 | 163160.1 | 98.64
24.8 | 0 | 100 | 65.95 | 163059.8 | 98.62
26 | 0 | 100 | 65.77 | 162603.5 | 98.34
27.2 | 0 | 100 | 65.53 | 162025.3 | 98.19
28.4 | 0 | 100 | 65.35 | 161563.4 | 97.89
29.6 | 0 | 100 | 65.37 | 161620.9 | 97.78
30.8 | 0 | 100 | 65.16 | 161093.9 | 97.6
32 | 0 | 100 | 65.15 | 161072.1 | 97.65
33.2 | 0 | 100 | 64.96 | 160612.2 | 97.23
34.4 | 0 | 100 | 64.84 | 160296.2 | 97.09
35.6 | 0 | 100 | 64.68 | 159908.7 | 96.8
36.8 | 0 | 100 | 64.5 | 159457.9 | 96.4
38 | 0 | 100 | 64.41 | 159240 | 96.35
39.2 | 0 | 100 | 64.11 | 158492 | 95.88
40.4 | 0 | 100 | 63.95 | 158097.1 | 95.68
41.6 | 0 | 100 | 63.76 | 157629.8 | 95.48
42.8 | 0 | 100 | 63.53 | 157068.5 | 95.13
44.1 | 0 | 100 | 63.25 | 156370 | 94.74
45.2 | 0 | 100 | 62.83 | 155336.1 | 94.08
46.5 | 0 | 100 | 62.39 | 154243.3 | 93.3
47.7 | 0 | 100 | 61.6 | 152302.7 | 92.28
48.9 | 0 | 100 | 60.43 | 149412.2 | 90.43
50.1 | 0 | 100 | 58.2 | 143896.7 | 87.11
51.3 | 0 | 100 | 54.8 | 135473.6 | 82.02
52.5 | 0 | 100 | 49.64 | 122730.1 | 74.33
53.7 | 0 | 100 | 43.17 | 106719.9 | 64.63
54.9 | 0 | 100 | 35.47 | 87695.9 | 53.15
56.1 | 0 | 100 | 27.46 | 67899.5 | 41.14
57.3 | 0 | 100 | 20.45 | 50548.5 | 30.6
58.5 | 0 | 100 | 14.92 | 36880.5 | 22.33
59.7 | 0 | 100 | 11.66 | 28838.6 | 17.46
60.9 | 0 | 100 | 9.84 | 24324.8 | 14.73
62.1 | 0 | 100 | 8.73 | 21595.8 | 13.07
63.3 | 0 | 100 | 8.06 | 19929 | 12.05
64.5 | 0 | 100 | 7.49 | 18521.6 | 11.2
65.7 | 0 | 100 | 7.15 | 17680 | 10.69
66.9 | 0 | 100 | 6.77 | 16745.8 | 10.12
68.1 | 0 | 100 | 6.44 | 15931.4 | 9.62
69.3 | 0 | 100 | 6.21 | 15355 | 9.3
70.5 | 0 | 100 | 5.93 | 14661.7 | 8.87
71.7 | 0 | 100 | 5.67 | 14016.6 | 8.49
72.9 | 0 | 100 | 5.43 | 13434.6 | 8.14
74.1 | 0 | 100 | 5.3 | 13102.3 | 7.92
75.3 | 0 | 100 | 5.07 | 12539.5 | 7.59
76.5 | 0 | 100 | 4.87 | 12042.4 | 7.29
77.7 | 0 | 100 | 4.71 | 11641 | 7.06
79 | 0 | 100 | 4.55 | 11248.6 | 6.81
80.2 | 0 | 100 | 4.4 | 10872.4 | 6.6
81.4 | 0 | 100 | 4.24 | 10484 | 6.35
82.6 | 0 | 100 | 4.06 | 10031.7 | 6.08
83.8 | 0 | 100 | 3.94 | 9737.7 | 5.9
85 | 0 | 100 | 3.79 | 9372.5 | 5.69
86.2 | 0 | 100 | 3.71 | 9171.7 | 5.56
87.4 | 0 | 100 | 3.54 | 8761.1 | 5.32
88.6 | 0 | 100 | 3.45 | 8525.8 | 5.17
89.8 | 0 | 100 | 3.32 | 8220.4 | 4.98
91 | 0 | 100 | 3.23 | 7997.7 | 4.85
92.2 | 0 | 100 | 3.12 | 7706.9 | 4.66
93.4 | 0 | 100 | 3.05 | 7538.5 | 4.56
94.6 | 0 | 100 | 2.9 | 7172.7 | 4.34
95.8 | 0 | 100 | 2.77 | 6855.5 | 4.15
97 | 0 | 100 | 2.76 | 6829.9 | 4.13
98.2 | 0 | 100 | 2.65 | 6563.1 | 3.97
100 | 0 | 100 | 2.55 | 6315.7 | 3.82
Measurement time: | 12/08/2022 19:23:34
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Position | Centre
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
100 | 0 | 100 | 2.7 | 6669 | 3.82
99.4 | 0 | 100 | 2.6 | 6417 | 3.87
98.1 | 0 | 100 | 2.7 | 6672.2 | 4.03
97 | 0 | 100 | 2.8 | 6917.8 | 4.17
95.9 | 0 | 100 | 2.87 | 7087.7 | 4.28
94.7 | 0 | 100 | 2.95 | 7286.8 | 4.4
93.4 | 0 | 100 | 3.07 | 7599.2 | 4.59
92.3 | 0 | 100 | 3.14 | 7775.4 | 4.69
91 | 0 | 100 | 3.26 | 8064.2 | 4.86
89.8 | 0 | 100 | 3.39 | 8389.3 | 5.06
88.6 | 0 | 100 | 3.5 | 8661.8 | 5.22
87.4 | 0 | 100 | 3.62 | 8945.4 | 5.4
86.2 | 0 | 100 | 3.77 | 9321.1 | 5.63
85 | 0 | 100 | 3.87 | 9567.9 | 5.78
83.8 | 0 | 100 | 3.98 | 9838.7 | 5.93
82.6 | 0 | 100 | 4.11 | 10150 | 6.13
81.4 | 0 | 100 | 4.3 | 10622.2 | 6.41
80.2 | 0 | 100 | 4.45 | 11009.3 | 6.65
79 | 0 | 100 | 4.6 | 11367.9 | 6.87
77.8 | 0 | 100 | 4.78 | 11819.5 | 7.13
76.6 | 0 | 100 | 4.93 | 12197.4 | 7.38
75.4 | 0 | 100 | 5.11 | 12640.6 | 7.64
74.2 | 0 | 100 | 5.3 | 13115.3 | 7.93
73 | 0 | 100 | 5.49 | 13567.8 | 8.21
71.8 | 0 | 100 | 5.77 | 14276.3 | 8.64
70.6 | 0 | 100 | 5.93 | 14671.6 | 8.86
69.4 | 0 | 100 | 6.19 | 15315.6 | 9.26
68.2 | 0 | 100 | 6.51 | 16095.7 | 9.73
67 | 0 | 100 | 6.82 | 16850.5 | 10.18
65.8 | 0 | 100 | 7.17 | 17727.2 | 10.71
64.6 | 0 | 100 | 7.57 | 18722.9 | 11.32
63.4 | 0 | 100 | 8.1 | 20022.4 | 12.09
62.2 | 0 | 100 | 8.78 | 21702.4 | 13.11
60.9 | 0 | 100 | 9.91 | 24499.1 | 14.83
59.8 | 0 | 100 | 11.69 | 28897.6 | 17.5
58.5 | 0 | 100 | 15.05 | 37201.4 | 22.48
57.3 | 0 | 100 | 20.72 | 51227.6 | 31.02
56.2 | 0 | 100 | 27.69 | 68448.8 | 41.33
54.9 | 0 | 100 | 35.9 | 88754.5 | 53.62
53.8 | 0 | 100 | 43.61 | 107831.1 | 65.1
52.5 | 0 | 100 | 49.8 | 123110.9 | 74.4
51.3 | 0 | 100 | 55.03 | 136041.4 | 82.2
50.2 | 0 | 100 | 58.33 | 144206.9 | 87.25
48.9 | 0 | 100 | 60.52 | 149634.6 | 90.55
47.7 | 0 | 100 | 61.73 | 152616 | 92.2
46.5 | 0 | 100 | 62.38 | 154225.5 | 93.49
45.3 | 0 | 100 | 62.87 | 155426 | 94.2
44.1 | 0 | 100 | 63.27 | 156426 | 94.69
42.9 | 0 | 100 | 63.55 | 157112.8 | 94.97
41.7 | 0 | 100 | 63.88 | 157931.2 | 95.59
40.5 | 0 | 100 | 63.96 | 158129.7 | 95.4
39.3 | 0 | 100 | 64.01 | 158264.6 | 95.77
38.1 | 0 | 100 | 64.15 | 158589.8 | 95.93
36.8 | 0 | 100 | 64.45 | 159353.6 | 96.44
35.7 | 0 | 100 | 64.58 | 159657.2 | 96.74
34.5 | 0 | 100 | 64.84 | 160314.3 | 96.99
33.3 | 0 | 100 | 64.94 | 160565.4 | 97.23
32.1 | 0 | 100 | 65.24 | 161288.9 | 97.35
30.8 | 0 | 100 | 65.24 | 161286.6 | 97.59
29.7 | 0 | 100 | 65.46 | 161848.9 | 97.92
28.4 | 0 | 100 | 65.6 | 162177.8 | 98.05
27.2 | 0 | 100 | 65.65 | 162312.1 | 98.26
26.1 | 0 | 100 | 65.75 | 162554.9 | 98.42
24.8 | 0 | 100 | 65.86 | 162833 | 98.66
23.6 | 0 | 100 | 65.98 | 163118.3 | 98.64
22.4 | 0 | 100 | 65.9 | 162932.3 | 98.59
21.2 | 0 | 100 | 66.05 | 163291.9 | 98.91
20 | 0 | 100 | 66.11 | 163435.3 | 99.01
18.8 | 0 | 100 | 66.3 | 163914.4 | 99.17
17.6 | 0 | 100 | 66.08 | 163379.3 | 98.75
16.4 | 0 | 100 | 66.26 | 163808 | 99.13
15.2 | 0 | 100 | 66.24 | 163779.7 | 99.15
14 | 0 | 100 | 66.27 | 163845 | 99.16
12.8 | 0 | 100 | 66.3 | 163920.5 | 99.25
11.6 | 0 | 100 | 66.31 | 163939.8 | 99.25
10.4 | 0 | 100 | 66.33 | 163984.2 | 99.26
9.1 | 0 | 100 | 66.26 | 163809.7 | 99.13
8 | 0 | 100 | 66.3 | 163920.2 | 99.38
6.7 | 0 | 100 | 66.38 | 164111.3 | 99.27
5.6 | 0 | 100 | 66.27 | 163854.2 | 99.13
4.4 | 0 | 100 | 66.22 | 163727.4 | 99.2
3.1 | 0 | 100 | 66.31 | 163946.7 | 99.2
2 | 0 | 100 | 66.18 | 163615 | 99.02
0.7 | 0 | 100 | 66.35 | 164050.7 | 99.16
-0.4 | 0 | 100 | 66.36 | 164074.3 | 99.25
-1.7 | 0 | 100 | 66.29 | 163897.6 | 99.29
-2.9 | 0 | 100 | 66.37 | 164093.1 | 99.33
-4 | 0 | 100 | 66.31 | 163953.4 | 99.34
-5.3 | 0 | 100 | 66.37 | 164099.4 | 99.56
-6.4 | 0 | 100 | 66.46 | 164301.7 | 99.62
-7.7 | 0 | 100 | 66.5 | 164411.9 | 99.5
-8.9 | 0 | 100 | 66.5 | 164400.1 | 99.53
-10.1 | 0 | 100 | 66.48 | 164351.3 | 99.5
-11.3 | 0 | 100 | 66.61 | 164693.5 | 99.54
-12.5 | 0 | 100 | 66.67 | 164827.3 | 99.77
-13.7 | 0 | 100 | 66.64 | 164759.7 | 99.97
-14.9 | 0 | 100 | 66.66 | 164796.5 | 99.82
-16.1 | 0 | 100 | 66.67 | 164839 | 99.75
-17.3 | 0 | 100 | 66.79 | 165127.2 | 99.8
-18.5 | 0 | 100 | 66.79 | 165127 | 100
-19.7 | 0 | 100 | 66.53 | 164487 | 99.76
-20.9 | 0 | 100 | 66.61 | 164685.3 | 99.74
-22.1 | 0 | 100 | 66.63 | 164727.3 | 99.81
-23.3 | 0 | 100 | 66.57 | 164591.8 | 99.67
-24.5 | 0 | 100 | 66.45 | 164285.7 | 99.45
-25.7 | 0 | 100 | 66.24 | 163761.4 | 99.38
-27 | 0 | 100 | 66.36 | 164060.7 | 99.34
-28.1 | 0 | 100 | 66.13 | 163501.1 | 99.06
-29.4 | 0 | 100 | 66.14 | 163511.2 | 99.03
-30.5 | 0 | 100 | 66.16 | 163571.9 | 99.1
-31.7 | 0 | 100 | 65.81 | 162706 | 98.67
-32.9 | 0 | 100 | 65.75 | 162548.1 | 98.43
-34.1 | 0 | 100 | 65.71 | 162461.3 | 98.44
-35.4 | 0 | 100 | 65.67 | 162360.6 | 98.36
-36.5 | 0 | 100 | 65.52 | 161976.6 | 98.02
-37.8 | 0 | 100 | 65.2 | 161200.3 | 97.57
-39 | 0 | 100 | 65.11 | 160986.8 | 97.47
-40.2 | 0 | 100 | 64.97 | 160620.5 | 97.27
-41.4 | 0 | 100 | 64.85 | 160338.5 | 97.19
-42.5 | 0 | 100 | 64.69 | 159941.4 | 96.94
-43.8 | 0 | 100 | 64.47 | 159399.3 | 96.59
-45 | 0 | 100 | 64.1 | 158472.6 | 96.14
-46.2 | 0 | 100 | 63.72 | 157531.7 | 95.45
-47.4 | 0 | 100 | 63.3 | 156511 | 94.81
-48.6 | 0 | 100 | 62.53 | 154603.4 | 93.69
-49.8 | 0 | 100 | 61.38 | 151763.6 | 91.98
-51 | 0 | 100 | 58.9 | 145620.6 | 88.11
-52.2 | 0 | 100 | 54.77 | 135418.6 | 81.99
-53.4 | 0 | 100 | 48.2 | 119163.5 | 72.13
-54.6 | 0 | 100 | 40.52 | 100179.3 | 60.6
-55.9 | 0 | 100 | 32.25 | 79727.1 | 48.28
-57.1 | 0 | 100 | 24.66 | 60964.1 | 36.93
-58.2 | 0 | 100 | 18.9 | 46735 | 28.32
-59.5 | 0 | 100 | 14.34 | 35463.9 | 21.48
-60.6 | 0 | 100 | 11.47 | 28345.9 | 17.15
-61.9 | 0 | 100 | 9.75 | 24112 | 14.64
-63.1 | 0 | 100 | 8.75 | 21635.2 | 13.12
-64.3 | 0 | 100 | 8.06 | 19936.5 | 12.07
-65.5 | 0 | 100 | 7.54 | 18642 | 11.29
-66.7 | 0 | 100 | 7.14 | 17660.2 | 10.71
-67.9 | 0 | 100 | 6.84 | 16910.7 | 10.24
-69.1 | 0 | 100 | 6.54 | 16180.7 | 9.82
-70.3 | 0 | 100 | 6.29 | 15555.3 | 9.42
-71.5 | 0 | 100 | 6 | 14825.9 | 8.99
-72.7 | 0 | 100 | 5.78 | 14294.1 | 8.66
-73.9 | 0 | 100 | 5.55 | 13711.2 | 8.32
-75.1 | 0 | 100 | 5.34 | 13204.5 | 7.99
-76.3 | 0 | 100 | 5.14 | 12699.1 | 7.69
-77.5 | 0 | 100 | 4.96 | 12273.8 | 7.43
-78.7 | 0 | 100 | 4.82 | 11921.7 | 7.22
-80 | 0 | 100 | 4.59 | 11340.7 | 6.86
-81.1 | 0 | 100 | 4.45 | 11003.3 | 6.66
-82.3 | 0 | 100 | 4.32 | 10687.7 | 6.48
-83.6 | 0 | 100 | 4.12 | 10193.9 | 6.17
-84.7 | 0 | 100 | 4 | 9896.3 | 5.99
-86 | 0 | 100 | 3.86 | 9533.2 | 5.78
-87.2 | 0 | 100 | 3.75 | 9269.3 | 5.61
-88.4 | 0 | 100 | 3.62 | 8942.4 | 5.42
-89.6 | 0 | 100 | 3.49 | 8621 | 5.22
-90.8 | 0 | 100 | 3.35 | 8284.7 | 5.02
-92 | 0 | 100 | 3.24 | 8019.5 | 4.86
-93.2 | 0 | 100 | 3.17 | 7836.6 | 4.75
-94.4 | 0 | 100 | 3.04 | 7525.3 | 4.56
-95.6 | 0 | 100 | 2.98 | 7367.3 | 4.47
-96.8 | 0 | 100 | 2.83 | 6998.3 | 4.24
-98 | 0 | 100 | 2.78 | 6866.6 | 4.16
-100 | 0 | 100 | 2.63 | 6499.4 | 3.94
Measurement time: | 12/08/2022 19:28:12
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Position | Hors champ
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-98.7 | 0 | 100 | 2.69 | 6649.1 | 4.02
-96.6 | 0 | 100 | 2.88 | 7130.8 | 4.32
-95.4 | 0 | 100 | 3.01 | 7430.4 | 4.51
-94.2 | 0 | 100 | 3.06 | 7560.4 | 4.57
-93 | 0 | 100 | 3.24 | 8022.2 | 4.86
-91.8 | 0 | 100 | 3.29 | 8134.3 | 4.92
-90.5 | 0 | 100 | 3.4 | 8413.5 | 5.1
-89.3 | 0 | 100 | 3.53 | 8717.3 | 5.28
-88.1 | 0 | 100 | 3.62 | 8950.1 | 5.4
-86.9 | 0 | 100 | 3.76 | 9305.6 | 5.63
-85.7 | 0 | 100 | 3.92 | 9703.3 | 5.87
-84.5 | 0 | 100 | 4.06 | 10038.9 | 6.09
-83.3 | 0 | 100 | 4.19 | 10367.3 | 6.26
-82.1 | 0 | 100 | 4.28 | 10592.5 | 6.41
-80.9 | 0 | 100 | 4.53 | 11187.7 | 6.77
-79.7 | 0 | 100 | 4.68 | 11571.9 | 6.99
-78.5 | 0 | 100 | 4.9 | 12107.9 | 7.33
-77.3 | 0 | 100 | 5.07 | 12524.7 | 7.59
-76.1 | 0 | 100 | 5.21 | 12882.9 | 7.79
-74.9 | 0 | 100 | 5.39 | 13330.9 | 8.06
-73.7 | 0 | 100 | 5.58 | 13806.8 | 8.36
-72.5 | 0 | 100 | 5.82 | 14389.7 | 8.7
-71.3 | 0 | 100 | 6.06 | 14985.8 | 9.06
-70.1 | 0 | 100 | 6.32 | 15626.4 | 9.43
-68.9 | 0 | 100 | 6.53 | 16138.8 | 9.75
-67.7 | 0 | 100 | 6.86 | 16965.2 | 10.27
-66.5 | 0 | 100 | 7.23 | 17881.5 | 10.82
-65.3 | 0 | 100 | 7.68 | 18978.9 | 11.47
-64.1 | 0 | 100 | 8.22 | 20315.6 | 12.27
-62.9 | 0 | 100 | 8.97 | 22171 | 13.41
-61.6 | 0 | 100 | 10.13 | 25044.7 | 15.13
-60.5 | 0 | 100 | 12.09 | 29879.3 | 18.05
-59.2 | 0 | 100 | 15.34 | 37936.3 | 22.96
-58 | 0 | 100 | 20.3 | 50191 | 30.36
-56.8 | 0 | 100 | 26.5 | 65509.3 | 39.72
-55.6 | 0 | 100 | 34.1 | 84308.9 | 50.9
-54.4 | 0 | 100 | 42.39 | 104793.9 | 63.79
-53.2 | 0 | 100 | 49.91 | 123397.3 | 74.82
-52 | 0 | 100 | 55.95 | 138327.1 | 83.53
-50.8 | 0 | 100 | 59.7 | 147589.5 | 89.64
-49.6 | 0 | 100 | 61.74 | 152651.7 | 92.3
-48.4 | 0 | 100 | 62.77 | 155177.9 | 93.58
-47.2 | 0 | 100 | 63.43 | 156817.4 | 94.66
-46 | 0 | 100 | 64.05 | 158360 | 95.44
-44.8 | 0 | 100 | 64.27 | 158903.6 | 96.02
-43.6 | 0 | 100 | 64.54 | 159564.7 | 96.55
-42.4 | 0 | 100 | 64.72 | 160020.3 | 96.9
-41.2 | 0 | 100 | 64.9 | 160457.5 | 97.08
-40 | 0 | 100 | 65.07 | 160882.8 | 97.56
-38.8 | 0 | 100 | 65.2 | 161195.4 | 97.75
-37.6 | 0 | 100 | 65.49 | 161912.5 | 97.41
-36.4 | 0 | 100 | 65.52 | 161979.5 | 97.67
-35.2 | 0 | 100 | 65.64 | 162288.6 | 98.18
-34 | 0 | 100 | 65.77 | 162606.3 | 98.23
-32.8 | 0 | 100 | 65.8 | 162683.1 | 98.78
-31.6 | 0 | 100 | 66 | 163165.5 | 98.2
-30.4 | 0 | 100 | 66.14 | 163524.9 | 98.75
-29.1 | 0 | 100 | 66.23 | 163744.9 | 98.93
-28 | 0 | 100 | 66.3 | 163916.9 | 99.12
-26.8 | 0 | 100 | 66.32 | 163954.1 | 99.16
-25.5 | 0 | 100 | 66.59 | 164632.8 | 99.72
-24.3 | 0 | 100 | 66.71 | 164922.6 | 99.73
-23.1 | 0 | 100 | 66.47 | 164347 | 99.53
-21.9 | 0 | 100 | 66.57 | 164580 | 99.48
-20.7 | 0 | 100 | 66.64 | 164763.2 | 99.82
-19.5 | 0 | 100 | 66.69 | 164891.3 | 99.88
-18.3 | 0 | 100 | 66.67 | 164823.6 | 99.43
-17.1 | 0 | 100 | 66.72 | 164963.4 | 99.8
-15.9 | 0 | 100 | 66.74 | 164992.5 | 99.69
-14.7 | 0 | 100 | 66.6 | 164650 | 99.59
-13.5 | 0 | 100 | 66.69 | 164888.4 | 99.97
-12.3 | 0 | 100 | 66.68 | 164852.9 | 99.82
-11.1 | 0 | 100 | 66.61 | 164694.4 | 100
-9.9 | 0 | 100 | 66.53 | 164483.9 | 99.44
-8.7 | 0 | 100 | 66.56 | 164554.1 | 99.73
-7.5 | 0 | 100 | 66.51 | 164444.8 | 99.32
-6.3 | 0 | 100 | 66.39 | 164148.1 | 99.73
-5.1 | 0 | 100 | 66.61 | 164693.8 | 99.36
-3.9 | 0 | 100 | 66.35 | 164045.5 | 99.12
-2.7 | 0 | 100 | 66.3 | 163922.7 | 99.24
-1.5 | 0 | 100 | 66.25 | 163801.7 | 99.14
-0.3 | 0 | 100 | 66.26 | 163820.4 | 99.43
0.9 | 0 | 100 | 66.3 | 163926.7 | 98.96
2.1 | 0 | 100 | 66.25 | 163781.8 | 99.01
3.4 | 0 | 100 | 66.35 | 164028.4 | 99.31
4.6 | 0 | 100 | 66.31 | 163941.9 | 99.15
5.8 | 0 | 100 | 66.35 | 164032.3 | 99.39
7 | 0 | 100 | 66.31 | 163952.8 | 99.36
8.2 | 0 | 100 | 66.31 | 163950.3 | 99.4
9.4 | 0 | 100 | 66.37 | 164093.6 | 98.86
10.6 | 0 | 100 | 66.26 | 163821.2 | 99.35
11.8 | 0 | 100 | 66.3 | 163917.5 | 99.11
13 | 0 | 100 | 66.32 | 163962.7 | 99.29
14.2 | 0 | 100 | 66.16 | 163569 | 99.07
15.4 | 0 | 100 | 66.17 | 163604.9 | 98.98
16.6 | 0 | 100 | 66.27 | 163839.7 | 98.99
17.8 | 0 | 100 | 66.12 | 163483.2 | 98.94
19 | 0 | 100 | 66.11 | 163447.3 | 98.51
20.2 | 0 | 100 | 66.02 | 163222.8 | 99
21.4 | 0 | 100 | 65.99 | 163155.2 | 98.62
22.6 | 0 | 100 | 65.97 | 163095.4 | 98.83
23.8 | 0 | 100 | 65.98 | 163133.8 | 98.7
25 | 0 | 100 | 65.68 | 162393.5 | 98.19
26.2 | 0 | 100 | 65.74 | 162525.4 | 97.97
27.4 | 0 | 100 | 65.5 | 161938.6 | 97.88
28.7 | 0 | 100 | 65.5 | 161942.8 | 97.9
29.8 | 0 | 100 | 65.23 | 161280.6 | 97.69
31.1 | 0 | 100 | 65.21 | 161223.3 | 97.79
32.3 | 0 | 100 | 65.01 | 160736.9 | 97.8
33.5 | 0 | 100 | 64.94 | 160558.8 | 97.31
34.7 | 0 | 100 | 64.77 | 160131 | 96.7
35.9 | 0 | 100 | 64.51 | 159480.6 | 96.74
37.1 | 0 | 100 | 64.41 | 159239.3 | 96.73
38.3 | 0 | 100 | 64.16 | 158624.7 | 96.25
39.5 | 0 | 100 | 64 | 158221.1 | 95.9
40.7 | 0 | 100 | 63.76 | 157631 | 95.74
41.9 | 0 | 100 | 63.65 | 157360 | 95.55
43.1 | 0 | 100 | 63.46 | 156885.5 | 95.06
44.3 | 0 | 100 | 63.09 | 155984.6 | 94.33
45.5 | 0 | 100 | 62.75 | 155147.4 | 93.72
46.7 | 0 | 100 | 62.29 | 153991 | 93.33
47.9 | 0 | 100 | 61.33 | 151618.8 | 91.94
49.1 | 0 | 100 | 60.04 | 148442.1 | 89.86
50.3 | 0 | 100 | 57.65 | 142534.4 | 86.43
51.5 | 0 | 100 | 53.74 | 132872.9 | 80.46
52.7 | 0 | 100 | 48.33 | 119481.6 | 72.14
53.9 | 0 | 100 | 41.48 | 102563.9 | 61.94
55.1 | 0 | 100 | 33.75 | 83454.1 | 50.57
56.4 | 0 | 100 | 25.8 | 63780.3 | 38.73
57.6 | 0 | 100 | 18.86 | 46616.4 | 28.34
58.8 | 0 | 100 | 13.93 | 34447.3 | 20.92
60 | 0 | 100 | 11.07 | 27376.7 | 16.6
61.2 | 0 | 100 | 9.53 | 23565.7 | 14.27
62.4 | 0 | 100 | 8.55 | 21148 | 12.8
63.6 | 0 | 100 | 7.92 | 19571.5 | 11.82
64.8 | 0 | 100 | 7.42 | 18335.6 | 11.12
66 | 0 | 100 | 7.02 | 17358.3 | 10.5
67.2 | 0 | 100 | 6.71 | 16590.3 | 10.07
68.4 | 0 | 100 | 6.39 | 15789.1 | 9.55
69.6 | 0 | 100 | 6.13 | 15147.1 | 9.16
70.8 | 0 | 100 | 5.89 | 14563.6 | 8.83
72 | 0 | 100 | 5.59 | 13815.3 | 8.38
73.2 | 0 | 100 | 5.43 | 13434.1 | 8.14
74.4 | 0 | 100 | 5.2 | 12860.6 | 7.77
75.6 | 0 | 100 | 5.03 | 12428.4 | 7.53
76.8 | 0 | 100 | 4.87 | 12029 | 7.27
78 | 0 | 100 | 4.69 | 11606.2 | 7.03
79.2 | 0 | 100 | 4.57 | 11306.1 | 6.82
80.5 | 0 | 100 | 4.36 | 10772.3 | 6.51
81.6 | 0 | 100 | 4.22 | 10426.6 | 6.31
82.9 | 0 | 100 | 4.08 | 10079.6 | 6.11
84 | 0 | 100 | 3.95 | 9775.1 | 5.92
85.2 | 0 | 100 | 3.76 | 9286.2 | 5.62
86.5 | 0 | 100 | 3.64 | 8990.5 | 5.42
87.6 | 0 | 100 | 3.58 | 8841.3 | 5.36
88.9 | 0 | 100 | 3.45 | 8519.8 | 5.16
90.1 | 0 | 100 | 3.3 | 8156.8 | 4.94
91.3 | 0 | 100 | 3.2 | 7920.6 | 4.8
92.5 | 0 | 100 | 3.13 | 7731.8 | 4.7
93.7 | 0 | 100 | 3.04 | 7507.9 | 4.56
94.9 | 0 | 100 | 2.95 | 7300.6 | 4.42
96.1 | 0 | 100 | 2.85 | 7040.3 | 4.26
97.3 | 0 | 100 | 2.74 | 6780.5 | 4.1
98.5 | 0 | 100 | 2.64 | 6538.4 | 3.95
100 | 0 | 100 | 2.53 | 6244.3 | 3.79