## 11_ Profil détecteur
Profil
Measurement time: | 12/08/2022 18:12:50
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-100 | 0 | 100 | 2.43 | 6017.8 | 3.78
-99.1 | 0 | 100 | 2.98 | 7363.3 | 4.43
-97.7 | 0 | 100 | 3.03 | 7502 | 4.51
-96.6 | 0 | 100 | 3.16 | 7801.4 | 4.69
-95.4 | 0 | 100 | 3.24 | 8009.7 | 4.81
-94.1 | 0 | 100 | 3.32 | 8198.1 | 4.92
-93 | 0 | 100 | 3.43 | 8467.9 | 5.08
-91.8 | 0 | 100 | 3.57 | 8821.7 | 5.3
-90.6 | 0 | 100 | 3.68 | 9097.9 | 5.47
-89.4 | 0 | 100 | 3.76 | 9301.1 | 5.59
-88.1 | 0 | 100 | 3.92 | 9693.8 | 5.82
-86.9 | 0 | 100 | 4.02 | 9930.5 | 5.97
-85.7 | 0 | 100 | 4.18 | 10341.6 | 6.21
-84.5 | 0 | 100 | 4.32 | 10689.8 | 6.44
-83.3 | 0 | 100 | 4.43 | 10940.6 | 6.6
-82.1 | 0 | 100 | 4.62 | 11423.2 | 6.87
-80.9 | 0 | 100 | 4.72 | 11681.8 | 7.02
-79.7 | 0 | 100 | 4.92 | 12167.6 | 7.31
-78.5 | 0 | 100 | 5.09 | 12589.6 | 7.56
-77.3 | 0 | 100 | 5.25 | 12972.6 | 7.79
-76.1 | 0 | 100 | 5.45 | 13463.5 | 8.1
-74.9 | 0 | 100 | 5.65 | 13957.7 | 8.39
-73.7 | 0 | 100 | 5.8 | 14339.6 | 8.63
-72.5 | 0 | 100 | 6.06 | 14991.6 | 9.01
-71.3 | 0 | 100 | 6.33 | 15644.6 | 9.4
-70.1 | 0 | 100 | 6.54 | 16178.4 | 9.72
-68.9 | 0 | 100 | 6.88 | 17008.1 | 10.22
-67.7 | 0 | 100 | 7.14 | 17659.6 | 10.6
-66.5 | 0 | 100 | 7.49 | 18510.3 | 11.14
-65.3 | 0 | 100 | 7.92 | 19573.7 | 11.78
-64.1 | 0 | 100 | 8.44 | 20854.5 | 12.54
-62.9 | 0 | 100 | 9.15 | 22615.7 | 13.61
-61.7 | 0 | 100 | 10.29 | 25446.7 | 15.29
-60.5 | 0 | 100 | 12.2 | 30162.1 | 18.14
-59.3 | 0 | 100 | 15.3 | 37836.1 | 22.75
-58 | 0 | 100 | 20.1 | 49689.6 | 29.9
-56.8 | 0 | 100 | 26.33 | 65108.2 | 39.15
-55.6 | 0 | 100 | 33.92 | 83864.5 | 50.51
-54.4 | 0 | 100 | 42.21 | 104351.1 | 62.76
-53.2 | 0 | 100 | 49.84 | 123212 | 74.1
-52 | 0 | 100 | 55.97 | 138366.6 | 83.2
-50.8 | 0 | 100 | 59.72 | 147654.1 | 88.93
-49.6 | 0 | 100 | 61.87 | 152956.1 | 92.23
-48.4 | 0 | 100 | 63.12 | 156064.7 | 93.91
-47.2 | 0 | 100 | 63.78 | 157694.1 | 94.91
-46 | 0 | 100 | 64.21 | 158744.5 | 95.5
-44.8 | 0 | 100 | 64.44 | 159316.2 | 96.1
-43.6 | 0 | 100 | 64.63 | 159781.4 | 96.33
-42.4 | 0 | 100 | 64.97 | 160634.4 | 96.73
-41.2 | 0 | 100 | 65.12 | 161010.6 | 96.94
-40 | 0 | 100 | 65.34 | 161543.3 | 97.23
-38.8 | 0 | 100 | 65.32 | 161498.9 | 97.26
-37.6 | 0 | 100 | 65.44 | 161799.3 | 97.57
-36.4 | 0 | 100 | 65.73 | 162515.8 | 97.83
-35.2 | 0 | 100 | 65.95 | 163057.7 | 97.99
-34 | 0 | 100 | 66.13 | 163495.7 | 98.34
-32.8 | 0 | 100 | 66.15 | 163542.7 | 98.48
-31.6 | 0 | 100 | 66.36 | 164057.9 | 98.66
-30.4 | 0 | 100 | 66.53 | 164473.8 | 98.97
-29.2 | 0 | 100 | 66.57 | 164588.7 | 99.17
-28 | 0 | 100 | 66.74 | 165015.9 | 99.27
-26.8 | 0 | 100 | 66.82 | 165191.2 | 99.49
-25.5 | 0 | 100 | 66.71 | 164924.1 | 99.31
-24.4 | 0 | 100 | 66.96 | 165550.5 | 99.67
-23.1 | 0 | 100 | 66.93 | 165476.1 | 99.79
-21.9 | 0 | 100 | 66.96 | 165553 | 99.61
-20.7 | 0 | 100 | 66.98 | 165608.6 | 99.68
-19.5 | 0 | 100 | 67 | 165652 | 99.66
-18.3 | 0 | 100 | 67.08 | 165847.8 | 99.88
-17.1 | 0 | 100 | 67.04 | 165748.2 | 99.74
-15.9 | 0 | 100 | 66.99 | 165617.2 | 99.87
-14.7 | 0 | 100 | 67.1 | 165890.4 | 99.93
-13.5 | 0 | 100 | 66.98 | 165599.5 | 99.8
-12.3 | 0 | 100 | 66.93 | 165483 | 99.86
-11.1 | 0 | 100 | 66.99 | 165627.2 | 100
-9.9 | 0 | 100 | 66.89 | 165373 | 99.72
-8.7 | 0 | 100 | 66.99 | 165628.2 | 99.66
-7.5 | 0 | 100 | 66.89 | 165363.8 | 99.56
-6.3 | 0 | 100 | 66.8 | 165157.8 | 99.51
-5.1 | 0 | 100 | 66.98 | 165588.2 | 99.75
-3.8 | 0 | 100 | 66.83 | 165232.3 | 99.55
-2.7 | 0 | 100 | 66.71 | 164931.2 | 99.43
-1.4 | 0 | 100 | 66.65 | 164781.7 | 99.11
-0.3 | 0 | 100 | 66.62 | 164718.3 | 99.28
1 | 0 | 100 | 66.68 | 164850.7 | 99.24
2.2 | 0 | 100 | 66.68 | 164859.5 | 99.29
3.4 | 0 | 100 | 66.53 | 164476.8 | 99.08
4.6 | 0 | 100 | 66.48 | 164354.3 | 99.1
5.8 | 0 | 100 | 66.62 | 164704.4 | 99.24
7 | 0 | 100 | 66.46 | 164323.9 | 99.09
8.2 | 0 | 100 | 66.64 | 164766.1 | 99.27
9.4 | 0 | 100 | 66.65 | 164776.5 | 99.38
10.6 | 0 | 100 | 66.52 | 164451.7 | 99.15
11.8 | 0 | 100 | 66.58 | 164614.9 | 99.24
13 | 0 | 100 | 66.58 | 164600.8 | 99.29
14.2 | 0 | 100 | 66.55 | 164545.8 | 99.3
15.4 | 0 | 100 | 66.52 | 164468 | 99.29
16.6 | 0 | 100 | 66.5 | 164414.9 | 99.05
17.8 | 0 | 100 | 66.44 | 164259.1 | 98.88
19 | 0 | 100 | 66.51 | 164446 | 98.84
20.2 | 0 | 100 | 66.36 | 164075.3 | 98.83
21.4 | 0 | 100 | 66.28 | 163856.4 | 98.77
22.6 | 0 | 100 | 66.1 | 163433.9 | 98.45
23.8 | 0 | 100 | 66.01 | 163189.2 | 98.38
25 | 0 | 100 | 65.99 | 163152.9 | 98.27
26.2 | 0 | 100 | 66.05 | 163310.6 | 98.48
27.4 | 0 | 100 | 65.85 | 162796.9 | 98.16
28.7 | 0 | 100 | 65.75 | 162550.3 | 97.94
29.8 | 0 | 100 | 65.63 | 162264.2 | 97.81
31 | 0 | 100 | 65.44 | 161779.4 | 97.41
32.3 | 0 | 100 | 65.28 | 161400.1 | 97.26
33.4 | 0 | 100 | 65.22 | 161236.3 | 97.24
34.7 | 0 | 100 | 65.01 | 160737.9 | 97.07
35.9 | 0 | 100 | 64.78 | 160150.6 | 96.5
37.1 | 0 | 100 | 64.71 | 159980.3 | 96.4
38.3 | 0 | 100 | 64.61 | 159749.8 | 96.24
39.5 | 0 | 100 | 64.35 | 159099 | 96
40.7 | 0 | 100 | 64.18 | 158674.8 | 95.71
41.9 | 0 | 100 | 63.85 | 157854.8 | 95.26
43.1 | 0 | 100 | 63.7 | 157493.9 | 94.9
44.3 | 0 | 100 | 63.36 | 156657.5 | 94.59
45.5 | 0 | 100 | 62.85 | 155375 | 93.58
46.7 | 0 | 100 | 62.32 | 154069.3 | 92.74
47.9 | 0 | 100 | 61.3 | 151555.5 | 91.37
49.1 | 0 | 100 | 59.78 | 147802.8 | 89.09
50.3 | 0 | 100 | 56.9 | 140672.5 | 84.87
51.5 | 0 | 100 | 52.65 | 130164.1 | 78.45
52.7 | 0 | 100 | 47.05 | 116319.4 | 70.12
53.9 | 0 | 100 | 39.68 | 98110 | 59.23
55.1 | 0 | 100 | 31.62 | 78172.9 | 47.16
56.3 | 0 | 100 | 23.99 | 59312.8 | 35.79
57.6 | 0 | 100 | 17.64 | 43602.6 | 26.32
58.8 | 0 | 100 | 13.28 | 32838.7 | 19.77
60 | 0 | 100 | 10.87 | 26864.8 | 16.21
61.2 | 0 | 100 | 9.49 | 23454.1 | 14.14
62.4 | 0 | 100 | 8.66 | 21415.8 | 12.9
63.6 | 0 | 100 | 8.03 | 19856.9 | 11.96
64.8 | 0 | 100 | 7.56 | 18702.2 | 11.26
66 | 0 | 100 | 7.19 | 17781 | 10.73
67.2 | 0 | 100 | 6.86 | 16957.3 | 10.22
68.4 | 0 | 100 | 6.54 | 16175 | 9.75
69.6 | 0 | 100 | 6.37 | 15741.7 | 9.49
70.8 | 0 | 100 | 6.06 | 14971 | 9.02
72 | 0 | 100 | 5.86 | 14493.3 | 8.74
73.2 | 0 | 100 | 5.65 | 13979.3 | 8.44
74.4 | 0 | 100 | 5.48 | 13539.2 | 8.17
75.6 | 0 | 100 | 5.28 | 13061.8 | 7.87
76.8 | 0 | 100 | 5.08 | 12551.8 | 7.55
78.1 | 0 | 100 | 4.88 | 12069.3 | 7.28
79.3 | 0 | 100 | 4.79 | 11833 | 7.14
80.5 | 0 | 100 | 4.57 | 11288.3 | 6.81
81.7 | 0 | 100 | 4.41 | 10906.6 | 6.58
82.9 | 0 | 100 | 4.26 | 10542.5 | 6.37
84.1 | 0 | 100 | 4.18 | 10343.8 | 6.24
85.3 | 0 | 100 | 4.04 | 9992.8 | 6.03
86.5 | 0 | 100 | 3.92 | 9685.3 | 5.83
87.7 | 0 | 100 | 3.84 | 9498.3 | 5.73
88.9 | 0 | 100 | 3.67 | 9064.5 | 5.47
90.1 | 0 | 100 | 3.51 | 8684.6 | 5.24
91.3 | 0 | 100 | 3.43 | 8491.5 | 5.11
92.5 | 0 | 100 | 3.34 | 8265.1 | 4.99
93.7 | 0 | 100 | 3.25 | 8032.2 | 4.85
94.9 | 0 | 100 | 3.18 | 7874.3 | 4.75
96.1 | 0 | 100 | 3.04 | 7527.2 | 4.55
97.3 | 0 | 100 | 2.93 | 7240.6 | 4.37
98.5 | 0 | 100 | 2.93 | 7255.9 | 4.37
100 | 0 | 100 | 2.81 | 6937.7 | 4.18
Measurement time: | 12/08/2022 20:26:32
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | FC 65-P
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
155 | 0 | 100 | 2.17 | 13220.5 | 1.46
154.1 | 0 | 100 | 2.29 | 13965.7 | 1.52
152.6 | 0 | 100 | 2.38 | 14502.1 | 1.58
151.5 | 0 | 100 | 2.39 | 14532.6 | 1.58
150.3 | 0 | 100 | 2.44 | 14854.5 | 1.61
149.1 | 0 | 100 | 2.5 | 15202.7 | 1.65
147.9 | 0 | 100 | 2.53 | 15405.8 | 1.67
146.7 | 0 | 100 | 2.6 | 15813.2 | 1.72
145.5 | 0 | 100 | 2.65 | 16122.7 | 1.75
144.3 | 0 | 100 | 2.68 | 16336 | 1.77
143.1 | 0 | 100 | 2.76 | 16829.3 | 1.83
141.9 | 0 | 100 | 2.82 | 17194.5 | 1.87
140.6 | 0 | 100 | 2.88 | 17530.9 | 1.9
139.5 | 0 | 100 | 2.93 | 17833.5 | 1.94
138.2 | 0 | 100 | 3.01 | 18298.6 | 1.99
137 | 0 | 100 | 3.06 | 18625.3 | 2.02
135.9 | 0 | 100 | 3.13 | 19042.3 | 2.07
134.6 | 0 | 100 | 3.2 | 19460 | 2.12
133.4 | 0 | 100 | 3.26 | 19867.1 | 2.16
132.2 | 0 | 100 | 3.33 | 20280.8 | 2.2
131 | 0 | 100 | 3.41 | 20780.7 | 2.26
129.8 | 0 | 100 | 3.51 | 21375.6 | 2.32
128.6 | 0 | 100 | 3.56 | 21672.2 | 2.35
127.4 | 0 | 100 | 3.64 | 22181.5 | 2.41
126.2 | 0 | 100 | 3.76 | 22905.7 | 2.49
125 | 0 | 100 | 3.85 | 23449.4 | 2.55
123.8 | 0 | 100 | 3.93 | 23923.3 | 2.6
122.6 | 0 | 100 | 4.02 | 24471.9 | 2.66
121.4 | 0 | 100 | 4.12 | 25073.1 | 2.73
120.2 | 0 | 100 | 4.24 | 25800.6 | 2.8
119 | 0 | 100 | 4.34 | 26385.8 | 2.87
117.8 | 0 | 100 | 4.45 | 27075.7 | 2.94
116.6 | 0 | 100 | 4.56 | 27771.5 | 3.02
115.4 | 0 | 100 | 4.67 | 28431.7 | 3.09
114.2 | 0 | 100 | 4.78 | 29120.2 | 3.17
113 | 0 | 100 | 4.92 | 29940.3 | 3.25
111.8 | 0 | 100 | 5.04 | 30688.5 | 3.33
110.6 | 0 | 100 | 5.19 | 31570.8 | 3.43
109.4 | 0 | 100 | 5.31 | 32320.5 | 3.52
108.2 | 0 | 100 | 5.46 | 33248.5 | 3.62
107 | 0 | 100 | 5.62 | 34193.1 | 3.72
105.8 | 0 | 100 | 5.76 | 35074.4 | 3.81
104.5 | 0 | 100 | 5.94 | 36180.4 | 3.94
103.4 | 0 | 100 | 6.07 | 36938.1 | 4.02
102.1 | 0 | 100 | 6.28 | 38254.6 | 4.16
100.9 | 0 | 100 | 6.47 | 39382.4 | 4.28
99.7 | 0 | 100 | 6.61 | 40233.5 | 4.37
98.5 | 0 | 100 | 6.82 | 41497.2 | 4.52
97.3 | 0 | 100 | 7.04 | 42835.8 | 4.66
96.1 | 0 | 100 | 7.22 | 43951.6 | 4.78
94.9 | 0 | 100 | 7.43 | 45206.9 | 4.92
93.7 | 0 | 100 | 7.63 | 46414 | 5.05
92.5 | 0 | 100 | 7.9 | 48113.6 | 5.23
91.3 | 0 | 100 | 8.13 | 49501.7 | 5.39
90.1 | 0 | 100 | 8.41 | 51195.6 | 5.56
88.9 | 0 | 100 | 8.65 | 52622.2 | 5.72
87.7 | 0 | 100 | 8.97 | 54591.5 | 5.94
86.5 | 0 | 100 | 9.25 | 56331.3 | 6.13
85.3 | 0 | 100 | 9.57 | 58222.3 | 6.33
84.1 | 0 | 100 | 9.89 | 60225.8 | 6.55
82.9 | 0 | 100 | 10.21 | 62129.1 | 6.76
81.7 | 0 | 100 | 10.57 | 64315.4 | 7
80.5 | 0 | 100 | 11 | 66931.2 | 7.28
79.3 | 0 | 100 | 11.35 | 69071 | 7.52
78 | 0 | 100 | 11.79 | 71778.8 | 7.8
76.9 | 0 | 100 | 12.27 | 74670.9 | 8.12
75.7 | 0 | 100 | 12.76 | 77638.8 | 8.44
74.4 | 0 | 100 | 13.32 | 81073.6 | 8.8
73.3 | 0 | 100 | 13.92 | 84708 | 9.21
72 | 0 | 100 | 14.71 | 89563.1 | 9.75
70.8 | 0 | 100 | 15.7 | 95588.1 | 10.4
69.6 | 0 | 100 | 17.06 | 103846.8 | 11.3
68.4 | 0 | 100 | 19.38 | 117970.7 | 12.83
67.3 | 0 | 100 | 22.77 | 138612.3 | 15.08
66 | 0 | 100 | 27.38 | 166679.1 | 18.13
64.8 | 0 | 100 | 32.63 | 198600.8 | 21.59
63.6 | 0 | 100 | 38.06 | 231656.3 | 25.23
62.4 | 0 | 100 | 44.33 | 269826.7 | 29.41
61.2 | 0 | 100 | 50.39 | 306732.9 | 33.4
60 | 0 | 100 | 56.87 | 346143.5 | 37.66
58.8 | 0 | 100 | 63.37 | 385740.8 | 41.95
57.6 | 0 | 100 | 69.85 | 425127.2 | 46.29
56.4 | 0 | 100 | 76.63 | 466448.5 | 50.73
55.2 | 0 | 100 | 83.09 | 505715.3 | 54.97
54 | 0 | 100 | 89.73 | 546149.6 | 59.43
52.8 | 0 | 100 | 96.25 | 585832.3 | 63.82
51.6 | 0 | 100 | 102.49 | 623816.5 | 67.85
50.3 | 0 | 100 | 109.31 | 665340 | 72.35
49.2 | 0 | 100 | 115.46 | 702759.5 | 76.35
47.9 | 0 | 100 | 121.45 | 739203.5 | 80.49
46.7 | 0 | 100 | 127.3 | 774854.8 | 84.37
45.6 | 0 | 100 | 132.21 | 804728.3 | 87.7
44.3 | 0 | 100 | 136.76 | 832392.4 | 90.79
43.2 | 0 | 100 | 140.23 | 853559 | 92.9
41.9 | 0 | 100 | 142.73 | 868732.8 | 94.48
40.7 | 0 | 100 | 144.39 | 878876.8 | 95.64
39.5 | 0 | 100 | 145.3 | 884389.1 | 96.04
38.3 | 0 | 100 | 145.91 | 888103.4 | 96.6
37.1 | 0 | 100 | 146.36 | 890870.7 | 96.85
35.9 | 0 | 100 | 146.66 | 892655.7 | 97.12
34.7 | 0 | 100 | 147.14 | 895602.4 | 97.43
33.5 | 0 | 100 | 147.5 | 897774.9 | 97.58
32.3 | 0 | 100 | 147.56 | 898121.6 | 97.72
31.1 | 0 | 100 | 147.95 | 900505.5 | 97.86
29.9 | 0 | 100 | 148.35 | 902953.1 | 98.15
28.7 | 0 | 100 | 148.34 | 902877.8 | 98.23
27.5 | 0 | 100 | 148.43 | 903464 | 98.31
26.3 | 0 | 100 | 148.75 | 905411.5 | 98.5
25.1 | 0 | 100 | 149.08 | 907400.6 | 98.36
23.8 | 0 | 100 | 149.31 | 908830.5 | 98.7
22.7 | 0 | 100 | 149.43 | 909504.6 | 98.97
21.5 | 0 | 100 | 149.58 | 910434.4 | 99.05
20.2 | 0 | 100 | 149.65 | 910874.4 | 99.17
19.1 | 0 | 100 | 149.8 | 911806 | 99.1
17.8 | 0 | 100 | 150.19 | 914146.6 | 99.32
16.7 | 0 | 100 | 149.87 | 912204.7 | 99.42
15.4 | 0 | 100 | 150.1 | 913633.2 | 99.42
14.2 | 0 | 100 | 150.16 | 913964.4 | 99.54
13 | 0 | 100 | 150.05 | 913320 | 99.5
11.8 | 0 | 100 | 150.32 | 914948.5 | 99.46
10.6 | 0 | 100 | 150.21 | 914305.1 | 99.6
9.4 | 0 | 100 | 150.45 | 915755.9 | 99.53
8.2 | 0 | 100 | 150.34 | 915082.5 | 99.67
7 | 0 | 100 | 150.58 | 916529.3 | 99.56
5.7 | 0 | 100 | 150.45 | 915745.9 | 99.4
4.6 | 0 | 100 | 150.22 | 914344.9 | 99.53
3.4 | 0 | 100 | 150.08 | 913464.6 | 99.34
2.2 | 0 | 100 | 150.28 | 914723.5 | 99.43
1 | 0 | 100 | 150.45 | 915736.2 | 99.61
-0.3 | 0 | 100 | 150.52 | 916171 | 99.62
-1.4 | 0 | 100 | 150.36 | 915199.6 | 99.57
-2.7 | 0 | 100 | 150.35 | 915112.1 | 99.56
-3.9 | 0 | 100 | 150.64 | 916868.2 | 99.77
-5 | 0 | 100 | 150.67 | 917063.9 | 99.81
-6.3 | 0 | 100 | 150.71 | 917345.4 | 99.77
-7.5 | 0 | 100 | 150.68 | 917152.1 | 99.7
-8.7 | 0 | 100 | 150.82 | 917987.6 | 99.81
-9.9 | 0 | 100 | 150.7 | 917256.4 | 99.79
-11.1 | 0 | 100 | 150.94 | 918738.7 | 99.81
-12.3 | 0 | 100 | 150.89 | 918429.4 | 99.94
-13.5 | 0 | 100 | 150.92 | 918576.3 | 99.81
-14.7 | 0 | 100 | 151 | 919100 | 100
-15.9 | 0 | 100 | 150.62 | 916792 | 99.91
-17.1 | 0 | 100 | 150.77 | 917685.9 | 99.73
-18.3 | 0 | 100 | 150.82 | 917976.5 | 99.76
-19.5 | 0 | 100 | 150.57 | 916488.5 | 99.76
-20.7 | 0 | 100 | 150.69 | 917223.8 | 99.75
-21.9 | 0 | 100 | 150.25 | 914528.2 | 99.63
-23.1 | 0 | 100 | 150.32 | 914939.7 | 99.45
-24.4 | 0 | 100 | 150.3 | 914802.7 | 99.35
-25.5 | 0 | 100 | 150.12 | 913707.7 | 99.27
-26.8 | 0 | 100 | 149.77 | 911612.7 | 99.1
-28 | 0 | 100 | 149.85 | 912059.2 | 99.29
-29.1 | 0 | 100 | 149.68 | 911041.7 | 99.18
-30.4 | 0 | 100 | 149.36 | 909117.2 | 98.92
-31.5 | 0 | 100 | 149.12 | 907639.4 | 98.8
-32.8 | 0 | 100 | 148.77 | 905492.7 | 98.63
-34 | 0 | 100 | 148.83 | 905886.1 | 98.37
-35.1 | 0 | 100 | 148.43 | 903453.5 | 98.2
-36.4 | 0 | 100 | 147.86 | 899997.8 | 97.87
-37.5 | 0 | 100 | 147.34 | 896787.4 | 97.42
-38.8 | 0 | 100 | 146.78 | 893391.6 | 97.18
-40 | 0 | 100 | 145.75 | 887162.6 | 96.55
-41.2 | 0 | 100 | 144.33 | 878470.9 | 95.67
-42.4 | 0 | 100 | 141.97 | 864132.7 | 93.92
-43.6 | 0 | 100 | 138.02 | 840092.7 | 91.38
-44.8 | 0 | 100 | 133.13 | 810296.2 | 88.13
-46 | 0 | 100 | 127.6 | 776652.8 | 84.65
-47.2 | 0 | 100 | 122.08 | 743043.1 | 80.85
-48.4 | 0 | 100 | 115.71 | 704301.9 | 76.62
-49.6 | 0 | 100 | 109.41 | 665972.1 | 72.37
-50.8 | 0 | 100 | 102.79 | 625680.3 | 68
-52 | 0 | 100 | 95.9 | 583700 | 63.46
-53.2 | 0 | 100 | 89.45 | 544448.3 | 59.27
-54.4 | 0 | 100 | 82.54 | 502401.2 | 54.6
-55.6 | 0 | 100 | 76.11 | 463266.2 | 50.37
-56.8 | 0 | 100 | 69.44 | 422653.1 | 45.98
-58 | 0 | 100 | 62.52 | 380539.6 | 41.49
-59.2 | 0 | 100 | 56.2 | 342063.3 | 37.14
-60.4 | 0 | 100 | 49.44 | 300940.3 | 32.72
-61.6 | 0 | 100 | 43.34 | 263796.5 | 28.69
-62.8 | 0 | 100 | 37.3 | 227007.1 | 24.7
-64 | 0 | 100 | 31.63 | 192528 | 20.93
-65.2 | 0 | 100 | 26.76 | 162899.9 | 17.71
-66.5 | 0 | 100 | 22.43 | 136545.6 | 14.85
-67.6 | 0 | 100 | 19.3 | 117475.6 | 12.76
-68.9 | 0 | 100 | 17.09 | 103993.2 | 11.29
-70.1 | 0 | 100 | 15.65 | 95275.5 | 10.36
-71.3 | 0 | 100 | 14.71 | 89561.3 | 9.73
-72.5 | 0 | 100 | 13.9 | 84578.5 | 9.21
-73.7 | 0 | 100 | 13.33 | 81143.8 | 8.83
-74.9 | 0 | 100 | 12.78 | 77785.7 | 8.46
-76.1 | 0 | 100 | 12.31 | 74951.7 | 8.16
-77.3 | 0 | 100 | 11.9 | 72437.7 | 7.87
-78.5 | 0 | 100 | 11.49 | 69943.4 | 7.61
-79.7 | 0 | 100 | 11.09 | 67497.3 | 7.34
-80.9 | 0 | 100 | 10.69 | 65054.9 | 7.07
-82.1 | 0 | 100 | 10.34 | 62961.6 | 6.84
-83.3 | 0 | 100 | 10.06 | 61212 | 6.65
-84.5 | 0 | 100 | 9.7 | 59053.6 | 6.43
-85.7 | 0 | 100 | 9.4 | 57184.9 | 6.22
-86.9 | 0 | 100 | 9.11 | 55475.5 | 6.03
-88.1 | 0 | 100 | 8.8 | 53548.2 | 5.83
-89.3 | 0 | 100 | 8.54 | 51951.4 | 5.64
-90.6 | 0 | 100 | 8.26 | 50269.4 | 5.47
-91.7 | 0 | 100 | 8.01 | 48784.2 | 5.3
-92.9 | 0 | 100 | 7.8 | 47467.7 | 5.15
-94.1 | 0 | 100 | 7.53 | 45810.1 | 4.98
-95.3 | 0 | 100 | 7.31 | 44481.4 | 4.84
-96.6 | 0 | 100 | 7.13 | 43399.7 | 4.72
-97.7 | 0 | 100 | 6.92 | 42107.1 | 4.58
-99 | 0 | 100 | 6.72 | 40919.8 | 4.45
-100.2 | 0 | 100 | 6.56 | 39922 | 4.34
-101.3 | 0 | 100 | 6.34 | 38567.2 | 4.19
-102.6 | 0 | 100 | 6.14 | 37386.5 | 4.06
-103.8 | 0 | 100 | 6.02 | 36612.3 | 3.98
-105 | 0 | 100 | 5.86 | 35644.7 | 3.88
-106.2 | 0 | 100 | 5.72 | 34843.3 | 3.79
-107.4 | 0 | 100 | 5.55 | 33796.5 | 3.67
-108.6 | 0 | 100 | 5.37 | 32674.8 | 3.55
-109.8 | 0 | 100 | 5.25 | 31933 | 3.47
-111 | 0 | 100 | 5.09 | 30975.5 | 3.36
-112.2 | 0 | 100 | 4.97 | 30261.1 | 3.29
-113.4 | 0 | 100 | 4.87 | 29625.1 | 3.22
-114.6 | 0 | 100 | 4.74 | 28844.9 | 3.13
-115.8 | 0 | 100 | 4.6 | 28013.1 | 3.04
-117 | 0 | 100 | 4.51 | 27426.3 | 2.98
-118.2 | 0 | 100 | 4.38 | 26662.3 | 2.9
-119.4 | 0 | 100 | 4.27 | 25988 | 2.83
-120.7 | 0 | 100 | 4.17 | 25352.1 | 2.76
-121.8 | 0 | 100 | 4.06 | 24712.1 | 2.69
-123.1 | 0 | 100 | 3.98 | 24207.3 | 2.63
-124.3 | 0 | 100 | 3.88 | 23640 | 2.57
-125.4 | 0 | 100 | 3.79 | 23052.9 | 2.5
-126.7 | 0 | 100 | 3.69 | 22456 | 2.44
-127.8 | 0 | 100 | 3.62 | 22024.6 | 2.39
-129.1 | 0 | 100 | 3.54 | 21522.1 | 2.34
-130.3 | 0 | 100 | 3.43 | 20852.5 | 2.27
-131.5 | 0 | 100 | 3.35 | 20419.1 | 2.22
-132.7 | 0 | 100 | 3.29 | 20036.6 | 2.18
-133.9 | 0 | 100 | 3.22 | 19583.4 | 2.13
-135.1 | 0 | 100 | 3.14 | 19093.8 | 2.08
-136.3 | 0 | 100 | 3.08 | 18752.3 | 2.04
-137.5 | 0 | 100 | 3.01 | 18314.4 | 1.99
-138.7 | 0 | 100 | 2.93 | 17821.9 | 1.94
-139.9 | 0 | 100 | 2.87 | 17449.6 | 1.9
-141.1 | 0 | 100 | 2.82 | 17149 | 1.86
-142.3 | 0 | 100 | 2.75 | 16763 | 1.82
-143.5 | 0 | 100 | 2.69 | 16395 | 1.78
-144.7 | 0 | 100 | 2.62 | 15918.2 | 1.73
-145.9 | 0 | 100 | 2.6 | 15798.7 | 1.72
-147.1 | 0 | 100 | 2.52 | 15365.9 | 1.67
-148.3 | 0 | 100 | 2.47 | 15009.9 | 1.63
-149.5 | 0 | 100 | 2.41 | 14697.2 | 1.6
-150.7 | 0 | 100 | 2.39 | 14566.4 | 1.58
-151.9 | 0 | 100 | 2.33 | 14173 | 1.54
-153.1 | 0 | 100 | 2.27 | 13830.6 | 1.5
-155 | 0 | 100 | 2.23 | 13557.5 | 1.47
Rendement
Measurement time: | 12/08/2022 18:00:16
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 19.24 | 47561.7 | 20.48
0 | 0 | 309.1 | 20.74 | 51265.8 | 20.6
0 | 0 | 307.6 | 20.83 | 51504.6 | 20.67
0 | 0 | 306.6 | 21.02 | 51974.8 | 20.89
0 | 0 | 305.3 | 21.15 | 52292.6 | 21.01
0 | 0 | 304.1 | 21.35 | 52782.9 | 21.19
0 | 0 | 302.9 | 21.46 | 53066 | 21.33
0 | 0 | 301.7 | 21.62 | 53451.6 | 21.46
0 | 0 | 300.5 | 21.73 | 53713.5 | 21.55
0 | 0 | 299.3 | 21.83 | 53983.6 | 21.68
0 | 0 | 298.1 | 22.08 | 54597.4 | 21.97
0 | 0 | 296.9 | 22.18 | 54836.5 | 22.02
0 | 0 | 295.7 | 22.34 | 55238 | 22.21
0 | 0 | 294.5 | 22.52 | 55683.7 | 22.35
0 | 0 | 293.3 | 22.64 | 55983.6 | 22.46
0 | 0 | 292.1 | 22.78 | 56332 | 22.67
0 | 0 | 290.9 | 22.95 | 56744.8 | 22.82
0 | 0 | 289.7 | 23.07 | 57049 | 22.94
0 | 0 | 288.5 | 23.29 | 57574 | 23.13
0 | 0 | 287.3 | 23.45 | 57968 | 23.31
0 | 0 | 286.1 | 23.53 | 58174.1 | 23.42
0 | 0 | 284.9 | 23.7 | 58600 | 23.54
0 | 0 | 283.7 | 23.93 | 59151.6 | 23.76
0 | 0 | 282.5 | 24.01 | 59365.5 | 23.88
0 | 0 | 281.2 | 24.25 | 59960.5 | 24.07
0 | 0 | 280.1 | 24.39 | 60289 | 24.23
0 | 0 | 278.8 | 24.46 | 60477.1 | 24.31
0 | 0 | 277.7 | 24.68 | 61005.5 | 24.5
0 | 0 | 276.4 | 24.86 | 61469.1 | 24.72
0 | 0 | 275.2 | 25.02 | 61854 | 24.83
0 | 0 | 274 | 25.22 | 62360 | 25.07
0 | 0 | 272.8 | 25.32 | 62610.7 | 25.19
0 | 0 | 271.6 | 25.54 | 63143.9 | 25.37
0 | 0 | 270.4 | 25.66 | 63445.9 | 25.55
0 | 0 | 269.2 | 25.84 | 63885.9 | 25.73
0 | 0 | 268 | 26.05 | 64400.7 | 25.89
0 | 0 | 266.8 | 26.3 | 65024.5 | 26.08
0 | 0 | 265.6 | 26.42 | 65327.7 | 26.21
0 | 0 | 264.4 | 26.53 | 65584.6 | 26.36
0 | 0 | 263.2 | 26.75 | 66146.4 | 26.61
0 | 0 | 262 | 26.97 | 66670.8 | 26.79
0 | 0 | 260.8 | 27.05 | 66881.8 | 26.88
0 | 0 | 259.6 | 27.34 | 67602.7 | 27.15
0 | 0 | 258.4 | 27.47 | 67924.8 | 27.26
0 | 0 | 257.2 | 27.65 | 68368.3 | 27.45
0 | 0 | 256 | 27.85 | 68849.4 | 27.69
0 | 0 | 254.8 | 28.04 | 69317.7 | 27.86
0 | 0 | 253.6 | 28.19 | 69703.5 | 28.04
0 | 0 | 252.4 | 28.45 | 70348.8 | 28.25
0 | 0 | 251.2 | 28.64 | 70815.2 | 28.46
0 | 0 | 250 | 28.82 | 71244.8 | 28.62
0 | 0 | 248.8 | 29.06 | 71852.1 | 28.86
0 | 0 | 247.6 | 29.19 | 72175.6 | 29.03
0 | 0 | 246.4 | 29.43 | 72757.2 | 29.3
0 | 0 | 245.1 | 29.65 | 73312.1 | 29.45
0 | 0 | 244 | 29.9 | 73917.3 | 29.71
0 | 0 | 242.7 | 30.03 | 74247.9 | 29.82
0 | 0 | 241.6 | 30.24 | 74771.8 | 30.06
0 | 0 | 240.3 | 30.32 | 74956.5 | 30.08
0 | 0 | 239.1 | 30.57 | 75581.7 | 30.35
0 | 0 | 237.9 | 30.83 | 76219.8 | 30.63
0 | 0 | 236.7 | 31.01 | 76674 | 30.79
0 | 0 | 235.5 | 31.17 | 77051.9 | 30.99
0 | 0 | 234.3 | 31.43 | 77707.8 | 31.23
0 | 0 | 233.1 | 31.52 | 77936.1 | 31.35
0 | 0 | 231.9 | 31.82 | 78671.4 | 31.59
0 | 0 | 230.7 | 32.12 | 79421.2 | 31.9
0 | 0 | 229.5 | 32.33 | 79929.1 | 32.1
0 | 0 | 228.3 | 32.5 | 80362.2 | 32.32
0 | 0 | 227.1 | 32.75 | 80967.7 | 32.54
0 | 0 | 225.9 | 32.97 | 81509.4 | 32.75
0 | 0 | 224.7 | 33.21 | 82105.1 | 32.97
0 | 0 | 223.5 | 33.45 | 82702.7 | 33.24
0 | 0 | 222.2 | 33.64 | 83165.5 | 33.38
0 | 0 | 221 | 33.91 | 83846.6 | 33.69
0 | 0 | 219.9 | 34.1 | 84313.7 | 33.85
0 | 0 | 218.6 | 34.31 | 84814.5 | 34.1
0 | 0 | 217.5 | 34.65 | 85665 | 34.4
0 | 0 | 216.2 | 34.78 | 85995 | 34.58
0 | 0 | 215 | 35.02 | 86580.8 | 34.76
0 | 0 | 213.8 | 35.24 | 87132.5 | 34.98
0 | 0 | 212.6 | 35.45 | 87656.8 | 35.26
0 | 0 | 211.4 | 35.79 | 88497.1 | 35.59
0 | 0 | 210.2 | 35.96 | 88911.9 | 35.76
0 | 0 | 209 | 36.22 | 89556.4 | 35.98
0 | 0 | 207.8 | 36.44 | 90090.1 | 36.21
0 | 0 | 206.6 | 36.83 | 91054.7 | 36.49
0 | 0 | 205.4 | 37 | 91482 | 36.81
0 | 0 | 204.2 | 37.19 | 91958.6 | 36.95
0 | 0 | 203 | 37.47 | 92627.4 | 37.21
0 | 0 | 201.8 | 37.77 | 93378.5 | 37.5 | D20 | Z | D
0 | 0 | 200.6 | 38.03 | 94030.3 | 37.81 |  | 200.6 | 37.81
0 | 0 | 199.4 | 38.27 | 94618.1 | 38.03 |  | 199.4 | 38.03
0 | 0 | 198.2 | 38.55 | 95320.8 | 38.28 |  | 200 | 37.92
0 | 0 | 197 | 38.79 | 95913.9 | 38.59
0 | 0 | 195.8 | 39.1 | 96666.5 | 38.8
0 | 0 | 194.6 | 39.25 | 97033 | 38.94
0 | 0 | 193.4 | 39.64 | 98008.2 | 39.35
0 | 0 | 192.2 | 39.93 | 98723.7 | 39.68
0 | 0 | 191 | 40.15 | 99276.6 | 39.92
0 | 0 | 189.8 | 40.52 | 100176.2 | 40.22
0 | 0 | 188.5 | 40.73 | 100696.1 | 40.48
0 | 0 | 187.4 | 41.03 | 101429 | 40.76
0 | 0 | 186.1 | 41.28 | 102046.6 | 40.98
0 | 0 | 184.9 | 41.5 | 102609.4 | 41.24
0 | 0 | 183.7 | 41.87 | 103520.2 | 41.53
0 | 0 | 182.5 | 42.14 | 104174.2 | 41.87
0 | 0 | 181.3 | 42.38 | 104788.9 | 42.1
0 | 0 | 180.1 | 42.66 | 105475.2 | 42.38
0 | 0 | 178.9 | 42.95 | 106175.7 | 42.69
0 | 0 | 177.7 | 43.3 | 107043.9 | 43.02
0 | 0 | 176.5 | 43.56 | 107694.7 | 43.3
0 | 0 | 175.3 | 43.95 | 108666.8 | 43.64
0 | 0 | 174.1 | 44.18 | 109239.1 | 43.86
0 | 0 | 172.9 | 44.42 | 109830.9 | 44.13
0 | 0 | 171.7 | 44.82 | 110804.1 | 44.5
0 | 0 | 170.5 | 45.09 | 111469.7 | 44.85
0 | 0 | 169.3 | 45.38 | 112194.3 | 45.07
0 | 0 | 168.1 | 45.79 | 113207.6 | 45.41
0 | 0 | 166.9 | 46.02 | 113773.1 | 45.71
0 | 0 | 165.7 | 46.26 | 114381.3 | 45.98
0 | 0 | 164.4 | 46.66 | 115368.9 | 46.4
0 | 0 | 163.3 | 46.97 | 116120.2 | 46.7
0 | 0 | 162 | 47.22 | 116749.8 | 46.99
0 | 0 | 160.9 | 47.61 | 117715.4 | 47.29
0 | 0 | 159.6 | 47.99 | 118644.5 | 47.77
0 | 0 | 158.4 | 48.28 | 119363.4 | 47.94
0 | 0 | 157.2 | 48.55 | 120029.6 | 48.28
0 | 0 | 156 | 48.92 | 120939.2 | 48.58
0 | 0 | 154.8 | 49.25 | 121772.5 | 48.95
0 | 0 | 153.6 | 49.62 | 122682.4 | 49.35
0 | 0 | 152.4 | 50.03 | 123703.9 | 49.6 | R50 | Z | D
0 | 0 | 151.2 | 50.22 | 124168.6 | 49.91 |  | 151.2 | 49.91
0 | 0 | 150 | 50.46 | 124759.1 | 50.18 |  | 150 | 50.18
0 | 0 | 148.8 | 50.87 | 125765.6 | 50.56 |  | 150.8 | 50
0 | 0 | 147.6 | 51.25 | 126705.6 | 50.86
0 | 0 | 146.4 | 51.59 | 127539 | 51.15
0 | 0 | 145.2 | 51.97 | 128493.7 | 51.62
0 | 0 | 144 | 52.33 | 129378.5 | 52.01
0 | 0 | 142.8 | 52.72 | 130336.2 | 52.47
0 | 0 | 141.6 | 53.01 | 131055.7 | 52.71
0 | 0 | 140.4 | 53.4 | 132017.8 | 52.97
0 | 0 | 139.2 | 53.66 | 132656.7 | 53.35
0 | 0 | 138 | 54.05 | 133641 | 53.65
0 | 0 | 136.8 | 54.43 | 134578.9 | 54.14
0 | 0 | 135.6 | 54.75 | 135369.7 | 54.45
0 | 0 | 134.4 | 55.14 | 136318.4 | 54.76
0 | 0 | 133.2 | 55.49 | 137188.5 | 55.1
0 | 0 | 131.9 | 55.85 | 138072 | 55.48
0 | 0 | 130.8 | 56.25 | 139075.6 | 55.89
0 | 0 | 129.6 | 56.71 | 140213 | 56.4
0 | 0 | 128.3 | 57.02 | 140961.3 | 56.57
0 | 0 | 127.1 | 57.39 | 141899.8 | 56.96
0 | 0 | 126 | 57.87 | 143076.8 | 57.42
0 | 0 | 124.7 | 58.23 | 143964.5 | 57.84
0 | 0 | 123.5 | 58.57 | 144802.3 | 58.19
0 | 0 | 122.3 | 59.13 | 146193.7 | 58.73
0 | 0 | 121.1 | 59.48 | 147066 | 59.09
0 | 0 | 119.9 | 59.85 | 147964.3 | 59.48
0 | 0 | 118.7 | 60.24 | 148934.6 | 59.89
0 | 0 | 117.5 | 60.6 | 149836.2 | 60.29
0 | 0 | 116.3 | 60.88 | 150522.4 | 60.55
0 | 0 | 115.1 | 61.27 | 151475.4 | 60.89
0 | 0 | 113.9 | 61.79 | 152758.1 | 61.33
0 | 0 | 112.7 | 62.09 | 153506.7 | 61.68
0 | 0 | 111.5 | 62.54 | 154614.8 | 62.13
0 | 0 | 110.3 | 63.01 | 155792.1 | 62.56
0 | 0 | 109.1 | 63.51 | 157019.6 | 63
0 | 0 | 107.9 | 64.02 | 158289.7 | 63.56
0 | 0 | 106.7 | 64.51 | 159496.6 | 63.84
0 | 0 | 105.5 | 64.81 | 160233.8 | 64.36
0 | 0 | 104.2 | 65.18 | 161158.2 | 64.88
0 | 0 | 103.1 | 65.66 | 162345.9 | 65.19
0 | 0 | 101.9 | 66.12 | 163470.3 | 65.63 | D10 | Z | D
0 | 0 | 100.7 | 66.56 | 164553.7 | 66.23 |  | 100.7 | 66.23
0 | 0 | 99.5 | 66.8 | 165151.1 | 66.43 |  | 99.5 | 66.43
0 | 0 | 98.3 | 67.23 | 166216.7 | 66.84 |  | 100 | 66.3467
0 | 0 | 97.1 | 67.74 | 167469.8 | 67.24
0 | 0 | 95.8 | 68.06 | 168266.5 | 67.73
0 | 0 | 94.7 | 68.39 | 169076.6 | 68.15
0 | 0 | 93.5 | 68.97 | 170521.4 | 68.6
0 | 0 | 92.2 | 69.61 | 172091.4 | 68.98
0 | 0 | 91.1 | 69.94 | 172923.4 | 69.54
0 | 0 | 89.8 | 70.5 | 174294.8 | 70.01
0 | 0 | 88.6 | 71.05 | 175650.2 | 70.44
0 | 0 | 87.4 | 71.54 | 176880.2 | 71.06
0 | 0 | 86.2 | 71.94 | 177867.6 | 71.55
0 | 0 | 85 | 72.45 | 179111.9 | 71.91
0 | 0 | 83.8 | 72.91 | 180252.2 | 72.35
0 | 0 | 82.6 | 73.37 | 181404 | 72.9
0 | 0 | 81.4 | 73.82 | 182515.6 | 73.36
0 | 0 | 80.2 | 74.32 | 183742 | 73.79
0 | 0 | 79 | 74.68 | 184629.3 | 74.18
0 | 0 | 77.8 | 75.26 | 186064.7 | 74.7
0 | 0 | 76.6 | 75.69 | 187144.5 | 75.09
0 | 0 | 75.4 | 75.94 | 187756.6 | 75.41
0 | 0 | 74.1 | 76.42 | 188938 | 75.96
0 | 0 | 73 | 76.9 | 190125.5 | 76.55
0 | 0 | 71.7 | 77.54 | 191710 | 77.05
0 | 0 | 70.5 | 78.1 | 193090.9 | 77.55
0 | 0 | 69.3 | 78.72 | 194624 | 77.98
0 | 0 | 68.1 | 79.04 | 195410 | 78.39
0 | 0 | 66.9 | 79.59 | 196786.3 | 79.02
0 | 0 | 65.7 | 80.07 | 197953.7 | 79.48
0 | 0 | 64.5 | 80.58 | 199215.5 | 79.89
0 | 0 | 63.3 | 81.05 | 200396.4 | 80.58
0 | 0 | 62 | 81.7 | 201999.9 | 81.11
0 | 0 | 60.9 | 82.02 | 202779.4 | 81.45
0 | 0 | 59.6 | 82.47 | 203895.1 | 82.01
0 | 0 | 58.5 | 83.17 | 205630.1 | 82.51
0 | 0 | 57.2 | 83.78 | 207122.8 | 83.14
0 | 0 | 56 | 84.05 | 207801.8 | 83.42
0 | 0 | 54.8 | 84.53 | 208993.7 | 83.93
0 | 0 | 53.6 | 84.95 | 210038.5 | 84.43
0 | 0 | 52.4 | 85.63 | 211702.6 | 85.05
0 | 0 | 51.2 | 86.21 | 213142.3 | 85.4
0 | 0 | 50 | 86.73 | 214435.1 | 86.13
0 | 0 | 48.8 | 87.06 | 215249.2 | 86.55
0 | 0 | 47.6 | 87.57 | 216492.8 | 86.98
0 | 0 | 46.4 | 88.17 | 217993.2 | 87.38
0 | 0 | 45.2 | 88.51 | 218825 | 87.97
0 | 0 | 44 | 89.08 | 220237.2 | 88.49
0 | 0 | 42.8 | 89.42 | 221077.5 | 88.88
0 | 0 | 41.6 | 89.98 | 222465.5 | 89.46
0 | 0 | 40.4 | 90.66 | 224139.6 | 89.96
0 | 0 | 39.1 | 91.34 | 225820.5 | 90.54
0 | 0 | 38 | 91.63 | 226542.2 | 90.99
0 | 0 | 36.7 | 92.09 | 227684.5 | 91.53
0 | 0 | 35.5 | 92.62 | 228982 | 91.93
0 | 0 | 34.3 | 93.2 | 230429.1 | 92.41
0 | 0 | 33.1 | 93.86 | 232066.6 | 93
0 | 0 | 31.9 | 94.45 | 233507.6 | 93.77
0 | 0 | 30.7 | 94.91 | 234651 | 94.29
0 | 0 | 29.5 | 95.33 | 235691.5 | 94.68
0 | 0 | 28.3 | 95.98 | 237296.9 | 95.29
0 | 0 | 27.1 | 96.25 | 237958.1 | 95.63
0 | 0 | 25.9 | 96.81 | 239336 | 96.15
0 | 0 | 24.7 | 97.32 | 240618.6 | 96.6
0 | 0 | 23.5 | 98.01 | 242304.1 | 97.24
0 | 0 | 22.3 | 98.33 | 243106.8 | 97.82
0 | 0 | 21.1 | 99 | 244773.2 | 98.39
0 | 0 | 19.8 | 99.28 | 245457.3 | 98.79
0 | 0 | 18.6 | 99.83 | 246816.3 | 99.35
0 | 0 | 17.4 | 99.99 | 247202.3 | 99.27
0 | 0 | 16.2 | 100.74 | 249072.3 | 99.76
0 | 0 | 15 | 100.58 | 248663 | 99.79
0 | 0 | 13.8 | 100.78 | 249172.3 | 99.99
0 | 0 | 12.6 | 100.54 | 248559 | 100
0 | 0 | 11.4 | 100.04 | 247344.5 | 99.33
0 | 0 | 10.2 | 99.02 | 244807.6 | 98.17
0 | 0 | 9 | 97.79 | 241773.5 | 97.12
0 | 0 | 7.8 | 95.33 | 235679.2 | 94.65
0 | 0 | 6.6 | 91.96 | 227360.3 | 91.27
0 | 0 | 5.4 | 86.89 | 214830.9 | 86.3
0 | 0 | 4.2 | 79.72 | 197103.6 | 79.19
0 | 0 | 3 | 69.42 | 171637.6 | 68.95
0 | 0 | 1.7 | 59.81 | 147860.3 | 59.41
0 | 0 | 0.6 | 54.3 | 134255.9 | 53.92
0 | 0 | -0.5 | 50.98 | 126046.2 | 50.65
Measurement time: | 12/08/2022 20:27:53
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | FC 65-P
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 43.65 | 265700.6 | 20.53
0 | 0 | 309.1 | 46.61 | 283690.7 | 20.64
0 | 0 | 307.6 | 47.15 | 287015.7 | 20.85
0 | 0 | 306.5 | 47.41 | 288596.1 | 20.97
0 | 0 | 305.3 | 47.77 | 290762.1 | 21.11
0 | 0 | 304.1 | 48.05 | 292479.6 | 21.27
0 | 0 | 302.9 | 48.39 | 294521.6 | 21.43
0 | 0 | 301.7 | 48.69 | 296376.8 | 21.56
0 | 0 | 300.5 | 49.05 | 298552.2 | 21.72
0 | 0 | 299.3 | 49.28 | 299976.2 | 21.83
0 | 0 | 298.1 | 49.62 | 302006.1 | 21.99
0 | 0 | 296.9 | 49.9 | 303750.5 | 22.09
0 | 0 | 295.7 | 50.3 | 306185.6 | 22.26
0 | 0 | 294.5 | 50.7 | 308577 | 22.45
0 | 0 | 293.3 | 51.07 | 310865.9 | 22.58
0 | 0 | 292 | 51.35 | 312524.1 | 22.72
0 | 0 | 290.9 | 51.56 | 313804.6 | 22.83
0 | 0 | 289.7 | 52.04 | 316739.2 | 23.02
0 | 0 | 288.5 | 52.49 | 319478.5 | 23.22
0 | 0 | 287.3 | 52.74 | 320982.1 | 23.34
0 | 0 | 286 | 53.14 | 323434.7 | 23.5
0 | 0 | 284.9 | 53.46 | 325418.8 | 23.65
0 | 0 | 283.7 | 53.93 | 328267.4 | 23.88
0 | 0 | 282.4 | 54.24 | 330160.8 | 24.02
0 | 0 | 281.3 | 54.53 | 331911.8 | 24.15
0 | 0 | 280 | 55.02 | 334898.7 | 24.36
0 | 0 | 278.8 | 55.25 | 336311.2 | 24.46
0 | 0 | 277.6 | 55.71 | 339103.8 | 24.62
0 | 0 | 276.4 | 56.14 | 341692.6 | 24.83
0 | 0 | 275.2 | 56.5 | 343882.1 | 25.04
0 | 0 | 274 | 56.9 | 346360.9 | 25.15
0 | 0 | 272.8 | 57.1 | 347563.9 | 25.31
0 | 0 | 271.6 | 57.65 | 350904.1 | 25.49
0 | 0 | 270.4 | 58.03 | 353201.3 | 25.65
0 | 0 | 269.2 | 58.32 | 355000.3 | 25.83
0 | 0 | 268 | 58.77 | 357707.7 | 26.04
0 | 0 | 266.8 | 59.11 | 359804.6 | 26.16
0 | 0 | 265.6 | 59.56 | 362536.8 | 26.36
0 | 0 | 264.4 | 59.88 | 364442.8 | 26.51
0 | 0 | 263.2 | 60.2 | 366426.3 | 26.66
0 | 0 | 262 | 60.69 | 369376.8 | 26.86
0 | 0 | 260.8 | 61.13 | 372073.5 | 27.03
0 | 0 | 259.6 | 61.67 | 375359.7 | 27.27
0 | 0 | 258.4 | 62.05 | 377663.2 | 27.47
0 | 0 | 257.2 | 62.36 | 379542.3 | 27.62
0 | 0 | 256 | 62.86 | 382625.2 | 27.8
0 | 0 | 254.8 | 63.31 | 385322.2 | 28.01
0 | 0 | 253.6 | 63.68 | 387570.7 | 28.2
0 | 0 | 252.4 | 64.2 | 390789.1 | 28.39
0 | 0 | 251.2 | 64.51 | 392671.7 | 28.54
0 | 0 | 250 | 64.99 | 395579.1 | 28.76
0 | 0 | 248.8 | 65.34 | 397682.9 | 28.95
0 | 0 | 247.6 | 65.78 | 400364.7 | 29.15
0 | 0 | 246.3 | 66.41 | 404202.3 | 29.36
0 | 0 | 245.1 | 66.74 | 406200.9 | 29.55
0 | 0 | 243.9 | 67.27 | 409470.8 | 29.71
0 | 0 | 242.8 | 67.71 | 412125.9 | 29.98
0 | 0 | 241.5 | 68.16 | 414872 | 30.19
0 | 0 | 240.3 | 68.7 | 418185.4 | 30.36
0 | 0 | 239.1 | 69.09 | 420539.1 | 30.56
0 | 0 | 237.9 | 69.54 | 423272.9 | 30.8
0 | 0 | 236.7 | 70.01 | 426116.9 | 31
0 | 0 | 235.5 | 70.4 | 428494.6 | 31.21
0 | 0 | 234.3 | 70.99 | 432075.9 | 31.43
0 | 0 | 233.1 | 71.48 | 435080.9 | 31.65
0 | 0 | 231.9 | 72.05 | 438557.9 | 31.88
0 | 0 | 230.7 | 72.37 | 440521.4 | 32.05
0 | 0 | 229.5 | 72.78 | 443018.6 | 32.25
0 | 0 | 228.3 | 73.47 | 447185.1 | 32.54
0 | 0 | 227.1 | 73.87 | 449602.9 | 32.68
0 | 0 | 225.9 | 74.35 | 452540.3 | 32.93
0 | 0 | 224.7 | 74.85 | 455605.6 | 33.12
0 | 0 | 223.4 | 75.41 | 458970.6 | 33.4
0 | 0 | 222.3 | 75.98 | 462462 | 33.62
0 | 0 | 221.1 | 76.5 | 465609.1 | 33.88
0 | 0 | 219.8 | 77.11 | 469318.1 | 34.12
0 | 0 | 218.7 | 77.41 | 471166.7 | 34.33
0 | 0 | 217.4 | 77.97 | 474552.1 | 34.59
0 | 0 | 216.2 | 78.67 | 478815.8 | 34.81
0 | 0 | 215 | 79.13 | 481658.9 | 35.09
0 | 0 | 213.8 | 79.7 | 485108.1 | 35.31
0 | 0 | 212.6 | 80.07 | 487381.8 | 35.47
0 | 0 | 211.4 | 80.77 | 491639.2 | 35.76
0 | 0 | 210.2 | 81.26 | 494629.8 | 35.98
0 | 0 | 209 | 81.79 | 497834.6 | 36.24
0 | 0 | 207.8 | 82.42 | 501681.7 | 36.5
0 | 0 | 206.6 | 82.98 | 505089.4 | 36.73
0 | 0 | 205.4 | 83.44 | 507897.8 | 37
0 | 0 | 204.2 | 84.14 | 512135 | 37.23
0 | 0 | 203 | 84.6 | 514908.1 | 37.4
0 | 0 | 201.8 | 85.3 | 519175.8 | 37.73 | D20 | Z | D
0 | 0 | 200.6 | 85.8 | 522247.4 | 38.04 |  | 200.6 | 38.04
0 | 0 | 199.4 | 86.37 | 525698.5 | 38.23 |  | 199.4 | 38.23
0 | 0 | 198.2 | 87.13 | 530342.8 | 38.52 |  | 200 | 38.135
0 | 0 | 197 | 87.51 | 532665.2 | 38.74
0 | 0 | 195.7 | 88.11 | 536274.9 | 38.96
0 | 0 | 194.5 | 88.64 | 539543.3 | 39.25
0 | 0 | 193.3 | 89.37 | 543941 | 39.56
0 | 0 | 192.1 | 89.77 | 546397.5 | 39.8
0 | 0 | 191 | 90.63 | 551611.4 | 40.1
0 | 0 | 189.7 | 91.11 | 554558.5 | 40.36
0 | 0 | 188.6 | 91.73 | 558350 | 40.64
0 | 0 | 187.3 | 92.35 | 562088.8 | 40.87
0 | 0 | 186.1 | 93.06 | 566424.1 | 41.16
0 | 0 | 185 | 93.72 | 570434 | 41.51
0 | 0 | 183.7 | 94.36 | 574323.1 | 41.77
0 | 0 | 182.5 | 94.76 | 576802.8 | 41.99
0 | 0 | 181.3 | 95.57 | 581677.8 | 42.36
0 | 0 | 180.1 | 96.4 | 586727.3 | 42.66
0 | 0 | 179 | 97.03 | 590563.3 | 42.92
0 | 0 | 177.7 | 97.73 | 594877.4 | 43.3
0 | 0 | 176.5 | 98.22 | 597823.2 | 43.49
0 | 0 | 175.3 | 98.9 | 601987.3 | 43.78
0 | 0 | 174.1 | 99.63 | 606385.8 | 44.09
0 | 0 | 172.9 | 100.25 | 610211.4 | 44.41
0 | 0 | 171.7 | 100.87 | 613985.9 | 44.64
0 | 0 | 170.5 | 101.6 | 618384.1 | 45.04
0 | 0 | 169.3 | 102.39 | 623232.5 | 45.36
0 | 0 | 168.1 | 103.2 | 628156.5 | 45.61
0 | 0 | 166.9 | 103.95 | 632728.3 | 45.93
0 | 0 | 165.7 | 104.56 | 636424.6 | 46.31
0 | 0 | 164.5 | 104.97 | 638934.2 | 46.45
0 | 0 | 163.3 | 105.83 | 644135.2 | 46.87
0 | 0 | 162.1 | 106.66 | 649182.4 | 47.22
0 | 0 | 160.9 | 107.27 | 652889.8 | 47.47
0 | 0 | 159.7 | 107.83 | 656357.3 | 47.71
0 | 0 | 158.5 | 108.39 | 659706.2 | 48.09
0 | 0 | 157.3 | 109.18 | 664555.5 | 48.42
0 | 0 | 156.1 | 110.26 | 671091.5 | 48.81
0 | 0 | 154.9 | 111.15 | 676509.4 | 49.14
0 | 0 | 153.7 | 112.11 | 682347.8 | 49.54 | R50 | Z | D
0 | 0 | 152.5 | 112.63 | 685568.3 | 49.85 |  | 152.5 | 49.85
0 | 0 | 151.2 | 113.43 | 690412.1 | 50.21 |  | 151.2 | 50.21
0 | 0 | 150.1 | 114.19 | 695034.2 | 50.54 |  | 151.9583 | 50
0 | 0 | 148.8 | 115.03 | 700162.1 | 50.89
0 | 0 | 147.6 | 115.87 | 705241.7 | 51.25
0 | 0 | 146.4 | 116.5 | 709100.8 | 51.52
0 | 0 | 145.2 | 117.25 | 713693.2 | 51.95
0 | 0 | 144 | 118.2 | 719458.5 | 52.29
0 | 0 | 142.8 | 118.85 | 723392.8 | 52.6
0 | 0 | 141.6 | 119.63 | 728131.5 | 52.94
0 | 0 | 140.4 | 120.42 | 732954.9 | 53.33
0 | 0 | 139.2 | 121.23 | 737903.4 | 53.7
0 | 0 | 138 | 122.19 | 743743.3 | 54.06
0 | 0 | 136.8 | 122.85 | 747759.7 | 54.35
0 | 0 | 135.6 | 123.6 | 752308.3 | 54.74
0 | 0 | 134.4 | 124.14 | 755626.9 | 55.01
0 | 0 | 133.2 | 125.34 | 762925.1 | 55.48
0 | 0 | 132 | 126.3 | 768756.2 | 55.88
0 | 0 | 130.8 | 127.21 | 774257.1 | 56.3
0 | 0 | 129.6 | 128.11 | 779773.1 | 56.72
0 | 0 | 128.4 | 128.92 | 784724.3 | 57.14
0 | 0 | 127.2 | 129.9 | 790666.1 | 57.54
0 | 0 | 126 | 130.57 | 794747.8 | 57.87
0 | 0 | 124.8 | 131.69 | 801528.1 | 58.31
0 | 0 | 123.6 | 132.39 | 805821.7 | 58.57
0 | 0 | 122.4 | 133.16 | 810503.7 | 58.98
0 | 0 | 121.2 | 134.13 | 816409.9 | 59.37
0 | 0 | 120 | 134.77 | 820315.7 | 59.7
0 | 0 | 118.8 | 135.55 | 825068.6 | 60.1
0 | 0 | 117.6 | 136.8 | 832644.1 | 60.71
0 | 0 | 116.4 | 137.78 | 838600.8 | 60.96
0 | 0 | 115.2 | 138.85 | 845116.9 | 61.41
0 | 0 | 113.9 | 139.71 | 850393.4 | 61.82
0 | 0 | 112.7 | 140.56 | 855546.7 | 62.21
0 | 0 | 111.5 | 141.52 | 861374.6 | 62.71
0 | 0 | 110.3 | 142.58 | 867834.4 | 63.13
0 | 0 | 109.1 | 143.24 | 871831.6 | 63.46
0 | 0 | 107.9 | 143.86 | 875657.7 | 63.67
0 | 0 | 106.7 | 145.1 | 883173.5 | 64.25
0 | 0 | 105.5 | 146.01 | 888735.8 | 64.68
0 | 0 | 104.3 | 147.08 | 895199.8 | 65.06
0 | 0 | 103.1 | 148.39 | 903186.1 | 65.64
0 | 0 | 101.9 | 149.16 | 907915.1 | 66 | D10 | Z | D
0 | 0 | 100.7 | 150.39 | 915393.9 | 66.54 |  | 100.7 | 66.54
0 | 0 | 99.5 | 151 | 919093.5 | 66.87 |  | 99.5 | 66.87
0 | 0 | 98.3 | 152.03 | 925375.3 | 67.29 |  | 100 | 66.7325
0 | 0 | 97.1 | 153.23 | 932658 | 67.91
0 | 0 | 95.9 | 154.37 | 939631.6 | 68.36
0 | 0 | 94.7 | 155.31 | 945296.6 | 68.66
0 | 0 | 93.5 | 156.24 | 950970.5 | 69.17
0 | 0 | 92.3 | 157.28 | 957316 | 69.57
0 | 0 | 91.1 | 158.18 | 962795.8 | 70.12
0 | 0 | 89.9 | 159.39 | 970132.5 | 70.58
0 | 0 | 88.7 | 160.39 | 976266.9 | 70.85
0 | 0 | 87.4 | 161.48 | 982880.4 | 71.5
0 | 0 | 86.3 | 162.4 | 988460.1 | 71.89
0 | 0 | 85 | 163.35 | 994233.2 | 72.34
0 | 0 | 83.8 | 164.5 | 1001285.7 | 72.78
0 | 0 | 82.6 | 165.65 | 1008229.6 | 73.34
0 | 0 | 81.4 | 166.68 | 1014555 | 73.7
0 | 0 | 80.2 | 167.6 | 1020146.2 | 74.24
0 | 0 | 79 | 168.65 | 1026545.3 | 74.64
0 | 0 | 77.8 | 169.76 | 1033298.2 | 75.15
0 | 0 | 76.6 | 170.89 | 1040147.8 | 75.6
0 | 0 | 75.4 | 171.87 | 1046127.7 | 76.08
0 | 0 | 74.2 | 173.17 | 1054041.5 | 76.69
0 | 0 | 73 | 173.99 | 1059003.5 | 77.09
0 | 0 | 71.8 | 175.39 | 1067549.7 | 77.64
0 | 0 | 70.6 | 176.29 | 1073036.3 | 78.03
0 | 0 | 69.3 | 177.33 | 1079331.1 | 78.58
0 | 0 | 68.2 | 178.44 | 1086087.9 | 78.97
0 | 0 | 66.9 | 179.76 | 1094156.4 | 79.67
0 | 0 | 65.7 | 180.96 | 1101471.5 | 80.11
0 | 0 | 64.5 | 181.99 | 1107730.9 | 80.45
0 | 0 | 63.3 | 183.04 | 1114095.3 | 80.83
0 | 0 | 62.1 | 184.14 | 1120778.3 | 81.41
0 | 0 | 60.9 | 185.37 | 1128278.6 | 82.03
0 | 0 | 59.7 | 186.39 | 1134491.5 | 82.41
0 | 0 | 58.5 | 187.66 | 1142202 | 82.99
0 | 0 | 57.3 | 188.81 | 1149245.5 | 83.56
0 | 0 | 56.1 | 189.86 | 1155647.3 | 84.06
0 | 0 | 54.8 | 190.96 | 1162286.4 | 84.56
0 | 0 | 53.7 | 192.23 | 1170064.8 | 85.07
0 | 0 | 52.4 | 193.35 | 1176836.8 | 85.67
0 | 0 | 51.2 | 194.6 | 1184451.7 | 86.2
0 | 0 | 50 | 195.55 | 1190267.4 | 86.62
0 | 0 | 48.8 | 196.99 | 1198986.7 | 87.26
0 | 0 | 47.6 | 198.22 | 1206520.7 | 87.67
0 | 0 | 46.4 | 199.17 | 1212262.1 | 88.18
0 | 0 | 45.2 | 200.52 | 1220507.8 | 88.65
0 | 0 | 44 | 201.4 | 1225855.2 | 89.05
0 | 0 | 42.8 | 202.81 | 1234442.8 | 89.64
0 | 0 | 41.6 | 203.84 | 1240715.3 | 90.15
0 | 0 | 40.3 | 204.81 | 1246621 | 90.58
0 | 0 | 39.1 | 206.11 | 1254525.9 | 91.14
0 | 0 | 37.9 | 207.04 | 1260159.4 | 91.61
0 | 0 | 36.7 | 208.36 | 1268228.6 | 92.35
0 | 0 | 35.5 | 209.61 | 1275840.2 | 92.75
0 | 0 | 34.3 | 210.67 | 1282267.9 | 93.21
0 | 0 | 33.1 | 211.98 | 1290269.3 | 93.72
0 | 0 | 31.9 | 213.02 | 1296559.7 | 94.28
0 | 0 | 30.7 | 214.18 | 1303645.8 | 94.78
0 | 0 | 29.5 | 215.16 | 1309604.7 | 95.2
0 | 0 | 28.3 | 216.54 | 1317996.7 | 95.73
0 | 0 | 27 | 217.61 | 1324502.9 | 96.32
0 | 0 | 25.9 | 218.68 | 1331020.6 | 96.93
0 | 0 | 24.6 | 220.18 | 1340191.7 | 97.32
0 | 0 | 23.4 | 221.26 | 1346741.6 | 97.82
0 | 0 | 22.3 | 222.11 | 1351914.2 | 98.37
0 | 0 | 21 | 222.97 | 1357177.6 | 98.69
0 | 0 | 19.8 | 223.94 | 1363073.6 | 99.24
0 | 0 | 18.6 | 224.79 | 1368201.8 | 99.45
0 | 0 | 17.4 | 225.63 | 1373348.1 | 99.9
0 | 0 | 16.2 | 226.15 | 1376494.6 | 99.97
0 | 0 | 15 | 225.96 | 1375323.3 | 100
0 | 0 | 13.8 | 225.15 | 1370431.4 | 99.63
0 | 0 | 12.6 | 223.21 | 1358636.9 | 98.91
0 | 0 | 11.4 | 221.67 | 1349205.4 | 98.08
0 | 0 | 10.2 | 217.51 | 1323928 | 96.32
0 | 0 | 9 | 211.99 | 1290304.4 | 93.75
0 | 0 | 7.8 | 203.3 | 1237430 | 89.96
0 | 0 | 6.5 | 190.19 | 1157604.4 | 84.14
0 | 0 | 5.4 | 171.35 | 1042958.7 | 75.83
0 | 0 | 4.2 | 144.24 | 877952.6 | 63.81
0 | 0 | 2.9 | 123.51 | 751769.5 | 54.68
0 | 0 | 1.7 | 113.35 | 689949.4 | 50.24
0 | 0 | 0.5 | 106.61 | 648896.4 | 47.15
0 | 0 | -0.5 | 102.77 | 625541.5 | 45.53