## 12082022 - RDM énergie_C2
Clinac 2
Measurement time: | 12/08/2022 18:00:16
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 19.24 | 47561.7 | 20.48
0 | 0 | 309.1 | 20.74 | 51265.8 | 20.6
0 | 0 | 307.6 | 20.83 | 51504.6 | 20.67
0 | 0 | 306.6 | 21.02 | 51974.8 | 20.89
0 | 0 | 305.3 | 21.15 | 52292.6 | 21.01
0 | 0 | 304.1 | 21.35 | 52782.9 | 21.19
0 | 0 | 302.9 | 21.46 | 53066 | 21.33
0 | 0 | 301.7 | 21.62 | 53451.6 | 21.46
0 | 0 | 300.5 | 21.73 | 53713.5 | 21.55
0 | 0 | 299.3 | 21.83 | 53983.6 | 21.68
0 | 0 | 298.1 | 22.08 | 54597.4 | 21.97
0 | 0 | 296.9 | 22.18 | 54836.5 | 22.02
0 | 0 | 295.7 | 22.34 | 55238 | 22.21
0 | 0 | 294.5 | 22.52 | 55683.7 | 22.35
0 | 0 | 293.3 | 22.64 | 55983.6 | 22.46
0 | 0 | 292.1 | 22.78 | 56332 | 22.67
0 | 0 | 290.9 | 22.95 | 56744.8 | 22.82
0 | 0 | 289.7 | 23.07 | 57049 | 22.94
0 | 0 | 288.5 | 23.29 | 57574 | 23.13
0 | 0 | 287.3 | 23.45 | 57968 | 23.31
0 | 0 | 286.1 | 23.53 | 58174.1 | 23.42
0 | 0 | 284.9 | 23.7 | 58600 | 23.54
0 | 0 | 283.7 | 23.93 | 59151.6 | 23.76
0 | 0 | 282.5 | 24.01 | 59365.5 | 23.88
0 | 0 | 281.2 | 24.25 | 59960.5 | 24.07
0 | 0 | 280.1 | 24.39 | 60289 | 24.23
0 | 0 | 278.8 | 24.46 | 60477.1 | 24.31
0 | 0 | 277.7 | 24.68 | 61005.5 | 24.5
0 | 0 | 276.4 | 24.86 | 61469.1 | 24.72
0 | 0 | 275.2 | 25.02 | 61854 | 24.83
0 | 0 | 274 | 25.22 | 62360 | 25.07
0 | 0 | 272.8 | 25.32 | 62610.7 | 25.19
0 | 0 | 271.6 | 25.54 | 63143.9 | 25.37
0 | 0 | 270.4 | 25.66 | 63445.9 | 25.55
0 | 0 | 269.2 | 25.84 | 63885.9 | 25.73
0 | 0 | 268 | 26.05 | 64400.7 | 25.89
0 | 0 | 266.8 | 26.3 | 65024.5 | 26.08
0 | 0 | 265.6 | 26.42 | 65327.7 | 26.21
0 | 0 | 264.4 | 26.53 | 65584.6 | 26.36
0 | 0 | 263.2 | 26.75 | 66146.4 | 26.61
0 | 0 | 262 | 26.97 | 66670.8 | 26.79
0 | 0 | 260.8 | 27.05 | 66881.8 | 26.88
0 | 0 | 259.6 | 27.34 | 67602.7 | 27.15
0 | 0 | 258.4 | 27.47 | 67924.8 | 27.26
0 | 0 | 257.2 | 27.65 | 68368.3 | 27.45
0 | 0 | 256 | 27.85 | 68849.4 | 27.69
0 | 0 | 254.8 | 28.04 | 69317.7 | 27.86
0 | 0 | 253.6 | 28.19 | 69703.5 | 28.04
0 | 0 | 252.4 | 28.45 | 70348.8 | 28.25
0 | 0 | 251.2 | 28.64 | 70815.2 | 28.46
0 | 0 | 250 | 28.82 | 71244.8 | 28.62
0 | 0 | 248.8 | 29.06 | 71852.1 | 28.86
0 | 0 | 247.6 | 29.19 | 72175.6 | 29.03
0 | 0 | 246.4 | 29.43 | 72757.2 | 29.3
0 | 0 | 245.1 | 29.65 | 73312.1 | 29.45
0 | 0 | 244 | 29.9 | 73917.3 | 29.71
0 | 0 | 242.7 | 30.03 | 74247.9 | 29.82
0 | 0 | 241.6 | 30.24 | 74771.8 | 30.06
0 | 0 | 240.3 | 30.32 | 74956.5 | 30.08
0 | 0 | 239.1 | 30.57 | 75581.7 | 30.35
0 | 0 | 237.9 | 30.83 | 76219.8 | 30.63
0 | 0 | 236.7 | 31.01 | 76674 | 30.79
0 | 0 | 235.5 | 31.17 | 77051.9 | 30.99
0 | 0 | 234.3 | 31.43 | 77707.8 | 31.23
0 | 0 | 233.1 | 31.52 | 77936.1 | 31.35
0 | 0 | 231.9 | 31.82 | 78671.4 | 31.59
0 | 0 | 230.7 | 32.12 | 79421.2 | 31.9
0 | 0 | 229.5 | 32.33 | 79929.1 | 32.1
0 | 0 | 228.3 | 32.5 | 80362.2 | 32.32
0 | 0 | 227.1 | 32.75 | 80967.7 | 32.54
0 | 0 | 225.9 | 32.97 | 81509.4 | 32.75
0 | 0 | 224.7 | 33.21 | 82105.1 | 32.97
0 | 0 | 223.5 | 33.45 | 82702.7 | 33.24
0 | 0 | 222.2 | 33.64 | 83165.5 | 33.38
0 | 0 | 221 | 33.91 | 83846.6 | 33.69
0 | 0 | 219.9 | 34.1 | 84313.7 | 33.85
0 | 0 | 218.6 | 34.31 | 84814.5 | 34.1
0 | 0 | 217.5 | 34.65 | 85665 | 34.4
0 | 0 | 216.2 | 34.78 | 85995 | 34.58
0 | 0 | 215 | 35.02 | 86580.8 | 34.76
0 | 0 | 213.8 | 35.24 | 87132.5 | 34.98
0 | 0 | 212.6 | 35.45 | 87656.8 | 35.26
0 | 0 | 211.4 | 35.79 | 88497.1 | 35.59
0 | 0 | 210.2 | 35.96 | 88911.9 | 35.76
0 | 0 | 209 | 36.22 | 89556.4 | 35.98
0 | 0 | 207.8 | 36.44 | 90090.1 | 36.21
0 | 0 | 206.6 | 36.83 | 91054.7 | 36.49
0 | 0 | 205.4 | 37 | 91482 | 36.81
0 | 0 | 204.2 | 37.19 | 91958.6 | 36.95
0 | 0 | 203 | 37.47 | 92627.4 | 37.21
0 | 0 | 201.8 | 37.77 | 93378.5 | 37.5 |  | Z | D
0 | 0 | 200.6 | 38.03 | 94030.3 | 37.81 | D20 | 200.6 | 37.81
0 | 0 | 199.4 | 38.27 | 94618.1 | 38.03 |  | 199.4 | 38.03
0 | 0 | 198.2 | 38.55 | 95320.8 | 38.28 |  | 200 | 37.92
0 | 0 | 197 | 38.79 | 95913.9 | 38.59
0 | 0 | 195.8 | 39.1 | 96666.5 | 38.8
0 | 0 | 194.6 | 39.25 | 97033 | 38.94
0 | 0 | 193.4 | 39.64 | 98008.2 | 39.35
0 | 0 | 192.2 | 39.93 | 98723.7 | 39.68
0 | 0 | 191 | 40.15 | 99276.6 | 39.92
0 | 0 | 189.8 | 40.52 | 100176.2 | 40.22
0 | 0 | 188.5 | 40.73 | 100696.1 | 40.48
0 | 0 | 187.4 | 41.03 | 101429 | 40.76
0 | 0 | 186.1 | 41.28 | 102046.6 | 40.98
0 | 0 | 184.9 | 41.5 | 102609.4 | 41.24
0 | 0 | 183.7 | 41.87 | 103520.2 | 41.53
0 | 0 | 182.5 | 42.14 | 104174.2 | 41.87
0 | 0 | 181.3 | 42.38 | 104788.9 | 42.1
0 | 0 | 180.1 | 42.66 | 105475.2 | 42.38
0 | 0 | 178.9 | 42.95 | 106175.7 | 42.69
0 | 0 | 177.7 | 43.3 | 107043.9 | 43.02
0 | 0 | 176.5 | 43.56 | 107694.7 | 43.3
0 | 0 | 175.3 | 43.95 | 108666.8 | 43.64
0 | 0 | 174.1 | 44.18 | 109239.1 | 43.86
0 | 0 | 172.9 | 44.42 | 109830.9 | 44.13
0 | 0 | 171.7 | 44.82 | 110804.1 | 44.5
0 | 0 | 170.5 | 45.09 | 111469.7 | 44.85
0 | 0 | 169.3 | 45.38 | 112194.3 | 45.07
0 | 0 | 168.1 | 45.79 | 113207.6 | 45.41
0 | 0 | 166.9 | 46.02 | 113773.1 | 45.71
0 | 0 | 165.7 | 46.26 | 114381.3 | 45.98
0 | 0 | 164.4 | 46.66 | 115368.9 | 46.4
0 | 0 | 163.3 | 46.97 | 116120.2 | 46.7
0 | 0 | 162 | 47.22 | 116749.8 | 46.99
0 | 0 | 160.9 | 47.61 | 117715.4 | 47.29
0 | 0 | 159.6 | 47.99 | 118644.5 | 47.77
0 | 0 | 158.4 | 48.28 | 119363.4 | 47.94
0 | 0 | 157.2 | 48.55 | 120029.6 | 48.28
0 | 0 | 156 | 48.92 | 120939.2 | 48.58
0 | 0 | 154.8 | 49.25 | 121772.5 | 48.95
0 | 0 | 153.6 | 49.62 | 122682.4 | 49.35
0 | 0 | 152.4 | 50.03 | 123703.9 | 49.6 |  | Z | D
0 | 0 | 151.2 | 50.22 | 124168.6 | 49.91 | R50 | 151.2 | 49.91
0 | 0 | 150 | 50.46 | 124759.1 | 50.18 |  | 150 | 50.18
0 | 0 | 148.8 | 50.87 | 125765.6 | 50.56 |  | 150.8 | 50
0 | 0 | 147.6 | 51.25 | 126705.6 | 50.86
0 | 0 | 146.4 | 51.59 | 127539 | 51.15
0 | 0 | 145.2 | 51.97 | 128493.7 | 51.62
0 | 0 | 144 | 52.33 | 129378.5 | 52.01
0 | 0 | 142.8 | 52.72 | 130336.2 | 52.47
0 | 0 | 141.6 | 53.01 | 131055.7 | 52.71
0 | 0 | 140.4 | 53.4 | 132017.8 | 52.97
0 | 0 | 139.2 | 53.66 | 132656.7 | 53.35
0 | 0 | 138 | 54.05 | 133641 | 53.65
0 | 0 | 136.8 | 54.43 | 134578.9 | 54.14
0 | 0 | 135.6 | 54.75 | 135369.7 | 54.45
0 | 0 | 134.4 | 55.14 | 136318.4 | 54.76
0 | 0 | 133.2 | 55.49 | 137188.5 | 55.1
0 | 0 | 131.9 | 55.85 | 138072 | 55.48
0 | 0 | 130.8 | 56.25 | 139075.6 | 55.89
0 | 0 | 129.6 | 56.71 | 140213 | 56.4
0 | 0 | 128.3 | 57.02 | 140961.3 | 56.57
0 | 0 | 127.1 | 57.39 | 141899.8 | 56.96
0 | 0 | 126 | 57.87 | 143076.8 | 57.42
0 | 0 | 124.7 | 58.23 | 143964.5 | 57.84
0 | 0 | 123.5 | 58.57 | 144802.3 | 58.19
0 | 0 | 122.3 | 59.13 | 146193.7 | 58.73
0 | 0 | 121.1 | 59.48 | 147066 | 59.09
0 | 0 | 119.9 | 59.85 | 147964.3 | 59.48
0 | 0 | 118.7 | 60.24 | 148934.6 | 59.89
0 | 0 | 117.5 | 60.6 | 149836.2 | 60.29
0 | 0 | 116.3 | 60.88 | 150522.4 | 60.55
0 | 0 | 115.1 | 61.27 | 151475.4 | 60.89
0 | 0 | 113.9 | 61.79 | 152758.1 | 61.33
0 | 0 | 112.7 | 62.09 | 153506.7 | 61.68
0 | 0 | 111.5 | 62.54 | 154614.8 | 62.13
0 | 0 | 110.3 | 63.01 | 155792.1 | 62.56
0 | 0 | 109.1 | 63.51 | 157019.6 | 63
0 | 0 | 107.9 | 64.02 | 158289.7 | 63.56
0 | 0 | 106.7 | 64.51 | 159496.6 | 63.84
0 | 0 | 105.5 | 64.81 | 160233.8 | 64.36
0 | 0 | 104.2 | 65.18 | 161158.2 | 64.88
0 | 0 | 103.1 | 65.66 | 162345.9 | 65.19
0 | 0 | 101.9 | 66.12 | 163470.3 | 65.63 |  | Z | D
0 | 0 | 100.7 | 66.56 | 164553.7 | 66.23 | D10 | 100.7 | 66.23
0 | 0 | 99.5 | 66.8 | 165151.1 | 66.43 |  | 99.5 | 66.43
0 | 0 | 98.3 | 67.23 | 166216.7 | 66.84 |  | 100 | 66.3467
0 | 0 | 97.1 | 67.74 | 167469.8 | 67.24
0 | 0 | 95.8 | 68.06 | 168266.5 | 67.73
0 | 0 | 94.7 | 68.39 | 169076.6 | 68.15
0 | 0 | 93.5 | 68.97 | 170521.4 | 68.6
0 | 0 | 92.2 | 69.61 | 172091.4 | 68.98
0 | 0 | 91.1 | 69.94 | 172923.4 | 69.54
0 | 0 | 89.8 | 70.5 | 174294.8 | 70.01
0 | 0 | 88.6 | 71.05 | 175650.2 | 70.44
0 | 0 | 87.4 | 71.54 | 176880.2 | 71.06
0 | 0 | 86.2 | 71.94 | 177867.6 | 71.55
0 | 0 | 85 | 72.45 | 179111.9 | 71.91
0 | 0 | 83.8 | 72.91 | 180252.2 | 72.35
0 | 0 | 82.6 | 73.37 | 181404 | 72.9
0 | 0 | 81.4 | 73.82 | 182515.6 | 73.36
0 | 0 | 80.2 | 74.32 | 183742 | 73.79
0 | 0 | 79 | 74.68 | 184629.3 | 74.18
0 | 0 | 77.8 | 75.26 | 186064.7 | 74.7
0 | 0 | 76.6 | 75.69 | 187144.5 | 75.09
0 | 0 | 75.4 | 75.94 | 187756.6 | 75.41
0 | 0 | 74.1 | 76.42 | 188938 | 75.96
0 | 0 | 73 | 76.9 | 190125.5 | 76.55
0 | 0 | 71.7 | 77.54 | 191710 | 77.05
0 | 0 | 70.5 | 78.1 | 193090.9 | 77.55
0 | 0 | 69.3 | 78.72 | 194624 | 77.98
0 | 0 | 68.1 | 79.04 | 195410 | 78.39
0 | 0 | 66.9 | 79.59 | 196786.3 | 79.02
0 | 0 | 65.7 | 80.07 | 197953.7 | 79.48
0 | 0 | 64.5 | 80.58 | 199215.5 | 79.89
0 | 0 | 63.3 | 81.05 | 200396.4 | 80.58
0 | 0 | 62 | 81.7 | 201999.9 | 81.11
0 | 0 | 60.9 | 82.02 | 202779.4 | 81.45
0 | 0 | 59.6 | 82.47 | 203895.1 | 82.01
0 | 0 | 58.5 | 83.17 | 205630.1 | 82.51
0 | 0 | 57.2 | 83.78 | 207122.8 | 83.14
0 | 0 | 56 | 84.05 | 207801.8 | 83.42
0 | 0 | 54.8 | 84.53 | 208993.7 | 83.93
0 | 0 | 53.6 | 84.95 | 210038.5 | 84.43
0 | 0 | 52.4 | 85.63 | 211702.6 | 85.05
0 | 0 | 51.2 | 86.21 | 213142.3 | 85.4
0 | 0 | 50 | 86.73 | 214435.1 | 86.13
0 | 0 | 48.8 | 87.06 | 215249.2 | 86.55
0 | 0 | 47.6 | 87.57 | 216492.8 | 86.98
0 | 0 | 46.4 | 88.17 | 217993.2 | 87.38
0 | 0 | 45.2 | 88.51 | 218825 | 87.97
0 | 0 | 44 | 89.08 | 220237.2 | 88.49
0 | 0 | 42.8 | 89.42 | 221077.5 | 88.88
0 | 0 | 41.6 | 89.98 | 222465.5 | 89.46
0 | 0 | 40.4 | 90.66 | 224139.6 | 89.96
0 | 0 | 39.1 | 91.34 | 225820.5 | 90.54
0 | 0 | 38 | 91.63 | 226542.2 | 90.99
0 | 0 | 36.7 | 92.09 | 227684.5 | 91.53
0 | 0 | 35.5 | 92.62 | 228982 | 91.93
0 | 0 | 34.3 | 93.2 | 230429.1 | 92.41
0 | 0 | 33.1 | 93.86 | 232066.6 | 93
0 | 0 | 31.9 | 94.45 | 233507.6 | 93.77
0 | 0 | 30.7 | 94.91 | 234651 | 94.29
0 | 0 | 29.5 | 95.33 | 235691.5 | 94.68
0 | 0 | 28.3 | 95.98 | 237296.9 | 95.29
0 | 0 | 27.1 | 96.25 | 237958.1 | 95.63
0 | 0 | 25.9 | 96.81 | 239336 | 96.15
0 | 0 | 24.7 | 97.32 | 240618.6 | 96.6
0 | 0 | 23.5 | 98.01 | 242304.1 | 97.24
0 | 0 | 22.3 | 98.33 | 243106.8 | 97.82
0 | 0 | 21.1 | 99 | 244773.2 | 98.39
0 | 0 | 19.8 | 99.28 | 245457.3 | 98.79
0 | 0 | 18.6 | 99.83 | 246816.3 | 99.35
0 | 0 | 17.4 | 99.99 | 247202.3 | 99.27
0 | 0 | 16.2 | 100.74 | 249072.3 | 99.76
0 | 0 | 15 | 100.58 | 248663 | 99.79
0 | 0 | 13.8 | 100.78 | 249172.3 | 99.99
0 | 0 | 12.6 | 100.54 | 248559 | 100 | z_max
0 | 0 | 11.4 | 100.04 | 247344.5 | 99.33
0 | 0 | 10.2 | 99.02 | 244807.6 | 98.17
0 | 0 | 9 | 97.79 | 241773.5 | 97.12
0 | 0 | 7.8 | 95.33 | 235679.2 | 94.65
0 | 0 | 6.6 | 91.96 | 227360.3 | 91.27
0 | 0 | 5.4 | 86.89 | 214830.9 | 86.3
0 | 0 | 4.2 | 79.72 | 197103.6 | 79.19
0 | 0 | 3 | 69.42 | 171637.6 | 68.95
0 | 0 | 1.7 | 59.81 | 147860.3 | 59.41
0 | 0 | 0.6 | 54.3 | 134255.9 | 53.92
0 | 0 | -0.5 | 50.98 | 126046.2 | 50.65
Measurement time: | 12/08/2022 18:08:37
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 33.1 | 81833.9 | 33.6
0 | 0 | 309.1 | 34.77 | 85965.8 | 33.86
0 | 0 | 307.6 | 35.14 | 86867.2 | 34.19
0 | 0 | 306.6 | 35.13 | 86861.2 | 34.24
0 | 0 | 305.3 | 35.24 | 87125.6 | 34.33
0 | 0 | 304.2 | 35.38 | 87477.2 | 34.52
0 | 0 | 302.9 | 35.77 | 88431.4 | 34.67
0 | 0 | 301.7 | 35.69 | 88236.2 | 34.83
0 | 0 | 300.5 | 35.81 | 88538.9 | 35.03
0 | 0 | 299.3 | 36.29 | 89714.5 | 35.26
0 | 0 | 298.2 | 36.3 | 89741.2 | 35.37
0 | 0 | 296.9 | 36.37 | 89909.5 | 35.49
0 | 0 | 295.7 | 36.65 | 90610.7 | 35.73
0 | 0 | 294.5 | 36.81 | 91002.6 | 35.89
0 | 0 | 293.3 | 37.02 | 91532 | 36.11
0 | 0 | 292.1 | 37.24 | 92073.8 | 36.22
0 | 0 | 290.9 | 37.22 | 92019.1 | 36.38
0 | 0 | 289.7 | 37.58 | 92909.5 | 36.55
0 | 0 | 288.5 | 37.62 | 93011.1 | 36.73
0 | 0 | 287.3 | 37.97 | 93882.8 | 36.92
0 | 0 | 286.1 | 37.97 | 93884.1 | 37.04
0 | 0 | 284.9 | 38.3 | 94692.1 | 37.31
0 | 0 | 283.7 | 38.39 | 94907.2 | 37.48
0 | 0 | 282.5 | 38.67 | 95602.5 | 37.49
0 | 0 | 281.3 | 38.79 | 95903.7 | 37.79
0 | 0 | 280.1 | 39.1 | 96663.6 | 38.05
0 | 0 | 278.9 | 39.2 | 96917.8 | 38.26
0 | 0 | 277.7 | 39.33 | 97229.6 | 38.41
0 | 0 | 276.5 | 39.57 | 97820.6 | 38.54
0 | 0 | 275.2 | 39.79 | 98383.3 | 38.67
0 | 0 | 274.1 | 39.92 | 98698.7 | 38.95
0 | 0 | 272.8 | 40.23 | 99458.6 | 39.24
0 | 0 | 271.6 | 40.47 | 100065.5 | 39.41
0 | 0 | 270.4 | 40.53 | 100198 | 39.53
0 | 0 | 269.2 | 40.75 | 100740.1 | 39.78
0 | 0 | 268 | 40.83 | 100949.8 | 39.95
0 | 0 | 266.8 | 41.2 | 101872.6 | 40.18
0 | 0 | 265.6 | 41.44 | 102466.2 | 40.27
0 | 0 | 264.4 | 41.48 | 102555.3 | 40.55
0 | 0 | 263.2 | 41.73 | 103173.7 | 40.79
0 | 0 | 262 | 41.8 | 103346.9 | 40.88
0 | 0 | 260.8 | 42.34 | 104690.9 | 41.15
0 | 0 | 259.6 | 42.28 | 104524.2 | 41.23
0 | 0 | 258.4 | 42.52 | 105128.6 | 41.43
0 | 0 | 257.2 | 42.7 | 105580.4 | 41.63
0 | 0 | 256 | 42.91 | 106100.2 | 41.91
0 | 0 | 254.8 | 43.13 | 106644.1 | 42.11
0 | 0 | 253.6 | 43.35 | 107175.1 | 42.24
0 | 0 | 252.4 | 43.66 | 107930.9 | 42.49
0 | 0 | 251.2 | 43.79 | 108254.8 | 42.69
0 | 0 | 250 | 44.01 | 108796.4 | 42.94
0 | 0 | 248.8 | 44.22 | 109320.4 | 43.14
0 | 0 | 247.6 | 44.34 | 109623.4 | 43.41
0 | 0 | 246.4 | 44.62 | 110311.2 | 43.65
0 | 0 | 245.2 | 45.02 | 111302.9 | 43.86
0 | 0 | 244 | 45.16 | 111655.1 | 44.01
0 | 0 | 242.7 | 45.19 | 111717 | 44.19
0 | 0 | 241.6 | 45.55 | 112607.3 | 44.5
0 | 0 | 240.4 | 45.78 | 113175.7 | 44.69
0 | 0 | 239.1 | 46.12 | 114034.2 | 44.98
0 | 0 | 238 | 46.3 | 114477.2 | 45.1
0 | 0 | 236.7 | 46.45 | 114836.5 | 45.43
0 | 0 | 235.5 | 46.76 | 115595.3 | 45.64
0 | 0 | 234.3 | 46.96 | 116109.8 | 45.78
0 | 0 | 233.1 | 47.15 | 116568.6 | 45.89
0 | 0 | 231.9 | 47.3 | 116937.3 | 46.25
0 | 0 | 230.7 | 47.51 | 117469.4 | 46.53
0 | 0 | 229.5 | 47.89 | 118410.4 | 46.69
0 | 0 | 228.3 | 48.21 | 119182.8 | 47.01
0 | 0 | 227.1 | 48.32 | 119459.3 | 47.15
0 | 0 | 225.9 | 48.46 | 119810.1 | 47.28
0 | 0 | 224.7 | 48.7 | 120409.2 | 47.73
0 | 0 | 223.5 | 49.19 | 121612.5 | 48.03
0 | 0 | 222.3 | 49.26 | 121778.5 | 48.13
0 | 0 | 221.1 | 49.43 | 122205.7 | 48.38
0 | 0 | 219.9 | 49.69 | 122855.5 | 48.69
0 | 0 | 218.7 | 49.99 | 123599.2 | 48.83
0 | 0 | 217.5 | 50.5 | 124851.1 | 49.18
0 | 0 | 216.3 | 50.43 | 124683.7 | 49.35
0 | 0 | 215.1 | 50.86 | 125741.9 | 49.67 |  | Z | D
0 | 0 | 213.9 | 51.04 | 126201.1 | 49.76 | R50 | 213.9 | 49.76
0 | 0 | 212.6 | 51.42 | 127121 | 50.07 |  | 212.6 | 50.07
0 | 0 | 211.5 | 51.59 | 127555.2 | 50.47 |  | 212.8935 | 50
0 | 0 | 210.2 | 51.87 | 128241.4 | 50.55
0 | 0 | 209 | 51.9 | 128327.1 | 50.86
0 | 0 | 207.9 | 52.28 | 129256.6 | 50.93
0 | 0 | 206.6 | 52.55 | 129922.6 | 51.35
0 | 0 | 205.4 | 52.72 | 130336.3 | 51.58
0 | 0 | 204.2 | 52.84 | 130634.9 | 51.9
0 | 0 | 203 | 53.25 | 131647.5 | 52.15
0 | 0 | 201.8 | 53.68 | 132719 | 52.34 |  | Z | D
0 | 0 | 200.6 | 53.82 | 133070.9 | 52.63 | D20 | 200.6 | 52.63
0 | 0 | 199.4 | 54.28 | 134208.1 | 52.99 |  | 199.4 | 52.99
0 | 0 | 198.2 | 54.39 | 134472.9 | 53.13 |  | 200 | 52.81
0 | 0 | 197 | 54.53 | 134812.3 | 53.31
0 | 0 | 195.8 | 54.82 | 135523.3 | 53.4
0 | 0 | 194.6 | 55.28 | 136675.6 | 53.86
0 | 0 | 193.4 | 55.44 | 137059.9 | 54.27
0 | 0 | 192.2 | 55.94 | 138313.7 | 54.42
0 | 0 | 191 | 55.89 | 138173 | 54.63
0 | 0 | 189.8 | 56.37 | 139370.8 | 54.85
0 | 0 | 188.5 | 56.54 | 139782.9 | 55.28
0 | 0 | 187.4 | 56.76 | 140334.3 | 55.51
0 | 0 | 186.2 | 57.11 | 141202.1 | 55.66
0 | 0 | 185 | 57.36 | 141825.2 | 55.97
0 | 0 | 183.8 | 57.51 | 142187.8 | 56.27
0 | 0 | 182.6 | 57.95 | 143262.1 | 56.66
0 | 0 | 181.4 | 58.36 | 144276.7 | 57
0 | 0 | 180.1 | 58.67 | 145054.6 | 57.29
0 | 0 | 178.9 | 58.77 | 145298.2 | 57.48
0 | 0 | 177.8 | 59.17 | 146297.5 | 57.81
0 | 0 | 176.5 | 59.38 | 146817.3 | 58.12
0 | 0 | 175.3 | 59.84 | 147933.2 | 58.26
0 | 0 | 174.1 | 60.06 | 148496.6 | 58.65
0 | 0 | 172.9 | 60.06 | 148491.9 | 58.88
0 | 0 | 171.7 | 60.57 | 149746.6 | 59.32
0 | 0 | 170.5 | 60.97 | 150738.4 | 59.52
0 | 0 | 169.3 | 61.38 | 151744.6 | 59.9
0 | 0 | 168.1 | 61.46 | 151941.3 | 60.12
0 | 0 | 166.9 | 61.85 | 152921.9 | 60.38
0 | 0 | 165.7 | 61.84 | 152889.6 | 60.61
0 | 0 | 164.5 | 62.51 | 154552.7 | 60.86
0 | 0 | 163.3 | 62.51 | 154555.2 | 61.28
0 | 0 | 162.1 | 62.99 | 155721.9 | 61.58
0 | 0 | 160.9 | 63.37 | 156677.2 | 61.78
0 | 0 | 159.7 | 63.84 | 157839.8 | 62.2
0 | 0 | 158.5 | 63.96 | 158125.9 | 62.42
0 | 0 | 157.3 | 64.15 | 158597.7 | 62.75
0 | 0 | 156.1 | 64.65 | 159834.5 | 63.14
0 | 0 | 154.9 | 64.74 | 160059.3 | 63.38
0 | 0 | 153.6 | 65.28 | 161406 | 63.55
0 | 0 | 152.4 | 65.49 | 161908.2 | 63.89
0 | 0 | 151.3 | 65.75 | 162559.3 | 64.34
0 | 0 | 150 | 66.03 | 163240.5 | 64.61
0 | 0 | 148.8 | 66.62 | 164701.8 | 65.13
0 | 0 | 147.6 | 66.91 | 165421.5 | 65.34
0 | 0 | 146.4 | 67.3 | 166391.9 | 65.69
0 | 0 | 145.2 | 67.48 | 166823.8 | 66.01
0 | 0 | 144 | 67.74 | 167478.5 | 66.09
0 | 0 | 142.8 | 68.22 | 168672.6 | 66.45
0 | 0 | 141.6 | 68.57 | 169526.7 | 66.87
0 | 0 | 140.4 | 68.72 | 169893.8 | 67.4
0 | 0 | 139.2 | 69.14 | 170932.3 | 67.53
0 | 0 | 138 | 69.66 | 172223.8 | 67.83
0 | 0 | 136.8 | 69.83 | 172652.2 | 68.15
0 | 0 | 135.6 | 70.25 | 173688.4 | 68.49
0 | 0 | 134.4 | 70.56 | 174440.5 | 69.08
0 | 0 | 133.2 | 71.13 | 175854.2 | 69.38
0 | 0 | 132 | 71.28 | 176236.4 | 69.61
0 | 0 | 130.8 | 71.66 | 177173.7 | 70.09
0 | 0 | 129.6 | 71.98 | 177967.1 | 70.42
0 | 0 | 128.4 | 72.48 | 179197.9 | 70.78
0 | 0 | 127.2 | 73.11 | 180757.5 | 71.13
0 | 0 | 126 | 72.71 | 179764.5 | 71.32
0 | 0 | 124.8 | 73.56 | 181862.9 | 71.68
0 | 0 | 123.6 | 73.67 | 182138.4 | 72.07
0 | 0 | 122.3 | 74.12 | 183260.5 | 72.59
0 | 0 | 121.1 | 74.59 | 184417.5 | 72.88
0 | 0 | 119.9 | 74.86 | 185087.5 | 73.09
0 | 0 | 118.7 | 75.32 | 186226.4 | 73.51
0 | 0 | 117.6 | 75.32 | 186214.2 | 73.74
0 | 0 | 116.3 | 76.08 | 188090.1 | 74.31
0 | 0 | 115.1 | 76.29 | 188617.7 | 74.61
0 | 0 | 113.9 | 76.78 | 189822.8 | 74.99
0 | 0 | 112.7 | 77.11 | 190648.5 | 75.32
0 | 0 | 111.5 | 77.32 | 191166.2 | 75.6
0 | 0 | 110.3 | 77.99 | 192828.2 | 75.93
0 | 0 | 109.1 | 78.05 | 192974.6 | 76.55
0 | 0 | 107.9 | 78.49 | 194044.7 | 76.75
0 | 0 | 106.7 | 79.09 | 195531.7 | 77.14
0 | 0 | 105.5 | 79.45 | 196422.9 | 77.31
0 | 0 | 104.3 | 79.64 | 196899.2 | 77.78
0 | 0 | 103.1 | 79.85 | 197410 | 78.43
0 | 0 | 101.9 | 80.69 | 199491.6 | 78.52 |  | Z | D
0 | 0 | 100.7 | 80.88 | 199964.6 | 79.08 | D10 | 100.7 | 79.08
0 | 0 | 99.5 | 81.22 | 200802.8 | 79.39 |  | 99.5 | 79.39
0 | 0 | 98.3 | 81.66 | 201899.7 | 79.8 |  | 100 | 79.2608
0 | 0 | 97.1 | 82.06 | 202874.7 | 80.12
0 | 0 | 95.9 | 82.34 | 203578.6 | 80.47
0 | 0 | 94.7 | 82.86 | 204863.3 | 81.08
0 | 0 | 93.5 | 83.28 | 205906.9 | 81.37
0 | 0 | 92.2 | 83.68 | 206889.1 | 82.04
0 | 0 | 91.1 | 84.36 | 208557.7 | 82.21
0 | 0 | 89.9 | 84.53 | 208986.4 | 82.73
0 | 0 | 88.7 | 84.97 | 210064.3 | 82.89
0 | 0 | 87.5 | 85.23 | 210707.1 | 83.32
0 | 0 | 86.2 | 85.86 | 212269.2 | 83.93
0 | 0 | 85.1 | 86.28 | 213309.7 | 84.21
0 | 0 | 83.8 | 86.51 | 213875.1 | 84.55
0 | 0 | 82.6 | 87.07 | 215264.3 | 85.1
0 | 0 | 81.4 | 87.27 | 215761.8 | 85.35
0 | 0 | 80.2 | 87.72 | 216866.4 | 85.79
0 | 0 | 79 | 88.27 | 218232.7 | 86.04
0 | 0 | 77.8 | 88.51 | 218832.3 | 86.46
0 | 0 | 76.6 | 88.76 | 219437.7 | 86.88
0 | 0 | 75.4 | 89.51 | 221307.9 | 87.31
0 | 0 | 74.2 | 89.81 | 222037 | 87.8
0 | 0 | 73 | 90.55 | 223876.1 | 88.32
0 | 0 | 71.8 | 90.45 | 223614.5 | 88.58
0 | 0 | 70.5 | 91.08 | 225173 | 89.17
0 | 0 | 69.3 | 91.66 | 226607.8 | 89.55
0 | 0 | 68.1 | 91.97 | 227391.5 | 89.73
0 | 0 | 66.9 | 92.47 | 228615.1 | 90.37
0 | 0 | 65.7 | 92.99 | 229906.7 | 91.01
0 | 0 | 64.5 | 93.38 | 230873.5 | 91.29
0 | 0 | 63.3 | 93.72 | 231716.6 | 91.52
0 | 0 | 62.1 | 94.17 | 232830 | 91.95
0 | 0 | 60.9 | 94.7 | 234124.4 | 92.46
0 | 0 | 59.7 | 94.88 | 234581.7 | 92.85
0 | 0 | 58.5 | 95.4 | 235855.4 | 93.28
0 | 0 | 57.3 | 96.07 | 237512.2 | 93.54
0 | 0 | 56.1 | 96.47 | 238510.3 | 94.11
0 | 0 | 54.9 | 96.46 | 238493.9 | 94.7
0 | 0 | 53.6 | 97.28 | 240511.1 | 94.85
0 | 0 | 52.4 | 97.52 | 241096.7 | 95.35
0 | 0 | 51.2 | 97.99 | 242256.7 | 95.91
0 | 0 | 50 | 98.06 | 242449.4 | 96.14
0 | 0 | 48.8 | 98.69 | 243986.8 | 96.36
0 | 0 | 47.6 | 99.07 | 244945.8 | 96.78
0 | 0 | 46.4 | 99.32 | 245552.9 | 96.96
0 | 0 | 45.2 | 99.71 | 246528.1 | 97.43
0 | 0 | 44 | 100.15 | 247596.8 | 97.82
0 | 0 | 42.8 | 100.36 | 248131.4 | 98.12
0 | 0 | 41.6 | 100.88 | 249405.3 | 98.3
0 | 0 | 40.4 | 101.24 | 250306.6 | 98.7
0 | 0 | 39.2 | 100.95 | 249574.9 | 99.27
0 | 0 | 38 | 101.5 | 250941.2 | 99.27
0 | 0 | 36.8 | 102.1 | 252438.1 | 99.52
0 | 0 | 35.5 | 101.93 | 252018.7 | 99.72
0 | 0 | 34.3 | 102.32 | 252970.9 | 99.96
0 | 0 | 33.1 | 102.06 | 252335.1 | 100
0 | 0 | 31.9 | 102.23 | 252756.5 | 99.91
0 | 0 | 30.7 | 102.3 | 252930.6 | 99.7
0 | 0 | 29.5 | 102.27 | 252848.2 | 99.79
0 | 0 | 28.3 | 101.94 | 252031.8 | 99.74
0 | 0 | 27.1 | 101.62 | 251250.9 | 99.46
0 | 0 | 25.9 | 101.66 | 251350.4 | 99.28
0 | 0 | 24.7 | 100.87 | 249397.5 | 98.66
0 | 0 | 23.5 | 100.65 | 248854.2 | 98.45
0 | 0 | 22.3 | 99.78 | 246694.8 | 97.51
0 | 0 | 21 | 98.69 | 243992 | 96.83
0 | 0 | 19.8 | 98.07 | 242464.2 | 95.77
0 | 0 | 18.7 | 97 | 239813.6 | 94.61
0 | 0 | 17.4 | 95.7 | 236611 | 93.43
0 | 0 | 16.3 | 94 | 232402.7 | 91.95
0 | 0 | 15 | 92.08 | 227646.9 | 90
0 | 0 | 13.8 | 90.06 | 222648.8 | 87.87
0 | 0 | 12.6 | 87.56 | 216481.9 | 85.42
0 | 0 | 11.4 | 84.36 | 208571.4 | 82.84
0 | 0 | 10.2 | 81.48 | 201459.3 | 79.5
0 | 0 | 9 | 77.7 | 192105.4 | 75.86
0 | 0 | 7.8 | 72.85 | 180115.4 | 71.02
0 | 0 | 6.6 | 67.39 | 166609.8 | 65.82
0 | 0 | 5.4 | 60.8 | 150329.2 | 59.23
0 | 0 | 4.2 | 53.13 | 131355.7 | 51.76
0 | 0 | 3 | 44.31 | 109548.9 | 43.46
0 | 0 | 1.8 | 37.36 | 92377 | 36.44
0 | 0 | 0.6 | 33.46 | 82728.8 | 32.77
0 | 0 | -0.5 | 31.45 | 77754.3 | 30.64
TrueBeam
Measurement time: | 02/08/2022 19:59:54
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 306.6 | 35.9 | 52112.8 | 20.92
0 | 0 | 305.3 | 36.3 | 52694 | 21.17
0 | 0 | 304.1 | 36.46 | 52937.8 | 21.27
0 | 0 | 302.9 | 36.69 | 53262.2 | 21.34
0 | 0 | 301.7 | 37.04 | 53780.3 | 21.59
0 | 0 | 300.5 | 37.17 | 53956.3 | 21.64
0 | 0 | 299.3 | 37.49 | 54424.6 | 21.86
0 | 0 | 298.1 | 37.75 | 54811.3 | 22
0 | 0 | 296.9 | 38.05 | 55242 | 22.13
0 | 0 | 295.7 | 38.27 | 55552.8 | 22.32
0 | 0 | 294.5 | 38.52 | 55922.7 | 22.47
0 | 0 | 293.3 | 38.72 | 56211.9 | 22.52
0 | 0 | 292.1 | 39.09 | 56745.1 | 22.75
0 | 0 | 290.9 | 39.2 | 56910.8 | 22.81
0 | 0 | 289.7 | 39.64 | 57551.2 | 23.1
0 | 0 | 288.5 | 39.77 | 57732.3 | 23.12
0 | 0 | 287.3 | 40.01 | 58090.4 | 23.26
0 | 0 | 286.1 | 40.42 | 58685 | 23.59
0 | 0 | 284.9 | 40.6 | 58936 | 23.62
0 | 0 | 283.7 | 40.89 | 59368.1 | 23.77
0 | 0 | 282.4 | 41.18 | 59790.2 | 23.98
0 | 0 | 281.3 | 41.42 | 60126.6 | 24.11
0 | 0 | 280 | 41.64 | 60447.9 | 24.21
0 | 0 | 278.9 | 41.94 | 60885 | 24.41
0 | 0 | 277.6 | 42.26 | 61345.6 | 24.54
0 | 0 | 276.4 | 42.66 | 61939 | 24.8
0 | 0 | 275.3 | 42.83 | 62185.5 | 24.89
0 | 0 | 274 | 43.17 | 62678.4 | 25.11
0 | 0 | 272.9 | 43.41 | 63018.3 | 25.28
0 | 0 | 271.7 | 43.61 | 63313.8 | 25.4
0 | 0 | 270.4 | 44 | 63873.3 | 25.61
0 | 0 | 269.3 | 44.3 | 64311.4 | 25.72
0 | 0 | 268 | 44.51 | 64623.9 | 25.91
0 | 0 | 266.8 | 44.96 | 65265.1 | 26.19
0 | 0 | 265.6 | 45.07 | 65431.4 | 26.21
0 | 0 | 264.4 | 45.41 | 65922.1 | 26.36
0 | 0 | 263.2 | 45.75 | 66422.7 | 26.67
0 | 0 | 262 | 46.12 | 66960.4 | 26.87
0 | 0 | 260.9 | 46.47 | 67463.1 | 27.04
0 | 0 | 259.6 | 46.62 | 67676 | 27.15
0 | 0 | 258.4 | 47.03 | 68279.6 | 27.39
0 | 0 | 257.2 | 47.39 | 68806.7 | 27.61
0 | 0 | 256 | 47.56 | 69051.2 | 27.7
0 | 0 | 254.8 | 47.98 | 69662.7 | 27.87
0 | 0 | 253.6 | 48.3 | 70117.9 | 28.15
0 | 0 | 252.4 | 48.66 | 70643.6 | 28.37
0 | 0 | 251.2 | 48.92 | 71021.2 | 28.47
0 | 0 | 250 | 49.33 | 71619.5 | 28.71
0 | 0 | 248.8 | 49.62 | 72034.4 | 28.9
0 | 0 | 247.6 | 50.01 | 72598.3 | 29.08
0 | 0 | 246.4 | 50.14 | 72797 | 29.15
0 | 0 | 245.2 | 50.74 | 73659.4 | 29.52
0 | 0 | 244 | 51.06 | 74135.1 | 29.81
0 | 0 | 242.8 | 51.41 | 74635.4 | 29.95
0 | 0 | 241.6 | 51.6 | 74917.7 | 30
0 | 0 | 240.3 | 52.03 | 75530.9 | 30.3
0 | 0 | 239.2 | 52.4 | 76073.4 | 30.54
0 | 0 | 237.9 | 52.73 | 76556.9 | 30.72
0 | 0 | 236.8 | 53.02 | 76980.4 | 30.87
0 | 0 | 235.5 | 53.31 | 77401.2 | 31.06
0 | 0 | 234.3 | 53.74 | 78023.6 | 31.31
0 | 0 | 233.1 | 54.03 | 78443.9 | 31.47
0 | 0 | 231.9 | 54.46 | 79070.5 | 31.73
0 | 0 | 230.7 | 54.79 | 79543.7 | 31.88
0 | 0 | 229.5 | 55.3 | 80280.5 | 32.17
0 | 0 | 228.3 | 55.54 | 80627.6 | 32.35
0 | 0 | 227.1 | 56.12 | 81467.7 | 32.64
0 | 0 | 225.9 | 56.41 | 81896.4 | 32.86
0 | 0 | 224.7 | 56.71 | 82324.2 | 33.06
0 | 0 | 223.5 | 57.14 | 82953.9 | 33.29
0 | 0 | 222.3 | 57.44 | 83397.7 | 33.42
0 | 0 | 221.1 | 58.05 | 84273.8 | 33.84
0 | 0 | 219.9 | 58.41 | 84797.4 | 34.01
0 | 0 | 218.7 | 58.67 | 85182.3 | 34.15
0 | 0 | 217.5 | 59.08 | 85766.2 | 34.42
0 | 0 | 216.3 | 59.62 | 86552.6 | 34.65
0 | 0 | 215.1 | 59.95 | 87032.8 | 34.93
0 | 0 | 213.9 | 60.16 | 87335.7 | 35.03
0 | 0 | 212.7 | 60.72 | 88155.7 | 35.27
0 | 0 | 211.5 | 61.06 | 88647.7 | 35.55
0 | 0 | 210.3 | 61.44 | 89204.3 | 35.79
0 | 0 | 209.1 | 61.89 | 89853.6 | 36.02
0 | 0 | 207.9 | 62.33 | 90487.9 | 36.3
0 | 0 | 206.7 | 62.74 | 91085.1 | 36.54
0 | 0 | 205.5 | 63.15 | 91680.9 | 36.74
0 | 0 | 204.2 | 63.64 | 92393.1 | 37.05
0 | 0 | 203.1 | 64.01 | 92923.4 | 37.26
0 | 0 | 201.8 | 64.58 | 93761.3 | 37.65 |  | Z | D
0 | 0 | 200.6 | 64.9 | 94213.8 | 37.77 | D20 | 200.6 | 37.77
0 | 0 | 199.5 | 65.38 | 94913.9 | 38.05 |  | 199.5 | 38.05
0 | 0 | 198.2 | 65.8 | 95534.1 | 38.28 |  | 200 | 37.9227
0 | 0 | 197 | 66.13 | 96010.8 | 38.54
0 | 0 | 195.8 | 66.73 | 96880.1 | 38.84
0 | 0 | 194.6 | 67.01 | 97289.3 | 38.9
0 | 0 | 193.4 | 67.61 | 98160.7 | 39.33
0 | 0 | 192.2 | 68.21 | 99024.1 | 39.7
0 | 0 | 191 | 68.44 | 99360.7 | 39.87
0 | 0 | 189.8 | 69.04 | 100227.4 | 40.14
0 | 0 | 188.6 | 69.46 | 100842.7 | 40.47
0 | 0 | 187.4 | 69.96 | 101571.8 | 40.7
0 | 0 | 186.2 | 70.51 | 102360.5 | 41.02
0 | 0 | 185 | 70.9 | 102934.4 | 41.21
0 | 0 | 183.8 | 71.54 | 103854.3 | 41.63
0 | 0 | 182.6 | 72.14 | 104728.5 | 42.03
0 | 0 | 181.4 | 72.36 | 105056.8 | 42.15
0 | 0 | 180.2 | 72.86 | 105770.1 | 42.36
0 | 0 | 178.9 | 73.18 | 106237 | 42.62
0 | 0 | 177.7 | 74.03 | 107468.7 | 43.12
0 | 0 | 176.6 | 74.56 | 108246.6 | 43.39
0 | 0 | 175.3 | 74.82 | 108620.6 | 43.53
0 | 0 | 174.1 | 75.56 | 109692.9 | 43.96
0 | 0 | 173 | 75.96 | 110277.5 | 44.23
0 | 0 | 171.7 | 76.42 | 110939.8 | 44.45
0 | 0 | 170.5 | 77.01 | 111795.4 | 44.79
0 | 0 | 169.3 | 77.33 | 112271.3 | 45.03
0 | 0 | 168.1 | 77.88 | 113064.8 | 45.32
0 | 0 | 166.9 | 78.62 | 114144.4 | 45.71
0 | 0 | 165.7 | 79.13 | 114882.9 | 46.06
0 | 0 | 164.5 | 79.66 | 115654 | 46.38
0 | 0 | 163.3 | 80.27 | 116533.6 | 46.75
0 | 0 | 162.1 | 80.83 | 117345.8 | 47.04
0 | 0 | 160.9 | 81.39 | 118162.6 | 47.35
0 | 0 | 159.7 | 81.83 | 118796 | 47.69
0 | 0 | 158.5 | 82.46 | 119712.1 | 47.92
0 | 0 | 157.3 | 82.94 | 120410.9 | 48.31
0 | 0 | 156.1 | 83.28 | 120905.1 | 48.5
0 | 0 | 154.9 | 84.03 | 121986.5 | 48.99
0 | 0 | 153.7 | 84.52 | 122702.7 | 49.2
0 | 0 | 152.5 | 85.18 | 123662.3 | 49.49 |  | Z | D
0 | 0 | 151.3 | 85.87 | 124670.2 | 49.95 | R50 | 151.3 | 49.95
0 | 0 | 150.1 | 86.28 | 125254.4 | 50.19 |  | 150.1 | 50.19
0 | 0 | 148.9 | 86.78 | 125988.2 | 50.6 |  | 151.05 | 50
0 | 0 | 147.6 | 87.62 | 127203.7 | 50.94
0 | 0 | 146.5 | 88.01 | 127768.7 | 51.25
0 | 0 | 145.2 | 88.79 | 128898.3 | 51.65
0 | 0 | 144 | 89.25 | 129574.6 | 52.02
0 | 0 | 142.8 | 89.81 | 130382.3 | 52.23
0 | 0 | 141.6 | 90.56 | 131469 | 52.65
0 | 0 | 140.4 | 91.08 | 132222.4 | 53.02
0 | 0 | 139.2 | 91.9 | 133416.2 | 53.5
0 | 0 | 138 | 92.3 | 133993.4 | 53.78
0 | 0 | 136.8 | 92.91 | 134888.3 | 54.06
0 | 0 | 135.6 | 93.82 | 136212.8 | 54.59
0 | 0 | 134.4 | 94.23 | 136800.4 | 54.8
0 | 0 | 133.2 | 94.93 | 137817.8 | 55.23
0 | 0 | 132 | 95.59 | 138779.5 | 55.67
0 | 0 | 130.8 | 96.02 | 139402.8 | 55.92
0 | 0 | 129.6 | 96.9 | 140674.7 | 56.46
0 | 0 | 128.4 | 97.27 | 141221.4 | 56.54
0 | 0 | 127.2 | 97.96 | 142219.3 | 57.08
0 | 0 | 126 | 98.72 | 143318.6 | 57.5
0 | 0 | 124.8 | 99.41 | 144320.1 | 57.75
0 | 0 | 123.5 | 99.99 | 145163.7 | 58.15
0 | 0 | 122.4 | 100.76 | 146276.7 | 58.6
0 | 0 | 121.1 | 101.21 | 146937.7 | 58.93
0 | 0 | 120 | 101.95 | 148003.4 | 59.35
0 | 0 | 118.8 | 102.74 | 149159.3 | 59.72
0 | 0 | 117.5 | 103.56 | 150349.3 | 60.29
0 | 0 | 116.4 | 104.07 | 151091.1 | 60.66
0 | 0 | 115.1 | 104.9 | 152295.1 | 61.08
0 | 0 | 113.9 | 105.59 | 153300.8 | 61.41
0 | 0 | 112.7 | 106.29 | 154317.1 | 61.79
0 | 0 | 111.5 | 106.87 | 155146 | 62.24
0 | 0 | 110.3 | 107.26 | 155722 | 62.45
0 | 0 | 109.1 | 108.49 | 157501.4 | 63.02
0 | 0 | 107.9 | 109.12 | 158422.2 | 63.44
0 | 0 | 106.7 | 109.75 | 159332.4 | 63.87
0 | 0 | 105.5 | 110.64 | 160619.9 | 64.33
0 | 0 | 104.3 | 111.24 | 161492.2 | 64.76
0 | 0 | 103.1 | 111.99 | 162586.2 | 65.18
0 | 0 | 101.9 | 112.64 | 163534.1 | 65.52 |  | Z | D
0 | 0 | 100.7 | 113.48 | 164753.7 | 65.92 | D10 | 100.7 | 65.92
0 | 0 | 99.5 | 114.3 | 165939.6 | 66.58 |  | 99.5 | 66.58
0 | 0 | 98.3 | 114.97 | 166918.8 | 67.02 |  | 100 | 66.305
0 | 0 | 97.1 | 115.51 | 167701.3 | 67.26
0 | 0 | 95.9 | 116.57 | 169230 | 67.78
0 | 0 | 94.7 | 117.16 | 170094.4 | 68.06
0 | 0 | 93.5 | 118.05 | 171390.5 | 68.75
0 | 0 | 92.3 | 118.83 | 172510.5 | 69.24
0 | 0 | 91.1 | 119.55 | 173565.2 | 69.46
0 | 0 | 89.9 | 120.29 | 174633.6 | 70.11
0 | 0 | 88.7 | 120.95 | 175594.2 | 70.35
0 | 0 | 87.5 | 121.82 | 176860.3 | 70.94
0 | 0 | 86.3 | 122.81 | 178287.6 | 71.55
0 | 0 | 85.1 | 123.39 | 179131.8 | 71.81
0 | 0 | 83.8 | 124.32 | 180488.9 | 72.39
0 | 0 | 82.7 | 125.14 | 181673.7 | 72.83
0 | 0 | 81.4 | 125.95 | 182851.2 | 73.27
0 | 0 | 80.3 | 126.78 | 184061.4 | 73.8
0 | 0 | 79 | 127.52 | 185131 | 74.19
0 | 0 | 77.8 | 128.19 | 186103.7 | 74.68
0 | 0 | 76.6 | 128.91 | 187143.1 | 74.98
0 | 0 | 75.4 | 129.69 | 188278.1 | 75.46
0 | 0 | 74.2 | 130.73 | 189789.1 | 76.27
0 | 0 | 73 | 131.6 | 191058.6 | 76.53
0 | 0 | 71.8 | 132.07 | 191744.3 | 76.93
0 | 0 | 70.6 | 133.11 | 193250.1 | 77.56
0 | 0 | 69.3 | 134.08 | 194651.6 | 78.14
0 | 0 | 68.2 | 134.76 | 195644.1 | 78.41
0 | 0 | 66.9 | 135.62 | 196889.2 | 78.78
0 | 0 | 65.7 | 136.37 | 197985.3 | 79.36
0 | 0 | 64.5 | 137.26 | 199274.4 | 79.84
0 | 0 | 63.3 | 138.01 | 200355.7 | 80.41
0 | 0 | 62.1 | 139.17 | 202040 | 81.05
0 | 0 | 60.9 | 139.89 | 203095 | 81.56
0 | 0 | 59.7 | 140.75 | 204331.7 | 81.97
0 | 0 | 58.5 | 141.57 | 205527.5 | 82.23
0 | 0 | 57.3 | 142.46 | 206826.9 | 82.84
0 | 0 | 56.1 | 143.25 | 207967.5 | 83.21
0 | 0 | 54.9 | 144.28 | 209465.1 | 84.11
0 | 0 | 53.7 | 145.27 | 210902.9 | 84.54
0 | 0 | 52.4 | 145.99 | 211945.5 | 84.95
0 | 0 | 51.2 | 147.02 | 213442.1 | 85.43
0 | 0 | 50 | 147.54 | 214191.4 | 85.93
0 | 0 | 48.8 | 148.65 | 215813.2 | 86.38
0 | 0 | 47.6 | 149.43 | 216936.6 | 86.9
0 | 0 | 46.4 | 150.36 | 218287.7 | 87.5
0 | 0 | 45.2 | 151.51 | 219966.9 | 88.3
0 | 0 | 44 | 152.18 | 220929.3 | 88.57
0 | 0 | 42.8 | 152.99 | 222110.7 | 88.95
0 | 0 | 41.6 | 153.8 | 223290 | 89.42
0 | 0 | 40.4 | 154.85 | 224802.1 | 90.09
0 | 0 | 39.2 | 155.37 | 225569.1 | 90.44
0 | 0 | 38 | 156.65 | 227420.7 | 90.93
0 | 0 | 36.8 | 157.58 | 228765.7 | 91.6
0 | 0 | 35.6 | 158.22 | 229697.3 | 92.08
0 | 0 | 34.4 | 159.13 | 231016.3 | 92.56
0 | 0 | 33.2 | 159.99 | 232266.4 | 93.12
0 | 0 | 31.9 | 161.13 | 233922.3 | 93.75
0 | 0 | 30.7 | 161.59 | 234599.8 | 94.06
0 | 0 | 29.5 | 162.66 | 236144.5 | 94.61
0 | 0 | 28.3 | 163.73 | 237700.9 | 95.11
0 | 0 | 27.1 | 164.46 | 238763.4 | 95.86
0 | 0 | 25.9 | 165.55 | 240343.3 | 96.37
0 | 0 | 24.7 | 166.05 | 241071.7 | 96.55
0 | 0 | 23.5 | 166.94 | 242355 | 96.92
0 | 0 | 22.3 | 168.07 | 243994.7 | 97.84
0 | 0 | 21.1 | 169 | 245357.1 | 98.53
0 | 0 | 19.9 | 169.57 | 246183.8 | 98.76
0 | 0 | 18.6 | 170.16 | 247039.7 | 98.94
0 | 0 | 17.5 | 170.66 | 247760.6 | 99.46
0 | 0 | 16.3 | 171.42 | 248859.7 | 99.83
0 | 0 | 15 | 171.85 | 249495.2 | 100
0 | 0 | 13.8 | 171.71 | 249279.2 | 99.89
0 | 0 | 12.6 | 171.41 | 248847.4 | 99.83
0 | 0 | 11.4 | 170.57 | 247634.2 | 99.35
0 | 0 | 10.2 | 169.23 | 245687.1 | 98.44
0 | 0 | 9 | 166.83 | 242200.6 | 97.14
0 | 0 | 7.8 | 163.49 | 237352.5 | 95.12
0 | 0 | 6.6 | 158.18 | 229640.6 | 92.08
0 | 0 | 5.4 | 151.17 | 219472.1 | 87.82
0 | 0 | 4.2 | 141.14 | 204899.7 | 82.02
0 | 0 | 2.9 | 123.89 | 179867.1 | 72.25
0 | 0 | 1.8 | 106.67 | 154861.2 | 62.09
0 | 0 | 0.5 | 94.37 | 137009 | 54.83
0 | 0 | -0.5 | 87.75 | 127391.6 | 51.08
Measurement time: | 02/08/2022 20:01:19
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 306.6 | 111.74 | 162225.7 | 18.59
0 | 0 | 305.3 | 112.81 | 163771.1 | 18.72
0 | 0 | 304.1 | 113.2 | 164342.4 | 18.82
0 | 0 | 302.9 | 113.96 | 165444.8 | 18.96
0 | 0 | 301.7 | 114.83 | 166714.3 | 19.1
0 | 0 | 300.5 | 115.65 | 167901.9 | 19.23
0 | 0 | 299.3 | 116.57 | 169227.8 | 19.38
0 | 0 | 298.1 | 117.21 | 170159.3 | 19.48
0 | 0 | 296.9 | 117.93 | 171215.8 | 19.6
0 | 0 | 295.7 | 118.78 | 172437.6 | 19.73
0 | 0 | 294.5 | 119.35 | 173272.6 | 19.86
0 | 0 | 293.3 | 120.56 | 175029.7 | 20
0 | 0 | 292.1 | 121.35 | 176170.3 | 20.13
0 | 0 | 290.9 | 122.01 | 177133.4 | 20.25
0 | 0 | 289.7 | 122.91 | 178443.7 | 20.43
0 | 0 | 288.5 | 123.6 | 179440.1 | 20.58
0 | 0 | 287.3 | 124.86 | 181266.2 | 20.7
0 | 0 | 286.1 | 125.49 | 182183.3 | 20.9
0 | 0 | 284.9 | 126.62 | 183820.8 | 21.02
0 | 0 | 283.7 | 127.09 | 184510.2 | 21.15
0 | 0 | 282.4 | 128.22 | 186151.4 | 21.31
0 | 0 | 281.2 | 129.09 | 187406.6 | 21.43
0 | 0 | 280 | 129.92 | 188617.8 | 21.6
0 | 0 | 278.8 | 130.73 | 189794.5 | 21.72
0 | 0 | 277.6 | 131.71 | 191216 | 21.89
0 | 0 | 276.4 | 132.78 | 192775.1 | 22.07
0 | 0 | 275.2 | 133.83 | 194292.6 | 22.21
0 | 0 | 274 | 134.77 | 195654.4 | 22.36
0 | 0 | 272.8 | 135.38 | 196545.3 | 22.53
0 | 0 | 271.6 | 136.33 | 197921.6 | 22.66
0 | 0 | 270.4 | 137.47 | 199570.4 | 22.81
0 | 0 | 269.2 | 138.41 | 200937.5 | 23
0 | 0 | 268 | 139.34 | 202291.2 | 23.13
0 | 0 | 266.8 | 140.54 | 204039.6 | 23.33
0 | 0 | 265.6 | 141.23 | 205030.8 | 23.47
0 | 0 | 264.4 | 142.48 | 206843.9 | 23.65
0 | 0 | 263.2 | 143.55 | 208406 | 23.85
0 | 0 | 262 | 144.26 | 209438.6 | 23.99
0 | 0 | 260.8 | 145.4 | 211087.9 | 24.14
0 | 0 | 259.6 | 146.29 | 212379.4 | 24.31
0 | 0 | 258.4 | 147.3 | 213853.5 | 24.49
0 | 0 | 257.2 | 148.48 | 215554.3 | 24.64
0 | 0 | 256 | 149.66 | 217276.2 | 24.86
0 | 0 | 254.8 | 150.54 | 218553.8 | 24.99
0 | 0 | 253.5 | 151.77 | 220343.2 | 25.22
0 | 0 | 252.3 | 153.04 | 222184.4 | 25.38
0 | 0 | 251.2 | 153.9 | 223426.4 | 25.53
0 | 0 | 249.9 | 154.77 | 224688.3 | 25.74
0 | 0 | 248.7 | 155.78 | 226158.8 | 25.9
0 | 0 | 247.5 | 157.1 | 228081.8 | 26.07
0 | 0 | 246.3 | 158.43 | 230003.4 | 26.25
0 | 0 | 245.1 | 159.22 | 231152.8 | 26.47
0 | 0 | 243.9 | 160.09 | 232420.3 | 26.65
0 | 0 | 242.7 | 161.48 | 234437.2 | 26.78
0 | 0 | 241.5 | 163.37 | 237180 | 27.08
0 | 0 | 240.3 | 163.54 | 237423 | 27.23
0 | 0 | 239.1 | 164.98 | 239511.4 | 27.41
0 | 0 | 237.9 | 166.77 | 242109.3 | 27.6
0 | 0 | 236.7 | 166.94 | 242359.6 | 27.79
0 | 0 | 235.5 | 168.9 | 245208 | 28.03
0 | 0 | 234.3 | 169.91 | 246672 | 28.24
0 | 0 | 233.1 | 170.86 | 248055.7 | 28.41
0 | 0 | 231.9 | 172.35 | 250210.4 | 28.56
0 | 0 | 230.7 | 173.14 | 251360.6 | 28.81
0 | 0 | 229.5 | 174.85 | 253841.7 | 29.03
0 | 0 | 228.3 | 176.13 | 255705.6 | 29.26
0 | 0 | 227.1 | 177.04 | 257031 | 29.43
0 | 0 | 225.9 | 178.76 | 259515.2 | 29.68
0 | 0 | 224.7 | 180.3 | 261760 | 29.91
0 | 0 | 223.5 | 180.94 | 262682.2 | 30.06
0 | 0 | 222.2 | 182.82 | 265419.4 | 30.3
0 | 0 | 221.1 | 183.47 | 266363 | 30.5
0 | 0 | 219.8 | 185.15 | 268792.2 | 30.74
0 | 0 | 218.7 | 186.68 | 271015.7 | 30.98
0 | 0 | 217.4 | 187.46 | 272154.6 | 31.16
0 | 0 | 216.2 | 189.48 | 275087.5 | 31.4
0 | 0 | 215 | 189.97 | 275801.7 | 31.58
0 | 0 | 213.8 | 191.8 | 278453.1 | 31.88
0 | 0 | 212.6 | 193.91 | 281513.2 | 32.16
0 | 0 | 211.4 | 194.68 | 282629.8 | 32.34
0 | 0 | 210.2 | 195.62 | 283996.8 | 32.55
0 | 0 | 209 | 197.86 | 287249.9 | 32.77
0 | 0 | 207.8 | 198.64 | 288379.7 | 33.04
0 | 0 | 206.6 | 200.17 | 290606.7 | 33.3
0 | 0 | 205.4 | 201.46 | 292480.6 | 33.49
0 | 0 | 204.2 | 203.62 | 295605.8 | 33.78
0 | 0 | 203 | 204.71 | 297201.5 | 33.94
0 | 0 | 201.8 | 205.98 | 299038.8 | 34.24 |  | Z | D
0 | 0 | 200.6 | 207.24 | 300866 | 34.51 | D10 | 200.6 | 34.51
0 | 0 | 199.4 | 209.61 | 304311.9 | 34.79 |  | 199.4 | 34.79
0 | 0 | 198.2 | 210.61 | 305756.9 | 34.94 |  | 200 | 34.65
0 | 0 | 197 | 212.14 | 307988.8 | 35.24
0 | 0 | 195.7 | 214.39 | 311241.5 | 35.57
0 | 0 | 194.6 | 215.17 | 312384.6 | 35.8
0 | 0 | 193.3 | 217.03 | 315084 | 36.04
0 | 0 | 192.1 | 218.33 | 316972 | 36.29
0 | 0 | 191 | 220.54 | 320183.6 | 36.55
0 | 0 | 189.7 | 221.67 | 321823.2 | 36.83
0 | 0 | 188.6 | 224.02 | 325224.6 | 37.12
0 | 0 | 187.3 | 224.89 | 326488 | 37.35
0 | 0 | 186.1 | 226.67 | 329070.2 | 37.62
0 | 0 | 184.9 | 227.45 | 330212.9 | 37.82
0 | 0 | 183.7 | 230.62 | 334806.5 | 38.23
0 | 0 | 182.5 | 231.32 | 335823.2 | 38.5
0 | 0 | 181.3 | 233.59 | 339121.9 | 38.75
0 | 0 | 180.1 | 235.31 | 341622.5 | 39.02
0 | 0 | 178.9 | 236.5 | 343350.5 | 39.22
0 | 0 | 177.7 | 238.55 | 346320.9 | 39.65
0 | 0 | 176.5 | 240.69 | 349436.1 | 39.93
0 | 0 | 175.3 | 242.25 | 351694.6 | 40.23
0 | 0 | 174.1 | 243.97 | 354186.9 | 40.47
0 | 0 | 172.9 | 245.68 | 356676.8 | 40.74
0 | 0 | 171.7 | 247.15 | 358805.4 | 41.04
0 | 0 | 170.5 | 249.34 | 361984.9 | 41.44
0 | 0 | 169.3 | 251.05 | 364468.5 | 41.68
0 | 0 | 168.1 | 253.1 | 367445.2 | 41.97
0 | 0 | 166.8 | 254.38 | 369300.7 | 42.27
0 | 0 | 165.6 | 256.39 | 372222.3 | 42.61
0 | 0 | 164.4 | 259.19 | 376287.6 | 42.95
0 | 0 | 163.2 | 260.31 | 377908 | 43.18
0 | 0 | 162 | 262.37 | 380908.2 | 43.56
0 | 0 | 160.8 | 264.37 | 383812.7 | 43.83
0 | 0 | 159.6 | 265.68 | 385710.6 | 44.23
0 | 0 | 158.4 | 268.6 | 389956.2 | 44.55
0 | 0 | 157.2 | 269.92 | 391865.2 | 44.84
0 | 0 | 156 | 272.14 | 395091.5 | 45.17
0 | 0 | 154.8 | 274.02 | 397821.9 | 45.48
0 | 0 | 153.6 | 275.81 | 400412.8 | 45.85
0 | 0 | 152.4 | 278.22 | 403909 | 46.17
0 | 0 | 151.2 | 280.12 | 406679.9 | 46.51
0 | 0 | 150 | 282.06 | 409492.2 | 46.85
0 | 0 | 148.8 | 284.27 | 412703.5 | 47.16
0 | 0 | 147.6 | 285.91 | 415076.9 | 47.48
0 | 0 | 146.4 | 288.74 | 419183.7 | 47.92
0 | 0 | 145.2 | 290.27 | 421406.7 | 48.26
0 | 0 | 144 | 292.6 | 424799.3 | 48.57
0 | 0 | 142.8 | 294.57 | 427656 | 48.94
0 | 0 | 141.5 | 296.97 | 431131.3 | 49.33
0 | 0 | 140.3 | 299.49 | 434798.1 | 49.69 |  | Z | D
0 | 0 | 139.1 | 301.22 | 437311.3 | 49.99 | R50 | 139.1 | 49.99
0 | 0 | 137.9 | 303.48 | 440594.1 | 50.35 |  | 137.9 | 50.35
0 | 0 | 136.8 | 306.19 | 444520.5 | 50.83 |  | 139.0667 | 50
0 | 0 | 135.5 | 307.74 | 446770.8 | 51.15
0 | 0 | 134.4 | 310.34 | 450541.6 | 51.51
0 | 0 | 133.1 | 312.63 | 453878.3 | 51.88
0 | 0 | 131.9 | 314.62 | 456765.7 | 52.24
0 | 0 | 130.7 | 317.83 | 461425.7 | 52.74
0 | 0 | 129.5 | 320.1 | 464720.7 | 53.04
0 | 0 | 128.3 | 321.53 | 466798.5 | 53.43
0 | 0 | 127.1 | 324.39 | 470948.6 | 53.8
0 | 0 | 125.9 | 326.18 | 473549 | 54.21
0 | 0 | 124.7 | 329.03 | 477679.4 | 54.7
0 | 0 | 123.5 | 331.08 | 480662 | 54.92
0 | 0 | 122.3 | 333.97 | 484856.5 | 55.41
0 | 0 | 121.1 | 336.5 | 488532.2 | 55.83
0 | 0 | 119.9 | 338.76 | 491813 | 56.23
0 | 0 | 118.7 | 340.94 | 494977.4 | 56.6
0 | 0 | 117.5 | 343.25 | 498330.2 | 56.95
0 | 0 | 116.3 | 345.91 | 502186 | 57.4
0 | 0 | 115.1 | 348.3 | 505655.9 | 57.71
0 | 0 | 113.9 | 351.21 | 509880.9 | 58.32
0 | 0 | 112.7 | 353.49 | 513190.6 | 58.7
0 | 0 | 111.5 | 355.94 | 516741.9 | 59.13
0 | 0 | 110.3 | 358.66 | 520700.8 | 59.56
0 | 0 | 109.1 | 361.09 | 524229.2 | 59.98
0 | 0 | 107.9 | 363.79 | 528151.8 | 60.38
0 | 0 | 106.7 | 366.79 | 532496.2 | 60.88
0 | 0 | 105.4 | 369.47 | 536387.1 | 61.34
0 | 0 | 104.3 | 371.8 | 539772.6 | 61.69
0 | 0 | 103.1 | 374.21 | 543277.1 | 62.19
0 | 0 | 101.9 | 377.77 | 548446 | 62.68 |  | Z | D
0 | 0 | 100.7 | 380.02 | 551702.5 | 63.11 | D10 | 100.7 | 63.11
0 | 0 | 99.5 | 383.42 | 556642.5 | 63.55 |  | 99.5 | 63.55
0 | 0 | 98.3 | 385.7 | 559948.1 | 64.07 |  | 100 | 63.3667
0 | 0 | 97 | 387.79 | 562982.5 | 64.49
0 | 0 | 95.9 | 390.95 | 567578.1 | 64.9
0 | 0 | 94.7 | 394.27 | 572392 | 65.4
0 | 0 | 93.4 | 396.82 | 576096.7 | 65.91
0 | 0 | 92.3 | 399.79 | 580415.5 | 66.37
0 | 0 | 91 | 403.8 | 586229 | 66.89
0 | 0 | 89.9 | 405.74 | 589048 | 67.42
0 | 0 | 88.6 | 408.53 | 593103.1 | 67.83
0 | 0 | 87.4 | 411 | 596679.8 | 68.26
0 | 0 | 86.2 | 414.05 | 601116.4 | 68.73
0 | 0 | 85 | 416.58 | 604790.8 | 69.14
0 | 0 | 83.8 | 419.17 | 608547 | 69.65
0 | 0 | 82.6 | 423.5 | 614829.6 | 70.23
0 | 0 | 81.4 | 425.73 | 618065.8 | 70.69
0 | 0 | 80.2 | 429.31 | 623266.3 | 71.3
0 | 0 | 79 | 432.84 | 628396.2 | 71.79
0 | 0 | 77.8 | 435.05 | 631601 | 72.24
0 | 0 | 76.6 | 438.44 | 636522 | 72.84
0 | 0 | 75.4 | 440.76 | 639889.9 | 73.28
0 | 0 | 74.2 | 444.51 | 645328.9 | 73.74
0 | 0 | 73 | 446.51 | 648231.1 | 74.19
0 | 0 | 71.8 | 450.54 | 654087.1 | 74.77
0 | 0 | 70.5 | 453.7 | 658676.6 | 75.32
0 | 0 | 69.3 | 457.11 | 663630.7 | 75.74
0 | 0 | 68.1 | 459.17 | 666611.1 | 76.31
0 | 0 | 66.9 | 463.81 | 673356.4 | 76.87
0 | 0 | 65.7 | 465.6 | 675945.5 | 77.44
0 | 0 | 64.5 | 469.76 | 681988.5 | 77.98
0 | 0 | 63.3 | 472.64 | 686168.5 | 78.51
0 | 0 | 62.1 | 476.05 | 691125.1 | 79.01
0 | 0 | 60.9 | 478.65 | 694898.1 | 79.53
0 | 0 | 59.7 | 483.97 | 702625.7 | 80.06
0 | 0 | 58.4 | 484.93 | 704015 | 80.58
0 | 0 | 57.2 | 488.16 | 708705.9 | 81.17
0 | 0 | 56.1 | 492.75 | 715360.9 | 81.68
0 | 0 | 54.8 | 494.52 | 717941 | 82.22
0 | 0 | 53.6 | 498.67 | 723956.5 | 82.74
0 | 0 | 52.4 | 501.95 | 728730.9 | 83.37
0 | 0 | 51.2 | 504.82 | 732892.5 | 83.84
0 | 0 | 50 | 508.67 | 738486.5 | 84.42
0 | 0 | 48.8 | 511.99 | 743296.8 | 85.01
0 | 0 | 47.6 | 516.07 | 749220.7 | 85.53
0 | 0 | 46.4 | 517.84 | 751799.1 | 86.11
0 | 0 | 45.2 | 522.29 | 758253.7 | 86.58
0 | 0 | 43.9 | 524.7 | 761753.9 | 87.14
0 | 0 | 42.7 | 528.39 | 767106.6 | 87.77
0 | 0 | 41.6 | 531.7 | 771919.7 | 88.27
0 | 0 | 40.3 | 534.95 | 776635 | 88.87
0 | 0 | 39.1 | 539.29 | 782932.7 | 89.51
0 | 0 | 37.9 | 541.25 | 785778.4 | 89.95
0 | 0 | 36.7 | 545.04 | 791282.1 | 90.42
0 | 0 | 35.5 | 548.76 | 796676.8 | 91.03
0 | 0 | 34.3 | 552.58 | 802220.7 | 91.64
0 | 0 | 33.1 | 555.23 | 806077.5 | 92.31
0 | 0 | 31.9 | 558.68 | 811082.2 | 92.73
0 | 0 | 30.7 | 561.76 | 815555.4 | 93.32
0 | 0 | 29.5 | 565.22 | 820571.2 | 93.88
0 | 0 | 28.3 | 568.99 | 826047.5 | 94.51
0 | 0 | 27 | 573.36 | 832395 | 95.14
0 | 0 | 25.8 | 575.93 | 836120.4 | 95.55
0 | 0 | 24.6 | 578.07 | 839235.9 | 96.12
0 | 0 | 23.4 | 581.22 | 843809.6 | 96.63
0 | 0 | 22.2 | 585.29 | 849716.7 | 97.16
0 | 0 | 21 | 589.94 | 856462.4 | 97.7
0 | 0 | 19.8 | 590.36 | 857074.8 | 98.2
0 | 0 | 18.6 | 594.68 | 863355 | 98.66
0 | 0 | 17.4 | 596.07 | 865368.5 | 99.12
0 | 0 | 16.2 | 598.77 | 869288.2 | 99.52
0 | 0 | 15 | 601.69 | 873517.7 | 99.8
0 | 0 | 13.8 | 602.3 | 874405.4 | 100
0 | 0 | 12.6 | 602.86 | 875227.2 | 99.92
0 | 0 | 11.4 | 600.07 | 871167.4 | 99.76
0 | 0 | 10.2 | 597.5 | 867437.9 | 99.3
0 | 0 | 8.9 | 591.63 | 858914 | 98.35
0 | 0 | 7.7 | 582.52 | 845697.6 | 96.73
0 | 0 | 6.5 | 569.69 | 827073.8 | 94.58
0 | 0 | 5.3 | 551.11 | 800087.8 | 91.3
0 | 0 | 4.1 | 518.67 | 752991 | 86.28
0 | 0 | 2.9 | 474.03 | 688188.8 | 78.79
0 | 0 | 1.7 | 418.33 | 607320.3 | 69.51
0 | 0 | 0.5 | 379.58 | 551070.8 | 63.02
0 | 0 | -0.5 | 359.95 | 522565.5 | 59.74
Measurement time: | 02/08/2022 20:03:25
Radiation device: | Novalis TB Accelerator
Energy: | 10 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 49.29 | 71552.7 | 27.58
0 | 0 | 309.1 | 48.4 | 70263.2 | 27.91
0 | 0 | 307.6 | 48.74 | 70762.8 | 28.13
0 | 0 | 306.6 | 49.05 | 71204.4 | 28.35
0 | 0 | 305.3 | 49.3 | 71569 | 28.42
0 | 0 | 304.1 | 49.59 | 71996.4 | 28.56
0 | 0 | 302.9 | 49.95 | 72513.1 | 28.8
0 | 0 | 301.7 | 50.24 | 72933.7 | 28.92
0 | 0 | 300.5 | 50.38 | 73137.4 | 29.03
0 | 0 | 299.3 | 50.68 | 73576.9 | 29.25
0 | 0 | 298.1 | 50.95 | 73974.9 | 29.34
0 | 0 | 296.9 | 51.27 | 74427.4 | 29.57
0 | 0 | 295.7 | 51.51 | 74783 | 29.72
0 | 0 | 294.5 | 51.81 | 75220.4 | 29.86
0 | 0 | 293.3 | 52.08 | 75612.5 | 30.02
0 | 0 | 292.1 | 52.4 | 76079.7 | 30.14
0 | 0 | 290.9 | 52.55 | 76295.2 | 30.31
0 | 0 | 289.7 | 52.99 | 76922.9 | 30.55
0 | 0 | 288.5 | 53.24 | 77295.2 | 30.66
0 | 0 | 287.3 | 53.56 | 77754.1 | 30.86
0 | 0 | 286.1 | 53.82 | 78140 | 31.01
0 | 0 | 284.9 | 54.13 | 78587.9 | 31.22
0 | 0 | 283.7 | 54.53 | 79172.7 | 31.41
0 | 0 | 282.5 | 54.65 | 79343.2 | 31.52
0 | 0 | 281.3 | 54.9 | 79698.8 | 31.67
0 | 0 | 280.1 | 55.34 | 80347.3 | 31.98
0 | 0 | 278.9 | 55.64 | 80783.8 | 32.07
0 | 0 | 277.6 | 55.96 | 81238.7 | 32.27
0 | 0 | 276.4 | 56.3 | 81729.3 | 32.43
0 | 0 | 275.3 | 56.6 | 82167.2 | 32.63
0 | 0 | 274.1 | 56.86 | 82552 | 32.75
0 | 0 | 272.8 | 57.29 | 83166.9 | 33
0 | 0 | 271.6 | 57.44 | 83393.1 | 33.11
0 | 0 | 270.4 | 57.98 | 84170.3 | 33.42
0 | 0 | 269.2 | 58.22 | 84519 | 33.51
0 | 0 | 268 | 58.49 | 84918.5 | 33.78
0 | 0 | 266.8 | 58.76 | 85303.4 | 33.85
0 | 0 | 265.6 | 59.19 | 85930.4 | 34.13
0 | 0 | 264.4 | 59.58 | 86490.9 | 34.3
0 | 0 | 263.2 | 59.83 | 86864.9 | 34.46
0 | 0 | 262 | 60.26 | 87479.4 | 34.72
0 | 0 | 260.8 | 60.57 | 87932.4 | 34.89
0 | 0 | 259.6 | 60.86 | 88357.4 | 35.04
0 | 0 | 258.4 | 61.28 | 88966.2 | 35.34
0 | 0 | 257.2 | 61.48 | 89249 | 35.49
0 | 0 | 256 | 61.85 | 89798.3 | 35.69
0 | 0 | 254.8 | 62.21 | 90320.2 | 35.9
0 | 0 | 253.6 | 62.6 | 90888 | 36.06
0 | 0 | 252.4 | 62.86 | 91255.7 | 36.26
0 | 0 | 251.2 | 63.3 | 91902 | 36.46
0 | 0 | 250 | 63.74 | 92542.2 | 36.73
0 | 0 | 248.7 | 63.9 | 92775.8 | 36.89
0 | 0 | 247.5 | 64.32 | 93379.8 | 37.05
0 | 0 | 246.3 | 64.67 | 93888.3 | 37.27
0 | 0 | 245.1 | 64.98 | 94335.6 | 37.42
0 | 0 | 243.9 | 65.36 | 94890.9 | 37.65
0 | 0 | 242.7 | 65.83 | 95574.8 | 37.92
0 | 0 | 241.5 | 66.09 | 95955 | 38.23
0 | 0 | 240.3 | 66.32 | 96275.3 | 38.25
0 | 0 | 239.1 | 66.91 | 97143.9 | 38.6
0 | 0 | 237.9 | 67.19 | 97538.9 | 38.78
0 | 0 | 236.7 | 67.65 | 98217.7 | 38.94
0 | 0 | 235.5 | 68.04 | 98781.3 | 39.22
0 | 0 | 234.3 | 68.41 | 99319.7 | 39.41
0 | 0 | 233.1 | 68.88 | 99995 | 39.69
0 | 0 | 231.9 | 69.09 | 100310.9 | 39.79
0 | 0 | 230.7 | 69.52 | 100924.2 | 40
0 | 0 | 229.5 | 69.91 | 101501 | 40.35
0 | 0 | 228.3 | 70.32 | 102094.3 | 40.46
0 | 0 | 227.1 | 70.63 | 102538.7 | 40.69
0 | 0 | 225.9 | 71.23 | 103410.2 | 41.05
0 | 0 | 224.6 | 71.69 | 104083.2 | 41.29
0 | 0 | 223.4 | 71.84 | 104300.4 | 41.38
0 | 0 | 222.2 | 72.3 | 104966.5 | 41.64
0 | 0 | 221 | 72.61 | 105416.1 | 41.85
0 | 0 | 219.8 | 73.19 | 106260.6 | 42.14
0 | 0 | 218.6 | 73.45 | 106640.2 | 42.38
0 | 0 | 217.4 | 74.04 | 107495.2 | 42.65
0 | 0 | 216.2 | 74.34 | 107932 | 42.83
0 | 0 | 215 | 74.84 | 108649.1 | 43.09
0 | 0 | 213.8 | 75.24 | 109227.3 | 43.44
0 | 0 | 212.6 | 75.59 | 109734.3 | 43.6
0 | 0 | 211.4 | 76.13 | 110517.5 | 43.93
0 | 0 | 210.2 | 76.45 | 110984 | 44.03
0 | 0 | 209 | 76.98 | 111753.8 | 44.34
0 | 0 | 207.8 | 77.34 | 112280.7 | 44.54
0 | 0 | 206.6 | 77.81 | 112962.9 | 44.85
0 | 0 | 205.4 | 78.16 | 113468.3 | 45.1
0 | 0 | 204.2 | 78.61 | 114130.8 | 45.3
0 | 0 | 203 | 79.11 | 114848.4 | 45.67
0 | 0 | 201.8 | 79.45 | 115344.1 | 45.87 |  | Z | D
0 | 0 | 200.6 | 80.06 | 116234.6 | 46.15 | D20 | 200.6 | 46.15
0 | 0 | 199.4 | 80.45 | 116791.4 | 46.34 |  | 199.4 | 46.34
0 | 0 | 198.1 | 80.91 | 117459.3 | 46.61 |  | 200 | 46.245
0 | 0 | 197 | 81.38 | 118152.5 | 46.92
0 | 0 | 195.7 | 81.86 | 118837.7 | 47.09
0 | 0 | 194.6 | 82.38 | 119597.5 | 47.44
0 | 0 | 193.4 | 82.65 | 119986.9 | 47.59
0 | 0 | 192.1 | 83.29 | 120923.2 | 48
0 | 0 | 191 | 83.58 | 121340.6 | 48.18
0 | 0 | 189.7 | 83.97 | 121909.8 | 48.42
0 | 0 | 188.5 | 84.53 | 122719 | 48.76
0 | 0 | 187.3 | 85.16 | 123639.6 | 49.04
0 | 0 | 186.1 | 85.45 | 124060.2 | 49.29
0 | 0 | 184.9 | 86.13 | 125039.1 | 49.7 |  | Z | D
0 | 0 | 183.7 | 86.59 | 125709.2 | 49.8 | R50 | 183.7 | 49.8
0 | 0 | 182.5 | 87.07 | 126404.3 | 50.21 |  | 182.5 | 50.21
0 | 0 | 181.3 | 87.43 | 126936.7 | 50.4 |  | 183.1146 | 50
0 | 0 | 180.1 | 88.01 | 127772.5 | 50.65
0 | 0 | 178.9 | 88.54 | 128540.3 | 51.09
0 | 0 | 177.7 | 89.01 | 129221.2 | 51.32
0 | 0 | 176.5 | 89.37 | 129740.7 | 51.58
0 | 0 | 175.3 | 89.94 | 130568.3 | 51.85
0 | 0 | 174.1 | 90.38 | 131212.8 | 52.1
0 | 0 | 172.9 | 90.98 | 132085.8 | 52.44
0 | 0 | 171.7 | 91.56 | 132926 | 52.71
0 | 0 | 170.5 | 92.11 | 133716.8 | 53.17
0 | 0 | 169.3 | 92.46 | 134227.2 | 53.29
0 | 0 | 168.1 | 93.12 | 135196.4 | 53.65
0 | 0 | 166.9 | 93.52 | 135772.3 | 53.91
0 | 0 | 165.6 | 94.18 | 136736.4 | 54.29
0 | 0 | 164.5 | 94.74 | 137539.1 | 54.63
0 | 0 | 163.2 | 95.1 | 138071.5 | 54.72
0 | 0 | 162 | 95.69 | 138918.2 | 55.14
0 | 0 | 160.8 | 96.26 | 139754.2 | 55.52
0 | 0 | 159.6 | 96.77 | 140495.6 | 55.73
0 | 0 | 158.4 | 97.32 | 141292.7 | 56.12
0 | 0 | 157.2 | 97.92 | 142157.9 | 56.44
0 | 0 | 156 | 98.32 | 142741.7 | 56.62
0 | 0 | 154.8 | 98.9 | 143580.5 | 57.04
0 | 0 | 153.6 | 99.6 | 144600.4 | 57.4
0 | 0 | 152.4 | 99.93 | 145082.4 | 57.55
0 | 0 | 151.2 | 100.59 | 146030.6 | 57.85
0 | 0 | 150 | 101.17 | 146880.3 | 58.32
0 | 0 | 148.8 | 101.77 | 147751.5 | 58.63
0 | 0 | 147.6 | 102.3 | 148513.3 | 59.02
0 | 0 | 146.4 | 102.77 | 149206.9 | 59.23
0 | 0 | 145.2 | 103.42 | 150147.6 | 59.53
0 | 0 | 144 | 104.11 | 151142.6 | 59.95
0 | 0 | 142.8 | 104.57 | 151817.7 | 60.17
0 | 0 | 141.6 | 105.27 | 152834.2 | 60.65
0 | 0 | 140.4 | 105.8 | 153598.7 | 60.96
0 | 0 | 139.2 | 106.52 | 154638.3 | 61.35
0 | 0 | 138 | 107.09 | 155465 | 61.63
0 | 0 | 136.8 | 107.49 | 156049 | 61.99
0 | 0 | 135.5 | 108.3 | 157222.9 | 62.43
0 | 0 | 134.4 | 108.73 | 157851.7 | 62.78
0 | 0 | 133.1 | 109.41 | 158841.6 | 63.17
0 | 0 | 131.9 | 110.06 | 159787.1 | 63.48
0 | 0 | 130.7 | 110.61 | 160589.2 | 63.68
0 | 0 | 129.5 | 111.36 | 161670.9 | 64.17
0 | 0 | 128.3 | 111.87 | 162418.1 | 64.59
0 | 0 | 127.1 | 112.61 | 163486.9 | 64.86
0 | 0 | 125.9 | 113.11 | 164216.1 | 65.18
0 | 0 | 124.7 | 113.71 | 165089.5 | 65.65
0 | 0 | 123.5 | 114.36 | 166027.1 | 65.89
0 | 0 | 122.3 | 115.12 | 167133.2 | 66.27
0 | 0 | 121.1 | 115.68 | 167947.5 | 66.64
0 | 0 | 119.9 | 116.26 | 168790.1 | 66.97
0 | 0 | 118.7 | 117.01 | 169876.7 | 67.41
0 | 0 | 117.5 | 117.63 | 170775.5 | 67.7
0 | 0 | 116.3 | 118.27 | 171704.4 | 68.01
0 | 0 | 115.1 | 119.01 | 172779.6 | 68.47
0 | 0 | 113.9 | 119.59 | 173625.9 | 68.89
0 | 0 | 112.7 | 120.37 | 174755.2 | 69.22
0 | 0 | 111.4 | 120.87 | 175480.5 | 69.62
0 | 0 | 110.3 | 121.46 | 176332.6 | 69.96
0 | 0 | 109 | 122.1 | 177269 | 70.33
0 | 0 | 107.8 | 122.8 | 178273.4 | 70.88
0 | 0 | 106.7 | 123.54 | 179353.7 | 71.07
0 | 0 | 105.4 | 124.25 | 180385.6 | 71.6
0 | 0 | 104.3 | 125.04 | 181533.4 | 71.95
0 | 0 | 103 | 125.76 | 182576.7 | 72.47
0 | 0 | 101.8 | 126.23 | 183258.2 | 72.83 |  | Z | D
0 | 0 | 100.6 | 126.93 | 184278.3 | 73.3 | D10 | 100.6 | 73.3
0 | 0 | 99.4 | 127.51 | 185110.9 | 73.44 |  | 99.4 | 73.44
0 | 0 | 98.2 | 128.19 | 186100.5 | 73.81 |  | 100 | 73.37
0 | 0 | 97 | 129.17 | 187532 | 74.43
0 | 0 | 95.8 | 129.74 | 188351.4 | 74.83
0 | 0 | 94.6 | 130.35 | 189247.2 | 75.04
0 | 0 | 93.4 | 131.2 | 190474.1 | 75.56
0 | 0 | 92.2 | 131.97 | 191587.2 | 76.03
0 | 0 | 91 | 132.71 | 192661.1 | 76.38
0 | 0 | 89.8 | 133.32 | 193546 | 76.91
0 | 0 | 88.6 | 134.1 | 194685.9 | 77.18
0 | 0 | 87.4 | 134.78 | 195676.2 | 77.54
0 | 0 | 86.2 | 135.39 | 196552.7 | 78.19
0 | 0 | 85 | 136.36 | 197966.7 | 78.55
0 | 0 | 83.8 | 137.04 | 198945.8 | 78.81
0 | 0 | 82.6 | 137.96 | 200286.9 | 79.54
0 | 0 | 81.4 | 138.52 | 201102 | 79.86
0 | 0 | 80.2 | 139.13 | 201983.2 | 80.19
0 | 0 | 78.9 | 140.09 | 203373.9 | 80.62
0 | 0 | 77.7 | 140.78 | 204380 | 81.04
0 | 0 | 76.5 | 141.38 | 205255.1 | 81.44
0 | 0 | 75.3 | 142.2 | 206450.6 | 82.06
0 | 0 | 74.1 | 143.01 | 207623.9 | 82.44
0 | 0 | 72.9 | 143.67 | 208575.7 | 82.75
0 | 0 | 71.7 | 144.46 | 209731.5 | 83.24
0 | 0 | 70.5 | 145.18 | 210763.6 | 83.59
0 | 0 | 69.3 | 145.85 | 211741.5 | 84.04
0 | 0 | 68.1 | 146.54 | 212751.6 | 84.45
0 | 0 | 66.9 | 147.48 | 214113.3 | 84.99
0 | 0 | 65.7 | 148.2 | 215156.2 | 85.34
0 | 0 | 64.5 | 149.16 | 216547.8 | 86.13
0 | 0 | 63.3 | 149.88 | 217593.4 | 86.35
0 | 0 | 62.1 | 150.44 | 218405.1 | 86.84
0 | 0 | 60.8 | 151.5 | 219941.2 | 87.29
0 | 0 | 59.7 | 152.06 | 220762 | 87.65
0 | 0 | 58.4 | 153.03 | 222167.9 | 88.16
0 | 0 | 57.2 | 153.61 | 223004.1 | 88.52
0 | 0 | 56 | 154.47 | 224262.7 | 88.91
0 | 0 | 54.8 | 155.18 | 225289.3 | 89.32
0 | 0 | 53.6 | 156.17 | 226728.4 | 89.97
0 | 0 | 52.4 | 157.01 | 227947.7 | 90.52
0 | 0 | 51.2 | 157.64 | 228865.1 | 90.79
0 | 0 | 50 | 158.57 | 230212.7 | 91.41
0 | 0 | 48.8 | 159.59 | 231688 | 91.95
0 | 0 | 47.6 | 160.03 | 232335.3 | 92.22
0 | 0 | 46.4 | 160.93 | 233640.4 | 92.87
0 | 0 | 45.1 | 161.79 | 234891.6 | 93.22
0 | 0 | 43.9 | 162.64 | 236111.5 | 93.66
0 | 0 | 42.7 | 163.55 | 237446.4 | 94.29
0 | 0 | 41.5 | 164.19 | 238370.9 | 94.69
0 | 0 | 40.3 | 165.2 | 239832.1 | 95.18
0 | 0 | 39.1 | 166.15 | 241216 | 95.7
0 | 0 | 37.9 | 166.51 | 241734.3 | 95.84
0 | 0 | 36.7 | 167.56 | 243258.7 | 96.4
0 | 0 | 35.5 | 168.14 | 244104.4 | 96.87
0 | 0 | 34.2 | 169.2 | 245641.2 | 97.45
0 | 0 | 33.1 | 169.77 | 246469 | 97.79
0 | 0 | 31.9 | 170.13 | 246988.9 | 98.11
0 | 0 | 30.6 | 171.08 | 248364.5 | 98.52
0 | 0 | 29.4 | 171.54 | 249035.2 | 98.79
0 | 0 | 28.2 | 172.22 | 250028.7 | 99.12
0 | 0 | 27 | 172.69 | 250715 | 99.48
0 | 0 | 25.8 | 172.92 | 251039.5 | 99.45
0 | 0 | 24.6 | 173.27 | 251545.2 | 99.75
0 | 0 | 23.4 | 173.25 | 251529.1 | 100
0 | 0 | 22.2 | 173.39 | 251730.5 | 99.86
0 | 0 | 21 | 173.24 | 251511.7 | 99.68
0 | 0 | 19.8 | 172.87 | 250965.2 | 99.58
0 | 0 | 18.6 | 172.03 | 249754.7 | 99.14
0 | 0 | 17.4 | 171.16 | 248482.3 | 98.58
0 | 0 | 16.2 | 169.65 | 246290.4 | 97.75
0 | 0 | 15 | 167.95 | 243830.2 | 96.62
0 | 0 | 13.8 | 165.5 | 240277.7 | 95.35
0 | 0 | 12.6 | 162.65 | 236132.2 | 93.69
0 | 0 | 11.4 | 159.1 | 230978.3 | 91.57
0 | 0 | 10.1 | 154.45 | 224235.5 | 88.89
0 | 0 | 8.9 | 148.87 | 216132.4 | 85.66
0 | 0 | 7.7 | 142.11 | 206307.2 | 82
0 | 0 | 6.5 | 133.31 | 193540.5 | 76.92
0 | 0 | 5.3 | 123.03 | 178607.1 | 70.92
0 | 0 | 4.1 | 109.63 | 159165.1 | 63.09
0 | 0 | 2.9 | 93.2 | 135307.1 | 53.76
0 | 0 | 1.7 | 76.35 | 110850.9 | 43.95
0 | 0 | 0.5 | 65.45 | 95023.8 | 37.6
0 | 0 | -0.5 | 60.46 | 87773.4 | 34.81