## 2_Profils énergie
Clinac 2
Measurement time: | 12/08/2022 18:12:50
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none |  | Valeur max | 67.1
Points [mm]:
Inline | Crossline | Depth | Normalized field |  | Current field | Ratio
-100 | 0 | 100 | 2.43 | 3.6215 | 6017.8 | 3.78
-99.1 | 0 | 100 | 2.98 | 4.4411 | 7363.3 | 4.43
-97.7 | 0 | 100 | 3.03 | 4.5156 | 7502 | 4.51
-96.6 | 0 | 100 | 3.16 | 4.7094 | 7801.4 | 4.69
-95.4 | 0 | 100 | 3.24 | 4.8286 | 8009.7 | 4.81
-94.1 | 0 | 100 | 3.32 | 4.9478 | 8198.1 | 4.92
-93 | 0 | 100 | 3.43 | 5.1118 | 8467.9 | 5.08
-91.8 | 0 | 100 | 3.57 | 5.3204 | 8821.7 | 5.3
-90.6 | 0 | 100 | 3.68 | 5.4844 | 9097.9 | 5.47
-89.4 | 0 | 100 | 3.76 | 5.6036 | 9301.1 | 5.59
-88.1 | 0 | 100 | 3.92 | 5.842 | 9693.8 | 5.82
-86.9 | 0 | 100 | 4.02 | 5.9911 | 9930.5 | 5.97
-85.7 | 0 | 100 | 4.18 | 6.2295 | 10341.6 | 6.21
-84.5 | 0 | 100 | 4.32 | 6.4382 | 10689.8 | 6.44
-83.3 | 0 | 100 | 4.43 | 6.6021 | 10940.6 | 6.6
-82.1 | 0 | 100 | 4.62 | 6.8852 | 11423.2 | 6.87
-80.9 | 0 | 100 | 4.72 | 7.0343 | 11681.8 | 7.02
-79.7 | 0 | 100 | 4.92 | 7.3323 | 12167.6 | 7.31
-78.5 | 0 | 100 | 5.09 | 7.5857 | 12589.6 | 7.56
-77.3 | 0 | 100 | 5.25 | 7.8241 | 12972.6 | 7.79
-76.1 | 0 | 100 | 5.45 | 8.1222 | 13463.5 | 8.1
-74.9 | 0 | 100 | 5.65 | 8.4203 | 13957.7 | 8.39
-73.7 | 0 | 100 | 5.8 | 8.6438 | 14339.6 | 8.63
-72.5 | 0 | 100 | 6.06 | 9.0313 | 14991.6 | 9.01
-71.3 | 0 | 100 | 6.33 | 9.4337 | 15644.6 | 9.4
-70.1 | 0 | 100 | 6.54 | 9.7466 | 16178.4 | 9.72
-68.9 | 0 | 100 | 6.88 | 10.2534 | 17008.1 | 10.22
-67.7 | 0 | 100 | 7.14 | 10.6408 | 17659.6 | 10.6
-66.5 | 0 | 100 | 7.49 | 11.1624 | 18510.3 | 11.14
-65.3 | 0 | 100 | 7.92 | 11.8033 | 19573.7 | 11.78
-64.1 | 0 | 100 | 8.44 | 12.5782 | 20854.5 | 12.54
-62.9 | 0 | 100 | 9.15 | 13.6364 | 22615.7 | 13.61
-61.7 | 0 | 100 | 10.29 | 15.3353 | 25446.7 | 15.29
-60.5 | 0 | 100 | 12.2 | 18.1818 | 30162.1 | 18.14
-59.3 | 0 | 100 | 15.3 | 22.8018 | 37836.1 | 22.75
-58 | 0 | 100 | 20.1 | 29.9553 | 49689.6 | 29.9
-56.8 | 0 | 100 | 26.33 | 39.2399 | 65108.2 | 39.15
-55.6 | 0 | 100 | 33.92 | 50.5514 | 83864.5 | 50.51
-54.4 | 0 | 100 | 42.21 | 62.9061 | 104351.1 | 62.76
-53.2 | 0 | 100 | 49.84 | 74.2772 | 123212 | 74.1
-52 | 0 | 100 | 55.97 | 83.4128 | 138366.6 | 83.2
-50.8 | 0 | 100 | 59.72 | 89.0015 | 147654.1 | 88.93
-49.6 | 0 | 100 | 61.87 | 92.2057 | 152956.1 | 92.23
-48.4 | 0 | 100 | 63.12 | 94.0686 | 156064.7 | 93.91
-47.2 | 0 | 100 | 63.78 | 95.0522 | 157694.1 | 94.91
-46 | 0 | 100 | 64.21 | 95.693 | 158744.5 | 95.5
-44.8 | 0 | 100 | 64.44 | 96.0358 | 159316.2 | 96.1
-43.6 | 0 | 100 | 64.63 | 96.3189 | 159781.4 | 96.33
-42.4 | 0 | 100 | 64.97 | 96.8256 | 160634.4 | 96.73
-41.2 | 0 | 100 | 65.12 | 97.0492 | 161010.6 | 96.94
-40 | 0 | 100 | 65.34 | 97.377 | 161543.3 | 97.23
-38.8 | 0 | 100 | 65.32 | 97.3472 | 161498.9 | 97.26
-37.6 | 0 | 100 | 65.44 | 97.5261 | 161799.3 | 97.57
-36.4 | 0 | 100 | 65.73 | 97.9583 | 162515.8 | 97.83
-35.2 | 0 | 100 | 65.95 | 98.2861 | 163057.7 | 97.99
-34 | 0 | 100 | 66.13 | 98.5544 | 163495.7 | 98.34
-32.8 | 0 | 100 | 66.15 | 98.5842 | 163542.7 | 98.48
-31.6 | 0 | 100 | 66.36 | 98.8972 | 164057.9 | 98.66
-30.4 | 0 | 100 | 66.53 | 99.1505 | 164473.8 | 98.97
-29.2 | 0 | 100 | 66.57 | 99.2101 | 164588.7 | 99.17
-28 | 0 | 100 | 66.74 | 99.4635 | 165015.9 | 99.27
-26.8 | 0 | 100 | 66.82 | 99.5827 | 165191.2 | 99.49
-25.5 | 0 | 100 | 66.71 | 99.4188 | 164924.1 | 99.31
-24.4 | 0 | 100 | 66.96 | 99.7914 | 165550.5 | 99.67
-23.1 | 0 | 100 | 66.93 | 99.7466 | 165476.1 | 99.79
-21.9 | 0 | 100 | 66.96 | 99.7914 | 165553 | 99.61
-20.7 | 0 | 100 | 66.98 | 99.8212 | 165608.6 | 99.68
-19.5 | 0 | 100 | 67 | 99.851 | 165652 | 99.66
-18.3 | 0 | 100 | 67.08 | 99.9702 | 165847.8 | 99.88
-17.1 | 0 | 100 | 67.04 | 99.9106 | 165748.2 | 99.74
-15.9 | 0 | 100 | 66.99 | 99.8361 | 165617.2 | 99.87
-14.7 | 0 | 100 | 67.1 | 100 | 165890.4 | 99.93
-13.5 | 0 | 100 | 66.98 | 99.8212 | 165599.5 | 99.8
-12.3 | 0 | 100 | 66.93 | 99.7466 | 165483 | 99.86
-11.1 | 0 | 100 | 66.99 | 99.8361 | 165627.2 | 100
-9.9 | 0 | 100 | 66.89 | 99.687 | 165373 | 99.72
-8.7 | 0 | 100 | 66.99 | 99.8361 | 165628.2 | 99.66
-7.5 | 0 | 100 | 66.89 | 99.687 | 165363.8 | 99.56
-6.3 | 0 | 100 | 66.8 | 99.5529 | 165157.8 | 99.51
-5.1 | 0 | 100 | 66.98 | 99.8212 | 165588.2 | 99.75
-3.8 | 0 | 100 | 66.83 | 99.5976 | 165232.3 | 99.55
-2.7 | 0 | 100 | 66.71 | 99.4188 | 164931.2 | 99.43
-1.4 | 0 | 100 | 66.65 | 99.3294 | 164781.7 | 99.11
-0.3 | 0 | 100 | 66.62 | 99.2846 | 164718.3 | 99.28
1 | 0 | 100 | 66.68 | 99.3741 | 164850.7 | 99.24
2.2 | 0 | 100 | 66.68 | 99.3741 | 164859.5 | 99.29
3.4 | 0 | 100 | 66.53 | 99.1505 | 164476.8 | 99.08
4.6 | 0 | 100 | 66.48 | 99.076 | 164354.3 | 99.1
5.8 | 0 | 100 | 66.62 | 99.2846 | 164704.4 | 99.24
7 | 0 | 100 | 66.46 | 99.0462 | 164323.9 | 99.09
8.2 | 0 | 100 | 66.64 | 99.3145 | 164766.1 | 99.27
9.4 | 0 | 100 | 66.65 | 99.3294 | 164776.5 | 99.38
10.6 | 0 | 100 | 66.52 | 99.1356 | 164451.7 | 99.15
11.8 | 0 | 100 | 66.58 | 99.225 | 164614.9 | 99.24
13 | 0 | 100 | 66.58 | 99.225 | 164600.8 | 99.29
14.2 | 0 | 100 | 66.55 | 99.1803 | 164545.8 | 99.3
15.4 | 0 | 100 | 66.52 | 99.1356 | 164468 | 99.29
16.6 | 0 | 100 | 66.5 | 99.1058 | 164414.9 | 99.05
17.8 | 0 | 100 | 66.44 | 99.0164 | 164259.1 | 98.88
19 | 0 | 100 | 66.51 | 99.1207 | 164446 | 98.84
20.2 | 0 | 100 | 66.36 | 98.8972 | 164075.3 | 98.83
21.4 | 0 | 100 | 66.28 | 98.7779 | 163856.4 | 98.77
22.6 | 0 | 100 | 66.1 | 98.5097 | 163433.9 | 98.45
23.8 | 0 | 100 | 66.01 | 98.3756 | 163189.2 | 98.38
25 | 0 | 100 | 65.99 | 98.3458 | 163152.9 | 98.27
26.2 | 0 | 100 | 66.05 | 98.4352 | 163310.6 | 98.48
27.4 | 0 | 100 | 65.85 | 98.1371 | 162796.9 | 98.16
28.7 | 0 | 100 | 65.75 | 97.9881 | 162550.3 | 97.94
29.8 | 0 | 100 | 65.63 | 97.8092 | 162264.2 | 97.81
31 | 0 | 100 | 65.44 | 97.5261 | 161779.4 | 97.41
32.3 | 0 | 100 | 65.28 | 97.2876 | 161400.1 | 97.26
33.4 | 0 | 100 | 65.22 | 97.1982 | 161236.3 | 97.24
34.7 | 0 | 100 | 65.01 | 96.8852 | 160737.9 | 97.07
35.9 | 0 | 100 | 64.78 | 96.5425 | 160150.6 | 96.5
37.1 | 0 | 100 | 64.71 | 96.4382 | 159980.3 | 96.4
38.3 | 0 | 100 | 64.61 | 96.2891 | 159749.8 | 96.24
39.5 | 0 | 100 | 64.35 | 95.9016 | 159099 | 96
40.7 | 0 | 100 | 64.18 | 95.6483 | 158674.8 | 95.71
41.9 | 0 | 100 | 63.85 | 95.1565 | 157854.8 | 95.26
43.1 | 0 | 100 | 63.7 | 94.9329 | 157493.9 | 94.9
44.3 | 0 | 100 | 63.36 | 94.4262 | 156657.5 | 94.59
45.5 | 0 | 100 | 62.85 | 93.6662 | 155375 | 93.58
46.7 | 0 | 100 | 62.32 | 92.8763 | 154069.3 | 92.74
47.9 | 0 | 100 | 61.3 | 91.3562 | 151555.5 | 91.37
49.1 | 0 | 100 | 59.78 | 89.0909 | 147802.8 | 89.09
50.3 | 0 | 100 | 56.9 | 84.7988 | 140672.5 | 84.87
51.5 | 0 | 100 | 52.65 | 78.465 | 130164.1 | 78.45
52.7 | 0 | 100 | 47.05 | 70.1192 | 116319.4 | 70.12
53.9 | 0 | 100 | 39.68 | 59.1356 | 98110 | 59.23
55.1 | 0 | 100 | 31.62 | 47.1237 | 78172.9 | 47.16
56.3 | 0 | 100 | 23.99 | 35.7526 | 59312.8 | 35.79
57.6 | 0 | 100 | 17.64 | 26.2891 | 43602.6 | 26.32
58.8 | 0 | 100 | 13.28 | 19.7914 | 32838.7 | 19.77
60 | 0 | 100 | 10.87 | 16.1997 | 26864.8 | 16.21
61.2 | 0 | 100 | 9.49 | 14.1431 | 23454.1 | 14.14
62.4 | 0 | 100 | 8.66 | 12.9061 | 21415.8 | 12.9
63.6 | 0 | 100 | 8.03 | 11.9672 | 19856.9 | 11.96
64.8 | 0 | 100 | 7.56 | 11.2668 | 18702.2 | 11.26
66 | 0 | 100 | 7.19 | 10.7154 | 17781 | 10.73
67.2 | 0 | 100 | 6.86 | 10.2235 | 16957.3 | 10.22
68.4 | 0 | 100 | 6.54 | 9.7466 | 16175 | 9.75
69.6 | 0 | 100 | 6.37 | 9.4933 | 15741.7 | 9.49
70.8 | 0 | 100 | 6.06 | 9.0313 | 14971 | 9.02
72 | 0 | 100 | 5.86 | 8.7332 | 14493.3 | 8.74
73.2 | 0 | 100 | 5.65 | 8.4203 | 13979.3 | 8.44
74.4 | 0 | 100 | 5.48 | 8.1669 | 13539.2 | 8.17
75.6 | 0 | 100 | 5.28 | 7.8689 | 13061.8 | 7.87
76.8 | 0 | 100 | 5.08 | 7.5708 | 12551.8 | 7.55
78.1 | 0 | 100 | 4.88 | 7.2727 | 12069.3 | 7.28
79.3 | 0 | 100 | 4.79 | 7.1386 | 11833 | 7.14
80.5 | 0 | 100 | 4.57 | 6.8107 | 11288.3 | 6.81
81.7 | 0 | 100 | 4.41 | 6.5723 | 10906.6 | 6.58
82.9 | 0 | 100 | 4.26 | 6.3487 | 10542.5 | 6.37
84.1 | 0 | 100 | 4.18 | 6.2295 | 10343.8 | 6.24
85.3 | 0 | 100 | 4.04 | 6.0209 | 9992.8 | 6.03
86.5 | 0 | 100 | 3.92 | 5.842 | 9685.3 | 5.83
87.7 | 0 | 100 | 3.84 | 5.7228 | 9498.3 | 5.73
88.9 | 0 | 100 | 3.67 | 5.4694 | 9064.5 | 5.47
90.1 | 0 | 100 | 3.51 | 5.231 | 8684.6 | 5.24
91.3 | 0 | 100 | 3.43 | 5.1118 | 8491.5 | 5.11
92.5 | 0 | 100 | 3.34 | 4.9776 | 8265.1 | 4.99
93.7 | 0 | 100 | 3.25 | 4.8435 | 8032.2 | 4.85
94.9 | 0 | 100 | 3.18 | 4.7392 | 7874.3 | 4.75
96.1 | 0 | 100 | 3.04 | 4.5306 | 7527.2 | 4.55
97.3 | 0 | 100 | 2.93 | 4.3666 | 7240.6 | 4.37
98.5 | 0 | 100 | 2.93 | 4.3666 | 7255.9 | 4.37
100 | 0 | 100 | 2.81 | 4.1878 | 6937.7 | 4.18
Measurement time: | 12/08/2022 18:18:10
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none |  | Valeur max | 82.11
Points [mm]:
Inline | Crossline | Depth | Normalized field |  | Current field | Ratio
-100 | 0 | 100 | 2.2 | 2.6793 | 5430.7 | 3.3
-99.1 | 0 | 100 | 2.41 | 2.9351 | 5946.4 | 2.93
-97.6 | 0 | 100 | 2.56 | 3.1178 | 6333.4 | 3.13
-96.6 | 0 | 100 | 2.61 | 3.1787 | 6453.4 | 3.18
-95.4 | 0 | 100 | 2.69 | 3.2761 | 6643.4 | 3.28
-94.1 | 0 | 100 | 2.81 | 3.4222 | 6954.2 | 3.43
-93 | 0 | 100 | 2.91 | 3.544 | 7200.5 | 3.55
-91.8 | 0 | 100 | 3.01 | 3.6658 | 7447.9 | 3.68
-90.5 | 0 | 100 | 3.17 | 3.8607 | 7830.4 | 3.85
-89.4 | 0 | 100 | 3.22 | 3.9216 | 7952 | 3.93
-88.1 | 0 | 100 | 3.33 | 4.0555 | 8220.9 | 4.06
-86.9 | 0 | 100 | 3.46 | 4.2139 | 8551.1 | 4.21
-85.7 | 0 | 100 | 3.56 | 4.3356 | 8809.9 | 4.34
-84.5 | 0 | 100 | 3.73 | 4.5427 | 9220.1 | 4.53
-83.3 | 0 | 100 | 3.84 | 4.6767 | 9502.5 | 4.7
-82.1 | 0 | 100 | 3.99 | 4.8593 | 9866.7 | 4.85
-80.9 | 0 | 100 | 4.2 | 5.1151 | 10381.3 | 5.12
-79.7 | 0 | 100 | 4.38 | 5.3343 | 10837.3 | 5.34
-78.5 | 0 | 100 | 4.51 | 5.4926 | 11161.1 | 5.51
-77.3 | 0 | 100 | 4.72 | 5.7484 | 11681.7 | 5.77
-76.1 | 0 | 100 | 4.94 | 6.0163 | 12215.4 | 5.99
-74.9 | 0 | 100 | 5.14 | 6.2599 | 12716.9 | 6.29
-73.7 | 0 | 100 | 5.48 | 6.674 | 13546.9 | 6.69
-72.5 | 0 | 100 | 5.86 | 7.1368 | 14482.4 | 7.14
-71.3 | 0 | 100 | 6.2 | 7.5508 | 15326.8 | 7.58
-70.1 | 0 | 100 | 6.72 | 8.1841 | 16608.4 | 8.18
-68.9 | 0 | 100 | 7.19 | 8.7565 | 17770.2 | 8.8
-67.7 | 0 | 100 | 7.79 | 9.4873 | 19263.2 | 9.53
-66.5 | 0 | 100 | 8.61 | 10.4859 | 21285.7 | 10.52
-65.3 | 0 | 100 | 9.61 | 11.7038 | 23758.6 | 11.74
-64.1 | 0 | 100 | 10.78 | 13.1287 | 26660.5 | 13.15
-62.9 | 0 | 100 | 12.18 | 14.8338 | 30104.8 | 14.9
-61.6 | 0 | 100 | 14.27 | 17.3791 | 35286.2 | 17.37
-60.4 | 0 | 100 | 17.01 | 20.7161 | 42052.5 | 20.75
-59.2 | 0 | 100 | 20.85 | 25.3928 | 51543 | 25.45
-58 | 0 | 100 | 26.21 | 31.9206 | 64789.7 | 32.18
-56.8 | 0 | 100 | 33.33 | 40.5919 | 82399.2 | 40.65
-55.6 | 0 | 100 | 42.07 | 51.2361 | 103999.5 | 51.38
-54.4 | 0 | 100 | 51.26 | 62.4284 | 126735.7 | 62.67
-53.2 | 0 | 100 | 59.56 | 72.5368 | 147250.7 | 72.67
-52 | 0 | 100 | 65.97 | 80.3434 | 163103.7 | 80.69
-50.8 | 0 | 100 | 70.08 | 85.3489 | 173256 | 85.8
-49.6 | 0 | 100 | 72.99 | 88.8929 | 180449.1 | 89.21
-48.4 | 0 | 100 | 75.22 | 91.6088 | 185961.1 | 91.61
-47.2 | 0 | 100 | 76.57 | 93.253 | 189298.5 | 93.52
-46 | 0 | 100 | 77.55 | 94.4465 | 191721 | 95.01
-44.8 | 0 | 100 | 78.64 | 95.774 | 194433.8 | 96.15
-43.6 | 0 | 100 | 79.19 | 96.4438 | 195785 | 96.72
-42.4 | 0 | 100 | 79.87 | 97.272 | 197473.5 | 97.35
-41.2 | 0 | 100 | 80.2 | 97.6739 | 198293.9 | 98.11
-40 | 0 | 100 | 80.53 | 98.0758 | 199093.8 | 98.35
-38.8 | 0 | 100 | 81 | 98.6482 | 200260.2 | 99.04
-37.6 | 0 | 100 | 81.15 | 98.8308 | 200625.7 | 98.97
-36.4 | 0 | 100 | 81.37 | 99.0988 | 201183.7 | 99.44
-35.2 | 0 | 100 | 81.67 | 99.4641 | 201910.8 | 99.66
-34 | 0 | 100 | 81.64 | 99.4276 | 201841.1 | 99.78
-32.8 | 0 | 100 | 81.51 | 99.2693 | 201517.3 | 99.73
-31.6 | 0 | 100 | 81.77 | 99.5859 | 202155.6 | 99.89
-30.4 | 0 | 100 | 81.89 | 99.7321 | 202452.6 | 99.99
-29.2 | 0 | 100 | 81.92 | 99.7686 | 202526.1 | 99.93
-28 | 0 | 100 | 81.66 | 99.452 | 201893.5 | 99.8
-26.7 | 0 | 100 | 82.11 | 100 | 203005.2 | 99.9
-25.5 | 0 | 100 | 81.57 | 99.3423 | 201675.9 | 100
-24.3 | 0 | 100 | 81.79 | 99.6103 | 202201.9 | 99.73
-23.1 | 0 | 100 | 81.48 | 99.2327 | 201453.6 | 99.83
-21.9 | 0 | 100 | 81.67 | 99.4641 | 201916.3 | 99.5
-20.7 | 0 | 100 | 81.68 | 99.4763 | 201931.8 | 99.78
-19.5 | 0 | 100 | 81.37 | 99.0988 | 201185.1 | 99.54
-18.3 | 0 | 100 | 81.63 | 99.4154 | 201806.3 | 99.81
-17.1 | 0 | 100 | 81.53 | 99.2936 | 201580.1 | 99.49
-15.9 | 0 | 100 | 81.4 | 99.1353 | 201253.9 | 99.78
-14.7 | 0 | 100 | 81.46 | 99.2084 | 201388 | 99.63
-13.5 | 0 | 100 | 81.34 | 99.0622 | 201106.2 | 99.53
-12.3 | 0 | 100 | 81.51 | 99.2693 | 201525.7 | 99.64
-11.1 | 0 | 100 | 81.55 | 99.318 | 201627.6 | 99.38
-9.9 | 0 | 100 | 81.5 | 99.2571 | 201489.5 | 99.31
-8.7 | 0 | 100 | 81.16 | 98.843 | 200650.3 | 99.37
-7.5 | 0 | 100 | 81.33 | 99.0501 | 201087.4 | 99.33
-6.3 | 0 | 100 | 81.01 | 98.6603 | 200293 | 98.99
-5.1 | 0 | 100 | 81.35 | 99.0744 | 201130.6 | 99.08
-3.9 | 0 | 100 | 80.91 | 98.5385 | 200032.2 | 99.22
-2.7 | 0 | 100 | 81.24 | 98.9404 | 200842.8 | 99.17
-1.4 | 0 | 100 | 80.88 | 98.502 | 199967.2 | 99.16
-0.3 | 0 | 100 | 81.26 | 98.9648 | 200908.9 | 99.11
0.9 | 0 | 100 | 80.9 | 98.5264 | 200007.8 | 98.95
2.1 | 0 | 100 | 81.11 | 98.7821 | 200521 | 98.89
3.4 | 0 | 100 | 80.82 | 98.4289 | 199812.9 | 99.13
4.6 | 0 | 100 | 81.2 | 98.8917 | 200754.2 | 99.16
5.8 | 0 | 100 | 80.9 | 98.5264 | 200001.6 | 99.01
7 | 0 | 100 | 81.17 | 98.8552 | 200687.6 | 99.02
8.2 | 0 | 100 | 80.82 | 98.4289 | 199814.1 | 98.98
9.4 | 0 | 100 | 81.38 | 99.1109 | 201211.6 | 99.29
10.6 | 0 | 100 | 80.95 | 98.5873 | 200146 | 99.28
11.8 | 0 | 100 | 81.25 | 98.9526 | 200877 | 99.22
13 | 0 | 100 | 81.05 | 98.709 | 200396.2 | 99.24
14.2 | 0 | 100 | 80.94 | 98.5751 | 200118.9 | 99.06
15.4 | 0 | 100 | 81.23 | 98.9283 | 200820 | 99.12
16.6 | 0 | 100 | 80.68 | 98.2584 | 199464.6 | 98.83
17.8 | 0 | 100 | 81.11 | 98.7821 | 200543 | 98.87
19 | 0 | 100 | 80.67 | 98.2463 | 199441.8 | 98.98
20.2 | 0 | 100 | 81.02 | 98.6725 | 200311.7 | 98.95
21.4 | 0 | 100 | 81.17 | 98.8552 | 200672.9 | 98.87
22.6 | 0 | 100 | 80.69 | 98.2706 | 199505.9 | 99.09
23.8 | 0 | 100 | 80.95 | 98.5873 | 200144.4 | 98.98
25 | 0 | 100 | 80.95 | 98.5873 | 200139.5 | 98.93
26.2 | 0 | 100 | 80.76 | 98.3559 | 199659 | 98.86
27.4 | 0 | 100 | 81.04 | 98.6969 | 200352.6 | 98.91
28.6 | 0 | 100 | 80.83 | 98.4411 | 199850.4 | 99.33
29.8 | 0 | 100 | 80.91 | 98.5385 | 200045.4 | 99.15
31 | 0 | 100 | 80.88 | 98.502 | 199957 | 99.23
32.2 | 0 | 100 | 81.06 | 98.7212 | 200416.4 | 99.34
33.4 | 0 | 100 | 80.95 | 98.5873 | 200139.7 | 99
34.6 | 0 | 100 | 80.58 | 98.1366 | 199232.6 | 98.91
35.8 | 0 | 100 | 80.79 | 98.3924 | 199737.2 | 98.69
37 | 0 | 100 | 80.34 | 97.8444 | 198628.2 | 98.52
38.2 | 0 | 100 | 80.24 | 97.7226 | 198383.3 | 98.22
39.4 | 0 | 100 | 79.57 | 96.9066 | 196717.2 | 97.87
40.6 | 0 | 100 | 79.38 | 96.6752 | 196254 | 97.38
41.8 | 0 | 100 | 78.87 | 96.0541 | 194995.2 | 96.49
43.1 | 0 | 100 | 78.62 | 95.7496 | 194385.9 | 95.85
44.3 | 0 | 100 | 77.26 | 94.0933 | 191019.5 | 94.79
45.5 | 0 | 100 | 76.38 | 93.0216 | 188840.1 | 93.79
46.7 | 0 | 100 | 75.28 | 91.6819 | 186117.3 | 92.07
47.9 | 0 | 100 | 73.47 | 89.4775 | 181640.9 | 90.1
49.1 | 0 | 100 | 71.25 | 86.7738 | 176145.3 | 87.04
50.3 | 0 | 100 | 67.56 | 82.2799 | 167034.2 | 83.01
51.5 | 0 | 100 | 63.09 | 76.836 | 155990.8 | 77.22
52.7 | 0 | 100 | 56.87 | 69.2607 | 140614.6 | 69.64
53.9 | 0 | 100 | 49.1 | 59.7978 | 121403.2 | 59.97
55.1 | 0 | 100 | 39.99 | 48.703 | 98870.1 | 49.05
56.3 | 0 | 100 | 31.56 | 38.4362 | 78035.8 | 38.53
57.5 | 0 | 100 | 24.16 | 29.4239 | 59728 | 29.62
58.7 | 0 | 100 | 18.83 | 22.9327 | 46544.6 | 23.06
59.9 | 0 | 100 | 15.51 | 18.8893 | 38343.7 | 18.99
61.1 | 0 | 100 | 13.09 | 15.942 | 32354.3 | 15.99
62.3 | 0 | 100 | 11.32 | 13.7864 | 27991.9 | 13.84
63.5 | 0 | 100 | 10.03 | 12.2153 | 24787 | 12.28
64.7 | 0 | 100 | 8.95 | 10.9 | 22134.5 | 10.95
65.9 | 0 | 100 | 8.04 | 9.7917 | 19871.6 | 9.87
67.1 | 0 | 100 | 7.41 | 9.0245 | 18316.6 | 9.06
68.3 | 0 | 100 | 6.85 | 8.3425 | 16932.1 | 8.4
69.5 | 0 | 100 | 6.39 | 7.7822 | 15792.6 | 7.81
70.8 | 0 | 100 | 5.9 | 7.1855 | 14583.4 | 7.22
72 | 0 | 100 | 5.58 | 6.7958 | 13788.3 | 6.82
73.2 | 0 | 100 | 5.3 | 6.4548 | 13104.8 | 6.49
74.3 | 0 | 100 | 4.99 | 6.0772 | 12335.8 | 6.11
75.5 | 0 | 100 | 4.75 | 5.7849 | 11742.4 | 5.8
76.8 | 0 | 100 | 4.51 | 5.4926 | 11162.1 | 5.55
78 | 0 | 100 | 4.35 | 5.2978 | 10745.7 | 5.33
79.2 | 0 | 100 | 4.18 | 5.0907 | 10331.6 | 5.12
80.4 | 0 | 100 | 3.99 | 4.8593 | 9855.1 | 4.87
81.6 | 0 | 100 | 3.88 | 4.7254 | 9597.3 | 4.75
82.8 | 0 | 100 | 3.69 | 4.494 | 9118.8 | 4.53
84 | 0 | 100 | 3.56 | 4.3356 | 8794.5 | 4.35
85.2 | 0 | 100 | 3.47 | 4.226 | 8577.2 | 4.25
86.4 | 0 | 100 | 3.32 | 4.0434 | 8213.5 | 4.06
87.6 | 0 | 100 | 3.25 | 3.9581 | 8046 | 3.99
88.8 | 0 | 100 | 3.1 | 3.7754 | 7665.1 | 3.8
90 | 0 | 100 | 2.99 | 3.6415 | 7388.1 | 3.66
91.2 | 0 | 100 | 2.9 | 3.5318 | 7170.3 | 3.54
92.4 | 0 | 100 | 2.78 | 3.3857 | 6868.8 | 3.41
93.6 | 0 | 100 | 2.71 | 3.3005 | 6696.2 | 3.32
94.8 | 0 | 100 | 2.62 | 3.1908 | 6487.5 | 3.22
96 | 0 | 100 | 2.53 | 3.0812 | 6260.5 | 3.1
97.2 | 0 | 100 | 2.43 | 2.9594 | 6016.2 | 2.99
98.4 | 0 | 100 | 2.4 | 2.9229 | 5923.3 | 2.94
100 | 0 | 100 | 2.25 | 2.7402 | 5564.8 | 2.76
TrueBeam
Measurement time: | 30/08/2022 16:29:49
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-155 | 0 | 100 | 97.82 | 3347 | 1.98
-154.1 | 0 | 100 | 104.92 | 3590.1 | 2.14
-152.6 | 0 | 100 | 107.93 | 3692.8 | 2.21
-151.5 | 0 | 100 | 107.71 | 3685.5 | 2.2
-150.4 | 0 | 100 | 110.71 | 3787.9 | 2.26
-149.1 | 0 | 100 | 113.19 | 3872.8 | 2.31
-147.9 | 0 | 100 | 112.91 | 3863.2 | 2.31
-146.7 | 0 | 100 | 115.21 | 3941.9 | 2.35
-145.5 | 0 | 100 | 118.72 | 4062 | 2.43
-144.3 | 0 | 100 | 121.1 | 4143.7 | 2.47
-143.1 | 0 | 100 | 121.05 | 4142 | 2.47
-141.9 | 0 | 100 | 121.73 | 4164.9 | 2.49
-140.7 | 0 | 100 | 124.22 | 4250.4 | 2.53
-139.5 | 0 | 100 | 125.3 | 4287.4 | 2.56
-138.3 | 0 | 100 | 125.42 | 4291.3 | 2.56
-137.1 | 0 | 100 | 130.8 | 4475.5 | 2.68
-135.9 | 0 | 100 | 131.48 | 4498.7 | 2.68
-134.7 | 0 | 100 | 132.49 | 4533.1 | 2.7
-133.5 | 0 | 100 | 135.51 | 4636.7 | 2.76
-132.3 | 0 | 100 | 136.41 | 4667.3 | 2.78
-131.1 | 0 | 100 | 141.39 | 4837.7 | 2.88
-129.9 | 0 | 100 | 142.63 | 4880.2 | 2.91
-128.7 | 0 | 100 | 145.56 | 4980.5 | 2.97
-127.5 | 0 | 100 | 151.43 | 5181.4 | 3.09
-126.3 | 0 | 100 | 152.81 | 5228.6 | 3.12
-125.1 | 0 | 100 | 152.26 | 5209.8 | 3.11
-123.9 | 0 | 100 | 155.96 | 5336.3 | 3.18
-122.7 | 0 | 100 | 161.76 | 5534.6 | 3.3
-121.5 | 0 | 100 | 161.02 | 5509.4 | 3.29
-120.3 | 0 | 100 | 166.78 | 5706.5 | 3.4
-119.1 | 0 | 100 | 166.77 | 5706 | 3.39
-117.8 | 0 | 100 | 170.75 | 5842.3 | 3.48
-116.7 | 0 | 100 | 175.5 | 6004.7 | 3.58
-115.4 | 0 | 100 | 178.38 | 6103.3 | 3.64
-114.2 | 0 | 100 | 185.21 | 6337.2 | 3.78
-113 | 0 | 100 | 189 | 6466.7 | 3.85
-111.8 | 0 | 100 | 189.41 | 6480.9 | 3.86
-110.6 | 0 | 100 | 195.32 | 6683 | 3.98
-109.4 | 0 | 100 | 199.33 | 6820.1 | 4.06
-108.2 | 0 | 100 | 203.52 | 6963.5 | 4.15
-107 | 0 | 100 | 211.38 | 7232.4 | 4.31
-105.8 | 0 | 100 | 215.4 | 7370.1 | 4.39
-104.6 | 0 | 100 | 218.59 | 7479.2 | 4.46
-103.4 | 0 | 100 | 225.64 | 7720.3 | 4.6
-102.2 | 0 | 100 | 232.07 | 7940.5 | 4.74
-101 | 0 | 100 | 236.14 | 8079.6 | 4.81
-99.8 | 0 | 100 | 241.8 | 8273.3 | 4.92
-98.6 | 0 | 100 | 246.67 | 8439.9 | 5.03
-97.4 | 0 | 100 | 252.61 | 8643.3 | 5.15
-96.2 | 0 | 100 | 258.81 | 8855.3 | 5.28
-95 | 0 | 100 | 264.84 | 9061.8 | 5.4
-93.8 | 0 | 100 | 273.5 | 9358.1 | 5.58
-92.6 | 0 | 100 | 280.57 | 9599.8 | 5.72
-91.3 | 0 | 100 | 291.65 | 9979 | 5.95
-90.2 | 0 | 100 | 299.22 | 10238.1 | 6.11
-89 | 0 | 100 | 304.87 | 10431.5 | 6.21
-87.7 | 0 | 100 | 311.46 | 10656.7 | 6.35
-86.6 | 0 | 100 | 325.18 | 11126.2 | 6.62
-85.3 | 0 | 100 | 333.67 | 11416.6 | 6.79
-84.1 | 0 | 100 | 344.21 | 11777.3 | 7.01
-82.9 | 0 | 100 | 356.11 | 12184.6 | 7.26
-81.7 | 0 | 100 | 364.08 | 12457.1 | 7.42
-80.5 | 0 | 100 | 375.78 | 12857.6 | 7.65
-79.3 | 0 | 100 | 388.87 | 13305.5 | 7.92
-78.1 | 0 | 100 | 401.24 | 13728.8 | 8.18
-76.9 | 0 | 100 | 413.18 | 14137.2 | 8.42
-75.7 | 0 | 100 | 426.2 | 14582.8 | 8.68
-74.5 | 0 | 100 | 440.18 | 15061 | 8.97
-73.3 | 0 | 100 | 453.4 | 15513.5 | 9.26
-72.1 | 0 | 100 | 472.25 | 16158.4 | 9.63
-70.9 | 0 | 100 | 486.25 | 16637.3 | 9.91
-69.7 | 0 | 100 | 502.98 | 17209.7 | 10.26
-68.5 | 0 | 100 | 525.28 | 17972.9 | 10.71
-67.3 | 0 | 100 | 546.75 | 18707.3 | 11.13
-66.1 | 0 | 100 | 576.98 | 19741.9 | 11.77
-64.9 | 0 | 100 | 610.91 | 20902.7 | 12.45
-63.7 | 0 | 100 | 657.46 | 22495.5 | 13.4
-62.5 | 0 | 100 | 730.37 | 24990.2 | 14.86
-61.3 | 0 | 100 | 844.53 | 28896.1 | 17.21
-60.1 | 0 | 100 | 1036.64 | 35469.3 | 21.12
-58.9 | 0 | 100 | 1328.49 | 45455.1 | 27.05
-57.6 | 0 | 100 | 1733.18 | 59301.8 | 35.27
-56.4 | 0 | 100 | 2226.35 | 76176.1 | 45.43
-55.2 | 0 | 100 | 2807.32 | 96054.1 | 57.29
-54 | 0 | 100 | 3371.3 | 115351.3 | 68.71
-52.8 | 0 | 100 | 3852.93 | 131830.5 | 78.55
-51.6 | 0 | 100 | 4207.02 | 143945.7 | 85.82
-50.4 | 0 | 100 | 4398.35 | 150492.5 | 89.65
-49.2 | 0 | 100 | 4505.77 | 154167.7 | 91.81
-48 | 0 | 100 | 4584.95 | 156877.1 | 93.34
-46.8 | 0 | 100 | 4638.44 | 158707.2 | 94.36
-45.6 | 0 | 100 | 4653.46 | 159220.9 | 94.94
-44.4 | 0 | 100 | 4681.66 | 160186 | 95.43
-43.2 | 0 | 100 | 4707.85 | 161082.2 | 95.95
-42 | 0 | 100 | 4717.17 | 161401.1 | 96.16
-40.8 | 0 | 100 | 4728.38 | 161784.4 | 96.33
-39.6 | 0 | 100 | 4750.88 | 162554.4 | 96.79
-38.4 | 0 | 100 | 4767 | 163105.9 | 97.2
-37.2 | 0 | 100 | 4775.2 | 163386.5 | 97.29
-36 | 0 | 100 | 4781.4 | 163598.6 | 97.47
-34.8 | 0 | 100 | 4798.42 | 164180.9 | 97.85
-33.6 | 0 | 100 | 4809.75 | 164568.8 | 98.01
-32.4 | 0 | 100 | 4813.32 | 164690.6 | 98.01
-31.2 | 0 | 100 | 4826.27 | 165133.7 | 98.5
-30 | 0 | 100 | 4835.45 | 165448 | 98.56
-28.8 | 0 | 100 | 4843.08 | 165709.2 | 98.73
-27.6 | 0 | 100 | 4856.81 | 166179 | 98.9
-26.4 | 0 | 100 | 4860.91 | 166319.2 | 99.01
-25.1 | 0 | 100 | 4864.16 | 166430.5 | 99.24
-24 | 0 | 100 | 4871.4 | 166677.9 | 99.1
-22.8 | 0 | 100 | 4874.44 | 166782.2 | 99.34
-21.6 | 0 | 100 | 4878.3 | 166914.2 | 99.59
-20.4 | 0 | 100 | 4893.58 | 167437.1 | 99.79
-19.1 | 0 | 100 | 4889.49 | 167297.1 | 99.69
-18 | 0 | 100 | 4892.96 | 167415.8 | 99.78
-16.7 | 0 | 100 | 4891.16 | 167354 | 99.74
-15.5 | 0 | 100 | 4891.44 | 167363.9 | 99.71
-14.3 | 0 | 100 | 4885.61 | 167164.2 | 99.7
-13.1 | 0 | 100 | 4899.41 | 167636.4 | 99.72
-11.9 | 0 | 100 | 4892.25 | 167391.3 | 99.67
-10.7 | 0 | 100 | 4899.41 | 167636.6 | 100
-9.5 | 0 | 100 | 4899.38 | 167635.5 | 99.77
-8.3 | 0 | 100 | 4892.85 | 167411.9 | 99.61
-7.1 | 0 | 100 | 4901.85 | 167720.1 | 99.78
-5.9 | 0 | 100 | 4900.64 | 167678.4 | 99.91
-4.7 | 0 | 100 | 4895.56 | 167504.7 | 99.72
-3.5 | 0 | 100 | 4892.79 | 167409.8 | 99.74
-2.3 | 0 | 100 | 4894.91 | 167482.6 | 99.73
-1.1 | 0 | 100 | 4896.02 | 167520.4 | 99.78
0.1 | 0 | 100 | 4888.64 | 167268 | 99.66
1.3 | 0 | 100 | 4891.42 | 167363.2 | 99.75
2.5 | 0 | 100 | 4891.53 | 167366.7 | 99.9
3.7 | 0 | 100 | 4887.86 | 167241.4 | 99.68
4.9 | 0 | 100 | 4902.29 | 167734.9 | 99.97
6.1 | 0 | 100 | 4890.69 | 167338.1 | 99.63
7.3 | 0 | 100 | 4898.39 | 167601.5 | 99.95
8.5 | 0 | 100 | 4888.96 | 167279 | 99.81
9.7 | 0 | 100 | 4890.84 | 167343.1 | 99.78
10.9 | 0 | 100 | 4896.41 | 167533.7 | 99.77
12.1 | 0 | 100 | 4888.9 | 167276.8 | 99.6
13.3 | 0 | 100 | 4891.49 | 167365.5 | 99.76
14.5 | 0 | 100 | 4892.41 | 167396.9 | 99.8
15.7 | 0 | 100 | 4885.76 | 167169.6 | 99.75
16.9 | 0 | 100 | 4888.39 | 167259.3 | 99.71
18.1 | 0 | 100 | 4885.12 | 167147.4 | 99.59
19.3 | 0 | 100 | 4888.34 | 167257.6 | 99.54
20.5 | 0 | 100 | 4882.39 | 167054.2 | 99.43
21.7 | 0 | 100 | 4878.13 | 166908.3 | 99.49
22.9 | 0 | 100 | 4875.19 | 166807.8 | 99.52
24.1 | 0 | 100 | 4874.35 | 166779.1 | 99.38
25.4 | 0 | 100 | 4857.2 | 166192.3 | 99.04
26.6 | 0 | 100 | 4854.68 | 166105.8 | 99.05
27.7 | 0 | 100 | 4843.69 | 165730.1 | 98.67
29 | 0 | 100 | 4844.71 | 165764.9 | 98.92
30.2 | 0 | 100 | 4830.07 | 165264.1 | 98.36
31.4 | 0 | 100 | 4822.46 | 165003.7 | 98.22
32.6 | 0 | 100 | 4813.12 | 164684.1 | 98.18
33.7 | 0 | 100 | 4800.17 | 164240.9 | 97.72
35 | 0 | 100 | 4790.88 | 163923.1 | 97.79
36.2 | 0 | 100 | 4777.32 | 163459.1 | 97.34
37.4 | 0 | 100 | 4762.1 | 162938.3 | 97.03
38.6 | 0 | 100 | 4755.28 | 162704.8 | 96.88
39.8 | 0 | 100 | 4739.59 | 162167.9 | 96.49
41 | 0 | 100 | 4723.54 | 161618.9 | 96.31
42.2 | 0 | 100 | 4711.5 | 161206.8 | 96.07
43.4 | 0 | 100 | 4682.66 | 160220.1 | 95.36
44.6 | 0 | 100 | 4670.4 | 159800.8 | 95.27
45.8 | 0 | 100 | 4635.93 | 158621.3 | 94.52
47 | 0 | 100 | 4587.42 | 156961.4 | 93.53
48.2 | 0 | 100 | 4517.63 | 154573.7 | 92.06
49.4 | 0 | 100 | 4411.79 | 150952.1 | 90.01
50.6 | 0 | 100 | 4216.55 | 144271.9 | 86.09
51.8 | 0 | 100 | 3930.02 | 134468.1 | 80.06
53 | 0 | 100 | 3538.58 | 121074.9 | 72.14
54.2 | 0 | 100 | 3036.19 | 103885 | 61.84
55.4 | 0 | 100 | 2468.48 | 84460.5 | 50.42
56.6 | 0 | 100 | 1899.72 | 65000 | 38.78
57.8 | 0 | 100 | 1423.45 | 48704.2 | 29.05
59.1 | 0 | 100 | 1067.12 | 36512.1 | 21.74
60.3 | 0 | 100 | 864.71 | 29586.5 | 17.63
61.5 | 0 | 100 | 744.18 | 25462.5 | 15.17
62.7 | 0 | 100 | 670.48 | 22940.8 | 13.69
63.9 | 0 | 100 | 624.53 | 21368.7 | 12.72
65.1 | 0 | 100 | 582.49 | 19930.4 | 11.89
66.3 | 0 | 100 | 548.99 | 18784 | 11.2
67.5 | 0 | 100 | 532.69 | 18226.4 | 10.87
68.7 | 0 | 100 | 512.93 | 17550.2 | 10.46
69.9 | 0 | 100 | 491.43 | 16814.6 | 10.02
71.1 | 0 | 100 | 471.34 | 16127.2 | 9.61
72.3 | 0 | 100 | 454.71 | 15558.3 | 9.26
73.5 | 0 | 100 | 445.31 | 15236.5 | 9.07
74.7 | 0 | 100 | 426.83 | 14604.4 | 8.71
75.9 | 0 | 100 | 416.8 | 14261 | 8.5
77.1 | 0 | 100 | 405.53 | 13875.5 | 8.27
78.3 | 0 | 100 | 393.82 | 13474.9 | 8.02
79.5 | 0 | 100 | 380.79 | 13028.9 | 7.77
80.7 | 0 | 100 | 372.67 | 12751.3 | 7.61
81.9 | 0 | 100 | 360.34 | 12329.2 | 7.36
83.1 | 0 | 100 | 350.11 | 11979.3 | 7.12
84.3 | 0 | 100 | 337.67 | 11553.6 | 6.88
85.5 | 0 | 100 | 326.95 | 11186.7 | 6.67
86.7 | 0 | 100 | 317.76 | 10872.5 | 6.49
87.9 | 0 | 100 | 310.09 | 10610 | 6.32
89.1 | 0 | 100 | 303.02 | 10368 | 6.17
90.3 | 0 | 100 | 292.8 | 10018.4 | 5.97
91.5 | 0 | 100 | 288.14 | 9858.8 | 5.89
92.7 | 0 | 100 | 279.53 | 9564.4 | 5.7
93.9 | 0 | 100 | 272.38 | 9319.8 | 5.55
95.2 | 0 | 100 | 264.57 | 9052.4 | 5.4
96.4 | 0 | 100 | 258.69 | 8851.4 | 5.28
97.6 | 0 | 100 | 250.42 | 8568.5 | 5.11
98.8 | 0 | 100 | 248.9 | 8516.3 | 5.07
100 | 0 | 100 | 241.16 | 8251.4 | 4.92
101.2 | 0 | 100 | 241.53 | 8264.1 | 4.93
102.4 | 0 | 100 | 232.7 | 7961.9 | 4.74
103.6 | 0 | 100 | 222.63 | 7617.6 | 4.55
104.8 | 0 | 100 | 222.61 | 7616.7 | 4.54
106 | 0 | 100 | 214.57 | 7341.8 | 4.38
107.2 | 0 | 100 | 212.39 | 7267.1 | 4.33
108.4 | 0 | 100 | 205.9 | 7045 | 4.2
109.6 | 0 | 100 | 205.64 | 7036.2 | 4.19
110.8 | 0 | 100 | 201.4 | 6890.9 | 4.11
112 | 0 | 100 | 194 | 6637.8 | 3.96
113.2 | 0 | 100 | 190.77 | 6527.3 | 3.89
114.4 | 0 | 100 | 187.99 | 6432.2 | 3.83
115.6 | 0 | 100 | 183.54 | 6279.8 | 3.74
116.8 | 0 | 100 | 180.95 | 6191.3 | 3.7
118 | 0 | 100 | 177.82 | 6084.4 | 3.62
119.3 | 0 | 100 | 170.82 | 5844.8 | 3.48
120.4 | 0 | 100 | 169.29 | 5792.2 | 3.45
121.7 | 0 | 100 | 163.48 | 5593.6 | 3.33
122.8 | 0 | 100 | 162.66 | 5565.4 | 3.32
124.1 | 0 | 100 | 158.79 | 5433.1 | 3.24
125.3 | 0 | 100 | 157.59 | 5392.1 | 3.22
126.5 | 0 | 100 | 156.67 | 5360.5 | 3.19
127.7 | 0 | 100 | 152.51 | 5218.1 | 3.11
128.9 | 0 | 100 | 150.41 | 5146.3 | 3.07
130.1 | 0 | 100 | 150.94 | 5164.6 | 3.08
131.3 | 0 | 100 | 147.29 | 5039.7 | 3
132.5 | 0 | 100 | 145.19 | 4967.8 | 2.96
133.7 | 0 | 100 | 142.24 | 4866.9 | 2.9
134.9 | 0 | 100 | 142.08 | 4861.4 | 2.9
136.1 | 0 | 100 | 140.65 | 4812.4 | 2.87
137.3 | 0 | 100 | 135.36 | 4631.4 | 2.76
138.5 | 0 | 100 | 134.92 | 4616.3 | 2.75
139.7 | 0 | 100 | 131.98 | 4515.9 | 2.69
140.9 | 0 | 100 | 131.19 | 4488.6 | 2.68
142.1 | 0 | 100 | 128.2 | 4386.6 | 2.61
143.3 | 0 | 100 | 126.95 | 4343.7 | 2.59
144.5 | 0 | 100 | 125.08 | 4279.6 | 2.55
145.7 | 0 | 100 | 124.4 | 4256.5 | 2.53
146.9 | 0 | 100 | 120.69 | 4129.3 | 2.46
148.1 | 0 | 100 | 120.16 | 4111.3 | 2.45
149.4 | 0 | 100 | 121.42 | 4154.5 | 2.48
150.6 | 0 | 100 | 120.12 | 4109.9 | 2.45
151.8 | 0 | 100 | 112.92 | 3863.6 | 2.3
153 | 0 | 100 | 112.6 | 3852.7 | 2.3
155 | 0 | 100 | 112.22 | 3839.8 | 2.29
Measurement time: | 30/08/2022 16:31:44
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
154.6 | 0 | 100 | 318.6 | 10901.2 | 1.96
153.4 | 0 | 100 | 325.9 | 11151 | 2
152.3 | 0 | 100 | 324.35 | 11097.8 | 1.99
151.2 | 0 | 100 | 330.19 | 11297.7 | 2.03
149.9 | 0 | 100 | 334.69 | 11451.6 | 2.05
148.7 | 0 | 100 | 336.9 | 11527.3 | 2.07
147.5 | 0 | 100 | 343.7 | 11759.8 | 2.11
146.3 | 0 | 100 | 346.29 | 11848.6 | 2.12
145.1 | 0 | 100 | 349.31 | 11951.7 | 2.15
143.9 | 0 | 100 | 356.26 | 12189.8 | 2.18
142.7 | 0 | 100 | 361.99 | 12385.8 | 2.22
141.5 | 0 | 100 | 364.6 | 12475 | 2.24
140.3 | 0 | 100 | 368.28 | 12601.1 | 2.26
139.1 | 0 | 100 | 377.06 | 12901.3 | 2.31
137.9 | 0 | 100 | 380.36 | 13014.3 | 2.33
136.7 | 0 | 100 | 383.09 | 13107.5 | 2.35
135.5 | 0 | 100 | 392.21 | 13419.7 | 2.4
134.2 | 0 | 100 | 393.02 | 13447.4 | 2.42
133 | 0 | 100 | 399.97 | 13685.3 | 2.45
131.8 | 0 | 100 | 409.75 | 14019.8 | 2.51
130.6 | 0 | 100 | 413.91 | 14162.1 | 2.54
129.4 | 0 | 100 | 422.79 | 14466 | 2.59
128.2 | 0 | 100 | 429.82 | 14706.6 | 2.64
127 | 0 | 100 | 434.72 | 14874.1 | 2.67
125.8 | 0 | 100 | 443.69 | 15181 | 2.71
124.6 | 0 | 100 | 450 | 15397.1 | 2.76
123.4 | 0 | 100 | 463.15 | 15847 | 2.84
122.2 | 0 | 100 | 467.07 | 15981 | 2.86
121 | 0 | 100 | 474.89 | 16248.6 | 2.92
119.8 | 0 | 100 | 485.9 | 16625.4 | 2.97
118.6 | 0 | 100 | 491.14 | 16804.7 | 3.01
117.4 | 0 | 100 | 499.69 | 17097.4 | 3.06
116.2 | 0 | 100 | 506.04 | 17314.5 | 3.11
115 | 0 | 100 | 520.84 | 17820.9 | 3.19
113.8 | 0 | 100 | 529.51 | 18117.5 | 3.25
112.6 | 0 | 100 | 539.16 | 18447.6 | 3.3
111.4 | 0 | 100 | 552.1 | 18890.5 | 3.39
110.2 | 0 | 100 | 560.15 | 19165.8 | 3.43
109 | 0 | 100 | 569.49 | 19485.5 | 3.5
107.8 | 0 | 100 | 584.9 | 20012.6 | 3.58
106.6 | 0 | 100 | 592.37 | 20268.3 | 3.64
105.3 | 0 | 100 | 613.34 | 20985.8 | 3.76
104.2 | 0 | 100 | 622.03 | 21283 | 3.81
103 | 0 | 100 | 637.72 | 21820.1 | 3.91
101.8 | 0 | 100 | 654.33 | 22388.4 | 4.01
100.6 | 0 | 100 | 668.65 | 22878.3 | 4.1
99.3 | 0 | 100 | 682.3 | 23345.2 | 4.18
98.2 | 0 | 100 | 699.87 | 23946.6 | 4.29
96.9 | 0 | 100 | 717.91 | 24563.6 | 4.41
95.8 | 0 | 100 | 735.97 | 25181.6 | 4.51
94.6 | 0 | 100 | 751.6 | 25716.4 | 4.61
93.3 | 0 | 100 | 774.62 | 26504 | 4.74
92.2 | 0 | 100 | 793.06 | 27135.1 | 4.87
90.9 | 0 | 100 | 820.53 | 28075 | 5.02
89.7 | 0 | 100 | 843.28 | 28853.4 | 5.18
88.5 | 0 | 100 | 863.99 | 29561.9 | 5.3
87.3 | 0 | 100 | 892.75 | 30546.2 | 5.47
86.1 | 0 | 100 | 916.22 | 31349.1 | 5.61
84.9 | 0 | 100 | 942.24 | 32239.5 | 5.79
83.7 | 0 | 100 | 975.59 | 33380.3 | 5.97
82.5 | 0 | 100 | 1000.41 | 34229.5 | 6.15
81.3 | 0 | 100 | 1036.77 | 35473.7 | 6.34
80.1 | 0 | 100 | 1068.27 | 36551.7 | 6.55
78.9 | 0 | 100 | 1102.46 | 37721.5 | 6.78
77.7 | 0 | 100 | 1141.41 | 39054 | 7
76.5 | 0 | 100 | 1185.59 | 40565.6 | 7.24
75.3 | 0 | 100 | 1215.18 | 41578 | 7.46
74.1 | 0 | 100 | 1260.81 | 43139.5 | 7.73
72.9 | 0 | 100 | 1311.29 | 44866.7 | 8.02
71.7 | 0 | 100 | 1354.17 | 46333.8 | 8.31
70.5 | 0 | 100 | 1404.91 | 48069.8 | 8.61
69.2 | 0 | 100 | 1462.77 | 50049.7 | 8.97
68.1 | 0 | 100 | 1520.24 | 52015.9 | 9.32
66.8 | 0 | 100 | 1598.71 | 54700.8 | 9.79
65.7 | 0 | 100 | 1682.03 | 57551.8 | 10.32
64.5 | 0 | 100 | 1788.65 | 61199.7 | 10.92
63.2 | 0 | 100 | 1924.64 | 65852.6 | 11.82
62.1 | 0 | 100 | 2136.12 | 73088.6 | 13.07
60.8 | 0 | 100 | 2465.61 | 84362.3 | 15.11
59.6 | 0 | 100 | 3059.83 | 104694.2 | 18.69
58.4 | 0 | 100 | 4040.4 | 138244.9 | 24.83
57.2 | 0 | 100 | 5572.24 | 190657.8 | 34.12
56 | 0 | 100 | 7170.94 | 245358.2 | 43.95
54.8 | 0 | 100 | 8975.05 | 307087.2 | 54.88
53.6 | 0 | 100 | 10516.23 | 359819.5 | 64.65
52.4 | 0 | 100 | 11752.95 | 402134.8 | 71.86
51.1 | 0 | 100 | 12707.74 | 434803.5 | 77.87
50 | 0 | 100 | 13295.11 | 454900.6 | 81.5
48.8 | 0 | 100 | 13732.22 | 469856.6 | 84.02
47.6 | 0 | 100 | 13938.16 | 476903.1 | 85.6
46.4 | 0 | 100 | 14158.81 | 484452.9 | 86.67
45.2 | 0 | 100 | 14294.72 | 489102.9 | 87.55
44 | 0 | 100 | 14429.67 | 493720.2 | 88.42
42.7 | 0 | 100 | 14565.42 | 498365.1 | 89.22
41.6 | 0 | 100 | 14638.04 | 500849.9 | 89.91
40.4 | 0 | 100 | 14779.18 | 505679.1 | 90.42
39.1 | 0 | 100 | 14863.5 | 508564.1 | 90.96
38 | 0 | 100 | 14908.5 | 510103.9 | 91.5
36.7 | 0 | 100 | 15031.39 | 514308.6 | 92.17
35.6 | 0 | 100 | 15108.81 | 516957.7 | 92.49
34.4 | 0 | 100 | 15178.31 | 519335.4 | 92.94
33.1 | 0 | 100 | 15240.29 | 521456.3 | 93.45
32 | 0 | 100 | 15382.43 | 526319.8 | 94.01
30.7 | 0 | 100 | 15400.77 | 526947 | 94.38
29.6 | 0 | 100 | 15477.42 | 529569.7 | 94.76
28.3 | 0 | 100 | 15546.43 | 531931 | 95.18
27.2 | 0 | 100 | 15603.82 | 533894.8 | 95.49
26 | 0 | 100 | 15624.31 | 534595.9 | 95.93
24.7 | 0 | 100 | 15710.7 | 537551.7 | 96.29
23.6 | 0 | 100 | 15777.01 | 539820.6 | 96.52
22.3 | 0 | 100 | 15833.97 | 541769.2 | 96.81
21.1 | 0 | 100 | 15879.01 | 543310.5 | 97.37
20 | 0 | 100 | 15903.42 | 544145.6 | 97.55
18.7 | 0 | 100 | 15947.48 | 545653.1 | 97.86
17.5 | 0 | 100 | 16044.85 | 548984.7 | 98.08
16.3 | 0 | 100 | 15999.03 | 547416.9 | 98.26
15.1 | 0 | 100 | 16100.96 | 550904.5 | 98.59
14 | 0 | 100 | 16142.05 | 552310.7 | 98.73
12.7 | 0 | 100 | 16142.54 | 552327.4 | 98.89
11.5 | 0 | 100 | 16157.22 | 552829.6 | 99.18
10.3 | 0 | 100 | 16190.09 | 553954.3 | 99.16
9.1 | 0 | 100 | 16201.82 | 554355.6 | 99.32
7.9 | 0 | 100 | 16248.75 | 555961.3 | 99.5
6.7 | 0 | 100 | 16253.24 | 556115 | 99.59
5.5 | 0 | 100 | 16255.89 | 556205.8 | 99.5
4.3 | 0 | 100 | 16275.05 | 556861.2 | 99.83
3.1 | 0 | 100 | 16301.21 | 557756.3 | 99.9
1.9 | 0 | 100 | 16347.68 | 559346.4 | 100
0.7 | 0 | 100 | 16317.23 | 558304.4 | 99.82
-0.5 | 0 | 100 | 16247.81 | 555929.2 | 99.9
-1.7 | 0 | 100 | 16311.4 | 558104.9 | 99.94
-2.9 | 0 | 100 | 16313.9 | 558190.3 | 99.93
-4.1 | 0 | 100 | 16337.91 | 559012.1 | 99.92
-5.4 | 0 | 100 | 16286.6 | 557256.6 | 99.78
-6.5 | 0 | 100 | 16232.15 | 555393.5 | 99.62
-7.8 | 0 | 100 | 16281.12 | 557069.1 | 99.78
-9 | 0 | 100 | 16247.96 | 555934.4 | 99.59
-10.2 | 0 | 100 | 16253.41 | 556120.8 | 99.49
-11.4 | 0 | 100 | 16210.69 | 554658.9 | 99.37
-12.6 | 0 | 100 | 16177.14 | 553511.3 | 98.97
-13.8 | 0 | 100 | 16155.12 | 552757.6 | 98.87
-15 | 0 | 100 | 16144.95 | 552409.8 | 98.75
-16.2 | 0 | 100 | 16037.16 | 548721.5 | 98.48
-17.4 | 0 | 100 | 16070.39 | 549858.8 | 98.24
-18.6 | 0 | 100 | 15965 | 546252.7 | 97.94
-19.8 | 0 | 100 | 15964.64 | 546240.3 | 97.79
-21 | 0 | 100 | 15946.61 | 545623.3 | 97.56
-22.2 | 0 | 100 | 15878.08 | 543278.7 | 97.2
-23.4 | 0 | 100 | 15839.56 | 541960.7 | 96.89
-24.6 | 0 | 100 | 15747.08 | 538796.5 | 96.66
-25.8 | 0 | 100 | 15715.06 | 537700.7 | 96.27
-27.1 | 0 | 100 | 15679.03 | 536467.9 | 95.97
-28.2 | 0 | 100 | 15638.82 | 535092.2 | 95.56
-29.5 | 0 | 100 | 15521.55 | 531079.6 | 95.13
-30.6 | 0 | 100 | 15442.4 | 528371.5 | 94.76
-31.9 | 0 | 100 | 15403.92 | 527055 | 94.28
-33.1 | 0 | 100 | 15315.79 | 524039.6 | 93.85
-34.2 | 0 | 100 | 15281.26 | 522858.1 | 93.41
-35.5 | 0 | 100 | 15157.85 | 518635.7 | 93.01
-36.6 | 0 | 100 | 15082.27 | 516049.6 | 92.52
-37.9 | 0 | 100 | 15005.39 | 513418.9 | 91.9
-39.1 | 0 | 100 | 14915.58 | 510346.3 | 91.34
-40.2 | 0 | 100 | 14826.03 | 507282.1 | 90.8
-41.5 | 0 | 100 | 14717.14 | 503556.5 | 90.15
-42.6 | 0 | 100 | 14638.11 | 500852.1 | 89.5
-43.9 | 0 | 100 | 14535.6 | 497344.8 | 88.9
-45.1 | 0 | 100 | 14382.32 | 492100.1 | 88.07
-46.3 | 0 | 100 | 14207.96 | 486134.6 | 87.23
-47.5 | 0 | 100 | 14082.64 | 481846.4 | 86.09
-48.7 | 0 | 100 | 13792.53 | 471920.1 | 84.61
-49.9 | 0 | 100 | 13476.46 | 461105.6 | 82.47
-51.1 | 0 | 100 | 12809.57 | 438287.5 | 78.45
-52.3 | 0 | 100 | 11859.62 | 405784.6 | 72.5
-53.5 | 0 | 100 | 10296.58 | 352304.2 | 63.12
-54.7 | 0 | 100 | 8651.14 | 296004.4 | 52.81
-55.9 | 0 | 100 | 6837.29 | 233942.4 | 41.9
-57.1 | 0 | 100 | 5273.87 | 180448.9 | 32.33
-58.3 | 0 | 100 | 4092.54 | 140029 | 25.03
-59.6 | 0 | 100 | 3113.59 | 106533.3 | 19.09
-60.7 | 0 | 100 | 2548.94 | 87213.6 | 15.56
-62 | 0 | 100 | 2179.06 | 74557.8 | 13.33
-63.1 | 0 | 100 | 1955.63 | 66913.2 | 11.98
-64.3 | 0 | 100 | 1809.15 | 61901.3 | 11.09
-65.6 | 0 | 100 | 1698.94 | 58130.4 | 10.39
-66.7 | 0 | 100 | 1612.44 | 55170.6 | 9.87
-68 | 0 | 100 | 1529.93 | 52347.7 | 9.37
-69.2 | 0 | 100 | 1470.42 | 50311.5 | 9
-70.4 | 0 | 100 | 1413.59 | 48367 | 8.65
-71.6 | 0 | 100 | 1359.85 | 46528.1 | 8.34
-72.8 | 0 | 100 | 1310.3 | 44832.8 | 8.02
-74 | 0 | 100 | 1260.38 | 43124.8 | 7.72
-75.2 | 0 | 100 | 1222.61 | 41832.2 | 7.48
-76.4 | 0 | 100 | 1178.37 | 40318.7 | 7.2
-77.6 | 0 | 100 | 1140.85 | 39035 | 6.99
-78.8 | 0 | 100 | 1098.92 | 37600.2 | 6.73
-80 | 0 | 100 | 1069.16 | 36582 | 6.54
-81.2 | 0 | 100 | 1025.01 | 35071.3 | 6.28
-82.4 | 0 | 100 | 997.54 | 34131.4 | 6.12
-83.6 | 0 | 100 | 970.54 | 33207.5 | 5.94
-84.8 | 0 | 100 | 940.85 | 32191.9 | 5.75
-86.1 | 0 | 100 | 904.14 | 30935.7 | 5.54
-87.2 | 0 | 100 | 877.59 | 30027.4 | 5.37
-88.4 | 0 | 100 | 854.43 | 29235 | 5.23
-89.7 | 0 | 100 | 830.2 | 28406 | 5.07
-90.8 | 0 | 100 | 805.57 | 27562.9 | 4.94
-92.1 | 0 | 100 | 786.22 | 26900.9 | 4.82
-93.2 | 0 | 100 | 758.41 | 25949.3 | 4.64
-94.4 | 0 | 100 | 739.58 | 25305.2 | 4.53
-95.7 | 0 | 100 | 720.66 | 24657.7 | 4.41
-96.8 | 0 | 100 | 703.41 | 24067.5 | 4.3
-98.1 | 0 | 100 | 680.56 | 23285.9 | 4.17
-99.3 | 0 | 100 | 665.03 | 22754.5 | 4.07
-100.5 | 0 | 100 | 649.76 | 22231.8 | 3.98
-101.7 | 0 | 100 | 633.62 | 21679.8 | 3.88
-102.9 | 0 | 100 | 616.93 | 21108.5 | 3.79
-104.1 | 0 | 100 | 603.1 | 20635.4 | 3.69
-105.3 | 0 | 100 | 595.13 | 20362.9 | 3.64
-106.5 | 0 | 100 | 578.73 | 19801.6 | 3.54
-107.7 | 0 | 100 | 564.96 | 19330.6 | 3.46
-108.9 | 0 | 100 | 553.64 | 18943 | 3.38
-110.1 | 0 | 100 | 539.36 | 18454.7 | 3.3
-111.3 | 0 | 100 | 528.4 | 18079.4 | 3.23
-112.5 | 0 | 100 | 517.08 | 17692.3 | 3.17
-113.7 | 0 | 100 | 507.09 | 17350.5 | 3.09
-114.9 | 0 | 100 | 496.71 | 16995.1 | 3.05
-116.1 | 0 | 100 | 486.39 | 16642.2 | 2.98
-117.3 | 0 | 100 | 479.62 | 16410.5 | 2.93
-118.5 | 0 | 100 | 466.56 | 15963.7 | 2.85
-119.7 | 0 | 100 | 457.37 | 15649.1 | 2.8
-120.9 | 0 | 100 | 452.04 | 15467 | 2.76
-122.1 | 0 | 100 | 442.1 | 15126.8 | 2.71
-123.3 | 0 | 100 | 437.67 | 14975.2 | 2.67
-124.5 | 0 | 100 | 425.01 | 14541.8 | 2.61
-125.7 | 0 | 100 | 419.57 | 14355.9 | 2.56
-126.9 | 0 | 100 | 407.93 | 13957.5 | 2.5
-128.1 | 0 | 100 | 404.67 | 13845.9 | 2.48
-129.3 | 0 | 100 | 399.59 | 13672.2 | 2.45
-130.5 | 0 | 100 | 394.17 | 13486.7 | 2.41
-131.8 | 0 | 100 | 381.27 | 13045.6 | 2.33
-132.9 | 0 | 100 | 381.01 | 13036.4 | 2.32
-134.1 | 0 | 100 | 371.7 | 12717.8 | 2.27
-135.3 | 0 | 100 | 367.79 | 12584.3 | 2.25
-136.5 | 0 | 100 | 361.42 | 12366.2 | 2.21
-137.8 | 0 | 100 | 355.74 | 12171.9 | 2.18
-138.9 | 0 | 100 | 348.14 | 11912 | 2.13
-140.1 | 0 | 100 | 346.14 | 11843.2 | 2.12
-141.3 | 0 | 100 | 343.16 | 11741.6 | 2.1
-142.5 | 0 | 100 | 336.93 | 11528.1 | 2.06
-143.8 | 0 | 100 | 331.94 | 11357.5 | 2.03
-144.9 | 0 | 100 | 326.51 | 11171.7 | 2
-146.2 | 0 | 100 | 319.41 | 10928.8 | 1.95
-147.3 | 0 | 100 | 317.62 | 10867.6 | 1.94
-148.5 | 0 | 100 | 311.15 | 10646.2 | 1.9
-149.8 | 0 | 100 | 308.99 | 10572.2 | 1.89
-150.9 | 0 | 100 | 308.36 | 10550.8 | 1.88
-152.2 | 0 | 100 | 299.73 | 10255.4 | 1.84
-153.3 | 0 | 100 | 302.09 | 10336.2 | 1.85
-155 | 0 | 100 | 295.69 | 10117.3 | 1.81