## 4_Profils profondeur mesure
Measurement time: | 12/08/2022 18:26:00
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
100 | 0 | 50 | 2.76 | 6826.9 | 2.99
99.1 | 0 | 50 | 2.38 | 5876.7 | 2.73
97.6 | 0 | 50 | 2.45 | 6064 | 2.81
96.5 | 0 | 50 | 2.53 | 6263.9 | 2.91
95.3 | 0 | 50 | 2.6 | 6416.5 | 2.97
94.1 | 0 | 50 | 2.7 | 6681.5 | 3.1
92.9 | 0 | 50 | 2.77 | 6858.3 | 3.18
91.7 | 0 | 50 | 2.89 | 7145.8 | 3.31
90.5 | 0 | 50 | 2.89 | 7140.2 | 3.31
89.3 | 0 | 50 | 3.04 | 7511.2 | 3.48
88.1 | 0 | 50 | 3.14 | 7774.3 | 3.61
86.9 | 0 | 50 | 3.21 | 7941.8 | 3.69
85.7 | 0 | 50 | 3.37 | 8321.5 | 3.86
84.5 | 0 | 50 | 3.47 | 8569.7 | 3.97
83.2 | 0 | 50 | 3.62 | 8950.4 | 4.15
82 | 0 | 50 | 3.7 | 9147.8 | 4.24
80.8 | 0 | 50 | 3.8 | 9398.2 | 4.36
79.6 | 0 | 50 | 3.95 | 9758.6 | 4.53
78.4 | 0 | 50 | 4.05 | 10009.6 | 4.64
77.3 | 0 | 50 | 4.23 | 10448.5 | 4.85
76 | 0 | 50 | 4.4 | 10874.4 | 5.05
74.8 | 0 | 50 | 4.53 | 11209.5 | 5.2
73.6 | 0 | 50 | 4.72 | 11679.1 | 5.42
72.4 | 0 | 50 | 4.93 | 12187.9 | 5.65
71.2 | 0 | 50 | 5.09 | 12594.3 | 5.84
70 | 0 | 50 | 5.36 | 13251 | 6.14
68.8 | 0 | 50 | 5.53 | 13663.4 | 6.35
67.6 | 0 | 50 | 5.75 | 14217.9 | 6.59
66.4 | 0 | 50 | 6.02 | 14890.1 | 6.91
65.2 | 0 | 50 | 6.36 | 15712.6 | 7.3
64 | 0 | 50 | 6.64 | 16408.9 | 7.61
62.8 | 0 | 50 | 7.04 | 17406.7 | 8.09
61.6 | 0 | 50 | 7.55 | 18668.2 | 8.66
60.3 | 0 | 50 | 8.16 | 20173.2 | 9.36
59.2 | 0 | 50 | 9.07 | 22416.5 | 10.4
57.9 | 0 | 50 | 10.49 | 25924.1 | 12.02
56.8 | 0 | 50 | 13.05 | 32270.5 | 14.99
55.5 | 0 | 50 | 17.81 | 44027.6 | 20.42
54.3 | 0 | 50 | 26.32 | 65060.7 | 30.13
53.2 | 0 | 50 | 36.16 | 89388.7 | 41.49
51.9 | 0 | 50 | 48.06 | 118827.8 | 55.12
50.7 | 0 | 50 | 58.69 | 145114.5 | 67.35
49.5 | 0 | 50 | 67.28 | 166340.6 | 77.22
48.3 | 0 | 50 | 74.35 | 183813.9 | 85.36
47.1 | 0 | 50 | 78.66 | 194476.4 | 90.12
45.9 | 0 | 50 | 81.2 | 200754.1 | 93.14
44.7 | 0 | 50 | 82.5 | 203966 | 94.81
43.5 | 0 | 50 | 83.26 | 205840.6 | 95.42
42.3 | 0 | 50 | 83.74 | 207024.1 | 95.85
41.1 | 0 | 50 | 84.12 | 207972.6 | 96.54
39.8 | 0 | 50 | 84.38 | 208628.1 | 96.8
38.7 | 0 | 50 | 84.51 | 208946.2 | 96.99
37.4 | 0 | 50 | 84.59 | 209137.5 | 97.09
36.2 | 0 | 50 | 85.03 | 210230.2 | 97.57
35.1 | 0 | 50 | 85.15 | 210513.3 | 97.55
33.8 | 0 | 50 | 85.22 | 210690.7 | 97.74
32.7 | 0 | 50 | 85.44 | 211244.2 | 97.98
31.4 | 0 | 50 | 85.63 | 211695.4 | 98.19
30.2 | 0 | 50 | 85.76 | 212027.9 | 98.37
29 | 0 | 50 | 85.76 | 212025.6 | 98.53
27.8 | 0 | 50 | 85.88 | 212317.6 | 98.57
26.6 | 0 | 50 | 86.1 | 212863.9 | 98.75
25.4 | 0 | 50 | 86.15 | 212991.1 | 98.81
24.2 | 0 | 50 | 86.29 | 213332.4 | 98.87
23 | 0 | 50 | 86.24 | 213211.2 | 98.93
21.8 | 0 | 50 | 86.36 | 213507.5 | 99.15
20.6 | 0 | 50 | 86.36 | 213519.2 | 99.05
19.4 | 0 | 50 | 86.27 | 213280.6 | 99.1
18.2 | 0 | 50 | 86.33 | 213438.5 | 99.07
17 | 0 | 50 | 86.34 | 213460.4 | 99.23
15.8 | 0 | 50 | 86.59 | 214069.1 | 99.17
14.6 | 0 | 50 | 86.24 | 213211.5 | 98.89
13.4 | 0 | 50 | 86.38 | 213550.2 | 99.24
12.2 | 0 | 50 | 86.27 | 213282.8 | 98.91
11 | 0 | 50 | 86.22 | 213166.8 | 98.94
9.8 | 0 | 50 | 86.07 | 212803.8 | 99
8.6 | 0 | 50 | 86.01 | 212656.1 | 98.73
7.4 | 0 | 50 | 86.24 | 213216.3 | 98.87
6.2 | 0 | 50 | 86.29 | 213345.4 | 98.97
5 | 0 | 50 | 86.35 | 213490.5 | 99.04
3.8 | 0 | 50 | 86.27 | 213281.9 | 99.11
2.6 | 0 | 50 | 86.4 | 213600.1 | 99.01
1.4 | 0 | 50 | 86.27 | 213298.7 | 99.02
0.2 | 0 | 50 | 86.44 | 213719.6 | 99.11
-1 | 0 | 50 | 86.34 | 213463.1 | 99.02
-2.3 | 0 | 50 | 86.46 | 213751 | 99.09
-3.4 | 0 | 50 | 86.32 | 213410.4 | 99.18
-4.6 | 0 | 50 | 86.45 | 213739.5 | 99.2
-5.9 | 0 | 50 | 86.46 | 213766.3 | 99.2
-7 | 0 | 50 | 86.61 | 214121.9 | 99.31
-8.3 | 0 | 50 | 86.59 | 214092.5 | 99.42
-9.5 | 0 | 50 | 86.68 | 214295.8 | 99.36
-10.7 | 0 | 50 | 86.83 | 214670.8 | 99.78
-11.9 | 0 | 50 | 86.92 | 214890.3 | 99.75
-13.1 | 0 | 50 | 86.93 | 214931.1 | 99.65
-14.3 | 0 | 50 | 86.97 | 215014 | 99.82
-15.5 | 0 | 50 | 87.04 | 215183.6 | 99.94
-16.7 | 0 | 50 | 87.17 | 215503.1 | 99.95
-17.9 | 0 | 50 | 86.91 | 214861.3 | 99.64
-19.1 | 0 | 50 | 87.15 | 215465.1 | 99.95
-20.3 | 0 | 50 | 87.06 | 215246.5 | 99.92
-21.5 | 0 | 50 | 87.1 | 215334.3 | 99.9
-22.7 | 0 | 50 | 87.1 | 215353.8 | 100
-23.9 | 0 | 50 | 87.03 | 215159.4 | 99.89
-25.1 | 0 | 50 | 87 | 215093.2 | 99.57
-26.3 | 0 | 50 | 87.04 | 215182.8 | 99.93
-27.5 | 0 | 50 | 86.82 | 214650.7 | 99.76
-28.7 | 0 | 50 | 86.82 | 214655.6 | 99.71
-29.9 | 0 | 50 | 86.86 | 214759.5 | 99.64
-31.1 | 0 | 50 | 86.73 | 214421 | 99.42
-32.3 | 0 | 50 | 86.57 | 214025.1 | 99.37
-33.5 | 0 | 50 | 86.34 | 213469.6 | 99.11
-34.7 | 0 | 50 | 86.38 | 213570.3 | 99.05
-36 | 0 | 50 | 86.25 | 213249.9 | 98.84
-37.1 | 0 | 50 | 86.07 | 212802.6 | 98.78
-38.4 | 0 | 50 | 86.08 | 212825.2 | 98.87
-39.5 | 0 | 50 | 85.73 | 211942.4 | 98.34
-40.8 | 0 | 50 | 85.45 | 211266.9 | 98.1
-41.9 | 0 | 50 | 85.32 | 210928.7 | 97.98
-43.1 | 0 | 50 | 85.05 | 210264.9 | 97.45
-44.4 | 0 | 50 | 84.57 | 209090 | 97.19
-45.5 | 0 | 50 | 83.93 | 207497.3 | 96.33
-46.7 | 0 | 50 | 82.73 | 204536.8 | 94.97
-48 | 0 | 50 | 80.47 | 198953.2 | 92.18
-49.1 | 0 | 50 | 76.11 | 188167.6 | 87.34
-50.4 | 0 | 50 | 68.2 | 168624.9 | 78.39
-51.6 | 0 | 50 | 57.93 | 143211 | 66.58
-52.8 | 0 | 50 | 46.25 | 114351.8 | 53.06
-54 | 0 | 50 | 35.06 | 86669.2 | 40.28
-55.2 | 0 | 50 | 25.85 | 63914.4 | 29.68
-56.4 | 0 | 50 | 18.51 | 45765.2 | 21.23
-57.6 | 0 | 50 | 13.69 | 33850.6 | 15.71
-58.8 | 0 | 50 | 10.99 | 27163.8 | 12.61
-60 | 0 | 50 | 9.36 | 23138.4 | 10.73
-61.2 | 0 | 50 | 8.46 | 20912.9 | 9.69
-62.5 | 0 | 50 | 7.75 | 19171.4 | 8.91
-63.6 | 0 | 50 | 7.22 | 17852.3 | 8.28
-64.9 | 0 | 50 | 6.86 | 16959.1 | 7.87
-66.1 | 0 | 50 | 6.56 | 16211.6 | 7.52
-67.2 | 0 | 50 | 6.18 | 15281.5 | 7.09
-68.5 | 0 | 50 | 5.94 | 14693.9 | 6.82
-69.6 | 0 | 50 | 5.73 | 14176.9 | 6.58
-70.9 | 0 | 50 | 5.48 | 13543.6 | 6.29
-72.1 | 0 | 50 | 5.31 | 13131.8 | 6.1
-73.3 | 0 | 50 | 5.06 | 12521 | 5.81
-74.5 | 0 | 50 | 4.88 | 12066.8 | 5.61
-75.6 | 0 | 50 | 4.73 | 11698.2 | 5.42
-76.9 | 0 | 50 | 4.5 | 11130.4 | 5.16
-78.1 | 0 | 50 | 4.34 | 10721.8 | 4.98
-79.3 | 0 | 50 | 4.23 | 10455.5 | 4.86
-80.5 | 0 | 50 | 4.05 | 10006 | 4.65
-81.7 | 0 | 50 | 3.87 | 9568.9 | 4.44
-82.9 | 0 | 50 | 3.79 | 9369 | 4.35
-84.1 | 0 | 50 | 3.68 | 9109.3 | 4.23
-85.3 | 0 | 50 | 3.53 | 8719.4 | 4.05
-86.6 | 0 | 50 | 3.41 | 8433.2 | 3.92
-87.7 | 0 | 50 | 3.27 | 8084.3 | 3.75
-88.9 | 0 | 50 | 3.23 | 7982.3 | 3.71
-90.1 | 0 | 50 | 3.18 | 7859.3 | 3.64
-91.3 | 0 | 50 | 3.02 | 7462.9 | 3.47
-92.5 | 0 | 50 | 2.87 | 7085.1 | 3.29
-93.7 | 0 | 50 | 2.8 | 6927.5 | 3.21
-95 | 0 | 50 | 2.7 | 6674 | 3.1
-96.1 | 0 | 50 | 2.65 | 6539.7 | 3.03
-97.3 | 0 | 50 | 2.52 | 6230.1 | 2.89
-98.6 | 0 | 50 | 2.49 | 6156.2 | 2.86
-100 | 0 | 50 | 2.41 | 5946.3 | 2.76
Measurement time: | 12/08/2022 18:26:54
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-100 | 0 | 100 | 2.52 | 6239.8 | 3.92
-99.1 | 0 | 100 | 2.81 | 6946.8 | 4.2
-97.7 | 0 | 100 | 2.95 | 7281.2 | 4.42
-96.6 | 0 | 100 | 3.07 | 7581 | 4.6
-95.4 | 0 | 100 | 3.12 | 7705.2 | 4.65
-94.2 | 0 | 100 | 3.25 | 8028.7 | 4.85
-93 | 0 | 100 | 3.36 | 8301.7 | 5.02
-91.7 | 0 | 100 | 3.47 | 8586.7 | 5.19
-90.6 | 0 | 100 | 3.57 | 8827.8 | 5.34
-89.4 | 0 | 100 | 3.66 | 9044.8 | 5.47
-88.2 | 0 | 100 | 3.73 | 9216.1 | 5.58
-86.9 | 0 | 100 | 3.86 | 9534.3 | 5.78
-85.7 | 0 | 100 | 4.01 | 9910.5 | 5.99
-84.5 | 0 | 100 | 4.22 | 10442.6 | 6.32
-83.3 | 0 | 100 | 4.32 | 10692.7 | 6.47
-82.1 | 0 | 100 | 4.47 | 11052.4 | 6.68
-80.9 | 0 | 100 | 4.64 | 11461.2 | 6.93
-79.7 | 0 | 100 | 4.8 | 11875.8 | 7.19
-78.5 | 0 | 100 | 4.98 | 12307.4 | 7.44
-77.3 | 0 | 100 | 5.19 | 12839.6 | 7.78
-76.1 | 0 | 100 | 5.32 | 13141.4 | 7.96
-74.9 | 0 | 100 | 5.51 | 13616.8 | 8.25
-73.7 | 0 | 100 | 5.73 | 14169.9 | 8.58
-72.5 | 0 | 100 | 5.97 | 14764.6 | 8.93
-71.3 | 0 | 100 | 6.18 | 15282.4 | 9.26
-70.1 | 0 | 100 | 6.43 | 15900.9 | 9.62
-68.9 | 0 | 100 | 6.7 | 16564.6 | 10.02
-67.7 | 0 | 100 | 6.97 | 17240.1 | 10.42
-66.5 | 0 | 100 | 7.38 | 18247.9 | 11.03
-65.3 | 0 | 100 | 7.79 | 19260.6 | 11.67
-64.1 | 0 | 100 | 8.28 | 20481 | 12.39
-62.9 | 0 | 100 | 9.05 | 22382.7 | 13.54
-61.7 | 0 | 100 | 10.13 | 25056.2 | 15.19
-60.5 | 0 | 100 | 12.05 | 29783.7 | 18.02
-59.3 | 0 | 100 | 15.07 | 37263 | 22.57
-58 | 0 | 100 | 19.89 | 49177.1 | 29.74
-56.8 | 0 | 100 | 25.98 | 64220.8 | 38.85
-55.6 | 0 | 100 | 33.52 | 82867.4 | 50.18
-54.4 | 0 | 100 | 41.72 | 103140 | 62.45
-53.2 | 0 | 100 | 49.39 | 122108.1 | 73.83
-52 | 0 | 100 | 55.56 | 137362.9 | 83.23
-50.8 | 0 | 100 | 59.42 | 146914.5 | 88.73
-49.6 | 0 | 100 | 61.64 | 152406.4 | 92.04
-48.4 | 0 | 100 | 62.88 | 155452 | 94.09
-47.2 | 0 | 100 | 63.44 | 156851.8 | 94.79
-46 | 0 | 100 | 63.96 | 158135 | 95.52
-44.8 | 0 | 100 | 64.28 | 158917.1 | 96.11
-43.6 | 0 | 100 | 64.52 | 159505.2 | 96.59
-42.4 | 0 | 100 | 64.84 | 160295.2 | 96.85
-41.2 | 0 | 100 | 65.17 | 161116.2 | 97.38
-40 | 0 | 100 | 65.14 | 161057.4 | 97.46
-38.8 | 0 | 100 | 65.23 | 161275.3 | 97.49
-37.6 | 0 | 100 | 65.53 | 162016.9 | 97.96
-36.4 | 0 | 100 | 65.51 | 161951.8 | 97.92
-35.2 | 0 | 100 | 65.88 | 162872.9 | 98.42
-34 | 0 | 100 | 65.91 | 162957.8 | 98.75
-32.8 | 0 | 100 | 65.88 | 162883.2 | 98.72
-31.6 | 0 | 100 | 66.2 | 163661.4 | 98.93
-30.4 | 0 | 100 | 66.29 | 163898.8 | 99.03
-29.2 | 0 | 100 | 66.3 | 163920.7 | 99.11
-28 | 0 | 100 | 66.56 | 164556.5 | 99.43
-26.8 | 0 | 100 | 66.52 | 164466.8 | 99.53
-25.5 | 0 | 100 | 66.52 | 164450.4 | 99.52
-24.4 | 0 | 100 | 66.63 | 164729 | 99.68
-23.2 | 0 | 100 | 66.75 | 165041.4 | 99.7
-21.9 | 0 | 100 | 66.75 | 165029.2 | 99.86
-20.8 | 0 | 100 | 66.81 | 165181.4 | 99.94
-19.5 | 0 | 100 | 66.8 | 165153.5 | 99.84
-18.3 | 0 | 100 | 66.79 | 165119.9 | 99.78
-17.1 | 0 | 100 | 66.8 | 165160.5 | 99.78
-15.9 | 0 | 100 | 66.78 | 165095 | 100
-14.7 | 0 | 100 | 66.66 | 164813.1 | 99.65
-13.5 | 0 | 100 | 66.75 | 165026.1 | 99.79
-12.3 | 0 | 100 | 66.62 | 164718.9 | 99.69
-11.1 | 0 | 100 | 66.85 | 165269.6 | 99.99
-9.9 | 0 | 100 | 66.76 | 165044.5 | 99.81
-8.7 | 0 | 100 | 66.71 | 164942.3 | 99.74
-7.5 | 0 | 100 | 66.64 | 164764.8 | 99.69
-6.3 | 0 | 100 | 66.52 | 164469.6 | 99.62
-5.1 | 0 | 100 | 66.56 | 164571.5 | 99.55
-3.9 | 0 | 100 | 66.59 | 164622.8 | 99.46
-2.7 | 0 | 100 | 66.53 | 164491.8 | 99.53
-1.5 | 0 | 100 | 66.52 | 164471.7 | 99.55
-0.3 | 0 | 100 | 66.37 | 164086.1 | 99.4
0.9 | 0 | 100 | 66.4 | 164155.7 | 99.28
2.2 | 0 | 100 | 66.38 | 164111.7 | 99.26
3.4 | 0 | 100 | 66.3 | 163916.4 | 99.3
4.6 | 0 | 100 | 66.28 | 163855.7 | 99.26
5.8 | 0 | 100 | 66.49 | 164388.2 | 99.39
7 | 0 | 100 | 66.51 | 164447.7 | 99.35
8.2 | 0 | 100 | 66.34 | 164014.5 | 99.33
9.4 | 0 | 100 | 66.43 | 164244.9 | 99.34
10.6 | 0 | 100 | 66.44 | 164258.1 | 99.6
11.8 | 0 | 100 | 66.32 | 163967.3 | 99.25
13 | 0 | 100 | 66.42 | 164220.5 | 99.59
14.2 | 0 | 100 | 66.4 | 164171.1 | 99.4
15.4 | 0 | 100 | 66.44 | 164262.4 | 99.35
16.6 | 0 | 100 | 66.37 | 164096.5 | 99.35
17.8 | 0 | 100 | 66.34 | 164005.3 | 99.22
19 | 0 | 100 | 66.48 | 164350.8 | 99.25
20.2 | 0 | 100 | 66.33 | 163995.3 | 99.14
21.4 | 0 | 100 | 66.26 | 163808.9 | 99.12
22.6 | 0 | 100 | 66.04 | 163267.1 | 98.87
23.8 | 0 | 100 | 66 | 163174.7 | 98.65
25 | 0 | 100 | 66 | 163174.6 | 98.69
26.3 | 0 | 100 | 65.9 | 162934.8 | 98.64
27.4 | 0 | 100 | 65.68 | 162396 | 98.23
28.7 | 0 | 100 | 65.55 | 162065.8 | 98.04
29.9 | 0 | 100 | 65.54 | 162028.1 | 97.99
31 | 0 | 100 | 65.26 | 161347.7 | 97.75
32.3 | 0 | 100 | 65.37 | 161621.1 | 97.7
33.5 | 0 | 100 | 65 | 160713.3 | 97.35
34.7 | 0 | 100 | 64.91 | 160484.4 | 97.11
35.9 | 0 | 100 | 64.72 | 160013.9 | 96.93
37.1 | 0 | 100 | 64.64 | 159822.6 | 96.63
38.3 | 0 | 100 | 64.53 | 159537.6 | 96.58
39.5 | 0 | 100 | 64.24 | 158828.3 | 96.04
40.7 | 0 | 100 | 63.83 | 157807 | 95.59
41.9 | 0 | 100 | 63.72 | 157547.5 | 95.45
43.1 | 0 | 100 | 63.51 | 157010.4 | 94.96
44.3 | 0 | 100 | 63.16 | 156159.4 | 94.44
45.5 | 0 | 100 | 62.82 | 155304 | 94.12
46.7 | 0 | 100 | 62.21 | 153807.9 | 93
47.9 | 0 | 100 | 61.3 | 151550.9 | 91.67
49.1 | 0 | 100 | 59.7 | 147592 | 89.18
50.3 | 0 | 100 | 56.88 | 140619.3 | 85.25
51.5 | 0 | 100 | 52.71 | 130307.7 | 78.81
52.7 | 0 | 100 | 46.88 | 115903.8 | 70.22
53.9 | 0 | 100 | 39.7 | 98160 | 59.32
55.1 | 0 | 100 | 31.84 | 78710.8 | 47.6
56.3 | 0 | 100 | 24.09 | 59560.8 | 36.1
57.6 | 0 | 100 | 17.58 | 43469.9 | 26.28
58.8 | 0 | 100 | 13.3 | 32893.5 | 19.89
59.9 | 0 | 100 | 10.83 | 26779 | 16.21
61.2 | 0 | 100 | 9.37 | 23176 | 14.04
62.4 | 0 | 100 | 8.51 | 21042.3 | 12.74
63.6 | 0 | 100 | 7.93 | 19602.4 | 11.86
64.8 | 0 | 100 | 7.51 | 18559.2 | 11.22
66 | 0 | 100 | 7.03 | 17371.3 | 10.52
67.2 | 0 | 100 | 6.8 | 16814.8 | 10.17
68.4 | 0 | 100 | 6.53 | 16133.7 | 9.76
69.6 | 0 | 100 | 6.25 | 15455.3 | 9.34
70.8 | 0 | 100 | 5.99 | 14798.2 | 8.97
72 | 0 | 100 | 5.77 | 14265.7 | 8.64
73.2 | 0 | 100 | 5.53 | 13675.3 | 8.27
74.4 | 0 | 100 | 5.34 | 13212 | 8
75.6 | 0 | 100 | 5.14 | 12711.2 | 7.71
76.8 | 0 | 100 | 4.95 | 12247 | 7.42
78 | 0 | 100 | 4.79 | 11853.7 | 7.17
79.2 | 0 | 100 | 4.6 | 11360.6 | 6.87
80.4 | 0 | 100 | 4.45 | 11010 | 6.65
81.6 | 0 | 100 | 4.28 | 10585.6 | 6.41
82.8 | 0 | 100 | 4.15 | 10252.6 | 6.2
84 | 0 | 100 | 4 | 9889.2 | 5.99
85.3 | 0 | 100 | 3.89 | 9626.7 | 5.83
86.4 | 0 | 100 | 3.81 | 9423.9 | 5.71
87.6 | 0 | 100 | 3.7 | 9154 | 5.54
88.9 | 0 | 100 | 3.51 | 8680.7 | 5.27
90 | 0 | 100 | 3.45 | 8535.6 | 5.17
91.3 | 0 | 100 | 3.29 | 8142.6 | 4.94
92.5 | 0 | 100 | 3.22 | 7958 | 4.82
93.7 | 0 | 100 | 3.11 | 7696.2 | 4.67
94.9 | 0 | 100 | 3.04 | 7504.1 | 4.55
96.1 | 0 | 100 | 2.97 | 7330.5 | 4.44
97.3 | 0 | 100 | 2.88 | 7119.2 | 4.31
98.5 | 0 | 100 | 2.76 | 6819.9 | 4.13
100 | 0 | 100 | 2.66 | 6580.2 | 3.98
Measurement time: | 12/08/2022 18:27:42
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
100 | 0 | 150 | 3.14 | 7759.3 | 6.55
99.1 | 0 | 150 | 2.83 | 6986.3 | 5.59
97.6 | 0 | 150 | 2.95 | 7285.4 | 5.84
96.5 | 0 | 150 | 3.04 | 7506.1 | 6.01
95.4 | 0 | 150 | 3.15 | 7784 | 6.23
94.1 | 0 | 150 | 3.22 | 7956.5 | 6.37
92.9 | 0 | 150 | 3.37 | 8326.8 | 6.67
91.7 | 0 | 150 | 3.4 | 8416.7 | 6.73
90.5 | 0 | 150 | 3.5 | 8649.6 | 6.9
89.3 | 0 | 150 | 3.62 | 8954.6 | 7.18
88.1 | 0 | 150 | 3.74 | 9256.1 | 7.42
86.9 | 0 | 150 | 3.82 | 9453.7 | 7.56
85.7 | 0 | 150 | 3.98 | 9835.8 | 7.86
84.5 | 0 | 150 | 4.11 | 10159.3 | 8.13
83.3 | 0 | 150 | 4.23 | 10457 | 8.38
82.1 | 0 | 150 | 4.38 | 10835.2 | 8.68
80.9 | 0 | 150 | 4.47 | 11062.9 | 8.86
79.7 | 0 | 150 | 4.6 | 11364 | 9.1
78.4 | 0 | 150 | 4.85 | 11985.9 | 9.6
77.3 | 0 | 150 | 4.98 | 12318.6 | 9.87
76 | 0 | 150 | 5.14 | 12699.1 | 10.16
74.8 | 0 | 150 | 5.3 | 13114.4 | 10.49
73.7 | 0 | 150 | 5.46 | 13487.7 | 10.8
72.4 | 0 | 150 | 5.69 | 14070.3 | 11.26
71.3 | 0 | 150 | 5.89 | 14568.8 | 11.65
70 | 0 | 150 | 6.18 | 15279.1 | 12.21
68.8 | 0 | 150 | 6.43 | 15891.2 | 12.7
67.6 | 0 | 150 | 6.69 | 16545.2 | 13.25
66.4 | 0 | 150 | 7.06 | 17449.4 | 13.98
65.2 | 0 | 150 | 7.5 | 18543.5 | 14.82
64 | 0 | 150 | 8.15 | 20147.3 | 16.13
62.8 | 0 | 150 | 9.05 | 22374.1 | 17.89
61.6 | 0 | 150 | 10.74 | 26543 | 21.26
60.4 | 0 | 150 | 13.78 | 34071.4 | 27.29
59.2 | 0 | 150 | 18.06 | 44648.1 | 35.73
58 | 0 | 150 | 23.64 | 58444.8 | 46.76
56.8 | 0 | 150 | 29.43 | 72763.1 | 58.22
55.6 | 0 | 150 | 34.43 | 85114 | 68.15
54.3 | 0 | 150 | 38.96 | 96319 | 76.98
53.2 | 0 | 150 | 41.96 | 103744.8 | 82.91
51.9 | 0 | 150 | 44.28 | 109466.3 | 87.69
50.8 | 0 | 150 | 45.57 | 112655.6 | 90.05
49.6 | 0 | 150 | 46.21 | 114255.1 | 91.29
48.3 | 0 | 150 | 46.95 | 116067.5 | 92.74
47.2 | 0 | 150 | 47.11 | 116464.9 | 93.37
45.9 | 0 | 150 | 47.33 | 117014.1 | 93.78
44.7 | 0 | 150 | 47.54 | 117532.5 | 94.04
43.5 | 0 | 150 | 47.77 | 118097.4 | 94.51
42.3 | 0 | 150 | 48.05 | 118803.7 | 95.22
41.2 | 0 | 150 | 48.21 | 119186.2 | 95.41
39.9 | 0 | 150 | 48.34 | 119503.2 | 95.72
38.7 | 0 | 150 | 48.53 | 119988.8 | 96.11
37.5 | 0 | 150 | 48.67 | 120320.3 | 96.17
36.3 | 0 | 150 | 48.72 | 120448 | 96.44
35.1 | 0 | 150 | 48.94 | 120999.1 | 96.82
33.9 | 0 | 150 | 49 | 121154.1 | 96.76
32.7 | 0 | 150 | 49.14 | 121485.1 | 97.24
31.5 | 0 | 150 | 49.3 | 121884.9 | 97.67
30.3 | 0 | 150 | 49.42 | 122184.2 | 97.85
29.1 | 0 | 150 | 49.59 | 122601.9 | 98.18
27.9 | 0 | 150 | 49.64 | 122732.8 | 98.25
26.7 | 0 | 150 | 49.76 | 123015.3 | 98.49
25.5 | 0 | 150 | 49.73 | 122942.4 | 98.5
24.2 | 0 | 150 | 49.83 | 123196.6 | 99.01
23.1 | 0 | 150 | 49.99 | 123603.6 | 98.91
21.8 | 0 | 150 | 50.05 | 123740.7 | 98.98
20.7 | 0 | 150 | 50.1 | 123853.2 | 99.08
19.5 | 0 | 150 | 50.17 | 124046.3 | 99.04
18.2 | 0 | 150 | 50.19 | 124082.1 | 99.09
17.1 | 0 | 150 | 50.12 | 123914 | 99.2
15.8 | 0 | 150 | 50.24 | 124206.9 | 99.17
14.7 | 0 | 150 | 50.05 | 123750.4 | 99.23
13.4 | 0 | 150 | 50.17 | 124048.4 | 99.29
12.2 | 0 | 150 | 50.22 | 124152.1 | 99.32
11 | 0 | 150 | 50.16 | 124006 | 99.23
9.8 | 0 | 150 | 50.29 | 124331.2 | 99.53
8.6 | 0 | 150 | 50.25 | 124226.1 | 99.65
7.4 | 0 | 150 | 50.18 | 124074.3 | 99.26
6.2 | 0 | 150 | 50.34 | 124450.1 | 99.43
5 | 0 | 150 | 50.26 | 124272 | 99.46
3.8 | 0 | 150 | 50.34 | 124452.9 | 99.55
2.6 | 0 | 150 | 50.22 | 124161.4 | 99.64
1.4 | 0 | 150 | 50.35 | 124473.6 | 99.71
0.2 | 0 | 150 | 50.33 | 124435.4 | 99.62
-1 | 0 | 150 | 50.41 | 124642.1 | 99.81
-2.2 | 0 | 150 | 50.3 | 124353.5 | 99.56
-3.4 | 0 | 150 | 50.37 | 124531.1 | 99.6
-4.6 | 0 | 150 | 50.41 | 124642.5 | 99.63
-5.8 | 0 | 150 | 50.46 | 124745 | 100
-7 | 0 | 150 | 50.43 | 124690.7 | 99.76
-8.3 | 0 | 150 | 50.4 | 124613.8 | 99.66
-9.4 | 0 | 150 | 50.36 | 124515.6 | 99.79
-10.7 | 0 | 150 | 50.52 | 124911.3 | 99.93
-11.8 | 0 | 150 | 50.47 | 124770.8 | 100
-13.1 | 0 | 150 | 50.44 | 124703.2 | 99.74
-14.3 | 0 | 150 | 50.49 | 124825.9 | 99.87
-15.5 | 0 | 150 | 50.57 | 125017 | 99.9
-16.7 | 0 | 150 | 50.46 | 124767.2 | 99.9
-17.9 | 0 | 150 | 50.56 | 125001.4 | 99.91
-19.1 | 0 | 150 | 50.55 | 124981.3 | 99.92
-20.3 | 0 | 150 | 50.46 | 124754.3 | 99.7
-21.5 | 0 | 150 | 50.39 | 124575.5 | 99.74
-22.7 | 0 | 150 | 50.32 | 124399.3 | 99.55
-23.9 | 0 | 150 | 50.34 | 124468.2 | 99.5
-25.1 | 0 | 150 | 50.36 | 124518.9 | 99.48
-26.3 | 0 | 150 | 50.11 | 123901.8 | 99.4
-27.5 | 0 | 150 | 50.16 | 124021.7 | 99.26
-28.7 | 0 | 150 | 50.05 | 123734 | 99.09
-29.9 | 0 | 150 | 49.97 | 123538.2 | 98.91
-31.1 | 0 | 150 | 49.94 | 123474.6 | 98.76
-32.4 | 0 | 150 | 49.85 | 123257.6 | 98.57
-33.5 | 0 | 150 | 49.61 | 122647.5 | 98.24
-34.7 | 0 | 150 | 49.56 | 122523.9 | 97.87
-36 | 0 | 150 | 49.55 | 122512.5 | 97.92
-37.1 | 0 | 150 | 49.29 | 121864.8 | 97.49
-38.4 | 0 | 150 | 49.22 | 121700.8 | 97.27
-39.5 | 0 | 150 | 48.99 | 121112.3 | 96.97
-40.8 | 0 | 150 | 48.98 | 121085 | 96.85
-42 | 0 | 150 | 48.82 | 120697.8 | 96.68
-43.1 | 0 | 150 | 48.7 | 120402.6 | 96.17
-44.4 | 0 | 150 | 48.51 | 119937.7 | 95.93
-45.6 | 0 | 150 | 48.29 | 119380.7 | 95.54
-46.8 | 0 | 150 | 48.03 | 118737.6 | 95.17
-48 | 0 | 150 | 47.75 | 118062.9 | 94.44
-49.2 | 0 | 150 | 47.5 | 117448.5 | 93.94
-50.4 | 0 | 150 | 47 | 116201.1 | 92.89
-51.6 | 0 | 150 | 46.18 | 114166.4 | 91.33
-52.8 | 0 | 150 | 44.88 | 110967.7 | 88.61
-54 | 0 | 150 | 42.29 | 104548 | 83.62
-55.2 | 0 | 150 | 38.35 | 94814.9 | 75.83
-56.5 | 0 | 150 | 33 | 81588.1 | 65.2
-57.6 | 0 | 150 | 27.38 | 67694.5 | 54.16
-58.8 | 0 | 150 | 21.92 | 54183.5 | 43.42
-60 | 0 | 150 | 17.08 | 42223.4 | 33.77
-61.2 | 0 | 150 | 13.49 | 33343.4 | 26.64
-62.5 | 0 | 150 | 10.79 | 26684.5 | 21.36
-63.7 | 0 | 150 | 9.25 | 22857.1 | 18.3
-64.9 | 0 | 150 | 8.18 | 20233.3 | 16.18
-66.1 | 0 | 150 | 7.64 | 18898.7 | 15.1
-67.3 | 0 | 150 | 7.16 | 17713.2 | 14.17
-68.5 | 0 | 150 | 6.8 | 16803.5 | 13.45
-69.7 | 0 | 150 | 6.49 | 16042.2 | 12.83
-70.9 | 0 | 150 | 6.24 | 15423.3 | 12.36
-72.1 | 0 | 150 | 5.99 | 14820 | 11.86
-73.3 | 0 | 150 | 5.82 | 14392 | 11.5
-74.5 | 0 | 150 | 5.57 | 13780.2 | 11.02
-75.7 | 0 | 150 | 5.41 | 13374.3 | 10.69
-76.9 | 0 | 150 | 5.19 | 12828.9 | 10.26
-78.1 | 0 | 150 | 5.11 | 12630.8 | 10.12
-79.3 | 0 | 150 | 4.88 | 12064.3 | 9.64
-80.5 | 0 | 150 | 4.74 | 11727.1 | 9.39
-81.7 | 0 | 150 | 4.59 | 11347.7 | 9.08
-83 | 0 | 150 | 4.45 | 10998.9 | 8.81
-84.2 | 0 | 150 | 4.3 | 10640.7 | 8.5
-85.3 | 0 | 150 | 4.21 | 10410.8 | 8.31
-86.6 | 0 | 150 | 4.03 | 9956.8 | 7.94
-87.7 | 0 | 150 | 3.96 | 9794.7 | 7.84
-89 | 0 | 150 | 3.83 | 9480.1 | 7.59
-90.2 | 0 | 150 | 3.69 | 9132 | 7.3
-91.3 | 0 | 150 | 3.61 | 8920 | 7.13
-92.6 | 0 | 150 | 3.45 | 8521.9 | 6.82
-93.7 | 0 | 150 | 3.35 | 8289.6 | 6.62
-95 | 0 | 150 | 3.28 | 8119.1 | 6.5
-96.2 | 0 | 150 | 3.15 | 7786.4 | 6.23
-97.4 | 0 | 150 | 3.09 | 7634.2 | 6.12
-98.6 | 0 | 150 | 3 | 7429.3 | 5.94
-100 | 0 | 150 | 2.91 | 7193 | 5.75