## 5_Profils vitesse déplacement c
Measurement time: | 12/08/2022 18:29:30
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Speed | 0,5 cm/s
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-155 | 0 | 100 | 0.75 | 1859 | 1.07
-154.1 | 0 | 100 | 0.83 | 2044.4 | 1.23
-152.6 | 0 | 100 | 0.9 | 2236.3 | 1.35
-151.5 | 0 | 100 | 0.91 | 2260.1 | 1.37
-150.3 | 0 | 100 | 0.93 | 2303.9 | 1.39
-149.1 | 0 | 100 | 0.95 | 2349 | 1.42
-147.9 | 0 | 100 | 0.9 | 2220.6 | 1.34
-146.7 | 0 | 100 | 0.99 | 2443.1 | 1.48
-145.5 | 0 | 100 | 1.02 | 2519.5 | 1.52
-144.3 | 0 | 100 | 1.05 | 2606.6 | 1.57
-143.1 | 0 | 100 | 1.04 | 2579.5 | 1.56
-141.9 | 0 | 100 | 1.07 | 2649.7 | 1.6
-140.7 | 0 | 100 | 1.07 | 2645.2 | 1.6
-139.5 | 0 | 100 | 1.11 | 2751.9 | 1.66
-138.3 | 0 | 100 | 1.15 | 2839.8 | 1.72
-137 | 0 | 100 | 1.22 | 3014.1 | 1.82
-135.9 | 0 | 100 | 1.24 | 3071 | 1.85
-134.7 | 0 | 100 | 1.28 | 3173.6 | 1.92
-133.4 | 0 | 100 | 1.28 | 3165 | 1.91
-132.3 | 0 | 100 | 1.3 | 3206.8 | 1.94
-131.1 | 0 | 100 | 1.33 | 3289.3 | 1.99
-129.8 | 0 | 100 | 1.38 | 3401.5 | 2.05
-128.6 | 0 | 100 | 1.38 | 3404.2 | 2.06
-127.4 | 0 | 100 | 1.45 | 3573.3 | 2.16
-126.2 | 0 | 100 | 1.47 | 3645 | 2.2
-125 | 0 | 100 | 1.52 | 3764.2 | 2.27
-123.8 | 0 | 100 | 1.56 | 3867.8 | 2.34
-122.6 | 0 | 100 | 1.56 | 3868.1 | 2.33
-121.4 | 0 | 100 | 1.66 | 4113.7 | 2.49
-120.2 | 0 | 100 | 1.71 | 4219.3 | 2.55
-119 | 0 | 100 | 1.75 | 4316.5 | 2.61
-117.8 | 0 | 100 | 1.81 | 4463.3 | 2.7
-116.6 | 0 | 100 | 1.87 | 4627.5 | 2.8
-115.4 | 0 | 100 | 1.9 | 4707.5 | 2.85
-114.2 | 0 | 100 | 1.94 | 4804.3 | 2.9
-113 | 0 | 100 | 2 | 4939.3 | 2.99
-111.8 | 0 | 100 | 2.05 | 5066.4 | 3.06
-110.6 | 0 | 100 | 2.09 | 5158.5 | 3.12
-109.4 | 0 | 100 | 2.17 | 5366.1 | 3.24
-108.2 | 0 | 100 | 2.24 | 5550.3 | 3.35
-107 | 0 | 100 | 2.3 | 5693.6 | 3.44
-105.7 | 0 | 100 | 2.34 | 5787.2 | 3.5
-104.6 | 0 | 100 | 2.43 | 5998 | 3.63
-103.3 | 0 | 100 | 2.5 | 6181 | 3.74
-102.1 | 0 | 100 | 2.59 | 6406.4 | 3.87
-100.9 | 0 | 100 | 2.72 | 6737.1 | 4.08
-99.7 | 0 | 100 | 2.76 | 6827.9 | 4.13
-98.5 | 0 | 100 | 2.89 | 7136.4 | 4.32
-97.3 | 0 | 100 | 2.93 | 7233.6 | 4.37
-96.1 | 0 | 100 | 3 | 7415.5 | 4.49
-94.9 | 0 | 100 | 3.13 | 7744.2 | 4.69
-93.7 | 0 | 100 | 3.23 | 7978.4 | 4.83
-92.5 | 0 | 100 | 3.34 | 8267.2 | 5
-91.3 | 0 | 100 | 3.46 | 8566 | 5.17
-90.1 | 0 | 100 | 3.57 | 8826.6 | 5.34
-88.9 | 0 | 100 | 3.66 | 9041.1 | 5.46
-87.7 | 0 | 100 | 3.76 | 9289.1 | 5.61
-86.5 | 0 | 100 | 3.93 | 9712.5 | 5.87
-85.3 | 0 | 100 | 4.05 | 10014.6 | 6.05
-84.1 | 0 | 100 | 4.2 | 10372.5 | 6.28
-82.9 | 0 | 100 | 4.33 | 10711.8 | 6.48
-81.7 | 0 | 100 | 4.54 | 11212.6 | 6.78
-80.5 | 0 | 100 | 4.62 | 11415.2 | 6.91
-79.3 | 0 | 100 | 4.83 | 11946.8 | 7.22
-78.1 | 0 | 100 | 5.05 | 12485.9 | 7.55
-76.9 | 0 | 100 | 5.16 | 12749 | 7.72
-75.7 | 0 | 100 | 5.37 | 13268.6 | 8.02
-74.4 | 0 | 100 | 5.61 | 13866 | 8.4
-73.2 | 0 | 100 | 5.83 | 14421.2 | 8.72
-72.1 | 0 | 100 | 6.01 | 14861 | 9.01
-70.8 | 0 | 100 | 6.25 | 15453 | 9.34
-69.6 | 0 | 100 | 6.52 | 16129.8 | 9.76
-68.4 | 0 | 100 | 6.76 | 16719.3 | 10.1
-67.2 | 0 | 100 | 7.13 | 17631.5 | 10.67
-66 | 0 | 100 | 7.49 | 18514.7 | 11.18
-64.8 | 0 | 100 | 7.91 | 19566.5 | 11.82
-63.6 | 0 | 100 | 8.56 | 21153.2 | 12.78
-62.4 | 0 | 100 | 9.32 | 23047.6 | 13.94
-61.2 | 0 | 100 | 10.64 | 26299.6 | 15.94
-60 | 0 | 100 | 12.97 | 32073.7 | 19.41
-58.8 | 0 | 100 | 16.66 | 41179.6 | 24.87
-57.6 | 0 | 100 | 22.03 | 54455 | 32.91
-56.4 | 0 | 100 | 28.65 | 70839.8 | 42.87
-55.2 | 0 | 100 | 36.72 | 90782.6 | 54.88
-54 | 0 | 100 | 44.75 | 110633.9 | 66.79
-52.8 | 0 | 100 | 51.92 | 128356.1 | 77.69
-51.6 | 0 | 100 | 57.26 | 141573.6 | 85.68
-50.4 | 0 | 100 | 60.43 | 149406.9 | 90.43
-49.2 | 0 | 100 | 62.11 | 153554.2 | 92.99
-48 | 0 | 100 | 62.98 | 155697 | 94.13
-46.8 | 0 | 100 | 63.73 | 157552 | 95.13
-45.5 | 0 | 100 | 64.09 | 158443.4 | 95.87
-44.4 | 0 | 100 | 64.39 | 159196.6 | 96.28
-43.1 | 0 | 100 | 64.55 | 159596.2 | 96.6
-42 | 0 | 100 | 64.87 | 160373.7 | 97.02
-40.7 | 0 | 100 | 65.07 | 160887.5 | 97.29
-39.5 | 0 | 100 | 65.05 | 160830.1 | 97.49
-38.3 | 0 | 100 | 65.37 | 161617.5 | 97.58
-37.1 | 0 | 100 | 65.4 | 161689.1 | 97.74
-35.9 | 0 | 100 | 65.51 | 161957.7 | 98
-34.7 | 0 | 100 | 65.68 | 162383.4 | 98.35
-33.5 | 0 | 100 | 65.91 | 162949.3 | 98.61
-32.3 | 0 | 100 | 66.05 | 163308.9 | 98.78
-31.1 | 0 | 100 | 66.09 | 163398.8 | 98.8
-29.9 | 0 | 100 | 66.14 | 163513.3 | 98.94
-28.7 | 0 | 100 | 66.23 | 163755.7 | 98.94
-27.5 | 0 | 100 | 66.38 | 164122.8 | 99.26
-26.3 | 0 | 100 | 66.53 | 164495.8 | 99.52
-25.1 | 0 | 100 | 66.54 | 164516.6 | 99.24
-23.9 | 0 | 100 | 66.54 | 164514.4 | 99.42
-22.7 | 0 | 100 | 66.61 | 164684 | 99.64
-21.5 | 0 | 100 | 66.64 | 164753.6 | 99.83
-20.3 | 0 | 100 | 66.77 | 165075.2 | 99.91
-19.1 | 0 | 100 | 66.69 | 164879.3 | 99.98
-17.9 | 0 | 100 | 66.64 | 164753.9 | 99.75
-16.7 | 0 | 100 | 66.75 | 165039.2 | 99.9
-15.5 | 0 | 100 | 66.7 | 164909.5 | 100
-14.3 | 0 | 100 | 66.87 | 165315.4 | 99.89
-13 | 0 | 100 | 66.71 | 164929.5 | 99.81
-11.9 | 0 | 100 | 66.69 | 164871.5 | 99.83
-10.6 | 0 | 100 | 66.61 | 164693.1 | 99.85
-9.5 | 0 | 100 | 66.67 | 164822.6 | 99.73
-8.3 | 0 | 100 | 66.59 | 164645.7 | 99.54
-7 | 0 | 100 | 66.47 | 164335.2 | 99.52
-5.8 | 0 | 100 | 66.5 | 164406.6 | 99.52
-4.6 | 0 | 100 | 66.38 | 164113.3 | 99.26
-3.4 | 0 | 100 | 66.45 | 164299.4 | 99.37
-2.2 | 0 | 100 | 66.35 | 164032.3 | 99.25
-1 | 0 | 100 | 66.44 | 164273.3 | 99.3
0.2 | 0 | 100 | 66.38 | 164111.1 | 99.3
1.4 | 0 | 100 | 66.53 | 164477.3 | 99.39
2.6 | 0 | 100 | 66.42 | 164218 | 99.36
3.8 | 0 | 100 | 66.53 | 164475.3 | 99.33
5 | 0 | 100 | 66.46 | 164303.5 | 99.4
6.2 | 0 | 100 | 66.3 | 163928 | 99.32
7.4 | 0 | 100 | 66.47 | 164345.4 | 99.5
8.6 | 0 | 100 | 66.35 | 164047.5 | 99.46
9.8 | 0 | 100 | 66.31 | 163941 | 99.38
11 | 0 | 100 | 66.35 | 164044.1 | 99.29
12.2 | 0 | 100 | 66.32 | 163954.4 | 99.2
13.4 | 0 | 100 | 66.35 | 164030.1 | 99.34
14.6 | 0 | 100 | 66.39 | 164145.4 | 99.33
15.8 | 0 | 100 | 66.38 | 164115.6 | 99.16
17 | 0 | 100 | 66.36 | 164070.2 | 99.21
18.2 | 0 | 100 | 66.31 | 163947.1 | 99.13
19.5 | 0 | 100 | 66.17 | 163596.9 | 99.11
20.6 | 0 | 100 | 66.21 | 163692.6 | 99.14
21.9 | 0 | 100 | 66.22 | 163730.6 | 99.05
23.1 | 0 | 100 | 66.09 | 163393.2 | 98.82
24.2 | 0 | 100 | 66.05 | 163305.1 | 98.9
25.5 | 0 | 100 | 65.92 | 162985.4 | 98.68
26.7 | 0 | 100 | 65.82 | 162736.2 | 98.52
27.9 | 0 | 100 | 65.59 | 162162.5 | 98.2
29.1 | 0 | 100 | 65.53 | 162008.1 | 97.9
30.3 | 0 | 100 | 65.39 | 161659.5 | 97.68
31.5 | 0 | 100 | 65.24 | 161295.3 | 97.58
32.7 | 0 | 100 | 65.15 | 161068.4 | 97.39
33.9 | 0 | 100 | 64.96 | 160613.8 | 96.99
35.1 | 0 | 100 | 64.73 | 160032.3 | 96.79
36.3 | 0 | 100 | 64.65 | 159840.5 | 96.6
37.5 | 0 | 100 | 64.4 | 159225.9 | 96.44
38.7 | 0 | 100 | 64.24 | 158821.6 | 96.13
39.9 | 0 | 100 | 64.11 | 158512 | 95.75
41.1 | 0 | 100 | 63.89 | 157949.4 | 95.36
42.3 | 0 | 100 | 63.74 | 157596.2 | 95.2
43.5 | 0 | 100 | 63.44 | 156849.3 | 94.99
44.7 | 0 | 100 | 62.97 | 155691.9 | 94.16
45.9 | 0 | 100 | 62.66 | 154905.7 | 93.83
47.1 | 0 | 100 | 61.92 | 153083.6 | 92.88
48.3 | 0 | 100 | 60.81 | 150335.1 | 91.06
49.5 | 0 | 100 | 58.76 | 145274.6 | 87.88
50.7 | 0 | 100 | 55.58 | 137412.7 | 83.3
51.9 | 0 | 100 | 50.96 | 125986.5 | 76.32
53.1 | 0 | 100 | 44.62 | 110323.4 | 66.85
54.3 | 0 | 100 | 37.03 | 91561.1 | 55.33
55.5 | 0 | 100 | 29.01 | 71727.2 | 43.41
56.7 | 0 | 100 | 21.71 | 53671.8 | 32.48
57.9 | 0 | 100 | 15.85 | 39177.3 | 23.69
59.2 | 0 | 100 | 12.13 | 29990.9 | 18.16
60.3 | 0 | 100 | 10.29 | 25432 | 15.39
61.6 | 0 | 100 | 9.03 | 22333.5 | 13.52
62.8 | 0 | 100 | 8.32 | 20557.7 | 12.42
64 | 0 | 100 | 7.73 | 19122.1 | 11.58
65.2 | 0 | 100 | 7.3 | 18042.3 | 10.93
66.4 | 0 | 100 | 6.99 | 17272.5 | 10.45
67.6 | 0 | 100 | 6.67 | 16488.3 | 9.97
68.8 | 0 | 100 | 6.37 | 15752.7 | 9.54
70 | 0 | 100 | 6.17 | 15254.1 | 9.23
71.2 | 0 | 100 | 5.87 | 14521.8 | 8.8
72.4 | 0 | 100 | 5.67 | 14013.9 | 8.49
73.6 | 0 | 100 | 5.43 | 13424.9 | 8.14
74.8 | 0 | 100 | 5.26 | 13002.6 | 7.89
76 | 0 | 100 | 5.07 | 12541 | 7.59
77.2 | 0 | 100 | 4.92 | 12152.8 | 7.35
78.4 | 0 | 100 | 4.71 | 11649.9 | 7.04
79.6 | 0 | 100 | 4.58 | 11325 | 6.84
80.8 | 0 | 100 | 4.45 | 11011.9 | 6.65
82 | 0 | 100 | 4.24 | 10493.3 | 6.35
83.3 | 0 | 100 | 4.12 | 10186.4 | 6.17
84.4 | 0 | 100 | 3.97 | 9805.1 | 5.94
85.6 | 0 | 100 | 3.85 | 9527.4 | 5.78
86.8 | 0 | 100 | 3.75 | 9263.1 | 5.61
88.1 | 0 | 100 | 3.61 | 8925.3 | 5.4
89.3 | 0 | 100 | 3.51 | 8667 | 5.25
90.5 | 0 | 100 | 3.37 | 8344.1 | 5.06
91.7 | 0 | 100 | 3.29 | 8131.2 | 4.92
92.9 | 0 | 100 | 3.17 | 7848.9 | 4.76
94.1 | 0 | 100 | 3.08 | 7611.8 | 4.61
95.3 | 0 | 100 | 2.98 | 7377.6 | 4.46
96.5 | 0 | 100 | 2.93 | 7235 | 4.38
97.7 | 0 | 100 | 2.8 | 6910.3 | 4.19
98.9 | 0 | 100 | 2.75 | 6802.5 | 4.12
100.1 | 0 | 100 | 2.64 | 6525.5 | 3.95
101.3 | 0 | 100 | 2.58 | 6374.7 | 3.85
102.5 | 0 | 100 | 2.54 | 6281.2 | 3.8
103.7 | 0 | 100 | 2.42 | 5990.9 | 3.63
104.9 | 0 | 100 | 2.38 | 5893.3 | 3.57
106.1 | 0 | 100 | 2.31 | 5707.4 | 3.46
107.3 | 0 | 100 | 2.24 | 5538 | 3.34
108.5 | 0 | 100 | 2.18 | 5384.6 | 3.26
109.8 | 0 | 100 | 2.1 | 5192.7 | 3.15
110.9 | 0 | 100 | 2.1 | 5191.9 | 3.14
112.2 | 0 | 100 | 2.02 | 4989 | 3.02
113.4 | 0 | 100 | 1.95 | 4816.7 | 2.92
114.6 | 0 | 100 | 1.91 | 4726.3 | 2.86
115.8 | 0 | 100 | 1.87 | 4621.7 | 2.8
117 | 0 | 100 | 1.81 | 4467.7 | 2.71
118.2 | 0 | 100 | 1.76 | 4341.9 | 2.63
119.4 | 0 | 100 | 1.68 | 4147.4 | 2.51
120.6 | 0 | 100 | 1.66 | 4099.5 | 2.48
121.8 | 0 | 100 | 1.62 | 4002.8 | 2.42
123 | 0 | 100 | 1.61 | 3973 | 2.4
124.2 | 0 | 100 | 1.56 | 3845.2 | 2.33
125.4 | 0 | 100 | 1.5 | 3699.8 | 2.24
126.6 | 0 | 100 | 1.48 | 3651.9 | 2.21
127.8 | 0 | 100 | 1.42 | 3516.6 | 2.13
129 | 0 | 100 | 1.44 | 3559.4 | 2.16
130.2 | 0 | 100 | 1.41 | 3485.3 | 2.11
131.4 | 0 | 100 | 1.3 | 3225 | 1.95
132.6 | 0 | 100 | 1.3 | 3213 | 1.95
133.8 | 0 | 100 | 1.27 | 3148.8 | 1.91
135 | 0 | 100 | 1.25 | 3085.9 | 1.87
136.2 | 0 | 100 | 1.18 | 2925.5 | 1.78
137.4 | 0 | 100 | 1.21 | 2988 | 1.81
138.6 | 0 | 100 | 1.14 | 2822.9 | 1.71
139.8 | 0 | 100 | 1.15 | 2844 | 1.72
141 | 0 | 100 | 1.09 | 2702.7 | 1.64
142.2 | 0 | 100 | 1.09 | 2698.7 | 1.63
143.4 | 0 | 100 | 1.01 | 2492.1 | 1.51
144.7 | 0 | 100 | 1.04 | 2573.8 | 1.56
145.9 | 0 | 100 | 1.01 | 2492.3 | 1.51
147 | 0 | 100 | 0.96 | 2382.3 | 1.44
148.2 | 0 | 100 | 0.99 | 2449.8 | 1.48
149.5 | 0 | 100 | 0.93 | 2287.8 | 1.39
150.7 | 0 | 100 | 0.95 | 2360.7 | 1.43
151.9 | 0 | 100 | 0.91 | 2240.4 | 1.36
153.1 | 0 | 100 | 0.89 | 2197.4 | 1.33
155 | 0 | 100 | 0.87 | 2148.5 | 1.3
Measurement time: | 12/08/2022 18:30:41
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Speed | 1 cm/s
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
155 | 0 | 100 | 0.69 | 1710.9 | 1.04
154.4 | 0 | 100 | 0.84 | 2082.6 | 1.26
152.9 | 0 | 100 | 0.87 | 2145.2 | 1.3
151.5 | 0 | 100 | 0.96 | 2372.5 | 1.43
150.2 | 0 | 100 | 0.98 | 2418.5 | 1.46
149.1 | 0 | 100 | 0.95 | 2353.1 | 1.42
148 | 0 | 100 | 0.94 | 2317.5 | 1.4
146.9 | 0 | 100 | 0.99 | 2457.8 | 1.48
145.6 | 0 | 100 | 1.04 | 2571 | 1.55
144.3 | 0 | 100 | 1.01 | 2508.1 | 1.51
143.1 | 0 | 100 | 1.04 | 2582.7 | 1.56
141.9 | 0 | 100 | 1.02 | 2524.4 | 1.53
140.7 | 0 | 100 | 1.05 | 2586.3 | 1.56
139.5 | 0 | 100 | 1.15 | 2844.6 | 1.72
138.3 | 0 | 100 | 1.1 | 2723.9 | 1.64
137.1 | 0 | 100 | 1.17 | 2882.5 | 1.74
135.9 | 0 | 100 | 1.18 | 2929.7 | 1.77
134.7 | 0 | 100 | 1.16 | 2871.8 | 1.74
133.5 | 0 | 100 | 1.21 | 3002.4 | 1.82
132.3 | 0 | 100 | 1.32 | 3265.1 | 1.97
131.1 | 0 | 100 | 1.35 | 3326 | 2.01
129.9 | 0 | 100 | 1.41 | 3494.5 | 2.11
128.7 | 0 | 100 | 1.42 | 3518.6 | 2.13
127.5 | 0 | 100 | 1.41 | 3482.2 | 2.11
126.3 | 0 | 100 | 1.47 | 3633.6 | 2.2
125.1 | 0 | 100 | 1.53 | 3786 | 2.29
123.9 | 0 | 100 | 1.48 | 3662.7 | 2.22
122.7 | 0 | 100 | 1.6 | 3967.7 | 2.4
121.5 | 0 | 100 | 1.64 | 4058.6 | 2.46
120.2 | 0 | 100 | 1.69 | 4182.3 | 2.52
119.1 | 0 | 100 | 1.67 | 4138.4 | 2.5
117.8 | 0 | 100 | 1.78 | 4391.8 | 2.65
116.6 | 0 | 100 | 1.78 | 4391.5 | 2.65
115.5 | 0 | 100 | 1.82 | 4503.4 | 2.72
114.2 | 0 | 100 | 1.92 | 4748.9 | 2.87
113 | 0 | 100 | 1.96 | 4842.5 | 2.92
111.8 | 0 | 100 | 2.04 | 5036.9 | 3.04
110.6 | 0 | 100 | 2.04 | 5041.4 | 3.05
109.4 | 0 | 100 | 2.11 | 5213.8 | 3.16
108.2 | 0 | 100 | 2.19 | 5402.4 | 3.26
107 | 0 | 100 | 2.21 | 5460.4 | 3.3
105.8 | 0 | 100 | 2.32 | 5746.3 | 3.47
104.6 | 0 | 100 | 2.35 | 5805.5 | 3.5
103.4 | 0 | 100 | 2.43 | 6005.2 | 3.64
102.2 | 0 | 100 | 2.55 | 6301.3 | 3.81
101 | 0 | 100 | 2.58 | 6375.6 | 3.86
99.8 | 0 | 100 | 2.69 | 6646.1 | 4.03
98.5 | 0 | 100 | 2.77 | 6839.8 | 4.13
97.3 | 0 | 100 | 2.83 | 6990 | 4.23
96.1 | 0 | 100 | 2.95 | 7289.6 | 4.41
94.9 | 0 | 100 | 3.05 | 7542 | 4.57
93.7 | 0 | 100 | 3.1 | 7657.2 | 4.62
92.5 | 0 | 100 | 3.24 | 8013.8 | 4.86
91.3 | 0 | 100 | 3.27 | 8096.8 | 4.9
90.1 | 0 | 100 | 3.38 | 8362.3 | 5.04
88.9 | 0 | 100 | 3.52 | 8701.9 | 5.26
87.7 | 0 | 100 | 3.62 | 8941.2 | 5.4
86.5 | 0 | 100 | 3.75 | 9259.2 | 5.58
85.2 | 0 | 100 | 3.8 | 9406.8 | 5.68
84 | 0 | 100 | 4 | 9900.3 | 6
82.8 | 0 | 100 | 4.21 | 10411.9 | 6.29
81.6 | 0 | 100 | 4.32 | 10673.7 | 6.45
80.4 | 0 | 100 | 4.44 | 10989.3 | 6.64
79.2 | 0 | 100 | 4.63 | 11457.1 | 6.93
78 | 0 | 100 | 4.89 | 12102.1 | 7.31
76.8 | 0 | 100 | 4.94 | 12217 | 7.39
75.6 | 0 | 100 | 5.11 | 12639.9 | 7.62
74.4 | 0 | 100 | 5.32 | 13145.6 | 7.96
73.2 | 0 | 100 | 5.54 | 13697.8 | 8.27
72 | 0 | 100 | 5.78 | 14285.8 | 8.64
70.8 | 0 | 100 | 5.96 | 14731.8 | 8.91
69.6 | 0 | 100 | 6.21 | 15352.8 | 9.28
68.4 | 0 | 100 | 6.46 | 15978.7 | 9.64
67.1 | 0 | 100 | 6.81 | 16839.7 | 10.16
65.9 | 0 | 100 | 7.16 | 17697.5 | 10.7
64.7 | 0 | 100 | 7.51 | 18556.7 | 11.22
63.5 | 0 | 100 | 7.97 | 19715.6 | 11.9
62.3 | 0 | 100 | 8.69 | 21475.9 | 12.98
61.1 | 0 | 100 | 9.64 | 23845.5 | 14.35
59.9 | 0 | 100 | 11.19 | 27656.6 | 16.72
58.7 | 0 | 100 | 13.98 | 34554 | 20.89
57.4 | 0 | 100 | 19.12 | 47259.5 | 28.57
56.3 | 0 | 100 | 25.75 | 63667.5 | 38.47
55 | 0 | 100 | 33.41 | 82602.3 | 49.96
53.8 | 0 | 100 | 41.6 | 102859.1 | 62.12
52.7 | 0 | 100 | 48.32 | 119460.8 | 72.2
51.4 | 0 | 100 | 53.82 | 133065.6 | 80.48
50.2 | 0 | 100 | 57.5 | 142158.3 | 85.84
49 | 0 | 100 | 59.95 | 148214.1 | 89.72
47.8 | 0 | 100 | 61.37 | 151721.8 | 91.88
46.6 | 0 | 100 | 62.3 | 154034.8 | 93.2
45.4 | 0 | 100 | 62.97 | 155687 | 94.06
44.2 | 0 | 100 | 63.21 | 156278.6 | 94.31
43 | 0 | 100 | 63.48 | 156944.3 | 94.79
41.8 | 0 | 100 | 63.87 | 157910.2 | 95.52
40.6 | 0 | 100 | 64.05 | 158351.8 | 95.73
39.3 | 0 | 100 | 64.31 | 158995 | 96.02
38.1 | 0 | 100 | 64.4 | 159228.1 | 96.21
36.9 | 0 | 100 | 64.46 | 159371.6 | 96.19
35.7 | 0 | 100 | 64.9 | 160464 | 96.85
34.5 | 0 | 100 | 65 | 160711.2 | 97.24
33.3 | 0 | 100 | 65.27 | 161366.5 | 97.42
32.1 | 0 | 100 | 65.07 | 160869.9 | 97.45
30.8 | 0 | 100 | 65.36 | 161604.4 | 97.66
29.7 | 0 | 100 | 65.56 | 162092.5 | 98.03
28.5 | 0 | 100 | 65.69 | 162403.1 | 98.33
27.2 | 0 | 100 | 65.7 | 162442.3 | 98.18
26.1 | 0 | 100 | 65.83 | 162745.8 | 98.53
24.8 | 0 | 100 | 66.04 | 163269.3 | 98.62
23.6 | 0 | 100 | 66 | 163169.6 | 98.76
22.4 | 0 | 100 | 66.13 | 163507.3 | 98.74
21.2 | 0 | 100 | 66.3 | 163906.4 | 99.06
20 | 0 | 100 | 66.23 | 163751.2 | 99.11
18.8 | 0 | 100 | 66.31 | 163933 | 99.16
17.6 | 0 | 100 | 66.45 | 164291.6 | 99.17
16.4 | 0 | 100 | 66.33 | 163999.6 | 99.32
15.1 | 0 | 100 | 66.42 | 164216 | 99.39
14 | 0 | 100 | 66.4 | 164153.8 | 99.05
12.7 | 0 | 100 | 66.32 | 163956.8 | 98.95
11.6 | 0 | 100 | 66.43 | 164229.7 | 99.13
10.4 | 0 | 100 | 66.42 | 164214.5 | 99.28
9.1 | 0 | 100 | 66.51 | 164424.2 | 99.12
7.9 | 0 | 100 | 66.53 | 164476.7 | 99.06
6.7 | 0 | 100 | 66.29 | 163881.1 | 99.14
5.5 | 0 | 100 | 66.37 | 164078.6 | 99.04
4.3 | 0 | 100 | 66.34 | 164011.8 | 99.2
3.1 | 0 | 100 | 66.47 | 164329.6 | 99.34
1.9 | 0 | 100 | 66.35 | 164036.1 | 99.07
0.6 | 0 | 100 | 66.35 | 164043.4 | 99.29
-0.5 | 0 | 100 | 66.4 | 164162 | 99.1
-1.7 | 0 | 100 | 66.42 | 164224.1 | 99.19
-3 | 0 | 100 | 66.47 | 164329.9 | 99.31
-4.1 | 0 | 100 | 66.51 | 164427.9 | 99.51
-5.4 | 0 | 100 | 66.47 | 164331.7 | 99.42
-6.5 | 0 | 100 | 66.49 | 164382.2 | 99.23
-7.8 | 0 | 100 | 66.69 | 164882.4 | 99.72
-9 | 0 | 100 | 66.61 | 164686.3 | 99.41
-10.2 | 0 | 100 | 66.65 | 164776.5 | 99.87
-11.4 | 0 | 100 | 66.64 | 164751.6 | 100
-12.6 | 0 | 100 | 66.66 | 164803.2 | 99.58
-13.8 | 0 | 100 | 66.78 | 165107 | 99.81
-15 | 0 | 100 | 66.83 | 165226 | 99.92
-16.2 | 0 | 100 | 66.87 | 165336.9 | 99.94
-17.5 | 0 | 100 | 66.96 | 165548.1 | 99.73
-18.6 | 0 | 100 | 66.72 | 164958.3 | 99.84
-19.9 | 0 | 100 | 66.95 | 165526.6 | 99.67
-21.1 | 0 | 100 | 66.77 | 165076.9 | 99.72
-22.3 | 0 | 100 | 66.7 | 164898.3 | 99.73
-23.5 | 0 | 100 | 66.58 | 164614 | 99.52
-24.7 | 0 | 100 | 66.7 | 164899.2 | 99.67
-25.9 | 0 | 100 | 66.68 | 164850.2 | 99.57
-27.1 | 0 | 100 | 66.38 | 164106.4 | 99.2
-28.3 | 0 | 100 | 66.41 | 164177.8 | 99.1
-29.5 | 0 | 100 | 66.16 | 163573.6 | 99.07
-30.7 | 0 | 100 | 66.26 | 163815.1 | 98.92
-31.9 | 0 | 100 | 65.89 | 162908.4 | 98.44
-33.1 | 0 | 100 | 66 | 163178.1 | 98.48
-34.3 | 0 | 100 | 65.69 | 162417.9 | 98.16
-35.6 | 0 | 100 | 65.64 | 162279 | 97.98
-36.8 | 0 | 100 | 65.48 | 161877.8 | 97.79
-38 | 0 | 100 | 65.54 | 162049.7 | 97.9
-39.2 | 0 | 100 | 65.16 | 161109.5 | 97.3
-40.4 | 0 | 100 | 65.1 | 160938.3 | 97.14
-41.6 | 0 | 100 | 64.88 | 160402.8 | 96.93
-42.7 | 0 | 100 | 64.6 | 159724.7 | 96.54
-44 | 0 | 100 | 64.36 | 159130.1 | 96.31
-45.2 | 0 | 100 | 64.11 | 158494.1 | 95.73
-46.4 | 0 | 100 | 63.86 | 157876.2 | 95.09
-47.6 | 0 | 100 | 63.06 | 155910.3 | 94.25
-48.8 | 0 | 100 | 62.42 | 154324.2 | 92.85
-50 | 0 | 100 | 60.76 | 150211.7 | 90.98
-51.2 | 0 | 100 | 57.74 | 142749.4 | 86.19
-52.4 | 0 | 100 | 52.87 | 130713.4 | 78.98
-53.7 | 0 | 100 | 45.59 | 112719.5 | 68.1
-54.9 | 0 | 100 | 37.52 | 92760.6 | 56.15
-56.1 | 0 | 100 | 29.42 | 72733.2 | 43.98
-57.3 | 0 | 100 | 22.4 | 55384.2 | 33.37
-58.5 | 0 | 100 | 17.16 | 42426.2 | 25.64
-59.7 | 0 | 100 | 13.22 | 32695.7 | 19.73
-60.9 | 0 | 100 | 10.84 | 26794.5 | 16.16
-62.2 | 0 | 100 | 9.5 | 23484.9 | 14.17
-63.4 | 0 | 100 | 8.58 | 21222.9 | 12.81
-64.6 | 0 | 100 | 7.94 | 19635.3 | 11.89
-65.8 | 0 | 100 | 7.47 | 18466.2 | 11.15
-67 | 0 | 100 | 7.12 | 17595.3 | 10.64
-68.2 | 0 | 100 | 6.8 | 16816.4 | 10.17
-69.4 | 0 | 100 | 6.55 | 16197.6 | 9.8
-70.6 | 0 | 100 | 6.32 | 15617.3 | 9.42
-71.8 | 0 | 100 | 6.06 | 14994.3 | 9.06
-73 | 0 | 100 | 5.86 | 14481 | 8.78
-74.3 | 0 | 100 | 5.67 | 14030.5 | 8.46
-75.5 | 0 | 100 | 5.45 | 13466.9 | 8.12
-76.7 | 0 | 100 | 5.21 | 12877.2 | 7.78
-77.9 | 0 | 100 | 5.02 | 12399.4 | 7.48
-79.1 | 0 | 100 | 4.87 | 12028.5 | 7.27
-80.3 | 0 | 100 | 4.67 | 11536.6 | 6.98
-81.5 | 0 | 100 | 4.54 | 11227.3 | 6.79
-82.7 | 0 | 100 | 4.38 | 10826.8 | 6.54
-83.9 | 0 | 100 | 4.19 | 10362.5 | 6.25
-85.1 | 0 | 100 | 4.01 | 9902.5 | 5.98
-86.4 | 0 | 100 | 3.88 | 9582.4 | 5.81
-87.6 | 0 | 100 | 3.75 | 9266.7 | 5.6
-88.8 | 0 | 100 | 3.65 | 9033.9 | 5.46
-90 | 0 | 100 | 3.53 | 8733.6 | 5.28
-91.2 | 0 | 100 | 3.45 | 8535.8 | 5.16
-92.4 | 0 | 100 | 3.33 | 8228.6 | 4.97
-93.6 | 0 | 100 | 3.2 | 7901.6 | 4.79
-94.8 | 0 | 100 | 3.19 | 7898.6 | 4.77
-96 | 0 | 100 | 2.98 | 7358.5 | 4.44
-97.2 | 0 | 100 | 3 | 7428.9 | 4.49
-98.4 | 0 | 100 | 2.79 | 6889.7 | 4.16
-99.6 | 0 | 100 | 2.75 | 6787.7 | 4.09
-100.8 | 0 | 100 | 2.69 | 6659.8 | 4.02
-102 | 0 | 100 | 2.63 | 6512.4 | 3.92
-103.2 | 0 | 100 | 2.52 | 6232 | 3.76
-104.5 | 0 | 100 | 2.44 | 6022.4 | 3.63
-105.7 | 0 | 100 | 2.34 | 5787.1 | 3.5
-106.9 | 0 | 100 | 2.31 | 5699.6 | 3.44
-108.1 | 0 | 100 | 2.23 | 5504.4 | 3.33
-109.3 | 0 | 100 | 2.22 | 5476.3 | 3.31
-110.5 | 0 | 100 | 2.14 | 5289.9 | 3.2
-111.7 | 0 | 100 | 2.08 | 5143.2 | 3.12
-112.9 | 0 | 100 | 1.97 | 4882.9 | 2.95
-114.2 | 0 | 100 | 1.93 | 4782 | 2.89
-115.3 | 0 | 100 | 1.83 | 4523.1 | 2.74
-116.5 | 0 | 100 | 1.82 | 4510.8 | 2.72
-117.7 | 0 | 100 | 1.77 | 4387 | 2.65
-119 | 0 | 100 | 1.74 | 4309.2 | 2.6
-120.2 | 0 | 100 | 1.66 | 4100.8 | 2.48
-121.3 | 0 | 100 | 1.67 | 4132.2 | 2.5
-122.6 | 0 | 100 | 1.58 | 3905.8 | 2.36
-123.8 | 0 | 100 | 1.6 | 3949.9 | 2.39
-125 | 0 | 100 | 1.56 | 3849 | 2.33
-126.2 | 0 | 100 | 1.44 | 3552.9 | 2.15
-127.4 | 0 | 100 | 1.42 | 3509.3 | 2.12
-128.6 | 0 | 100 | 1.32 | 3269.4 | 1.97
-129.8 | 0 | 100 | 1.36 | 3352.2 | 2.03
-131 | 0 | 100 | 1.31 | 3242 | 1.96
-132.2 | 0 | 100 | 1.28 | 3155.6 | 1.91
-133.4 | 0 | 100 | 1.28 | 3153.4 | 1.9
-134.6 | 0 | 100 | 1.24 | 3061.6 | 1.85
-135.9 | 0 | 100 | 1.21 | 2993.5 | 1.81
-137.1 | 0 | 100 | 1.2 | 2973.9 | 1.79
-138.3 | 0 | 100 | 1.15 | 2848.5 | 1.72
-139.4 | 0 | 100 | 1.16 | 2870.1 | 1.73
-140.7 | 0 | 100 | 1.09 | 2707.2 | 1.63
-141.9 | 0 | 100 | 1.11 | 2745.8 | 1.66
-143.1 | 0 | 100 | 1.07 | 2639.7 | 1.59
-144.3 | 0 | 100 | 0.96 | 2376.5 | 1.43
-145.5 | 0 | 100 | 0.91 | 2251.2 | 1.36
-146.7 | 0 | 100 | 0.93 | 2311.4 | 1.4
-147.9 | 0 | 100 | 0.97 | 2404.3 | 1.45
-149.2 | 0 | 100 | 0.91 | 2256.3 | 1.36
-150.3 | 0 | 100 | 0.94 | 2321.7 | 1.4
-151.5 | 0 | 100 | 0.87 | 2157 | 1.3
-152.8 | 0 | 100 | 0.82 | 2030.8 | 1.23
-154 | 0 | 100 | 0.88 | 2164.4 | 1.31
-155 | 0 | 100 | 0.85 | 2094.8 | 1.26