## 6_Profils DSP
Measurement time: | 12/08/2022 18:32:47
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-155 | 0 | 100 | 0.77 | 1903.4 | 1.15
-154.2 | 0 | 100 | 0.84 | 2079.2 | 1.25
-152.8 | 0 | 100 | 0.93 | 2292 | 1.38
-151.7 | 0 | 100 | 0.88 | 2169.1 | 1.31
-150.6 | 0 | 100 | 0.91 | 2243.8 | 1.35
-149.3 | 0 | 100 | 0.94 | 2322.5 | 1.4
-148.1 | 0 | 100 | 0.96 | 2377.7 | 1.43
-146.9 | 0 | 100 | 0.98 | 2424.2 | 1.46
-145.7 | 0 | 100 | 0.99 | 2449.3 | 1.48
-144.5 | 0 | 100 | 1.01 | 2499.9 | 1.51
-143.3 | 0 | 100 | 1.06 | 2622.5 | 1.58
-142.1 | 0 | 100 | 1.08 | 2663.1 | 1.6
-140.9 | 0 | 100 | 1.13 | 2785.1 | 1.68
-139.7 | 0 | 100 | 1.18 | 2928 | 1.77
-138.5 | 0 | 100 | 1.18 | 2921 | 1.77
-137.3 | 0 | 100 | 1.2 | 2956.6 | 1.78
-136.1 | 0 | 100 | 1.21 | 2999.6 | 1.81
-134.9 | 0 | 100 | 1.24 | 3066.2 | 1.85
-133.7 | 0 | 100 | 1.26 | 3121.8 | 1.88
-132.5 | 0 | 100 | 1.29 | 3186.7 | 1.93
-131.3 | 0 | 100 | 1.33 | 3294.8 | 1.99
-130.1 | 0 | 100 | 1.38 | 3409.9 | 2.06
-128.9 | 0 | 100 | 1.36 | 3352 | 2.02
-127.6 | 0 | 100 | 1.48 | 3670.9 | 2.21
-126.5 | 0 | 100 | 1.48 | 3669.4 | 2.21
-125.3 | 0 | 100 | 1.55 | 3833.2 | 2.31
-124 | 0 | 100 | 1.52 | 3756 | 2.27
-122.9 | 0 | 100 | 1.62 | 4001.3 | 2.42
-121.6 | 0 | 100 | 1.64 | 4062.3 | 2.45
-120.4 | 0 | 100 | 1.75 | 4322.9 | 2.61
-119.2 | 0 | 100 | 1.78 | 4400.9 | 2.65
-118 | 0 | 100 | 1.78 | 4402.2 | 2.65
-116.9 | 0 | 100 | 1.87 | 4619.4 | 2.78
-115.6 | 0 | 100 | 1.9 | 4709.3 | 2.84
-114.4 | 0 | 100 | 1.94 | 4787.4 | 2.89
-113.2 | 0 | 100 | 2 | 4955.6 | 3
-112 | 0 | 100 | 2.07 | 5124.3 | 3.09
-110.8 | 0 | 100 | 2.16 | 5337.3 | 3.22
-109.6 | 0 | 100 | 2.21 | 5464.2 | 3.3
-108.4 | 0 | 100 | 2.24 | 5547.2 | 3.35
-107.2 | 0 | 100 | 2.32 | 5725.1 | 3.46
-106 | 0 | 100 | 2.43 | 6003.6 | 3.63
-104.8 | 0 | 100 | 2.44 | 6026.9 | 3.64
-103.6 | 0 | 100 | 2.56 | 6329.9 | 3.82
-102.4 | 0 | 100 | 2.61 | 6442.9 | 3.89
-101.2 | 0 | 100 | 2.74 | 6781.8 | 4.1
-100 | 0 | 100 | 2.81 | 6944.8 | 4.19
-98.8 | 0 | 100 | 2.87 | 7095.8 | 4.29
-97.5 | 0 | 100 | 2.93 | 7239.1 | 4.37
-96.4 | 0 | 100 | 3.04 | 7519.9 | 4.54
-95.1 | 0 | 100 | 3.11 | 7698.4 | 4.65
-94 | 0 | 100 | 3.27 | 8081 | 4.89
-92.8 | 0 | 100 | 3.35 | 8275.3 | 5
-91.5 | 0 | 100 | 3.45 | 8520.1 | 5.15
-90.3 | 0 | 100 | 3.58 | 8853.2 | 5.36
-89.1 | 0 | 100 | 3.67 | 9079.7 | 5.49
-87.9 | 0 | 100 | 3.8 | 9402.1 | 5.68
-86.7 | 0 | 100 | 3.91 | 9671.4 | 5.84
-85.5 | 0 | 100 | 4.09 | 10121.9 | 6.11
-84.3 | 0 | 100 | 4.17 | 10318.2 | 6.22
-83.1 | 0 | 100 | 4.28 | 10576.1 | 6.38
-81.9 | 0 | 100 | 4.48 | 11081.6 | 6.7
-80.7 | 0 | 100 | 4.6 | 11382.5 | 6.88
-79.5 | 0 | 100 | 4.79 | 11838.1 | 7.17
-78.3 | 0 | 100 | 4.98 | 12323.6 | 7.44
-77.1 | 0 | 100 | 5.2 | 12846.7 | 7.76
-75.9 | 0 | 100 | 5.37 | 13271.2 | 8.01
-74.7 | 0 | 100 | 5.61 | 13861.1 | 8.37
-73.5 | 0 | 100 | 5.77 | 14258 | 8.61
-72.3 | 0 | 100 | 5.98 | 14786.5 | 8.94
-71.1 | 0 | 100 | 6.21 | 15361.2 | 9.32
-69.9 | 0 | 100 | 6.5 | 16072.4 | 9.72
-68.7 | 0 | 100 | 6.77 | 16736.7 | 10.11
-67.5 | 0 | 100 | 7.05 | 17433.5 | 10.53
-66.3 | 0 | 100 | 7.42 | 18345.2 | 11.07
-65.1 | 0 | 100 | 7.87 | 19466.7 | 11.77
-63.9 | 0 | 100 | 8.41 | 20780.1 | 12.56
-62.7 | 0 | 100 | 9.16 | 22643.6 | 13.7
-61.4 | 0 | 100 | 10.47 | 25882.5 | 15.64
-60.2 | 0 | 100 | 12.44 | 30748.1 | 18.6
-59 | 0 | 100 | 15.79 | 39049.5 | 23.61
-57.8 | 0 | 100 | 20.84 | 51523.7 | 31.15
-56.6 | 0 | 100 | 27.15 | 67119.4 | 40.56
-55.4 | 0 | 100 | 34.92 | 86344.2 | 52.22
-54.2 | 0 | 100 | 43.28 | 107001.9 | 64.62
-53 | 0 | 100 | 50.54 | 124948.7 | 75.43
-51.8 | 0 | 100 | 56.46 | 139585.1 | 84.41
-50.6 | 0 | 100 | 59.91 | 148109.5 | 89.48
-49.4 | 0 | 100 | 62.03 | 153355.6 | 92.83
-48.2 | 0 | 100 | 62.9 | 155519.9 | 94.19
-47 | 0 | 100 | 63.64 | 157342.8 | 95.08
-45.8 | 0 | 100 | 64.05 | 158364 | 95.64
-44.6 | 0 | 100 | 64.26 | 158879.4 | 96.04
-43.3 | 0 | 100 | 64.62 | 159769.7 | 96.7
-42.2 | 0 | 100 | 64.79 | 160176.9 | 96.84
-41 | 0 | 100 | 64.92 | 160509.4 | 97.15
-39.7 | 0 | 100 | 65.26 | 161357.8 | 97.45
-38.5 | 0 | 100 | 65.3 | 161450.2 | 97.56
-37.3 | 0 | 100 | 65.28 | 161405.3 | 97.62
-36.2 | 0 | 100 | 65.58 | 162136.7 | 98.12
-34.9 | 0 | 100 | 65.73 | 162506.5 | 98.23
-33.7 | 0 | 100 | 65.93 | 163004.1 | 98.62
-32.5 | 0 | 100 | 65.96 | 163075.6 | 98.57
-31.3 | 0 | 100 | 66.11 | 163443.1 | 98.83
-30.1 | 0 | 100 | 66.27 | 163830.8 | 99.12
-28.9 | 0 | 100 | 66.44 | 164272.3 | 99.29
-27.7 | 0 | 100 | 66.59 | 164629.1 | 99.56
-26.5 | 0 | 100 | 66.55 | 164536.9 | 99.46
-25.3 | 0 | 100 | 66.62 | 164697.3 | 99.71
-24.1 | 0 | 100 | 66.59 | 164637.6 | 99.8
-22.9 | 0 | 100 | 66.79 | 165124 | 99.89
-21.7 | 0 | 100 | 66.87 | 165336.7 | 99.93
-20.5 | 0 | 100 | 66.85 | 165286.7 | 99.92
-19.3 | 0 | 100 | 66.76 | 165042.3 | 99.7
-18.1 | 0 | 100 | 66.79 | 165134.2 | 99.79
-16.9 | 0 | 100 | 66.83 | 165235 | 99.95
-15.7 | 0 | 100 | 66.8 | 165151.3 | 99.8
-14.5 | 0 | 100 | 66.88 | 165343 | 100
-13.3 | 0 | 100 | 66.66 | 164817.9 | 99.77
-12.1 | 0 | 100 | 66.75 | 165039.8 | 99.93
-10.9 | 0 | 100 | 66.72 | 164953.1 | 99.81
-9.6 | 0 | 100 | 66.71 | 164932.1 | 99.83
-8.4 | 0 | 100 | 66.69 | 164872.4 | 99.84
-7.2 | 0 | 100 | 66.59 | 164641.3 | 99.64
-6 | 0 | 100 | 66.58 | 164618 | 99.73
-4.8 | 0 | 100 | 66.59 | 164645.2 | 99.53
-3.6 | 0 | 100 | 66.46 | 164313 | 99.37
-2.4 | 0 | 100 | 66.49 | 164398.8 | 99.57
-1.2 | 0 | 100 | 66.51 | 164439.3 | 99.53
0 | 0 | 100 | 66.44 | 164266.4 | 99.42
1.2 | 0 | 100 | 66.51 | 164434.5 | 99.35
2.4 | 0 | 100 | 66.37 | 164094.7 | 99.32
3.6 | 0 | 100 | 66.41 | 164193.4 | 99.44
4.8 | 0 | 100 | 66.44 | 164251.7 | 99.3
6 | 0 | 100 | 66.36 | 164075.4 | 99.29
7.2 | 0 | 100 | 66.47 | 164333.8 | 99.3
8.4 | 0 | 100 | 66.36 | 164063 | 99.23
9.6 | 0 | 100 | 66.35 | 164050.9 | 99.35
10.8 | 0 | 100 | 66.38 | 164120.9 | 99.32
12 | 0 | 100 | 66.38 | 164103.2 | 99.26
13.3 | 0 | 100 | 66.43 | 164247.5 | 99.46
14.5 | 0 | 100 | 66.32 | 163966.1 | 99.1
15.7 | 0 | 100 | 66.41 | 164177.9 | 99.32
16.9 | 0 | 100 | 66.33 | 163990.4 | 99.23
18.1 | 0 | 100 | 66.14 | 163532.2 | 98.92
19.3 | 0 | 100 | 66.13 | 163491.7 | 98.97
20.5 | 0 | 100 | 66.15 | 163538.7 | 98.93
21.7 | 0 | 100 | 66.11 | 163451.1 | 98.94
22.9 | 0 | 100 | 65.93 | 163010.8 | 98.59
24.1 | 0 | 100 | 65.92 | 162981.2 | 98.59
25.3 | 0 | 100 | 65.92 | 162975.5 | 98.36
26.5 | 0 | 100 | 65.75 | 162546.5 | 98.39
27.7 | 0 | 100 | 65.63 | 162261 | 98.19
28.9 | 0 | 100 | 65.49 | 161911.7 | 97.98
30.1 | 0 | 100 | 65.5 | 161951.4 | 98.01
31.3 | 0 | 100 | 65.29 | 161413.8 | 97.57
32.5 | 0 | 100 | 65.05 | 160824.4 | 97.26
33.7 | 0 | 100 | 64.91 | 160490.2 | 97.22
34.9 | 0 | 100 | 64.75 | 160094 | 96.82
36.1 | 0 | 100 | 64.6 | 159718.1 | 96.53
37.3 | 0 | 100 | 64.4 | 159215.4 | 96.41
38.5 | 0 | 100 | 64.23 | 158801.6 | 95.84
39.7 | 0 | 100 | 64.11 | 158514.6 | 96.01
40.9 | 0 | 100 | 63.93 | 158062.2 | 95.63
42.1 | 0 | 100 | 63.56 | 157145.9 | 95.24
43.3 | 0 | 100 | 63.36 | 156642.6 | 94.73
44.6 | 0 | 100 | 63.04 | 155847 | 94.37
45.7 | 0 | 100 | 62.68 | 154963.5 | 93.83
47 | 0 | 100 | 62.01 | 153320.2 | 92.81
48.2 | 0 | 100 | 60.96 | 150721.1 | 91.31
49.4 | 0 | 100 | 59.32 | 146669.5 | 88.71
50.6 | 0 | 100 | 56.03 | 138526.1 | 83.73
51.8 | 0 | 100 | 51.47 | 127262.7 | 76.91
53 | 0 | 100 | 45.48 | 112454 | 68.1
54.2 | 0 | 100 | 38.06 | 94107.3 | 57.01
55.4 | 0 | 100 | 30.09 | 74387.9 | 45.06
56.6 | 0 | 100 | 22.52 | 55665.3 | 33.67
57.8 | 0 | 100 | 16.42 | 40607.7 | 24.58
59 | 0 | 100 | 12.56 | 31063.6 | 18.77
60.2 | 0 | 100 | 10.43 | 25796.1 | 15.6
61.4 | 0 | 100 | 9.19 | 22710.9 | 13.74
62.6 | 0 | 100 | 8.37 | 20683.5 | 12.52
63.8 | 0 | 100 | 7.84 | 19373.8 | 11.73
65 | 0 | 100 | 7.41 | 18308.6 | 11.08
66.2 | 0 | 100 | 7.01 | 17342.6 | 10.5
67.4 | 0 | 100 | 6.69 | 16546.6 | 10.02
68.6 | 0 | 100 | 6.45 | 15940.7 | 9.66
69.8 | 0 | 100 | 6.14 | 15182.1 | 9.2
71 | 0 | 100 | 5.88 | 14538.6 | 8.81
72.2 | 0 | 100 | 5.66 | 14003.1 | 8.47
73.5 | 0 | 100 | 5.46 | 13509.7 | 8.18
74.7 | 0 | 100 | 5.22 | 12908.8 | 7.83
75.9 | 0 | 100 | 5.04 | 12456.3 | 7.53
77.1 | 0 | 100 | 4.89 | 12081 | 7.31
78.3 | 0 | 100 | 4.69 | 11607.5 | 7.02
79.5 | 0 | 100 | 4.51 | 11158.7 | 6.75
80.7 | 0 | 100 | 4.42 | 10933.8 | 6.61
81.9 | 0 | 100 | 4.25 | 10509.4 | 6.36
83.1 | 0 | 100 | 4.12 | 10181.3 | 6.16
84.3 | 0 | 100 | 3.98 | 9842.2 | 5.96
85.5 | 0 | 100 | 3.83 | 9478.5 | 5.75
86.7 | 0 | 100 | 3.74 | 9254.4 | 5.6
87.9 | 0 | 100 | 3.59 | 8874.5 | 5.37
89.1 | 0 | 100 | 3.49 | 8628.7 | 5.22
90.3 | 0 | 100 | 3.4 | 8402.8 | 5.09
91.5 | 0 | 100 | 3.23 | 7996.8 | 4.85
92.7 | 0 | 100 | 3.2 | 7901.6 | 4.78
93.9 | 0 | 100 | 3.1 | 7657.5 | 4.64
95.1 | 0 | 100 | 2.98 | 7376.5 | 4.46
96.3 | 0 | 100 | 2.91 | 7197.5 | 4.35
97.6 | 0 | 100 | 2.81 | 6936.2 | 4.2
98.8 | 0 | 100 | 2.71 | 6694.5 | 4.06
100 | 0 | 100 | 2.66 | 6577.3 | 3.98
101.2 | 0 | 100 | 2.54 | 6283.3 | 3.81
102.4 | 0 | 100 | 2.48 | 6122.8 | 3.71
103.6 | 0 | 100 | 2.4 | 5941.2 | 3.61
104.8 | 0 | 100 | 2.36 | 5836.4 | 3.53
106 | 0 | 100 | 2.32 | 5724 | 3.46
107.2 | 0 | 100 | 2.29 | 5652.2 | 3.42
108.4 | 0 | 100 | 2.15 | 5311 | 3.22
109.6 | 0 | 100 | 2.09 | 5174.3 | 3.14
110.8 | 0 | 100 | 2.05 | 5079.9 | 3.07
112 | 0 | 100 | 1.99 | 4930.3 | 2.98
113.2 | 0 | 100 | 1.95 | 4827.9 | 2.92
114.4 | 0 | 100 | 1.9 | 4701.3 | 2.85
115.6 | 0 | 100 | 1.87 | 4617 | 2.8
116.8 | 0 | 100 | 1.76 | 4356.3 | 2.64
118 | 0 | 100 | 1.73 | 4278.1 | 2.59
119.2 | 0 | 100 | 1.77 | 4376 | 2.65
120.4 | 0 | 100 | 1.68 | 4164.9 | 2.52
121.6 | 0 | 100 | 1.61 | 3979 | 2.41
122.8 | 0 | 100 | 1.53 | 3784.6 | 2.29
124 | 0 | 100 | 1.53 | 3790.5 | 2.3
125.2 | 0 | 100 | 1.48 | 3649.2 | 2.21
126.4 | 0 | 100 | 1.47 | 3626.4 | 2.2
127.6 | 0 | 100 | 1.42 | 3517.2 | 2.13
128.8 | 0 | 100 | 1.41 | 3482.1 | 2.11
130 | 0 | 100 | 1.34 | 3306.2 | 2
131.2 | 0 | 100 | 1.36 | 3355.4 | 2.03
132.4 | 0 | 100 | 1.27 | 3141 | 1.9
133.6 | 0 | 100 | 1.22 | 3021.6 | 1.83
134.8 | 0 | 100 | 1.25 | 3082.8 | 1.87
136 | 0 | 100 | 1.23 | 3036.3 | 1.84
137.2 | 0 | 100 | 1.17 | 2898 | 1.75
138.4 | 0 | 100 | 1.15 | 2846.3 | 1.72
139.6 | 0 | 100 | 1.13 | 2782.6 | 1.68
140.8 | 0 | 100 | 1.07 | 2647.4 | 1.6
142.1 | 0 | 100 | 1.05 | 2586.3 | 1.57
143.3 | 0 | 100 | 1.03 | 2536 | 1.53
144.5 | 0 | 100 | 1.04 | 2574 | 1.56
145.7 | 0 | 100 | 1.03 | 2544.7 | 1.54
146.9 | 0 | 100 | 1.01 | 2499.1 | 1.51
148.1 | 0 | 100 | 0.98 | 2433.7 | 1.47
149.3 | 0 | 100 | 0.96 | 2376 | 1.44
150.5 | 0 | 100 | 0.94 | 2329.6 | 1.41
151.7 | 0 | 100 | 0.93 | 2290.8 | 1.39
152.9 | 0 | 100 | 0.87 | 2151.2 | 1.3
155 | 0 | 100 | 0.85 | 2101.5 | 1.27
Measurement time: | 12/08/2022 18:38:22
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 110 x 110 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
153.8 | 0 | 100 | 1.08 | 2680.2 | 1.33
151.9 | 0 | 100 | 1.11 | 2740.7 | 1.36
150.8 | 0 | 100 | 1.15 | 2852.2 | 1.42
149.5 | 0 | 100 | 1.17 | 2880.8 | 1.43
148.3 | 0 | 100 | 1.19 | 2939.6 | 1.46
147.1 | 0 | 100 | 1.21 | 2983.1 | 1.48
145.9 | 0 | 100 | 1.22 | 3011.8 | 1.49
144.7 | 0 | 100 | 1.26 | 3118.3 | 1.55
143.6 | 0 | 100 | 1.32 | 3264.2 | 1.62
142.3 | 0 | 100 | 1.29 | 3191.3 | 1.58
141.1 | 0 | 100 | 1.33 | 3291.5 | 1.63
139.9 | 0 | 100 | 1.38 | 3416.6 | 1.69
138.7 | 0 | 100 | 1.4 | 3462.7 | 1.72
137.5 | 0 | 100 | 1.42 | 3512.6 | 1.74
136.3 | 0 | 100 | 1.49 | 3682 | 1.83
135.1 | 0 | 100 | 1.51 | 3741 | 1.85
133.9 | 0 | 100 | 1.55 | 3833.1 | 1.9
132.7 | 0 | 100 | 1.62 | 3995.6 | 1.98
131.5 | 0 | 100 | 1.65 | 4069.6 | 2.02
130.3 | 0 | 100 | 1.71 | 4229.5 | 2.1
129.1 | 0 | 100 | 1.67 | 4136.6 | 2.05
127.9 | 0 | 100 | 1.74 | 4305.5 | 2.13
126.7 | 0 | 100 | 1.8 | 4439.7 | 2.2
125.5 | 0 | 100 | 1.83 | 4532.3 | 2.25
124.3 | 0 | 100 | 1.85 | 4564 | 2.26
123.1 | 0 | 100 | 1.98 | 4883 | 2.42
121.9 | 0 | 100 | 1.99 | 4910.1 | 2.44
120.6 | 0 | 100 | 1.98 | 4899.9 | 2.43
119.5 | 0 | 100 | 2.06 | 5099.8 | 2.53
118.3 | 0 | 100 | 2.14 | 5284.5 | 2.63
117.1 | 0 | 100 | 2.18 | 5397.6 | 2.68
115.9 | 0 | 100 | 2.25 | 5556.5 | 2.76
114.6 | 0 | 100 | 2.29 | 5670.5 | 2.82
113.5 | 0 | 100 | 2.36 | 5837 | 2.9
112.3 | 0 | 100 | 2.46 | 6084.6 | 3.02
111 | 0 | 100 | 2.53 | 6262 | 3.11
109.9 | 0 | 100 | 2.56 | 6337 | 3.14
108.6 | 0 | 100 | 2.62 | 6488.6 | 3.22
107.5 | 0 | 100 | 2.72 | 6736.1 | 3.34
106.2 | 0 | 100 | 2.82 | 6968.3 | 3.46
105 | 0 | 100 | 2.88 | 7116.3 | 3.53
103.8 | 0 | 100 | 2.96 | 7320.7 | 3.63
102.6 | 0 | 100 | 3.08 | 7607.7 | 3.78
101.4 | 0 | 100 | 3.16 | 7806.2 | 3.87
100.2 | 0 | 100 | 3.22 | 7956.5 | 3.95
99 | 0 | 100 | 3.34 | 8245.7 | 4.09
97.8 | 0 | 100 | 3.42 | 8454.3 | 4.19
96.6 | 0 | 100 | 3.58 | 8849.1 | 4.39
95.4 | 0 | 100 | 3.7 | 9140.8 | 4.54
94.2 | 0 | 100 | 3.75 | 9259.4 | 4.6
93 | 0 | 100 | 3.94 | 9741.1 | 4.83
91.8 | 0 | 100 | 4.03 | 9973.7 | 4.96
90.5 | 0 | 100 | 4.12 | 10186.7 | 5.05
89.4 | 0 | 100 | 4.3 | 10643.3 | 5.28
88.1 | 0 | 100 | 4.46 | 11019.4 | 5.47
86.9 | 0 | 100 | 4.58 | 11333.4 | 5.63
85.8 | 0 | 100 | 4.76 | 11772.8 | 5.84
84.5 | 0 | 100 | 4.88 | 12061.2 | 5.98
83.4 | 0 | 100 | 5.06 | 12513.9 | 6.22
82.1 | 0 | 100 | 5.24 | 12947.6 | 6.44
80.9 | 0 | 100 | 5.45 | 13473.5 | 6.69
79.7 | 0 | 100 | 5.64 | 13949.4 | 6.92
78.5 | 0 | 100 | 5.86 | 14495.6 | 7.19
77.3 | 0 | 100 | 6.04 | 14936.9 | 7.42
76.1 | 0 | 100 | 6.26 | 15486.3 | 7.68
74.9 | 0 | 100 | 6.5 | 16059.9 | 7.97
73.7 | 0 | 100 | 6.76 | 16709.6 | 8.3
72.5 | 0 | 100 | 7.04 | 17398.3 | 8.63
71.4 | 0 | 100 | 7.33 | 18117.5 | 8.99
70.1 | 0 | 100 | 7.6 | 18781.8 | 9.32
68.9 | 0 | 100 | 7.92 | 19575.6 | 9.71
67.7 | 0 | 100 | 8.29 | 20501.1 | 10.17
66.5 | 0 | 100 | 8.66 | 21420.2 | 10.63
65.3 | 0 | 100 | 9.12 | 22554.3 | 11.2
64.1 | 0 | 100 | 9.67 | 23907.1 | 11.88
62.9 | 0 | 100 | 10.32 | 25514.6 | 12.67
61.7 | 0 | 100 | 11.29 | 27915.3 | 13.85
60.5 | 0 | 100 | 12.88 | 31851.2 | 15.8
59.3 | 0 | 100 | 15.26 | 37739.8 | 18.75
58.1 | 0 | 100 | 20.01 | 49465.3 | 24.54
56.9 | 0 | 100 | 27.5 | 67987.3 | 33.77
55.7 | 0 | 100 | 36.39 | 89980.4 | 44.72
54.4 | 0 | 100 | 46.81 | 115719.8 | 57.48
53.3 | 0 | 100 | 55.68 | 137664.9 | 68.37
52 | 0 | 100 | 63.2 | 156258.7 | 77.59
50.8 | 0 | 100 | 68.68 | 169789.1 | 84.44
49.7 | 0 | 100 | 72.13 | 178334.7 | 88.75
48.4 | 0 | 100 | 74.52 | 184246 | 91.64
47.3 | 0 | 100 | 75.68 | 187111.8 | 92.96
46 | 0 | 100 | 76.8 | 189874.3 | 94.14
44.8 | 0 | 100 | 77.26 | 191015.8 | 94.76
43.6 | 0 | 100 | 77.61 | 191872.4 | 95.19
42.4 | 0 | 100 | 78.04 | 192944.7 | 95.72
41.2 | 0 | 100 | 78.24 | 193425.6 | 95.86
40 | 0 | 100 | 78.3 | 193593.6 | 96.08
38.8 | 0 | 100 | 78.4 | 193824.8 | 96.44
37.6 | 0 | 100 | 78.85 | 194946.9 | 96.78
36.3 | 0 | 100 | 78.99 | 195282.7 | 97.01
35.2 | 0 | 100 | 79.12 | 195606.3 | 97.15
33.9 | 0 | 100 | 79.23 | 195872.2 | 97.37
32.8 | 0 | 100 | 79.46 | 196460 | 97.51
31.6 | 0 | 100 | 79.68 | 196996.9 | 97.87
30.3 | 0 | 100 | 79.72 | 197100.7 | 98.04
29.2 | 0 | 100 | 79.93 | 197602.9 | 98.19
27.9 | 0 | 100 | 80.06 | 197930.9 | 98.33
26.7 | 0 | 100 | 80.24 | 198390.3 | 98.56
25.5 | 0 | 100 | 80.32 | 198570.1 | 98.6
24.3 | 0 | 100 | 80.36 | 198666.9 | 98.69
23.1 | 0 | 100 | 80.66 | 199424 | 98.96
21.9 | 0 | 100 | 80.72 | 199565.2 | 99.2
20.7 | 0 | 100 | 80.89 | 200000.3 | 99.29
19.5 | 0 | 100 | 80.84 | 199872.3 | 99.25
18.3 | 0 | 100 | 80.86 | 199905.7 | 99.32
17.1 | 0 | 100 | 80.99 | 200238.6 | 99.47
15.9 | 0 | 100 | 80.94 | 200110.6 | 99.33
14.7 | 0 | 100 | 80.93 | 200079.2 | 99.45
13.5 | 0 | 100 | 81.03 | 200334.1 | 99.44
12.3 | 0 | 100 | 81.14 | 200595.3 | 99.58
11.1 | 0 | 100 | 80.79 | 199749.7 | 99.36
9.9 | 0 | 100 | 80.9 | 200023.9 | 99.41
8.7 | 0 | 100 | 80.9 | 200008 | 99.26
7.5 | 0 | 100 | 81.06 | 200410.3 | 99.41
6.3 | 0 | 100 | 80.81 | 199793.4 | 99.29
5.1 | 0 | 100 | 80.9 | 200023.7 | 99.45
3.9 | 0 | 100 | 80.79 | 199734.5 | 99.34
2.7 | 0 | 100 | 80.79 | 199730.1 | 99.22
1.5 | 0 | 100 | 81.06 | 200412.1 | 99.47
0.3 | 0 | 100 | 81.06 | 200400.3 | 99.39
-0.9 | 0 | 100 | 80.97 | 200193.1 | 99.5
-2.2 | 0 | 100 | 80.98 | 200202.6 | 99.57
-3.3 | 0 | 100 | 81.01 | 200296 | 99.7
-4.5 | 0 | 100 | 80.95 | 200127.8 | 99.45
-5.8 | 0 | 100 | 81.1 | 200505.6 | 99.64
-6.9 | 0 | 100 | 81.31 | 201039 | 99.76
-8.2 | 0 | 100 | 81.18 | 200704.6 | 99.72
-9.3 | 0 | 100 | 81.22 | 200801.5 | 99.75
-10.6 | 0 | 100 | 81.34 | 201109.8 | 99.85
-11.8 | 0 | 100 | 81.46 | 201405 | 99.92
-13 | 0 | 100 | 81.34 | 201109.5 | 99.93
-14.2 | 0 | 100 | 81.5 | 201484.4 | 100
-15.4 | 0 | 100 | 81.36 | 201160.6 | 99.87
-16.6 | 0 | 100 | 81.47 | 201422.5 | 99.95
-17.8 | 0 | 100 | 81.34 | 201090.8 | 99.92
-19 | 0 | 100 | 81.3 | 201005.8 | 99.98
-20.2 | 0 | 100 | 81.42 | 201288.7 | 99.9
-21.4 | 0 | 100 | 81.38 | 201200.9 | 99.9
-22.6 | 0 | 100 | 81.31 | 201028.1 | 99.82
-23.8 | 0 | 100 | 81.2 | 200747.7 | 99.8
-25 | 0 | 100 | 81.17 | 200690.1 | 99.65
-26.2 | 0 | 100 | 81.12 | 200551.1 | 99.69
-27.4 | 0 | 100 | 81.07 | 200426.7 | 99.64
-28.6 | 0 | 100 | 81.02 | 200299 | 99.55
-29.8 | 0 | 100 | 80.7 | 199514.1 | 99.18
-31 | 0 | 100 | 80.69 | 199485.3 | 99.02
-32.2 | 0 | 100 | 80.57 | 199197.3 | 98.92
-33.4 | 0 | 100 | 80.31 | 198563.2 | 98.48
-34.7 | 0 | 100 | 80.42 | 198834.7 | 98.75
-35.9 | 0 | 100 | 80.3 | 198538.2 | 98.54
-37 | 0 | 100 | 79.98 | 197726.7 | 98.23
-38.3 | 0 | 100 | 79.82 | 197333.8 | 98.08
-39.4 | 0 | 100 | 79.68 | 197004.9 | 97.84
-40.7 | 0 | 100 | 79.55 | 196674.5 | 97.75
-41.9 | 0 | 100 | 79.23 | 195883.8 | 97.27
-43 | 0 | 100 | 79.07 | 195484.1 | 97.06
-44.3 | 0 | 100 | 78.72 | 194624.7 | 96.84
-45.5 | 0 | 100 | 78.34 | 193683.7 | 96.15
-46.7 | 0 | 100 | 77.86 | 192504.3 | 95.55
-47.9 | 0 | 100 | 77.11 | 190650.2 | 94.58
-49.1 | 0 | 100 | 76.04 | 187994.8 | 93.39
-50.3 | 0 | 100 | 74.24 | 183541.2 | 91.06
-51.5 | 0 | 100 | 70.42 | 174100.4 | 86.52
-52.7 | 0 | 100 | 64.3 | 158962.5 | 78.85
-53.9 | 0 | 100 | 55.19 | 136458.5 | 67.76
-55.1 | 0 | 100 | 45.43 | 112312.7 | 55.8
-56.3 | 0 | 100 | 35.34 | 87365.3 | 43.44
-57.5 | 0 | 100 | 27 | 66764.7 | 33.15
-58.7 | 0 | 100 | 20.65 | 51052.5 | 25.36
-59.9 | 0 | 100 | 16 | 39568.7 | 19.66
-61.1 | 0 | 100 | 13.24 | 32729.2 | 16.26
-62.4 | 0 | 100 | 11.61 | 28702 | 14.25
-63.5 | 0 | 100 | 10.58 | 26148.2 | 13
-64.8 | 0 | 100 | 9.81 | 24252.4 | 12.05
-66 | 0 | 100 | 9.26 | 22884.3 | 11.36
-67.1 | 0 | 100 | 8.82 | 21816.9 | 10.85
-68.4 | 0 | 100 | 8.39 | 20737.8 | 10.31
-69.6 | 0 | 100 | 8.05 | 19891.9 | 9.87
-70.8 | 0 | 100 | 7.72 | 19092.2 | 9.48
-72 | 0 | 100 | 7.43 | 18375.1 | 9.14
-73.2 | 0 | 100 | 7.15 | 17689.4 | 8.8
-74.4 | 0 | 100 | 6.88 | 17012.1 | 8.44
-75.5 | 0 | 100 | 6.6 | 16317 | 8.11
-76.8 | 0 | 100 | 6.4 | 15813.6 | 7.85
-78 | 0 | 100 | 6.13 | 15162.2 | 7.52
-79.2 | 0 | 100 | 5.92 | 14638.2 | 7.28
-80.4 | 0 | 100 | 5.77 | 14267.7 | 7.07
-81.6 | 0 | 100 | 5.47 | 13534.6 | 6.72
-82.8 | 0 | 100 | 5.36 | 13258.3 | 6.59
-84 | 0 | 100 | 5.13 | 12684.2 | 6.31
-85.2 | 0 | 100 | 4.93 | 12196.9 | 6.04
-86.4 | 0 | 100 | 4.78 | 11823.8 | 5.88
-87.6 | 0 | 100 | 4.64 | 11470.6 | 5.71
-88.8 | 0 | 100 | 4.5 | 11130.2 | 5.53
-90 | 0 | 100 | 4.37 | 10815.4 | 5.37
-91.2 | 0 | 100 | 4.21 | 10407.6 | 5.17
-92.5 | 0 | 100 | 4.03 | 9968.6 | 4.95
-93.6 | 0 | 100 | 3.9 | 9644.5 | 4.78
-94.8 | 0 | 100 | 3.76 | 9288 | 4.6
-96 | 0 | 100 | 3.66 | 9047.7 | 4.5
-97.2 | 0 | 100 | 3.53 | 8717.3 | 4.33
-98.5 | 0 | 100 | 3.43 | 8486.3 | 4.22
-99.6 | 0 | 100 | 3.35 | 8290.4 | 4.11
-100.9 | 0 | 100 | 3.25 | 8034.5 | 3.99
-102.1 | 0 | 100 | 3.16 | 7818 | 3.89
-103.3 | 0 | 100 | 3.05 | 7545.5 | 3.75
-104.5 | 0 | 100 | 2.93 | 7242.8 | 3.6
-105.6 | 0 | 100 | 2.84 | 7013.5 | 3.49
-106.9 | 0 | 100 | 2.8 | 6931.3 | 3.44
-108.1 | 0 | 100 | 2.69 | 6653.8 | 3.31
-109.3 | 0 | 100 | 2.64 | 6534.8 | 3.25
-110.5 | 0 | 100 | 2.52 | 6240.4 | 3.1
-111.7 | 0 | 100 | 2.47 | 6107 | 3.03
-112.9 | 0 | 100 | 2.43 | 6019.9 | 2.99
-114.1 | 0 | 100 | 2.36 | 5828.7 | 2.89
-115.3 | 0 | 100 | 2.26 | 5586.5 | 2.77
-116.5 | 0 | 100 | 2.23 | 5505.7 | 2.73
-117.7 | 0 | 100 | 2.15 | 5305.6 | 2.63
-118.9 | 0 | 100 | 2.1 | 5184 | 2.58
-120.1 | 0 | 100 | 1.99 | 4928.5 | 2.44
-121.3 | 0 | 100 | 1.93 | 4767.7 | 2.36
-122.5 | 0 | 100 | 1.88 | 4642.6 | 2.3
-123.7 | 0 | 100 | 1.87 | 4635.4 | 2.3
-124.9 | 0 | 100 | 1.78 | 4396.4 | 2.18
-126.1 | 0 | 100 | 1.77 | 4385.8 | 2.18
-127.3 | 0 | 100 | 1.76 | 4344.9 | 2.16
-128.5 | 0 | 100 | 1.69 | 4167.7 | 2.07
-129.7 | 0 | 100 | 1.69 | 4175.5 | 2.07
-130.9 | 0 | 100 | 1.63 | 4039.6 | 2
-132.1 | 0 | 100 | 1.55 | 3823.1 | 1.9
-133.3 | 0 | 100 | 1.49 | 3678.6 | 1.83
-134.6 | 0 | 100 | 1.42 | 3521.6 | 1.75
-135.7 | 0 | 100 | 1.45 | 3572.8 | 1.77
-136.9 | 0 | 100 | 1.39 | 3440.9 | 1.71
-138.2 | 0 | 100 | 1.41 | 3485.2 | 1.73
-139.3 | 0 | 100 | 1.33 | 3297.9 | 1.64
-140.6 | 0 | 100 | 1.33 | 3278.6 | 1.63
-141.7 | 0 | 100 | 1.27 | 3142.8 | 1.56
-143 | 0 | 100 | 1.23 | 3029.1 | 1.51
-144.2 | 0 | 100 | 1.19 | 2951.6 | 1.47
-145.4 | 0 | 100 | 1.18 | 2907.2 | 1.44
-146.6 | 0 | 100 | 1.19 | 2931.2 | 1.46
-147.8 | 0 | 100 | 1.12 | 2767.4 | 1.38
-149 | 0 | 100 | 1.1 | 2725.5 | 1.35
-150.2 | 0 | 100 | 1.11 | 2745.7 | 1.36
-151.4 | 0 | 100 | 1.04 | 2570.8 | 1.28
-152.6 | 0 | 100 | 1.03 | 2545.9 | 1.26
-153.8 | 0 | 100 | 1.04 | 2564.1 | 1.28
-155 | 0 | 100 | 0.99 | 2453.6 | 1.22
Measurement time: | 12/08/2022 18:41:46
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1100 mm
Field size: | 90 x 90 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-153 | 0 | 100 | 0.56 | 1396.4 | 1.01
-151.5 | 0 | 100 | 0.59 | 1467.3 | 1.07
-150.5 | 0 | 100 | 0.62 | 1538.1 | 1.12
-149.2 | 0 | 100 | 0.62 | 1539.8 | 1.12
-148 | 0 | 100 | 0.62 | 1526.1 | 1.11
-146.8 | 0 | 100 | 0.61 | 1511.6 | 1.1
-145.6 | 0 | 100 | 0.66 | 1634.8 | 1.19
-144.4 | 0 | 100 | 0.71 | 1748 | 1.27
-143.2 | 0 | 100 | 0.72 | 1781.9 | 1.29
-142 | 0 | 100 | 0.72 | 1771.7 | 1.28
-140.8 | 0 | 100 | 0.74 | 1825.8 | 1.33
-139.6 | 0 | 100 | 0.73 | 1808 | 1.31
-138.4 | 0 | 100 | 0.77 | 1905.1 | 1.38
-137.1 | 0 | 100 | 0.82 | 2020.2 | 1.46
-135.9 | 0 | 100 | 0.84 | 2072.6 | 1.5
-134.8 | 0 | 100 | 0.87 | 2163.1 | 1.57
-133.5 | 0 | 100 | 0.88 | 2176 | 1.57
-132.3 | 0 | 100 | 0.88 | 2181.9 | 1.58
-131.1 | 0 | 100 | 0.88 | 2168.9 | 1.57
-129.9 | 0 | 100 | 0.89 | 2206.3 | 1.6
-128.8 | 0 | 100 | 0.99 | 2444.1 | 1.77
-127.5 | 0 | 100 | 1.01 | 2487.1 | 1.81
-126.3 | 0 | 100 | 1.02 | 2516.9 | 1.83
-125.1 | 0 | 100 | 1.08 | 2661.9 | 1.93
-123.9 | 0 | 100 | 1.16 | 2862.3 | 2.08
-122.7 | 0 | 100 | 1.16 | 2859.2 | 2.08
-121.5 | 0 | 100 | 1.18 | 2917.6 | 2.12
-120.3 | 0 | 100 | 1.21 | 2982 | 2.16
-119.1 | 0 | 100 | 1.28 | 3174 | 2.31
-117.9 | 0 | 100 | 1.3 | 3225.2 | 2.34
-116.7 | 0 | 100 | 1.33 | 3280.2 | 2.38
-115.5 | 0 | 100 | 1.37 | 3395.3 | 2.46
-114.3 | 0 | 100 | 1.43 | 3535.4 | 2.57
-113.1 | 0 | 100 | 1.49 | 3683.4 | 2.68
-111.9 | 0 | 100 | 1.47 | 3629.8 | 2.64
-110.7 | 0 | 100 | 1.54 | 3812.9 | 2.77
-109.5 | 0 | 100 | 1.65 | 4078.4 | 2.96
-108.3 | 0 | 100 | 1.64 | 4054.8 | 2.94
-107.1 | 0 | 100 | 1.69 | 4175.3 | 3.04
-105.8 | 0 | 100 | 1.77 | 4376.6 | 3.18
-104.7 | 0 | 100 | 1.82 | 4507.4 | 3.28
-103.5 | 0 | 100 | 1.83 | 4514.1 | 3.29
-102.2 | 0 | 100 | 1.92 | 4735.2 | 3.44
-101 | 0 | 100 | 2.01 | 4976.2 | 3.62
-99.8 | 0 | 100 | 2.05 | 5070.4 | 3.69
-98.6 | 0 | 100 | 2.16 | 5337.1 | 3.88
-97.4 | 0 | 100 | 2.24 | 5527.9 | 4.02
-96.2 | 0 | 100 | 2.29 | 5656.2 | 4.1
-95 | 0 | 100 | 2.35 | 5797.9 | 4.21
-93.8 | 0 | 100 | 2.43 | 6002.2 | 4.36
-92.6 | 0 | 100 | 2.55 | 6295.1 | 4.58
-91.4 | 0 | 100 | 2.59 | 6410.6 | 4.65
-90.2 | 0 | 100 | 2.71 | 6707.1 | 4.88
-89 | 0 | 100 | 2.73 | 6755.9 | 4.91
-87.8 | 0 | 100 | 2.9 | 7181.7 | 5.22
-86.6 | 0 | 100 | 3.01 | 7434.9 | 5.4
-85.4 | 0 | 100 | 3.07 | 7580 | 5.52
-84.2 | 0 | 100 | 3.2 | 7908.2 | 5.75
-83 | 0 | 100 | 3.33 | 8242.8 | 5.98
-81.8 | 0 | 100 | 3.47 | 8591.3 | 6.24
-80.6 | 0 | 100 | 3.57 | 8830.5 | 6.43
-79.4 | 0 | 100 | 3.72 | 9207.9 | 6.7
-78.2 | 0 | 100 | 3.84 | 9486.6 | 6.89
-77 | 0 | 100 | 3.95 | 9778 | 7.11
-75.8 | 0 | 100 | 4.13 | 10211 | 7.42
-74.6 | 0 | 100 | 4.29 | 10607.1 | 7.7
-73.3 | 0 | 100 | 4.43 | 10956.5 | 7.96
-72.2 | 0 | 100 | 4.63 | 11448.8 | 8.32
-71 | 0 | 100 | 4.76 | 11778.2 | 8.56
-69.7 | 0 | 100 | 5 | 12372.9 | 9
-68.5 | 0 | 100 | 5.19 | 12829.2 | 9.35
-67.3 | 0 | 100 | 5.38 | 13300.5 | 9.68
-66.2 | 0 | 100 | 5.68 | 14047.1 | 10.22
-64.9 | 0 | 100 | 5.95 | 14707.6 | 10.7
-63.7 | 0 | 100 | 6.36 | 15723.2 | 11.45
-62.5 | 0 | 100 | 6.84 | 16914.6 | 12.29
-61.3 | 0 | 100 | 7.55 | 18671.2 | 13.57
-60.1 | 0 | 100 | 8.74 | 21599.6 | 15.72
-58.9 | 0 | 100 | 10.73 | 26534.2 | 19.28
-57.7 | 0 | 100 | 13.91 | 34379.3 | 24.99
-56.5 | 0 | 100 | 18.3 | 45232 | 32.87
-55.3 | 0 | 100 | 24.07 | 59499.1 | 43.34
-54.1 | 0 | 100 | 30.54 | 75515 | 55.01
-52.9 | 0 | 100 | 37.08 | 91672.6 | 66.72
-51.7 | 0 | 100 | 43.01 | 106328.8 | 77.36
-50.5 | 0 | 100 | 47.33 | 117024.3 | 85.17
-49.3 | 0 | 100 | 49.93 | 123436.9 | 89.68
-48.1 | 0 | 100 | 51.42 | 127131 | 92.53
-46.9 | 0 | 100 | 52.13 | 128883.6 | 93.69
-45.7 | 0 | 100 | 52.63 | 130114.1 | 94.72
-44.5 | 0 | 100 | 52.99 | 131005.5 | 95.46
-43.3 | 0 | 100 | 53.22 | 131588 | 95.82
-42.1 | 0 | 100 | 53.45 | 132150.1 | 96.24
-40.9 | 0 | 100 | 53.69 | 132743.1 | 96.65
-39.7 | 0 | 100 | 53.79 | 132996 | 96.8
-38.5 | 0 | 100 | 54.03 | 133584 | 97.21
-37.2 | 0 | 100 | 54.14 | 133858.3 | 97.52
-36.1 | 0 | 100 | 54.43 | 134560 | 97.95
-34.8 | 0 | 100 | 54.35 | 134363.3 | 97.91
-33.6 | 0 | 100 | 54.51 | 134772.8 | 98.08
-32.4 | 0 | 100 | 54.79 | 135463.9 | 98.5
-31.2 | 0 | 100 | 54.78 | 135441.8 | 98.63
-30 | 0 | 100 | 54.94 | 135827.8 | 98.66
-28.8 | 0 | 100 | 55.04 | 136066.7 | 99.1
-27.6 | 0 | 100 | 55 | 135983.9 | 98.84
-26.4 | 0 | 100 | 55.19 | 136449 | 99.36
-25.2 | 0 | 100 | 55.2 | 136471.2 | 99.6
-24 | 0 | 100 | 55.23 | 136558.6 | 99.3
-22.8 | 0 | 100 | 55.29 | 136706.8 | 99.42
-21.6 | 0 | 100 | 55.49 | 137195 | 99.7
-20.4 | 0 | 100 | 55.36 | 136873.7 | 99.65
-19.2 | 0 | 100 | 55.39 | 136952.4 | 99.74
-18 | 0 | 100 | 55.55 | 137341.2 | 100
-16.8 | 0 | 100 | 55.41 | 137000.6 | 99.77
-15.6 | 0 | 100 | 55.54 | 137316.3 | 99.96
-14.4 | 0 | 100 | 55.33 | 136806 | 99.44
-13.2 | 0 | 100 | 55.52 | 137269 | 99.84
-12 | 0 | 100 | 55.3 | 136732 | 99.63
-10.8 | 0 | 100 | 55.21 | 136504.6 | 99.45
-9.6 | 0 | 100 | 55.29 | 136706.3 | 99.38
-8.4 | 0 | 100 | 55.13 | 136288.8 | 99.34
-7.1 | 0 | 100 | 55.25 | 136609 | 99.51
-5.9 | 0 | 100 | 55.27 | 136653.4 | 99.58
-4.7 | 0 | 100 | 55.15 | 136352.3 | 99.27
-3.5 | 0 | 100 | 55.09 | 136196.1 | 99.07
-2.3 | 0 | 100 | 55.13 | 136295.5 | 99.21
-1.1 | 0 | 100 | 55.03 | 136044.9 | 99.09
0.1 | 0 | 100 | 55.17 | 136391.8 | 99.3
1.3 | 0 | 100 | 55.06 | 136139.2 | 99.2
2.5 | 0 | 100 | 55.06 | 136139.6 | 99.03
3.7 | 0 | 100 | 54.96 | 135877 | 98.95
4.9 | 0 | 100 | 55.17 | 136400.2 | 99.32
6.1 | 0 | 100 | 55.05 | 136114.9 | 99.14
7.3 | 0 | 100 | 55.05 | 136090.7 | 99.07
8.5 | 0 | 100 | 55.05 | 136095.1 | 99.23
9.7 | 0 | 100 | 55.02 | 136027.2 | 98.99
10.9 | 0 | 100 | 55 | 135967.3 | 99.2
12.1 | 0 | 100 | 54.93 | 135804.1 | 98.92
13.3 | 0 | 100 | 54.95 | 135852.1 | 98.89
14.5 | 0 | 100 | 54.93 | 135808.1 | 98.99
15.7 | 0 | 100 | 54.94 | 135838.7 | 98.89
16.9 | 0 | 100 | 54.89 | 135701.9 | 98.9
18.1 | 0 | 100 | 54.95 | 135864.3 | 98.89
19.3 | 0 | 100 | 54.91 | 135751.4 | 98.98
20.5 | 0 | 100 | 54.83 | 135565.5 | 98.67
21.7 | 0 | 100 | 54.9 | 135722.4 | 98.8
23 | 0 | 100 | 54.69 | 135216.6 | 98.47
24.1 | 0 | 100 | 54.77 | 135404.5 | 98.63
25.4 | 0 | 100 | 54.63 | 135061.7 | 98.34
26.6 | 0 | 100 | 54.48 | 134684.4 | 98.1
27.8 | 0 | 100 | 54.37 | 134419 | 98.01
29 | 0 | 100 | 54.38 | 134446.8 | 97.86
30.2 | 0 | 100 | 54.32 | 134286.9 | 97.58
31.4 | 0 | 100 | 54.11 | 133788.1 | 97.33
32.6 | 0 | 100 | 53.97 | 133431.9 | 97.18
33.8 | 0 | 100 | 53.8 | 133007.2 | 96.89
35 | 0 | 100 | 53.65 | 132637.6 | 96.74
36.2 | 0 | 100 | 53.46 | 132168.4 | 96.07
37.4 | 0 | 100 | 53.5 | 132282.4 | 96.21
38.6 | 0 | 100 | 53.22 | 131572.3 | 95.87
39.8 | 0 | 100 | 52.97 | 130958.3 | 95.27
41 | 0 | 100 | 52.86 | 130680.7 | 95.3
42.2 | 0 | 100 | 52.68 | 130244.5 | 94.86
43.4 | 0 | 100 | 52.39 | 129522.9 | 94.28
44.6 | 0 | 100 | 51.93 | 128379.4 | 93.23
45.8 | 0 | 100 | 51.54 | 127435.9 | 92.87
47 | 0 | 100 | 50.75 | 125474.9 | 91.52
48.2 | 0 | 100 | 49.46 | 122285.4 | 88.96
49.4 | 0 | 100 | 47.11 | 116479.9 | 84.8
50.6 | 0 | 100 | 43.8 | 108300.9 | 78.96
51.8 | 0 | 100 | 39.02 | 96477.3 | 70.21
53 | 0 | 100 | 33.19 | 82059.3 | 59.74
54.2 | 0 | 100 | 26.82 | 66308.5 | 48.28
55.4 | 0 | 100 | 20.31 | 50201.5 | 36.57
56.6 | 0 | 100 | 14.94 | 36934.3 | 26.92
57.8 | 0 | 100 | 11.19 | 27675 | 20.08
59.1 | 0 | 100 | 8.81 | 21792.3 | 15.88
60.2 | 0 | 100 | 7.64 | 18892.5 | 13.75
61.5 | 0 | 100 | 6.87 | 16989.5 | 12.36
62.7 | 0 | 100 | 6.29 | 15556.9 | 11.32
63.9 | 0 | 100 | 5.98 | 14790.4 | 10.78
65.1 | 0 | 100 | 5.64 | 13935.9 | 10.15
66.3 | 0 | 100 | 5.35 | 13224.5 | 9.6
67.5 | 0 | 100 | 5.09 | 12594.1 | 9.18
68.7 | 0 | 100 | 4.87 | 12041.9 | 8.78
69.9 | 0 | 100 | 4.72 | 11659.5 | 8.48
71.1 | 0 | 100 | 4.5 | 11114.4 | 8.09
72.3 | 0 | 100 | 4.34 | 10725.2 | 7.81
73.5 | 0 | 100 | 4.16 | 10295.7 | 7.49
74.7 | 0 | 100 | 4 | 9882.4 | 7.2
75.9 | 0 | 100 | 3.89 | 9607.2 | 7
77.1 | 0 | 100 | 3.78 | 9354.2 | 6.82
78.3 | 0 | 100 | 3.61 | 8925.7 | 6.5
79.5 | 0 | 100 | 3.53 | 8721.7 | 6.35
80.7 | 0 | 100 | 3.35 | 8288.3 | 6.04
81.9 | 0 | 100 | 3.3 | 8161.6 | 5.94
83.1 | 0 | 100 | 3.12 | 7701.7 | 5.62
84.3 | 0 | 100 | 3.05 | 7544 | 5.5
85.5 | 0 | 100 | 2.92 | 7221 | 5.26
86.7 | 0 | 100 | 2.87 | 7106.8 | 5.18
88 | 0 | 100 | 2.71 | 6699 | 4.88
89.2 | 0 | 100 | 2.68 | 6620.9 | 4.83
90.4 | 0 | 100 | 2.57 | 6358.3 | 4.63
91.6 | 0 | 100 | 2.5 | 6189.7 | 4.51
92.8 | 0 | 100 | 2.42 | 5993.4 | 4.36
94 | 0 | 100 | 2.34 | 5773 | 4.21
95.2 | 0 | 100 | 2.24 | 5541 | 4.03
96.4 | 0 | 100 | 2.18 | 5388.3 | 3.93
97.6 | 0 | 100 | 2.1 | 5184.3 | 3.77
98.8 | 0 | 100 | 2.09 | 5158.3 | 3.75
100 | 0 | 100 | 1.98 | 4907.6 | 3.58
101.2 | 0 | 100 | 1.94 | 4786.3 | 3.49
102.4 | 0 | 100 | 1.85 | 4569.5 | 3.33
103.6 | 0 | 100 | 1.83 | 4527.5 | 3.29
104.8 | 0 | 100 | 1.76 | 4339.7 | 3.16
106 | 0 | 100 | 1.7 | 4213.1 | 3.07
107.2 | 0 | 100 | 1.67 | 4119.8 | 3
108.4 | 0 | 100 | 1.64 | 4052.3 | 2.95
109.6 | 0 | 100 | 1.59 | 3941.7 | 2.87
110.8 | 0 | 100 | 1.53 | 3788.3 | 2.76
112 | 0 | 100 | 1.51 | 3741 | 2.73
113.2 | 0 | 100 | 1.42 | 3511.2 | 2.56
114.4 | 0 | 100 | 1.4 | 3457.4 | 2.53
115.7 | 0 | 100 | 1.35 | 3330.9 | 2.42
116.9 | 0 | 100 | 1.34 | 3308.5 | 2.4
118.1 | 0 | 100 | 1.32 | 3251.9 | 2.37
119.3 | 0 | 100 | 1.25 | 3079.4 | 2.25
120.5 | 0 | 100 | 1.19 | 2947.5 | 2.15
121.7 | 0 | 100 | 1.2 | 2960.4 | 2.16
122.9 | 0 | 100 | 1.1 | 2726.9 | 1.99
124.1 | 0 | 100 | 1.08 | 2672.9 | 1.95
125.3 | 0 | 100 | 1.09 | 2695.8 | 1.96
126.5 | 0 | 100 | 1.08 | 2661.9 | 1.94
127.7 | 0 | 100 | 0.95 | 2359.6 | 1.72
128.9 | 0 | 100 | 0.98 | 2418.5 | 1.76
130.1 | 0 | 100 | 0.97 | 2386.6 | 1.74
131.3 | 0 | 100 | 0.93 | 2310.2 | 1.68
132.5 | 0 | 100 | 0.92 | 2271.3 | 1.65
133.7 | 0 | 100 | 0.9 | 2235.5 | 1.63
134.9 | 0 | 100 | 0.86 | 2120.1 | 1.54
136.1 | 0 | 100 | 0.83 | 2042 | 1.49
137.3 | 0 | 100 | 0.79 | 1947.4 | 1.42
138.5 | 0 | 100 | 0.79 | 1944.7 | 1.42
139.7 | 0 | 100 | 0.77 | 1911.4 | 1.39
140.9 | 0 | 100 | 0.77 | 1908.6 | 1.39
142.1 | 0 | 100 | 0.74 | 1832 | 1.34
143.3 | 0 | 100 | 0.71 | 1745.8 | 1.27
144.5 | 0 | 100 | 0.74 | 1824.7 | 1.33
145.7 | 0 | 100 | 0.67 | 1657.2 | 1.21
146.9 | 0 | 100 | 0.64 | 1585.7 | 1.15
148.1 | 0 | 100 | 0.63 | 1564.7 | 1.14
149.3 | 0 | 100 | 0.62 | 1532.9 | 1.12
150.5 | 0 | 100 | 0.56 | 1394 | 1.02
151.7 | 0 | 100 | 0.59 | 1466.3 | 1.07
153 | 0 | 100 | 0.58 | 1422.3 | 1.04
154 | 0 | 100 | 0.56 | 1381.4 | 1.01