## 7_Rendements DSP
Measurement time: | 12/08/2022 18:43:49
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1100 mm
Field size: | 90 x 90 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 17.36 | 42929.3 | 21.02
0 | 0 | 309.1 | 17.32 | 42809.3 | 21.08
0 | 0 | 307.6 | 17.53 | 43336.9 | 21.33
0 | 0 | 306.6 | 17.54 | 43377.3 | 21.31
0 | 0 | 305.3 | 17.72 | 43805.1 | 21.56
0 | 0 | 304.1 | 17.81 | 44030.7 | 21.65
0 | 0 | 302.9 | 17.85 | 44128.6 | 21.75
0 | 0 | 301.7 | 17.93 | 44328.1 | 21.86
0 | 0 | 300.5 | 18.1 | 44753.5 | 22.01
0 | 0 | 299.3 | 18.26 | 45141.2 | 22.22
0 | 0 | 298.1 | 18.37 | 45422.1 | 22.33
0 | 0 | 296.9 | 18.48 | 45690.5 | 22.51
0 | 0 | 295.7 | 18.62 | 46023.3 | 22.68
0 | 0 | 294.5 | 18.69 | 46219.5 | 22.78
0 | 0 | 293.3 | 18.88 | 46670.4 | 22.95
0 | 0 | 292.1 | 18.95 | 46848.1 | 23.05
0 | 0 | 290.9 | 19.14 | 47320.7 | 23.31
0 | 0 | 289.7 | 19.3 | 47707.1 | 23.49
0 | 0 | 288.5 | 19.39 | 47947.8 | 23.65
0 | 0 | 287.3 | 19.56 | 48368.3 | 23.77
0 | 0 | 286.1 | 19.56 | 48356.3 | 23.79
0 | 0 | 284.9 | 19.74 | 48801.6 | 24.01
0 | 0 | 283.7 | 19.93 | 49261.6 | 24.24
0 | 0 | 282.5 | 19.97 | 49364.5 | 24.33
0 | 0 | 281.2 | 20.18 | 49884.1 | 24.56
0 | 0 | 280.1 | 20.26 | 50087.2 | 24.72
0 | 0 | 278.8 | 20.49 | 50647.2 | 24.91
0 | 0 | 277.7 | 20.61 | 50962.4 | 25.08
0 | 0 | 276.4 | 20.78 | 51370.3 | 25.27
0 | 0 | 275.2 | 20.89 | 51637.6 | 25.37
0 | 0 | 274.1 | 21.01 | 51951.4 | 25.56
0 | 0 | 272.8 | 21.13 | 52232.8 | 25.73
0 | 0 | 271.6 | 21.29 | 52641.1 | 25.93
0 | 0 | 270.4 | 21.42 | 52949.5 | 26.06
0 | 0 | 269.2 | 21.63 | 53484.5 | 26.32
0 | 0 | 268.1 | 21.68 | 53605.1 | 26.39
0 | 0 | 266.8 | 21.88 | 54099.8 | 26.62
0 | 0 | 265.6 | 22 | 54383.4 | 26.81
0 | 0 | 264.4 | 22.16 | 54784.3 | 26.92
0 | 0 | 263.2 | 22.3 | 55129.8 | 27.12
0 | 0 | 262 | 22.44 | 55479.5 | 27.24
0 | 0 | 260.8 | 22.62 | 55918.5 | 27.47
0 | 0 | 259.6 | 22.8 | 56371.4 | 27.74
0 | 0 | 258.4 | 22.88 | 56579.6 | 27.85
0 | 0 | 257.2 | 23.08 | 57050.9 | 28.08
0 | 0 | 256 | 23.22 | 57403.8 | 28.26
0 | 0 | 254.8 | 23.35 | 57731.9 | 28.44
0 | 0 | 253.6 | 23.51 | 58135.6 | 28.65
0 | 0 | 252.4 | 23.72 | 58654.5 | 28.86
0 | 0 | 251.2 | 23.9 | 59085.2 | 29.09
0 | 0 | 250 | 24.03 | 59407.9 | 29.24
0 | 0 | 248.7 | 24.17 | 59757.6 | 29.36
0 | 0 | 247.6 | 24.25 | 59955.7 | 29.52
0 | 0 | 246.4 | 24.54 | 60677.8 | 29.89
0 | 0 | 245.1 | 24.66 | 60968.1 | 30
0 | 0 | 244 | 24.8 | 61308.2 | 30.24
0 | 0 | 242.7 | 25 | 61813.2 | 30.46
0 | 0 | 241.6 | 25.17 | 62233 | 30.62
0 | 0 | 240.3 | 25.36 | 62696.3 | 30.86
0 | 0 | 239.1 | 25.52 | 63100 | 31.04
0 | 0 | 238 | 25.69 | 63505.3 | 31.29
0 | 0 | 236.7 | 25.85 | 63906.1 | 31.5
0 | 0 | 235.5 | 26 | 64275.6 | 31.67
0 | 0 | 234.3 | 26.26 | 64933.2 | 31.98
0 | 0 | 233.1 | 26.41 | 65294 | 32.15
0 | 0 | 231.9 | 26.55 | 65645.1 | 32.33
0 | 0 | 230.7 | 26.71 | 66026.2 | 32.54
0 | 0 | 229.5 | 26.91 | 66536.1 | 32.7
0 | 0 | 228.3 | 27.2 | 67248 | 33.12
0 | 0 | 227.1 | 27.33 | 67562.7 | 33.28
0 | 0 | 225.9 | 27.47 | 67906.9 | 33.44
0 | 0 | 224.7 | 27.66 | 68378 | 33.69
0 | 0 | 223.5 | 27.91 | 68992.8 | 34.01
0 | 0 | 222.3 | 28.1 | 69474.6 | 34.23
0 | 0 | 221.1 | 28.28 | 69907.1 | 34.4
0 | 0 | 219.9 | 28.48 | 70409.6 | 34.71
0 | 0 | 218.6 | 28.68 | 70902.9 | 34.94
0 | 0 | 217.5 | 28.91 | 71476.9 | 35.18
0 | 0 | 216.3 | 29.04 | 71802 | 35.31
0 | 0 | 215.1 | 29.29 | 72415.4 | 35.64
0 | 0 | 213.9 | 29.34 | 72535.8 | 35.7
0 | 0 | 212.6 | 29.58 | 73142 | 36.04
0 | 0 | 211.4 | 29.81 | 73689.5 | 36.26
0 | 0 | 210.2 | 30.01 | 74202.9 | 36.55
0 | 0 | 209 | 30.26 | 74811.3 | 36.82
0 | 0 | 207.8 | 30.36 | 75055.1 | 36.88
0 | 0 | 206.6 | 30.53 | 75490.4 | 37.19
0 | 0 | 205.4 | 30.8 | 76145.5 | 37.47
0 | 0 | 204.2 | 30.99 | 76613.5 | 37.72
0 | 0 | 203 | 31.22 | 77196.1 | 37.97
0 | 0 | 201.8 | 31.44 | 77728.5 | 38.23 |  | Z | D
0 | 0 | 200.6 | 31.68 | 78322.8 | 38.55 | D20 | 200.6 | 38.55
0 | 0 | 199.4 | 31.9 | 78855.8 | 38.88 |  | 199.4 | 38.88
0 | 0 | 198.2 | 32.12 | 79407.9 | 39.03 |  | 200 | 38.715
0 | 0 | 197 | 32.24 | 79700.9 | 39.2
0 | 0 | 195.8 | 32.5 | 80346.8 | 39.66
0 | 0 | 194.6 | 32.74 | 80953.9 | 39.76
0 | 0 | 193.4 | 32.93 | 81406 | 40.18
0 | 0 | 192.2 | 33.15 | 81961.7 | 40.34
0 | 0 | 191 | 33.44 | 82671.4 | 40.74
0 | 0 | 189.8 | 33.61 | 83086.9 | 40.88
0 | 0 | 188.6 | 33.9 | 83808.4 | 41.25
0 | 0 | 187.4 | 34.16 | 84465.6 | 41.57
0 | 0 | 186.2 | 34.33 | 84886.5 | 41.72
0 | 0 | 185 | 34.52 | 85350.6 | 42.04
0 | 0 | 183.8 | 34.84 | 86124.8 | 42.32
0 | 0 | 182.5 | 35.07 | 86707 | 42.56
0 | 0 | 181.3 | 35.31 | 87300.4 | 43.01
0 | 0 | 180.1 | 35.43 | 87605.4 | 43.04
0 | 0 | 178.9 | 35.66 | 88174 | 43.46
0 | 0 | 177.7 | 36.01 | 89020.2 | 43.81
0 | 0 | 176.5 | 36.21 | 89514.9 | 44.04
0 | 0 | 175.3 | 36.44 | 90084.4 | 44.38
0 | 0 | 174.1 | 36.7 | 90736.4 | 44.69
0 | 0 | 172.9 | 36.89 | 91203.6 | 44.94
0 | 0 | 171.7 | 37.13 | 91805.2 | 45.29
0 | 0 | 170.5 | 37.36 | 92374.8 | 45.52
0 | 0 | 169.3 | 37.55 | 92842.5 | 45.8
0 | 0 | 168.1 | 37.88 | 93644.4 | 46.11
0 | 0 | 166.9 | 38.16 | 94332.7 | 46.46
0 | 0 | 165.7 | 38.52 | 95223.1 | 46.89
0 | 0 | 164.5 | 38.83 | 95998.6 | 47.13
0 | 0 | 163.3 | 38.97 | 96340.2 | 47.41
0 | 0 | 162.1 | 39.17 | 96832.9 | 47.66
0 | 0 | 160.9 | 39.59 | 97887.3 | 48.12
0 | 0 | 159.7 | 39.79 | 98378.1 | 48.4
0 | 0 | 158.4 | 40 | 98903.2 | 48.66
0 | 0 | 157.2 | 40.3 | 99645.9 | 49.1
0 | 0 | 156.1 | 40.66 | 100521.4 | 49.48 |  | Z | D
0 | 0 | 154.8 | 40.79 | 100836.9 | 49.72 | R50 | 154.8 | 49.72
0 | 0 | 153.6 | 41.09 | 101586.9 | 50.01 |  | 153.6 | 50.01
0 | 0 | 152.4 | 41.46 | 102493.6 | 50.55 |  | 153.6414 | 50
0 | 0 | 151.2 | 41.74 | 103189.8 | 50.78
0 | 0 | 150 | 41.96 | 103731.9 | 51.1
0 | 0 | 148.8 | 42.2 | 104328.1 | 51.37
0 | 0 | 147.6 | 42.43 | 104903.9 | 51.66
0 | 0 | 146.4 | 42.79 | 105794.6 | 52.14
0 | 0 | 145.2 | 43.01 | 106339.5 | 52.39
0 | 0 | 144 | 43.27 | 106970.7 | 52.64
0 | 0 | 142.8 | 43.65 | 107907.8 | 53.06
0 | 0 | 141.6 | 43.88 | 108481.2 | 53.35
0 | 0 | 140.4 | 44.19 | 109259.8 | 53.8
0 | 0 | 139.2 | 44.5 | 110024.4 | 54.17
0 | 0 | 138 | 44.84 | 110859.6 | 54.58
0 | 0 | 136.8 | 45.06 | 111411.7 | 54.88
0 | 0 | 135.6 | 45.36 | 112145.1 | 55.19
0 | 0 | 134.4 | 45.8 | 113236.7 | 55.66
0 | 0 | 133.2 | 45.98 | 113674.6 | 56.07
0 | 0 | 132 | 46.29 | 114454.7 | 56.44
0 | 0 | 130.8 | 46.68 | 115407.9 | 56.76
0 | 0 | 129.6 | 46.94 | 116049.4 | 57.09
0 | 0 | 128.4 | 47.36 | 117082.9 | 57.57
0 | 0 | 127.2 | 47.56 | 117589.5 | 57.95
0 | 0 | 126 | 47.9 | 118430.6 | 58.22
0 | 0 | 124.8 | 48.2 | 119161.4 | 58.72
0 | 0 | 123.6 | 48.53 | 119993.1 | 59.05
0 | 0 | 122.3 | 48.91 | 120933.1 | 59.54
0 | 0 | 121.2 | 49.26 | 121783.4 | 60.02
0 | 0 | 120 | 49.52 | 122438 | 60.25
0 | 0 | 118.8 | 49.83 | 123197.8 | 60.7
0 | 0 | 117.6 | 50.21 | 124128.3 | 61.15
0 | 0 | 116.3 | 50.43 | 124675.7 | 61.4
0 | 0 | 115.2 | 50.84 | 125684.7 | 61.96
0 | 0 | 113.9 | 51.15 | 126454.2 | 62.31
0 | 0 | 112.7 | 51.52 | 127379.3 | 62.78
0 | 0 | 111.5 | 51.84 | 128161.1 | 63.1
0 | 0 | 110.3 | 52.2 | 129046 | 63.53
0 | 0 | 109.1 | 52.52 | 129841.5 | 63.94
0 | 0 | 107.9 | 52.83 | 130616.4 | 64.26
0 | 0 | 106.7 | 53.18 | 131489 | 64.71
0 | 0 | 105.5 | 53.52 | 132309.7 | 65.21
0 | 0 | 104.3 | 53.85 | 133146.9 | 65.54
0 | 0 | 103.1 | 54.16 | 133906.2 | 65.89
0 | 0 | 101.9 | 54.6 | 134978.3 | 66.38 |  | Z | D
0 | 0 | 100.7 | 54.92 | 135788.7 | 66.9 | D10 | 100.7 | 66.9
0 | 0 | 99.5 | 55.3 | 136709.3 | 67.27 |  | 99.5 | 67.27
0 | 0 | 98.3 | 55.57 | 137376.7 | 67.5 |  | 100 | 67.1158
0 | 0 | 97.1 | 56.06 | 138607 | 68.2
0 | 0 | 95.9 | 56.39 | 139413 | 68.57
0 | 0 | 94.7 | 56.76 | 140322.5 | 69
0 | 0 | 93.5 | 57.12 | 141214.9 | 69.37
0 | 0 | 92.3 | 57.44 | 142019.2 | 69.81
0 | 0 | 91.1 | 57.84 | 143001.2 | 70.38
0 | 0 | 89.9 | 58.24 | 143997.4 | 70.94
0 | 0 | 88.7 | 58.52 | 144684.7 | 71.17
0 | 0 | 87.5 | 58.99 | 145838.9 | 71.77
0 | 0 | 86.3 | 59.29 | 146583.9 | 72.19
0 | 0 | 85.1 | 59.82 | 147900.2 | 72.79
0 | 0 | 83.8 | 60.11 | 148604.9 | 73.23
0 | 0 | 82.6 | 60.47 | 149498 | 73.6
0 | 0 | 81.5 | 60.84 | 150427.7 | 73.99
0 | 0 | 80.2 | 61.22 | 151353.6 | 74.55
0 | 0 | 79 | 61.56 | 152196.8 | 74.92
0 | 0 | 77.8 | 62.1 | 153525.3 | 75.84
0 | 0 | 76.6 | 62.37 | 154207 | 75.92
0 | 0 | 75.4 | 62.76 | 155164.7 | 76.39
0 | 0 | 74.2 | 63.2 | 156240.7 | 76.96
0 | 0 | 73 | 63.66 | 157386.1 | 77.46
0 | 0 | 71.8 | 64 | 158236.6 | 77.79
0 | 0 | 70.6 | 64.26 | 158879.5 | 78.16
0 | 0 | 69.3 | 64.81 | 160243 | 78.82
0 | 0 | 68.2 | 65.13 | 161021.6 | 79.26
0 | 0 | 66.9 | 65.5 | 161950.7 | 79.75
0 | 0 | 65.7 | 65.86 | 162825.2 | 80.3
0 | 0 | 64.5 | 66.28 | 163863.7 | 80.64
0 | 0 | 63.3 | 66.7 | 164914.8 | 81.08
0 | 0 | 62.1 | 67.2 | 166143.7 | 81.68
0 | 0 | 60.9 | 67.5 | 166894.5 | 82.09
0 | 0 | 59.7 | 67.97 | 168051.7 | 82.79
0 | 0 | 58.5 | 68.24 | 168712.4 | 83.09
0 | 0 | 57.3 | 68.76 | 170007.6 | 83.74
0 | 0 | 56 | 69.12 | 170893.1 | 84.29
0 | 0 | 54.9 | 69.45 | 171709.7 | 84.48
0 | 0 | 53.6 | 69.97 | 172995.8 | 85.23
0 | 0 | 52.5 | 70.24 | 173649.1 | 85.47
0 | 0 | 51.2 | 70.62 | 174590.9 | 86.09
0 | 0 | 50 | 71.12 | 175829.8 | 86.55
0 | 0 | 48.8 | 71.6 | 177020.8 | 87.22
0 | 0 | 47.6 | 72.1 | 178248.7 | 87.69
0 | 0 | 46.4 | 72.35 | 178881.1 | 88.12
0 | 0 | 45.2 | 72.77 | 179924.3 | 88.58
0 | 0 | 44 | 73.17 | 180897.8 | 89.23
0 | 0 | 42.8 | 73.58 | 181908.7 | 89.84
0 | 0 | 41.6 | 74.07 | 183131.5 | 90.17
0 | 0 | 40.4 | 74.54 | 184296.9 | 90.63
0 | 0 | 39.1 | 74.96 | 185323.2 | 91.33
0 | 0 | 38 | 75.39 | 186402.6 | 91.89
0 | 0 | 36.8 | 75.82 | 187444.6 | 92.06
0 | 0 | 35.5 | 76.14 | 188246.9 | 92.66
0 | 0 | 34.3 | 76.57 | 189307.4 | 93.17
0 | 0 | 33.1 | 76.96 | 190271.2 | 93.57
0 | 0 | 31.9 | 77.24 | 190964.1 | 93.9
0 | 0 | 30.7 | 77.76 | 192259.8 | 94.68
0 | 0 | 29.5 | 78.2 | 193334.9 | 95.17
0 | 0 | 28.3 | 78.34 | 193693.4 | 95.32
0 | 0 | 27.1 | 78.94 | 195162.7 | 95.98
0 | 0 | 25.9 | 79.35 | 196181.3 | 96.48
0 | 0 | 24.7 | 79.77 | 197211 | 97.01
0 | 0 | 23.5 | 80.23 | 198346.5 | 97.36
0 | 0 | 22.3 | 80.53 | 199108.9 | 98.07
0 | 0 | 21 | 80.91 | 200047.1 | 98.43
0 | 0 | 19.8 | 81.27 | 200924.2 | 99.12
0 | 0 | 18.6 | 81.75 | 202112 | 99.48
0 | 0 | 17.4 | 81.94 | 202581.9 | 99.42
0 | 0 | 16.2 | 82.15 | 203108.7 | 100
0 | 0 | 15 | 82.05 | 202857.2 | 99.99
0 | 0 | 13.8 | 82.19 | 203203.7 | 99.78
0 | 0 | 12.6 | 82.07 | 202894.8 | 99.75
0 | 0 | 11.4 | 81.6 | 201755.2 | 99.21
0 | 0 | 10.2 | 80.69 | 199483 | 98.06
0 | 0 | 9 | 79.42 | 196359.9 | 96.71
0 | 0 | 7.8 | 77.44 | 191447.4 | 94.4
0 | 0 | 6.6 | 74.59 | 184418.7 | 90.98
0 | 0 | 5.3 | 70.25 | 173674.9 | 85.46
0 | 0 | 4.1 | 63.95 | 158105.3 | 77.87
0 | 0 | 2.9 | 55.58 | 137411 | 67.6
0 | 0 | 1.7 | 47.68 | 117874.3 | 58.05
0 | 0 | 0.5 | 43.35 | 107179.3 | 52.8
0 | 0 | -0.5 | 40.75 | 100742.4 | 49.54
Measurement time: | 12/08/2022 18:48:09
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 110 x 110 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 22.31 | 55169 | 19.31
0 | 0 | 309.1 | 24.25 | 59966.2 | 19.4
0 | 0 | 307.6 | 24.45 | 60441.3 | 19.57
0 | 0 | 306.6 | 24.68 | 61014.7 | 19.74
0 | 0 | 305.3 | 24.83 | 61387.3 | 19.88
0 | 0 | 304.1 | 24.97 | 61741.2 | 19.98
0 | 0 | 302.9 | 25.17 | 62233.7 | 20.16
0 | 0 | 301.7 | 25.32 | 62604.4 | 20.26
0 | 0 | 300.5 | 25.43 | 62864.9 | 20.34
0 | 0 | 299.3 | 25.7 | 63551.6 | 20.56
0 | 0 | 298.1 | 25.9 | 64037.7 | 20.73
0 | 0 | 296.9 | 26.02 | 64332 | 20.82
0 | 0 | 295.7 | 26.21 | 64791.3 | 20.99
0 | 0 | 294.5 | 26.4 | 65267.4 | 21.11
0 | 0 | 293.3 | 26.53 | 65588.2 | 21.2
0 | 0 | 292.1 | 26.75 | 66132 | 21.41
0 | 0 | 290.9 | 26.96 | 66645.9 | 21.59
0 | 0 | 289.7 | 27.13 | 67073.2 | 21.69
0 | 0 | 288.5 | 27.24 | 67355 | 21.8
0 | 0 | 287.2 | 27.43 | 67821.2 | 22.04
0 | 0 | 286.1 | 27.69 | 68455.2 | 22.18
0 | 0 | 284.8 | 27.9 | 68986.2 | 22.33
0 | 0 | 283.6 | 28.08 | 69419 | 22.46
0 | 0 | 282.5 | 28.28 | 69912.5 | 22.61
0 | 0 | 281.2 | 28.41 | 70230.1 | 22.71
0 | 0 | 280.1 | 28.67 | 70888.2 | 22.95
0 | 0 | 278.8 | 28.83 | 71266.5 | 23.08
0 | 0 | 277.7 | 28.99 | 71684.1 | 23.2
0 | 0 | 276.4 | 29.21 | 72205.2 | 23.39
0 | 0 | 275.2 | 29.44 | 72780.6 | 23.58
0 | 0 | 274.1 | 29.63 | 73252.6 | 23.75
0 | 0 | 272.8 | 29.8 | 73682.3 | 23.87
0 | 0 | 271.6 | 30.06 | 74312.5 | 24.07
0 | 0 | 270.4 | 30.25 | 74787.9 | 24.21
0 | 0 | 269.2 | 30.48 | 75361.4 | 24.39
0 | 0 | 268 | 30.74 | 75989.8 | 24.62
0 | 0 | 266.8 | 30.87 | 76329.5 | 24.72
0 | 0 | 265.6 | 31.12 | 76937.6 | 24.92
0 | 0 | 264.4 | 31.29 | 77368.7 | 25.06
0 | 0 | 263.2 | 31.54 | 77977.6 | 25.24
0 | 0 | 262 | 31.71 | 78389 | 25.36
0 | 0 | 260.8 | 31.97 | 79038.5 | 25.59
0 | 0 | 259.6 | 32.15 | 79494.3 | 25.73
0 | 0 | 258.4 | 32.44 | 80194.7 | 25.95
0 | 0 | 257.2 | 32.7 | 80852.5 | 26.16
0 | 0 | 256 | 32.91 | 81356.6 | 26.37
0 | 0 | 254.8 | 33.07 | 81763.3 | 26.5
0 | 0 | 253.6 | 33.24 | 82177.2 | 26.65
0 | 0 | 252.4 | 33.56 | 82965 | 26.89
0 | 0 | 251.2 | 33.92 | 83853.5 | 27.16
0 | 0 | 250 | 34.06 | 84205.3 | 27.27
0 | 0 | 248.7 | 34.33 | 84876.7 | 27.49
0 | 0 | 247.6 | 34.51 | 85318.4 | 27.63
0 | 0 | 246.3 | 34.76 | 85932.9 | 27.86
0 | 0 | 245.1 | 35.09 | 86766.2 | 28.13
0 | 0 | 243.9 | 35.35 | 87403.3 | 28.28
0 | 0 | 242.7 | 35.55 | 87903.2 | 28.49
0 | 0 | 241.6 | 35.76 | 88416.2 | 28.64
0 | 0 | 240.3 | 36.05 | 89128 | 28.81
0 | 0 | 239.1 | 36.3 | 89741.5 | 29.1
0 | 0 | 238 | 36.59 | 90475.3 | 29.34
0 | 0 | 236.7 | 36.81 | 91006 | 29.45
0 | 0 | 235.5 | 37.04 | 91574.8 | 29.66
0 | 0 | 234.3 | 37.33 | 92282.7 | 29.91
0 | 0 | 233.1 | 37.57 | 92878.1 | 30.07
0 | 0 | 231.9 | 37.82 | 93493.7 | 30.3
0 | 0 | 230.7 | 38.06 | 94087.4 | 30.43
0 | 0 | 229.5 | 38.41 | 94956.6 | 30.79
0 | 0 | 228.3 | 38.56 | 95337.6 | 30.91
0 | 0 | 227.1 | 38.91 | 96206.6 | 31.2
0 | 0 | 225.9 | 39.2 | 96928.2 | 31.39
0 | 0 | 224.6 | 39.56 | 97799.9 | 31.67
0 | 0 | 223.5 | 39.64 | 97992.7 | 31.78
0 | 0 | 222.2 | 40.03 | 98961.4 | 32.08
0 | 0 | 221.1 | 40.39 | 99852.4 | 32.33
0 | 0 | 219.9 | 40.56 | 100286 | 32.49
0 | 0 | 218.6 | 40.9 | 101115.6 | 32.77
0 | 0 | 217.5 | 41.12 | 101663.9 | 32.97
0 | 0 | 216.2 | 41.28 | 102050.3 | 33.06
0 | 0 | 215.1 | 41.65 | 102973.7 | 33.4
0 | 0 | 213.8 | 41.93 | 103667.3 | 33.61
0 | 0 | 212.6 | 42.26 | 104480.2 | 33.87
0 | 0 | 211.4 | 42.52 | 105115.3 | 34.11
0 | 0 | 210.2 | 42.96 | 106209.3 | 34.37
0 | 0 | 209 | 43.25 | 106917.6 | 34.66
0 | 0 | 207.9 | 43.51 | 107578.4 | 34.81
0 | 0 | 206.6 | 43.85 | 108416.8 | 35.09
0 | 0 | 205.4 | 44.05 | 108895.6 | 35.23
0 | 0 | 204.2 | 44.39 | 109738 | 35.54
0 | 0 | 203 | 44.7 | 110516.3 | 35.81
0 | 0 | 201.8 | 45.01 | 111272 | 36 |  | Z | D
0 | 0 | 200.6 | 45.24 | 111839.1 | 36.24 | D20 | 200.6 | 36.24
0 | 0 | 199.4 | 45.65 | 112866.6 | 36.52 |  | 199.4 | 36.52
0 | 0 | 198.2 | 45.91 | 113516.7 | 36.85 |  | 200 | 36.38
0 | 0 | 197 | 46.21 | 114239 | 37.03
0 | 0 | 195.8 | 46.5 | 114955.9 | 37.32
0 | 0 | 194.6 | 46.75 | 115594.7 | 37.45
0 | 0 | 193.4 | 47.19 | 116661.1 | 37.86
0 | 0 | 192.2 | 47.56 | 117576.2 | 38.15
0 | 0 | 190.9 | 47.96 | 118580.9 | 38.42
0 | 0 | 189.8 | 48.21 | 119195.3 | 38.63
0 | 0 | 188.5 | 48.55 | 120037.1 | 38.98
0 | 0 | 187.3 | 49.02 | 121193.7 | 39.29
0 | 0 | 186.1 | 49.28 | 121833.4 | 39.47
0 | 0 | 184.9 | 49.68 | 122831.6 | 39.86
0 | 0 | 183.7 | 50.07 | 123784.6 | 40.11
0 | 0 | 182.5 | 50.33 | 124438.8 | 40.36
0 | 0 | 181.3 | 50.63 | 125185 | 40.58
0 | 0 | 180.1 | 51.09 | 126321.3 | 40.94
0 | 0 | 178.9 | 51.52 | 127374.3 | 41.27
0 | 0 | 177.7 | 51.88 | 128262.1 | 41.59
0 | 0 | 176.5 | 52.18 | 128995 | 41.85
0 | 0 | 175.3 | 52.48 | 129748.4 | 42.07
0 | 0 | 174.1 | 52.86 | 130677.4 | 42.4
0 | 0 | 172.9 | 53.22 | 131572.9 | 42.62
0 | 0 | 171.7 | 53.71 | 132786.2 | 43.02
0 | 0 | 170.5 | 53.96 | 133400.1 | 43.23
0 | 0 | 169.3 | 54.4 | 134495.8 | 43.6
0 | 0 | 168.1 | 54.71 | 135264.1 | 43.83
0 | 0 | 166.9 | 55.26 | 136625.1 | 44.3
0 | 0 | 165.7 | 55.56 | 137355 | 44.54
0 | 0 | 164.5 | 56.01 | 138484.8 | 44.88
0 | 0 | 163.3 | 56.41 | 139464.7 | 45.21
0 | 0 | 162.1 | 56.73 | 140249.3 | 45.54
0 | 0 | 160.9 | 57.25 | 141531.5 | 45.89
0 | 0 | 159.7 | 57.6 | 142397.6 | 46.13
0 | 0 | 158.4 | 58 | 143384.6 | 46.52
0 | 0 | 157.3 | 58.34 | 144228 | 46.72
0 | 0 | 156 | 58.8 | 145386.5 | 47.13
0 | 0 | 154.8 | 59.22 | 146423.6 | 47.38
0 | 0 | 153.6 | 59.69 | 147584.9 | 47.82
0 | 0 | 152.4 | 59.93 | 148174.2 | 48.01
0 | 0 | 151.2 | 60.28 | 149028.1 | 48.35
0 | 0 | 150 | 60.92 | 150612.2 | 48.8
0 | 0 | 148.8 | 61.39 | 151785.1 | 49.12
0 | 0 | 147.6 | 61.79 | 152773.7 | 49.4 |  | Z | D
0 | 0 | 146.4 | 62.19 | 153752 | 49.88 | R50 | 146.4 | 49.88
0 | 0 | 145.2 | 62.57 | 154704.3 | 50.12 |  | 145.2 | 50.12
0 | 0 | 144 | 63.06 | 155901.9 | 50.45 |  | 145.8 | 50
0 | 0 | 142.8 | 63.46 | 156890.6 | 50.87
0 | 0 | 141.6 | 63.8 | 157731.4 | 51.16
0 | 0 | 140.4 | 64.37 | 159136.9 | 51.63
0 | 0 | 139.2 | 64.78 | 160150.9 | 51.94
0 | 0 | 138 | 65.11 | 160970 | 52.24
0 | 0 | 136.8 | 65.64 | 162287.4 | 52.61
0 | 0 | 135.6 | 66.1 | 163413.4 | 52.93
0 | 0 | 134.4 | 66.68 | 164860.5 | 53.41
0 | 0 | 133.2 | 67.05 | 165776.9 | 53.69
0 | 0 | 132 | 67.48 | 166838.2 | 54.14
0 | 0 | 130.8 | 67.99 | 168101.4 | 54.58
0 | 0 | 129.6 | 68.51 | 169385.6 | 55
0 | 0 | 128.4 | 68.88 | 170289.4 | 55.26
0 | 0 | 127.2 | 69.65 | 172205.6 | 55.74
0 | 0 | 126 | 69.9 | 172826.5 | 55.98
0 | 0 | 124.8 | 70.53 | 174365.7 | 56.41
0 | 0 | 123.6 | 70.93 | 175373.8 | 56.75
0 | 0 | 122.3 | 71.55 | 176899.3 | 57.28
0 | 0 | 121.2 | 71.92 | 177810.5 | 57.62
0 | 0 | 119.9 | 72.25 | 178633.8 | 57.87
0 | 0 | 118.7 | 72.86 | 180129.8 | 58.37
0 | 0 | 117.5 | 73.33 | 181290.3 | 58.76
0 | 0 | 116.3 | 73.74 | 182315.7 | 59.16
0 | 0 | 115.1 | 74.25 | 183584 | 59.49
0 | 0 | 113.9 | 74.84 | 185039 | 59.88
0 | 0 | 112.7 | 75.31 | 186185 | 60.35
0 | 0 | 111.5 | 76.03 | 187977.8 | 60.88
0 | 0 | 110.3 | 76.2 | 188403.5 | 61.12
0 | 0 | 109.1 | 76.92 | 190164 | 61.62
0 | 0 | 107.9 | 77.35 | 191239.8 | 62.09
0 | 0 | 106.7 | 77.73 | 192185.7 | 62.35
0 | 0 | 105.5 | 78.5 | 194078.3 | 62.88
0 | 0 | 104.3 | 79.12 | 195607.7 | 63.39
0 | 0 | 103.1 | 79.63 | 196868.2 | 63.78
0 | 0 | 101.9 | 80.12 | 198077.1 | 64.19 |  | Z | D
0 | 0 | 100.7 | 80.69 | 199501.7 | 64.61 | D10 | 100.7 | 64.61
0 | 0 | 99.5 | 81.32 | 201046.7 | 65.11 |  | 99.5 | 65.11
0 | 0 | 98.3 | 81.77 | 202169.9 | 65.54 |  | 100 | 64.9017
0 | 0 | 97.1 | 82.51 | 203982.8 | 66.09
0 | 0 | 95.9 | 82.94 | 205067.6 | 66.53
0 | 0 | 94.7 | 83.54 | 206552.1 | 66.92
0 | 0 | 93.5 | 84.1 | 207919.6 | 67.43
0 | 0 | 92.3 | 84.57 | 209085.2 | 67.92
0 | 0 | 91.1 | 85.09 | 210377.5 | 68.2
0 | 0 | 89.9 | 85.61 | 211662.3 | 68.68
0 | 0 | 88.7 | 86.65 | 214217.4 | 69.34
0 | 0 | 87.5 | 87.13 | 215418.1 | 69.72
0 | 0 | 86.2 | 87.55 | 216451.3 | 70.14
0 | 0 | 85.1 | 88.2 | 218057.5 | 70.6
0 | 0 | 83.8 | 88.78 | 219488.9 | 71.17
0 | 0 | 82.6 | 89.31 | 220804.3 | 71.7
0 | 0 | 81.4 | 89.93 | 222335.9 | 72.14
0 | 0 | 80.2 | 90.56 | 223907.2 | 72.52
0 | 0 | 79 | 91.03 | 225059.5 | 72.84
0 | 0 | 77.8 | 91.51 | 226240.3 | 73.4
0 | 0 | 76.6 | 92.09 | 227685 | 73.87
0 | 0 | 75.4 | 92.77 | 229367.3 | 74.26
0 | 0 | 74.2 | 93.43 | 230999.9 | 74.9
0 | 0 | 73 | 93.95 | 232286.9 | 75.31
0 | 0 | 71.8 | 94.7 | 234129.8 | 75.87
0 | 0 | 70.5 | 95.29 | 235581.7 | 76.24
0 | 0 | 69.4 | 95.84 | 236952.7 | 76.72
0 | 0 | 68.1 | 96.42 | 238380.6 | 77.32
0 | 0 | 67 | 97.15 | 240185.8 | 77.79
0 | 0 | 65.7 | 97.88 | 241988.5 | 78.37
0 | 0 | 64.5 | 98.69 | 243992.2 | 78.96
0 | 0 | 63.3 | 99.19 | 245237.7 | 79.44
0 | 0 | 62.1 | 99.7 | 246487.9 | 79.96
0 | 0 | 60.9 | 100.36 | 248121.5 | 80.32
0 | 0 | 59.7 | 100.9 | 249472.3 | 80.85
0 | 0 | 58.5 | 101.7 | 251427.4 | 81.42
0 | 0 | 57.3 | 102.26 | 252815.6 | 81.98
0 | 0 | 56.1 | 102.96 | 254551.1 | 82.52
0 | 0 | 54.9 | 103.67 | 256296.3 | 83.07
0 | 0 | 53.7 | 104.25 | 257744.5 | 83.49
0 | 0 | 52.5 | 104.94 | 259459 | 84.05
0 | 0 | 51.3 | 105.58 | 261030.3 | 84.55
0 | 0 | 50 | 106.18 | 262510.5 | 85.03
0 | 0 | 48.9 | 106.77 | 263980.9 | 85.61
0 | 0 | 47.6 | 107.42 | 265579.2 | 86.1
0 | 0 | 46.5 | 108.29 | 267720.3 | 86.69
0 | 0 | 45.2 | 108.78 | 268937.4 | 87.14
0 | 0 | 44 | 109.32 | 270289 | 87.64
0 | 0 | 42.8 | 109.97 | 271875.4 | 88.08
0 | 0 | 41.6 | 110.89 | 274168.8 | 88.76
0 | 0 | 40.4 | 111.55 | 275789.5 | 89.25
0 | 0 | 39.2 | 112.09 | 277115.7 | 89.79
0 | 0 | 38 | 112.77 | 278814.9 | 90.42
0 | 0 | 36.8 | 113.49 | 280577.2 | 90.92
0 | 0 | 35.6 | 114.13 | 282162.1 | 91.42
0 | 0 | 34.4 | 114.75 | 283709.5 | 91.97
0 | 0 | 33.2 | 115.43 | 285375.6 | 92.46
0 | 0 | 32 | 115.93 | 286611.7 | 92.9
0 | 0 | 30.7 | 116.77 | 288704.7 | 93.5
0 | 0 | 29.5 | 117.52 | 290542.9 | 94.11
0 | 0 | 28.3 | 118.19 | 292205.5 | 94.67
0 | 0 | 27.1 | 118.78 | 293676.8 | 95.04
0 | 0 | 25.9 | 119.43 | 295263.2 | 95.68
0 | 0 | 24.7 | 119.99 | 296655.2 | 96.18
0 | 0 | 23.5 | 120.44 | 297774.7 | 96.67
0 | 0 | 22.3 | 121.39 | 300124.3 | 97.3
0 | 0 | 21.1 | 121.98 | 301570.6 | 97.71
0 | 0 | 19.9 | 122.82 | 303652.3 | 98.29
0 | 0 | 18.7 | 123.63 | 305666.2 | 98.85
0 | 0 | 17.5 | 123.91 | 306337.3 | 99.21
0 | 0 | 16.3 | 124.24 | 307157.8 | 99.56
0 | 0 | 15 | 124.85 | 308684.9 | 99.89
0 | 0 | 13.8 | 124.8 | 308556.9 | 100
0 | 0 | 12.6 | 124.43 | 307635.8 | 99.85
0 | 0 | 11.4 | 123.79 | 306059.1 | 99.19
0 | 0 | 10.2 | 122.97 | 304032.5 | 98.45
0 | 0 | 9 | 121.05 | 299270.5 | 97.16
0 | 0 | 7.8 | 118.54 | 293072.1 | 94.98
0 | 0 | 6.6 | 114.7 | 283586.8 | 91.93
0 | 0 | 5.4 | 108.63 | 268565.6 | 87.02
0 | 0 | 4.2 | 100.12 | 247542.8 | 80.23
0 | 0 | 3 | 87.27 | 215770.4 | 69.88
0 | 0 | 1.8 | 75.13 | 185742.8 | 60.31
0 | 0 | 0.6 | 68.88 | 170296.6 | 55.2
0 | 0 | -0.5 | 64.7 | 159954 | 51.88
Measurement time: | 12/08/2022 18:51:58
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 20.62 | 50985.5 | 20.43
0 | 0 | 309.1 | 20.48 | 50623 | 20.37
0 | 0 | 307.6 | 20.65 | 51042.1 | 20.54
0 | 0 | 306.6 | 20.81 | 51445.5 | 20.66
0 | 0 | 305.3 | 20.93 | 51753.1 | 20.82
0 | 0 | 304.1 | 21.06 | 52070.9 | 20.93
0 | 0 | 302.9 | 21.1 | 52163 | 20.96
0 | 0 | 301.7 | 21.3 | 52659.5 | 21.18
0 | 0 | 300.5 | 21.44 | 53019 | 21.33
0 | 0 | 299.3 | 21.58 | 53350.2 | 21.45
0 | 0 | 298.1 | 21.75 | 53779.2 | 21.62
0 | 0 | 296.9 | 21.87 | 54060.2 | 21.75
0 | 0 | 295.7 | 22.05 | 54514.3 | 21.92
0 | 0 | 294.5 | 22.28 | 55077.1 | 22.17
0 | 0 | 293.3 | 22.37 | 55299.1 | 22.22
0 | 0 | 292.1 | 22.41 | 55400.5 | 22.32
0 | 0 | 290.9 | 22.65 | 55993.4 | 22.5
0 | 0 | 289.7 | 22.78 | 56312.6 | 22.68
0 | 0 | 288.5 | 22.98 | 56807.9 | 22.84
0 | 0 | 287.3 | 23.08 | 57072.6 | 22.95
0 | 0 | 286.1 | 23.2 | 57352 | 23.1
0 | 0 | 284.9 | 23.47 | 58024.3 | 23.34
0 | 0 | 283.7 | 23.56 | 58247.1 | 23.45
0 | 0 | 282.5 | 23.72 | 58656 | 23.63
0 | 0 | 281.3 | 23.91 | 59113.3 | 23.79
0 | 0 | 280 | 24.12 | 59636.3 | 23.97
0 | 0 | 278.8 | 24.25 | 59949.3 | 24.1
0 | 0 | 277.7 | 24.42 | 60376 | 24.28
0 | 0 | 276.4 | 24.67 | 60986 | 24.55
0 | 0 | 275.2 | 24.76 | 61221.9 | 24.63
0 | 0 | 274.1 | 24.92 | 61615.9 | 24.81
0 | 0 | 272.8 | 25.08 | 61996.5 | 24.97
0 | 0 | 271.6 | 25.26 | 62445.9 | 25.13
0 | 0 | 270.4 | 25.43 | 62859.5 | 25.27
0 | 0 | 269.2 | 25.6 | 63282.3 | 25.51
0 | 0 | 268 | 25.75 | 63655.9 | 25.68
0 | 0 | 266.8 | 25.97 | 64199.3 | 25.83
0 | 0 | 265.6 | 26.09 | 64498.3 | 25.95
0 | 0 | 264.4 | 26.28 | 64978.1 | 26.13
0 | 0 | 263.2 | 26.42 | 65313.9 | 26.29
0 | 0 | 262 | 26.71 | 66043.8 | 26.54
0 | 0 | 260.8 | 26.79 | 66231.1 | 26.69
0 | 0 | 259.6 | 26.99 | 66732.3 | 26.88
0 | 0 | 258.4 | 27.18 | 67189.8 | 27.02
0 | 0 | 257.2 | 27.37 | 67671.8 | 27.26
0 | 0 | 256 | 27.59 | 68224.1 | 27.47
0 | 0 | 254.8 | 27.68 | 68429.5 | 27.57
0 | 0 | 253.6 | 27.95 | 69091.6 | 27.85
0 | 0 | 252.4 | 28.13 | 69543 | 28.01
0 | 0 | 251.2 | 28.31 | 69983.9 | 28.15
0 | 0 | 250 | 28.39 | 70193.5 | 28.27
0 | 0 | 248.8 | 28.68 | 70906.1 | 28.55
0 | 0 | 247.6 | 28.92 | 71494.6 | 28.82
0 | 0 | 246.3 | 29.1 | 71945.2 | 28.95
0 | 0 | 245.1 | 29.21 | 72222.7 | 29.11
0 | 0 | 243.9 | 29.42 | 72737.5 | 29.31
0 | 0 | 242.7 | 29.67 | 73346.5 | 29.56
0 | 0 | 241.5 | 29.9 | 73920 | 29.79
0 | 0 | 240.3 | 30.11 | 74438.1 | 29.96
0 | 0 | 239.1 | 30.28 | 74850.9 | 30.11
0 | 0 | 237.9 | 30.49 | 75372 | 30.33
0 | 0 | 236.7 | 30.76 | 76055.8 | 30.63
0 | 0 | 235.5 | 30.93 | 76473.8 | 30.8
0 | 0 | 234.3 | 31.16 | 77026.5 | 31.04
0 | 0 | 233.1 | 31.35 | 77505.4 | 31.23
0 | 0 | 231.9 | 31.59 | 78112.6 | 31.48
0 | 0 | 230.7 | 31.81 | 78638.5 | 31.69
0 | 0 | 229.5 | 32.02 | 79172.4 | 31.89
0 | 0 | 228.3 | 32.2 | 79618.2 | 32.04
0 | 0 | 227.1 | 32.5 | 80344.9 | 32.37
0 | 0 | 225.9 | 32.68 | 80798.7 | 32.58
0 | 0 | 224.7 | 32.88 | 81296 | 32.77
0 | 0 | 223.5 | 33.19 | 82052.6 | 33.05
0 | 0 | 222.3 | 33.33 | 82396.1 | 33.14
0 | 0 | 221.1 | 33.6 | 83075.3 | 33.46
0 | 0 | 219.9 | 33.81 | 83580.8 | 33.72
0 | 0 | 218.7 | 34 | 84055.1 | 33.83
0 | 0 | 217.5 | 34.29 | 84771.1 | 34.12
0 | 0 | 216.3 | 34.52 | 85351.2 | 34.42
0 | 0 | 215 | 34.77 | 85967.4 | 34.69
0 | 0 | 213.9 | 35.03 | 86597.9 | 34.84
0 | 0 | 212.7 | 35.29 | 87251.1 | 35.14
0 | 0 | 211.5 | 35.46 | 87661.4 | 35.31
0 | 0 | 210.2 | 35.76 | 88414.8 | 35.59
0 | 0 | 209 | 36.06 | 89161.7 | 35.88
0 | 0 | 207.8 | 36.27 | 89675.2 | 36.1
0 | 0 | 206.6 | 36.34 | 89838 | 36.31
0 | 0 | 205.4 | 36.74 | 90842.4 | 36.6
0 | 0 | 204.2 | 36.94 | 91322.4 | 36.8
0 | 0 | 203 | 37.2 | 91969.9 | 37.07
0 | 0 | 201.8 | 37.42 | 92514.5 | 37.31 |  | Z | D
0 | 0 | 200.6 | 37.61 | 92990.5 | 37.44 | D20 | 200.6 | 37.44
0 | 0 | 199.4 | 37.83 | 93533.1 | 37.73 |  | 199.4 | 37.73
0 | 0 | 198.2 | 38.18 | 94394.5 | 38.05 |  | 200 | 37.585
0 | 0 | 197 | 38.51 | 95207.4 | 38.32
0 | 0 | 195.8 | 38.72 | 95723.6 | 38.55
0 | 0 | 194.6 | 39.04 | 96520.3 | 38.9
0 | 0 | 193.4 | 39.28 | 97116.6 | 39.15
0 | 0 | 192.2 | 39.5 | 97654.1 | 39.3
0 | 0 | 191 | 39.83 | 98484.4 | 39.68
0 | 0 | 189.8 | 40.17 | 99307.9 | 39.95
0 | 0 | 188.5 | 40.47 | 100043.7 | 40.29
0 | 0 | 187.3 | 40.62 | 100418.2 | 40.5
0 | 0 | 186.1 | 40.98 | 101305 | 40.74
0 | 0 | 184.9 | 41.28 | 102063.5 | 41.12
0 | 0 | 183.8 | 41.68 | 103048.5 | 41.49
0 | 0 | 182.6 | 41.84 | 103438.8 | 41.67
0 | 0 | 181.3 | 42.07 | 104011.3 | 41.93
0 | 0 | 180.1 | 42.31 | 104611.2 | 42.17
0 | 0 | 178.9 | 42.69 | 105556.9 | 42.55
0 | 0 | 177.7 | 42.99 | 106274.2 | 42.83
0 | 0 | 176.5 | 43.26 | 106944.6 | 43.11
0 | 0 | 175.3 | 43.51 | 107565 | 43.35
0 | 0 | 174.1 | 43.78 | 108230.2 | 43.59
0 | 0 | 172.9 | 44.19 | 109247.9 | 43.96
0 | 0 | 171.7 | 44.48 | 109959.8 | 44.33
0 | 0 | 170.5 | 44.86 | 110910 | 44.65
0 | 0 | 169.3 | 45.07 | 111421.9 | 44.83
0 | 0 | 168.1 | 45.34 | 112092 | 45.2
0 | 0 | 166.9 | 45.72 | 113027.8 | 45.52
0 | 0 | 165.7 | 45.93 | 113556.4 | 45.71
0 | 0 | 164.5 | 46.25 | 114345.2 | 46.08
0 | 0 | 163.3 | 46.52 | 115004.8 | 46.38
0 | 0 | 162.1 | 46.8 | 115705.2 | 46.61
0 | 0 | 160.9 | 47.25 | 116827.8 | 47.11
0 | 0 | 159.7 | 47.66 | 117836 | 47.48
0 | 0 | 158.4 | 47.98 | 118612.3 | 47.76
0 | 0 | 157.3 | 48.34 | 119517 | 48.13
0 | 0 | 156.1 | 48.68 | 120354.3 | 48.49
0 | 0 | 154.8 | 48.98 | 121093 | 48.82
0 | 0 | 153.7 | 49.29 | 121856.8 | 49.07
0 | 0 | 152.4 | 49.59 | 122615.4 | 49.47 |  | Z | D
0 | 0 | 151.3 | 49.9 | 123363.6 | 49.67 | R50 | 151.3 | 49.67
0 | 0 | 150 | 50.22 | 124158 | 50.04 |  | 150 | 50.04
0 | 0 | 148.8 | 50.51 | 124879.1 | 50.33 |  | 150.1405 | 50
0 | 0 | 147.6 | 50.9 | 125839 | 50.78
0 | 0 | 146.4 | 51.35 | 126963.9 | 51.16
0 | 0 | 145.2 | 51.54 | 127434.7 | 51.4
0 | 0 | 144 | 52 | 128558.9 | 51.78
0 | 0 | 142.8 | 52.32 | 129356.2 | 52.15
0 | 0 | 141.6 | 52.76 | 130436.2 | 52.56
0 | 0 | 140.4 | 53.06 | 131177.6 | 52.8
0 | 0 | 139.2 | 53.43 | 132091.3 | 53.23
0 | 0 | 138 | 53.79 | 132981.6 | 53.64
0 | 0 | 136.8 | 54.07 | 133684.8 | 53.97
0 | 0 | 135.6 | 54.45 | 134625.9 | 54.31
0 | 0 | 134.4 | 54.87 | 135656.5 | 54.5
0 | 0 | 133.2 | 55.18 | 136424.4 | 54.95
0 | 0 | 132 | 55.74 | 137809.9 | 55.43
0 | 0 | 130.8 | 56.1 | 138705.1 | 55.84
0 | 0 | 129.6 | 56.4 | 139450.8 | 56.19
0 | 0 | 128.4 | 56.68 | 140124.9 | 56.5
0 | 0 | 127.2 | 57.13 | 141250.6 | 56.97
0 | 0 | 126 | 57.46 | 142067.4 | 57.3
0 | 0 | 124.8 | 57.91 | 143163.6 | 57.73
0 | 0 | 123.6 | 58.38 | 144327.6 | 58.22
0 | 0 | 122.4 | 58.68 | 145078.3 | 58.52
0 | 0 | 121.2 | 59.16 | 146271.3 | 58.87
0 | 0 | 120 | 59.4 | 146867.6 | 59.29
0 | 0 | 118.7 | 59.92 | 148151.2 | 59.71
0 | 0 | 117.5 | 60.37 | 149260.2 | 60.14
0 | 0 | 116.3 | 60.68 | 150031.6 | 60.66
0 | 0 | 115.1 | 61.03 | 150896.9 | 60.83
0 | 0 | 113.9 | 61.55 | 152176.7 | 61.33
0 | 0 | 112.7 | 62.14 | 153637.9 | 61.84
0 | 0 | 111.5 | 62.48 | 154471.5 | 62.09
0 | 0 | 110.3 | 62.87 | 155444.5 | 62.51
0 | 0 | 109.1 | 63.23 | 156317 | 62.93
0 | 0 | 107.9 | 63.57 | 157160.8 | 63.39
0 | 0 | 106.7 | 64.01 | 158263 | 63.82
0 | 0 | 105.5 | 64.56 | 159604.6 | 64.18
0 | 0 | 104.3 | 64.88 | 160406.1 | 64.54
0 | 0 | 103.1 | 65.21 | 161210.3 | 65.01
0 | 0 | 101.9 | 65.53 | 162001.2 | 65.31 |  | Z | D
0 | 0 | 100.7 | 66.16 | 163563.9 | 65.98 | D10 | 100.7 | 65.98
0 | 0 | 99.5 | 66.56 | 164558.7 | 66.38 |  | 99.5 | 66.38
0 | 0 | 98.3 | 67.02 | 165703.1 | 66.89 |  | 100 | 66.2133
0 | 0 | 97.1 | 67.59 | 167117 | 67.42
0 | 0 | 95.9 | 67.95 | 167994.1 | 67.68
0 | 0 | 94.7 | 68.47 | 169272.2 | 68.11
0 | 0 | 93.5 | 68.88 | 170307.1 | 68.68
0 | 0 | 92.3 | 69.27 | 171266.3 | 69
0 | 0 | 91.1 | 69.67 | 172252.7 | 69.41
0 | 0 | 89.9 | 70.16 | 173467.2 | 70.06
0 | 0 | 88.7 | 70.62 | 174608.9 | 70.41
0 | 0 | 87.5 | 71.16 | 175932.1 | 70.82
0 | 0 | 86.3 | 71.58 | 176974.7 | 71.37
0 | 0 | 85.1 | 72.11 | 178270.4 | 71.89
0 | 0 | 83.8 | 72.48 | 179189.5 | 72.44
0 | 0 | 82.7 | 72.93 | 180304.5 | 72.79
0 | 0 | 81.5 | 73.54 | 181815.1 | 73.21
0 | 0 | 80.2 | 74.06 | 183103.7 | 73.81
0 | 0 | 79 | 74.39 | 183907.3 | 74.15
0 | 0 | 77.8 | 74.79 | 184898.2 | 74.63
0 | 0 | 76.6 | 75.35 | 186298.8 | 75.18
0 | 0 | 75.4 | 75.77 | 187324.2 | 75.45
0 | 0 | 74.2 | 76.27 | 188573.3 | 76.01
0 | 0 | 73 | 76.67 | 189562.3 | 76.47
0 | 0 | 71.8 | 77.23 | 190927.7 | 77.01
0 | 0 | 70.6 | 77.78 | 192310.3 | 77.58
0 | 0 | 69.4 | 78.26 | 193494 | 78.02
0 | 0 | 68.2 | 78.76 | 194711 | 78.44
0 | 0 | 67 | 79.15 | 195677.2 | 78.83
0 | 0 | 65.7 | 79.68 | 197005.4 | 79.42
0 | 0 | 64.5 | 80.23 | 198366.7 | 79.94
0 | 0 | 63.3 | 80.83 | 199833.7 | 80.5
0 | 0 | 62.1 | 81.11 | 200531.3 | 80.84
0 | 0 | 60.9 | 81.65 | 201873.7 | 81.4
0 | 0 | 59.7 | 82.18 | 203177.8 | 82.03
0 | 0 | 58.5 | 82.67 | 204384.1 | 82.47
0 | 0 | 57.3 | 83.2 | 205704.3 | 82.91
0 | 0 | 56.1 | 83.68 | 206888.2 | 83.34
0 | 0 | 54.9 | 84.06 | 207823.1 | 83.89
0 | 0 | 53.7 | 84.56 | 209060.1 | 84.37
0 | 0 | 52.5 | 85.33 | 210956.4 | 85.03
0 | 0 | 51.3 | 85.74 | 211979.1 | 85.3
0 | 0 | 50 | 86.41 | 213634.3 | 86.09
0 | 0 | 48.8 | 86.83 | 214675.5 | 86.44
0 | 0 | 47.6 | 87.08 | 215297.2 | 86.93
0 | 0 | 46.4 | 87.56 | 216476.9 | 87.27
0 | 0 | 45.2 | 88.15 | 217948.6 | 87.95
0 | 0 | 44 | 88.63 | 219135.3 | 88.35
0 | 0 | 42.8 | 89.4 | 221038.4 | 89.13
0 | 0 | 41.6 | 89.76 | 221910.5 | 89.7
0 | 0 | 40.4 | 90.36 | 223399.4 | 90.29
0 | 0 | 39.2 | 91 | 224989.8 | 90.75
0 | 0 | 38 | 91.57 | 226388.7 | 91.23
0 | 0 | 36.8 | 92.01 | 227488.6 | 91.62
0 | 0 | 35.5 | 92.48 | 228639.2 | 92.25
0 | 0 | 34.3 | 93.16 | 230328.8 | 92.84
0 | 0 | 33.1 | 93.57 | 231346.5 | 93.3
0 | 0 | 31.9 | 94.17 | 232821.8 | 93.79
0 | 0 | 30.7 | 94.49 | 233618.4 | 94.18
0 | 0 | 29.5 | 94.99 | 234843.9 | 94.72
0 | 0 | 28.3 | 95.61 | 236391.8 | 95.39
0 | 0 | 27.1 | 96.04 | 237437.3 | 95.74
0 | 0 | 25.8 | 96.61 | 238864.8 | 96.2
0 | 0 | 24.7 | 97.06 | 239974.3 | 96.73
0 | 0 | 23.4 | 97.54 | 241147 | 97.28
0 | 0 | 22.2 | 98.2 | 242795.3 | 97.62
0 | 0 | 21.1 | 98.61 | 243808.9 | 98.3
0 | 0 | 19.8 | 99.21 | 245274.2 | 98.68
0 | 0 | 18.6 | 99.59 | 246211.4 | 99.28
0 | 0 | 17.4 | 99.63 | 246325 | 99.49
0 | 0 | 16.2 | 99.94 | 247084.9 | 99.69
0 | 0 | 15 | 100.28 | 247931.2 | 99.99
0 | 0 | 13.8 | 100.31 | 247992.5 | 100
0 | 0 | 12.6 | 100.02 | 247282.4 | 99.86
0 | 0 | 11.4 | 99.63 | 246311.4 | 99.28
0 | 0 | 10.2 | 98.67 | 243937.9 | 98.3
0 | 0 | 9 | 97.17 | 240231.4 | 96.82
0 | 0 | 7.8 | 94.81 | 234398.9 | 94.52
0 | 0 | 6.6 | 91.55 | 226344.7 | 91.32
0 | 0 | 5.4 | 86.21 | 213132.1 | 85.8
0 | 0 | 4.2 | 78.68 | 194527.2 | 78.53
0 | 0 | 2.9 | 68.76 | 170008.7 | 68.5
0 | 0 | 1.7 | 58.88 | 145570.7 | 58.67
0 | 0 | 0.6 | 53.86 | 133167.5 | 53.65
0 | 0 | -0.5 | 50.66 | 125253.9 | 50.52