## 8_Rendements taille de champ
Measurement time: | 12/08/2022 18:54:12
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 20.25 | 50053.1 | 20.2
0 | 0 | 309.1 | 20.35 | 50304.8 | 20.29
0 | 0 | 307.6 | 20.68 | 51118.5 | 20.61
0 | 0 | 306.6 | 20.74 | 51270.8 | 20.69
0 | 0 | 305.3 | 20.87 | 51599.4 | 20.8
0 | 0 | 304.1 | 20.95 | 51791.2 | 20.9
0 | 0 | 302.9 | 21.03 | 51986.8 | 20.94
0 | 0 | 301.7 | 21.24 | 52513.5 | 21.19
0 | 0 | 300.5 | 21.44 | 53017.3 | 21.4
0 | 0 | 299.3 | 21.58 | 53353.7 | 21.58
0 | 0 | 298.1 | 21.73 | 53713.2 | 21.68
0 | 0 | 296.9 | 21.8 | 53885.9 | 21.77
0 | 0 | 295.7 | 22.01 | 54426.3 | 21.98
0 | 0 | 294.5 | 22.16 | 54780.9 | 22.14
0 | 0 | 293.3 | 22.3 | 55137.8 | 22.27
0 | 0 | 292.1 | 22.49 | 55597.2 | 22.46
0 | 0 | 290.9 | 22.61 | 55905 | 22.58
0 | 0 | 289.7 | 22.71 | 56146.2 | 22.72
0 | 0 | 288.5 | 22.87 | 56546.2 | 22.84
0 | 0 | 287.3 | 23.04 | 56963.4 | 22.96
0 | 0 | 286.1 | 23.22 | 57407 | 23.2
0 | 0 | 284.8 | 23.37 | 57783.1 | 23.35
0 | 0 | 283.7 | 23.54 | 58196.3 | 23.49
0 | 0 | 282.5 | 23.68 | 58536.8 | 23.63
0 | 0 | 281.2 | 23.92 | 59134.2 | 23.9
0 | 0 | 280.1 | 24.05 | 59470.6 | 24.03
0 | 0 | 278.8 | 24.12 | 59635.2 | 24.12
0 | 0 | 277.6 | 24.29 | 60047.2 | 24.25
0 | 0 | 276.4 | 24.51 | 60594.3 | 24.49
0 | 0 | 275.2 | 24.69 | 61040.6 | 24.66
0 | 0 | 274 | 24.91 | 61579.7 | 24.86
0 | 0 | 272.8 | 24.94 | 61672.1 | 24.92
0 | 0 | 271.6 | 25.15 | 62189.2 | 25.11
0 | 0 | 270.4 | 25.34 | 62640 | 25.34
0 | 0 | 269.2 | 25.5 | 63048 | 25.49
0 | 0 | 268 | 25.67 | 63477.5 | 25.61
0 | 0 | 266.8 | 25.8 | 63781.9 | 25.76
0 | 0 | 265.6 | 26.05 | 64405.5 | 26.01
0 | 0 | 264.4 | 26.25 | 64888.7 | 26.27
0 | 0 | 263.2 | 26.41 | 65291.7 | 26.41
0 | 0 | 262 | 26.54 | 65620.8 | 26.51
0 | 0 | 260.8 | 26.74 | 66116.2 | 26.73
0 | 0 | 259.6 | 27.02 | 66809.6 | 26.95
0 | 0 | 258.4 | 27.18 | 67191 | 27.17
0 | 0 | 257.2 | 27.37 | 67672.3 | 27.34
0 | 0 | 256 | 27.56 | 68148.9 | 27.55
0 | 0 | 254.8 | 27.76 | 68631.2 | 27.71
0 | 0 | 253.6 | 27.86 | 68872.9 | 27.81
0 | 0 | 252.4 | 28.12 | 69534.2 | 28.11
0 | 0 | 251.2 | 28.21 | 69740.5 | 28.19
0 | 0 | 250 | 28.5 | 70452.4 | 28.49
0 | 0 | 248.8 | 28.64 | 70803.9 | 28.63
0 | 0 | 247.6 | 28.86 | 71345.3 | 28.82
0 | 0 | 246.4 | 29.03 | 71784.4 | 29.03
0 | 0 | 245.1 | 29.24 | 72295.9 | 29.2
0 | 0 | 243.9 | 29.45 | 72801.8 | 29.39
0 | 0 | 242.7 | 29.61 | 73211.8 | 29.59
0 | 0 | 241.5 | 29.82 | 73719.6 | 29.81
0 | 0 | 240.3 | 30.02 | 74228.3 | 29.95
0 | 0 | 239.1 | 30.34 | 75020.5 | 30.26
0 | 0 | 237.9 | 30.54 | 75516.3 | 30.55
0 | 0 | 236.7 | 30.64 | 75754.1 | 30.7
0 | 0 | 235.5 | 30.92 | 76443.9 | 30.88
0 | 0 | 234.3 | 31.16 | 77033.1 | 31.1
0 | 0 | 233.1 | 31.31 | 77412.2 | 31.26
0 | 0 | 231.9 | 31.51 | 77915.5 | 31.44
0 | 0 | 230.7 | 31.84 | 78709.4 | 31.77
0 | 0 | 229.5 | 31.89 | 78850.8 | 31.87
0 | 0 | 228.3 | 32.21 | 79644.5 | 32.21
0 | 0 | 227.1 | 32.46 | 80251.6 | 32.44
0 | 0 | 225.9 | 32.68 | 80800 | 32.62
0 | 0 | 224.7 | 32.89 | 81313.3 | 32.84
0 | 0 | 223.5 | 33.1 | 81845.3 | 33.08
0 | 0 | 222.2 | 33.41 | 82596 | 33.38
0 | 0 | 221.1 | 33.63 | 83139.7 | 33.6
0 | 0 | 219.9 | 33.75 | 83430.8 | 33.77
0 | 0 | 218.6 | 34 | 84061.1 | 34.03
0 | 0 | 217.5 | 34.29 | 84770.7 | 34.24
0 | 0 | 216.2 | 34.4 | 85052.4 | 34.38
0 | 0 | 215 | 34.81 | 86067.1 | 34.75
0 | 0 | 213.8 | 34.92 | 86331.2 | 34.89
0 | 0 | 212.6 | 35.18 | 86989.1 | 35.17
0 | 0 | 211.4 | 35.45 | 87644.1 | 35.42
0 | 0 | 210.2 | 35.65 | 88137.4 | 35.71
0 | 0 | 209 | 35.94 | 88851.9 | 35.87
0 | 0 | 207.8 | 36.14 | 89357 | 36.08
0 | 0 | 206.6 | 36.5 | 90236.1 | 36.43
0 | 0 | 205.4 | 36.52 | 90282.9 | 36.5
0 | 0 | 204.2 | 36.85 | 91117.6 | 36.85
0 | 0 | 203 | 37.22 | 92015 | 37.17
0 | 0 | 201.8 | 37.49 | 92681.5 | 37.41 | D20 | z | D
0 | 0 | 200.6 | 37.6 | 92949.6 | 37.57 |  | 200.6 | 37.57
0 | 0 | 199.4 | 37.9 | 93694.4 | 37.87 |  | 199.4 | 37.87
0 | 0 | 198.2 | 38.19 | 94409.6 | 38.12 |  | 200 | 37.72
0 | 0 | 197 | 38.59 | 95419.5 | 38.56
0 | 0 | 195.8 | 38.72 | 95724.9 | 38.65
0 | 0 | 194.6 | 39.04 | 96520.9 | 38.94
0 | 0 | 193.4 | 39.15 | 96800.1 | 39.13
0 | 0 | 192.2 | 39.54 | 97754.6 | 39.52
0 | 0 | 191 | 39.75 | 98263.7 | 39.67
0 | 0 | 189.8 | 40.12 | 99189.1 | 40.09
0 | 0 | 188.5 | 40.37 | 99801.8 | 40.34
0 | 0 | 187.3 | 40.68 | 100572.4 | 40.63
0 | 0 | 186.2 | 40.86 | 101012.7 | 40.86
0 | 0 | 185 | 41.19 | 101848.2 | 41.17
0 | 0 | 183.8 | 41.45 | 102486.4 | 41.44
0 | 0 | 182.5 | 41.77 | 103281.2 | 41.74
0 | 0 | 181.3 | 42.1 | 104074.1 | 41.98
0 | 0 | 180.1 | 42.49 | 105048.9 | 42.37
0 | 0 | 178.9 | 42.64 | 105416.5 | 42.61
0 | 0 | 177.7 | 43.03 | 106378.1 | 42.97
0 | 0 | 176.5 | 43.23 | 106883.2 | 43.22
0 | 0 | 175.3 | 43.58 | 107754.4 | 43.63
0 | 0 | 174.1 | 43.94 | 108624.5 | 43.91
0 | 0 | 172.9 | 44.25 | 109409 | 44.3
0 | 0 | 171.7 | 44.48 | 109972.2 | 44.44
0 | 0 | 170.5 | 44.74 | 110615 | 44.77
0 | 0 | 169.3 | 45.02 | 111300.5 | 45.03
0 | 0 | 168.1 | 45.32 | 112048.4 | 45.25
0 | 0 | 166.9 | 45.62 | 112799.9 | 45.61
0 | 0 | 165.7 | 46.02 | 113775.5 | 45.95
0 | 0 | 164.4 | 46.29 | 114432.8 | 46.25
0 | 0 | 163.3 | 46.6 | 115217.1 | 46.59
0 | 0 | 162 | 46.93 | 116017.9 | 46.86
0 | 0 | 160.9 | 47.24 | 116799.6 | 47.15
0 | 0 | 159.6 | 47.57 | 117610.8 | 47.55
0 | 0 | 158.4 | 47.96 | 118579.4 | 47.89
0 | 0 | 157.2 | 48.26 | 119304.3 | 48.18
0 | 0 | 156 | 48.57 | 120089.1 | 48.51
0 | 0 | 154.8 | 48.96 | 121053.6 | 48.92
0 | 0 | 153.6 | 49.31 | 121902.6 | 49.25
0 | 0 | 152.4 | 49.52 | 122437.6 | 49.57 | R50 | z | D
0 | 0 | 151.2 | 49.87 | 123289.8 | 49.8 |  | 151.2 | 49.8
0 | 0 | 150 | 50.15 | 123998.1 | 50.09 |  | 150 | 50.09
0 | 0 | 148.8 | 50.61 | 125129.7 | 50.54 |  | 151.1517 | 50
0 | 0 | 147.6 | 50.89 | 125806.2 | 50.84
0 | 0 | 146.4 | 51.23 | 126658.6 | 51.21
0 | 0 | 145.2 | 51.53 | 127404 | 51.43
0 | 0 | 144 | 51.93 | 128392.8 | 51.85
0 | 0 | 142.8 | 52.25 | 129185.3 | 52.2
0 | 0 | 141.6 | 52.66 | 130204.7 | 52.56
0 | 0 | 140.4 | 52.95 | 130915.7 | 52.85
0 | 0 | 139.2 | 53.3 | 131771.7 | 53.24
0 | 0 | 138 | 53.61 | 132550.8 | 53.61
0 | 0 | 136.8 | 54.07 | 133672.7 | 53.99
0 | 0 | 135.6 | 54.62 | 135028.8 | 54.5
0 | 0 | 134.4 | 54.92 | 135790.6 | 54.77
0 | 0 | 133.2 | 55.21 | 136504.2 | 55.13
0 | 0 | 132 | 55.64 | 137554.6 | 55.55
0 | 0 | 130.7 | 55.93 | 138289.3 | 55.87
0 | 0 | 129.6 | 56.27 | 139109.8 | 56.25
0 | 0 | 128.3 | 56.65 | 140062.1 | 56.64
0 | 0 | 127.1 | 56.89 | 140650.9 | 56.91
0 | 0 | 125.9 | 57.31 | 141698.2 | 57.34
0 | 0 | 124.7 | 57.72 | 142705.7 | 57.75
0 | 0 | 123.5 | 58.15 | 143766.1 | 58.15
0 | 0 | 122.3 | 58.53 | 144718.6 | 58.43
0 | 0 | 121.1 | 58.99 | 145855.7 | 58.86
0 | 0 | 119.9 | 59.48 | 147054 | 59.29
0 | 0 | 118.7 | 59.88 | 148033.6 | 59.73
0 | 0 | 117.5 | 60.2 | 148826.1 | 60.08
0 | 0 | 116.3 | 60.67 | 150005 | 60.55
0 | 0 | 115.1 | 60.96 | 150726.3 | 60.88
0 | 0 | 113.9 | 61.53 | 152114.6 | 61.44
0 | 0 | 112.7 | 61.98 | 153235.2 | 61.76
0 | 0 | 111.5 | 62.25 | 153895.6 | 62.19
0 | 0 | 110.3 | 62.68 | 154974.3 | 62.56
0 | 0 | 109.1 | 62.96 | 155670.3 | 62.86
0 | 0 | 107.9 | 63.47 | 156918.3 | 63.42
0 | 0 | 106.7 | 64.01 | 158263.3 | 63.91
0 | 0 | 105.5 | 64.45 | 159355.1 | 64.33
0 | 0 | 104.2 | 64.89 | 160432.7 | 64.84
0 | 0 | 103.1 | 65.28 | 161395.6 | 65.2
0 | 0 | 101.8 | 65.63 | 162249.8 | 65.6 | D10 | z | D
0 | 0 | 100.6 | 66.2 | 163681.8 | 66.08 |  | 100.6 | 66.08
0 | 0 | 99.5 | 66.52 | 164455.2 | 66.47 |  | 99.5 | 66.47
0 | 0 | 98.3 | 67.02 | 165686.6 | 66.97 |  | 100 | 66.2927
0 | 0 | 97.1 | 67.49 | 166855.4 | 67.4
0 | 0 | 95.9 | 67.83 | 167708.1 | 67.79
0 | 0 | 94.6 | 68.31 | 168893.5 | 68.22
0 | 0 | 93.5 | 68.85 | 170212.6 | 68.72
0 | 0 | 92.2 | 69.37 | 171504.3 | 69.18
0 | 0 | 91.1 | 69.77 | 172487.8 | 69.6
0 | 0 | 89.8 | 70.16 | 173467.3 | 70.15
0 | 0 | 88.6 | 70.49 | 174278.6 | 70.4
0 | 0 | 87.4 | 71.11 | 175804.9 | 71
0 | 0 | 86.2 | 71.62 | 177078.6 | 71.48
0 | 0 | 85 | 71.89 | 177741.1 | 71.71
0 | 0 | 83.8 | 72.4 | 178986.2 | 72.42
0 | 0 | 82.6 | 73.01 | 180507.3 | 73.04
0 | 0 | 81.4 | 73.43 | 181549.3 | 73.35
0 | 0 | 80.2 | 73.78 | 182410.2 | 73.81
0 | 0 | 79 | 74.37 | 183866.7 | 74.22
0 | 0 | 77.8 | 74.82 | 184989.5 | 74.7
0 | 0 | 76.6 | 75.39 | 186402.5 | 75.26
0 | 0 | 75.4 | 75.72 | 187210.5 | 75.62
0 | 0 | 74.1 | 76.35 | 188774.1 | 76.36
0 | 0 | 73 | 76.67 | 189557.2 | 76.62
0 | 0 | 71.7 | 77.22 | 190919.1 | 77.18
0 | 0 | 70.5 | 77.64 | 191948.1 | 77.47
0 | 0 | 69.3 | 78.36 | 193721.7 | 78.18
0 | 0 | 68.1 | 78.76 | 194733 | 78.51
0 | 0 | 66.9 | 79.21 | 195830.3 | 79.03
0 | 0 | 65.7 | 79.55 | 196672.4 | 79.53
0 | 0 | 64.5 | 79.98 | 197749.5 | 80.06
0 | 0 | 63.3 | 80.66 | 199418 | 80.6
0 | 0 | 62.1 | 81.34 | 201102.7 | 81.37
0 | 0 | 60.9 | 81.6 | 201736.2 | 81.5
0 | 0 | 59.7 | 82.29 | 203437.6 | 82.19
0 | 0 | 58.5 | 82.71 | 204477.6 | 82.47
0 | 0 | 57.3 | 83.37 | 206114.6 | 83.15
0 | 0 | 56 | 83.81 | 207212.5 | 83.78
0 | 0 | 54.8 | 84.26 | 208309.6 | 84.17
0 | 0 | 53.6 | 84.78 | 209605 | 84.68
0 | 0 | 52.4 | 85.32 | 210936.9 | 85.15
0 | 0 | 51.2 | 85.77 | 212064.9 | 85.74
0 | 0 | 50 | 86.37 | 213532.2 | 86.25
0 | 0 | 48.8 | 86.78 | 214542.2 | 86.68
0 | 0 | 47.6 | 87.38 | 216038.6 | 87.24
0 | 0 | 46.4 | 87.7 | 216829.8 | 87.66
0 | 0 | 45.2 | 88.28 | 218248.8 | 88.3
0 | 0 | 44 | 88.7 | 219295.7 | 88.8
0 | 0 | 42.8 | 89.1 | 220297.8 | 89.11
0 | 0 | 41.6 | 89.75 | 221884.7 | 89.7
0 | 0 | 40.4 | 90.26 | 223146.9 | 90.18
0 | 0 | 39.2 | 90.83 | 224551.5 | 90.71
0 | 0 | 38 | 91.48 | 226160.8 | 91.33
0 | 0 | 36.7 | 91.87 | 227127.9 | 91.71
0 | 0 | 35.5 | 92.38 | 228403 | 92.43
0 | 0 | 34.3 | 92.87 | 229605.9 | 92.82
0 | 0 | 33.1 | 93.42 | 230962.4 | 93.15
0 | 0 | 31.9 | 94.01 | 232418.5 | 93.9
0 | 0 | 30.7 | 94.39 | 233359.9 | 94.33
0 | 0 | 29.5 | 94.92 | 234681 | 94.98
0 | 0 | 28.3 | 95.54 | 236196.2 | 95.42
0 | 0 | 27.1 | 96.24 | 237945.4 | 95.87
0 | 0 | 25.9 | 96.58 | 238772.5 | 96.42
0 | 0 | 24.7 | 96.9 | 239569.2 | 96.88
0 | 0 | 23.5 | 97.44 | 240917.7 | 97.6
0 | 0 | 22.3 | 97.9 | 242037.5 | 97.85
0 | 0 | 21 | 98.4 | 243283.7 | 98.34
0 | 0 | 19.8 | 98.76 | 244173.8 | 98.7
0 | 0 | 18.6 | 99.23 | 245327.4 | 99.28
0 | 0 | 17.4 | 99.51 | 246015.6 | 99.66
0 | 0 | 16.2 | 99.83 | 246806.3 | 99.86
0 | 0 | 15 | 100.05 | 247363.4 | 100
0 | 0 | 13.8 | 100.23 | 247804.1 | 99.96
0 | 0 | 12.6 | 99.9 | 246994.4 | 99.72
0 | 0 | 11.4 | 99.64 | 246347.1 | 99.41
0 | 0 | 10.2 | 98.62 | 243834.6 | 98.58
0 | 0 | 9 | 97.23 | 240376.9 | 97.06
0 | 0 | 7.8 | 94.71 | 234144.2 | 94.75
0 | 0 | 6.6 | 91.68 | 226660.7 | 91.52
0 | 0 | 5.4 | 86.23 | 213188.9 | 86.14
0 | 0 | 4.2 | 78.73 | 194657.8 | 78.67
0 | 0 | 2.9 | 68.16 | 168516.7 | 68.12
0 | 0 | 1.7 | 58.98 | 145816.8 | 59.04
0 | 0 | 0.5 | 53.75 | 132881.1 | 53.64
0 | 0 | -0.5 | 50.5 | 124865 | 50.46
Measurement time: | 12/08/2022 18:59:41
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 16.29 | 40265.3 | 17.01
0 | 0 | 309.1 | 16.54 | 40882.8 | 17.43
0 | 0 | 307.6 | 16.76 | 41425.7 | 17.67
0 | 0 | 306.6 | 16.78 | 41489.7 | 17.73
0 | 0 | 305.3 | 16.94 | 41887.5 | 17.87
0 | 0 | 304.1 | 17.1 | 42273.9 | 18.03
0 | 0 | 302.9 | 17.17 | 42450.3 | 18.12
0 | 0 | 301.7 | 17.31 | 42805.8 | 18.27
0 | 0 | 300.5 | 17.44 | 43124.6 | 18.39
0 | 0 | 299.3 | 17.58 | 43474.5 | 18.57
0 | 0 | 298.1 | 17.72 | 43814 | 18.68
0 | 0 | 296.9 | 17.8 | 44017.9 | 18.77
0 | 0 | 295.7 | 17.89 | 44240.4 | 18.92
0 | 0 | 294.5 | 18.14 | 44844.3 | 19.18
0 | 0 | 293.3 | 18.17 | 44921.2 | 19.19
0 | 0 | 292.1 | 18.32 | 45304.9 | 19.36
0 | 0 | 290.9 | 18.5 | 45740.3 | 19.53
0 | 0 | 289.7 | 18.52 | 45787.4 | 19.57
0 | 0 | 288.5 | 18.7 | 46244.5 | 19.78
0 | 0 | 287.3 | 18.89 | 46691.6 | 19.94
0 | 0 | 286.1 | 19.02 | 47025.9 | 20.08
0 | 0 | 284.9 | 19.13 | 47296.9 | 20.2
0 | 0 | 283.7 | 19.25 | 47580.8 | 20.33
0 | 0 | 282.5 | 19.37 | 47897.8 | 20.45
0 | 0 | 281.3 | 19.55 | 48327 | 20.63
0 | 0 | 280.1 | 19.66 | 48612.1 | 20.75
0 | 0 | 278.9 | 19.92 | 49236.9 | 21
0 | 0 | 277.7 | 19.94 | 49290.9 | 21.03
0 | 0 | 276.5 | 20.05 | 49564.9 | 21.17
0 | 0 | 275.2 | 20.22 | 49999.6 | 21.34
0 | 0 | 274.1 | 20.41 | 50464.5 | 21.56
0 | 0 | 272.9 | 20.49 | 50670 | 21.65
0 | 0 | 271.7 | 20.69 | 51163.2 | 21.85
0 | 0 | 270.5 | 20.89 | 51642.5 | 22.07
0 | 0 | 269.2 | 21 | 51931.3 | 22.19
0 | 0 | 268.1 | 21.12 | 52207.6 | 22.31
0 | 0 | 266.8 | 21.33 | 52726.8 | 22.58
0 | 0 | 265.6 | 21.4 | 52897.2 | 22.59
0 | 0 | 264.4 | 21.48 | 53114.3 | 22.68
0 | 0 | 263.2 | 21.73 | 53713.7 | 22.93
0 | 0 | 262 | 21.93 | 54213.2 | 23.17
0 | 0 | 260.8 | 22.09 | 54617.4 | 23.36
0 | 0 | 259.6 | 22.21 | 54912.6 | 23.45
0 | 0 | 258.4 | 22.37 | 55304.1 | 23.67
0 | 0 | 257.2 | 22.56 | 55780.2 | 23.89
0 | 0 | 256 | 22.76 | 56278.7 | 24.04
0 | 0 | 254.8 | 22.89 | 56585.7 | 24.16
0 | 0 | 253.6 | 23.02 | 56921.6 | 24.35
0 | 0 | 252.4 | 23.17 | 57285.8 | 24.5
0 | 0 | 251.2 | 23.33 | 57689 | 24.65
0 | 0 | 250 | 23.47 | 58031.2 | 24.83
0 | 0 | 248.8 | 23.71 | 58613.1 | 25.05
0 | 0 | 247.6 | 23.87 | 59026.9 | 25.22
0 | 0 | 246.4 | 24.08 | 59527.7 | 25.45
0 | 0 | 245.2 | 24.22 | 59890.1 | 25.59
0 | 0 | 244 | 24.44 | 60427.2 | 25.8
0 | 0 | 242.7 | 24.57 | 60740.2 | 25.98
0 | 0 | 241.6 | 24.74 | 61170.1 | 26.13
0 | 0 | 240.3 | 24.89 | 61541.2 | 26.37
0 | 0 | 239.1 | 25.12 | 62097.5 | 26.53
0 | 0 | 237.9 | 25.33 | 62619.1 | 26.8
0 | 0 | 236.7 | 25.5 | 63035.9 | 26.95
0 | 0 | 235.5 | 25.64 | 63403 | 27.09
0 | 0 | 234.3 | 25.85 | 63910.4 | 27.32
0 | 0 | 233.1 | 26.08 | 64478.4 | 27.54
0 | 0 | 231.9 | 26.19 | 64759.1 | 27.69
0 | 0 | 230.7 | 26.48 | 65460.5 | 28.05
0 | 0 | 229.5 | 26.56 | 65675.7 | 28.16
0 | 0 | 228.3 | 26.82 | 66312.4 | 28.42
0 | 0 | 227.1 | 26.99 | 66724.4 | 28.56
0 | 0 | 225.9 | 27.23 | 67311.9 | 28.81
0 | 0 | 224.7 | 27.39 | 67716.6 | 28.93
0 | 0 | 223.5 | 27.64 | 68335.4 | 29.19
0 | 0 | 222.3 | 27.81 | 68745.4 | 29.43
0 | 0 | 221.1 | 28.09 | 69460.2 | 29.72
0 | 0 | 219.9 | 28.22 | 69779.5 | 29.85
0 | 0 | 218.6 | 28.43 | 70280.2 | 30.04
0 | 0 | 217.5 | 28.66 | 70853.7 | 30.33
0 | 0 | 216.3 | 28.85 | 71315.5 | 30.5
0 | 0 | 215.1 | 29.03 | 71769.5 | 30.75
0 | 0 | 213.9 | 29.23 | 72260.9 | 30.95
0 | 0 | 212.6 | 29.48 | 72875.4 | 31.16
0 | 0 | 211.5 | 29.72 | 73487.1 | 31.42
0 | 0 | 210.2 | 29.84 | 73780.5 | 31.59
0 | 0 | 209 | 30.18 | 74607.9 | 31.92
0 | 0 | 207.8 | 30.37 | 75091.1 | 32.13
0 | 0 | 206.6 | 30.62 | 75703 | 32.37
0 | 0 | 205.4 | 30.81 | 76183.5 | 32.58
0 | 0 | 204.2 | 31.03 | 76708 | 32.82
0 | 0 | 203 | 31.2 | 77146.9 | 33.02
0 | 0 | 201.8 | 31.56 | 78017.1 | 33.43 | D20 | z | D
0 | 0 | 200.6 | 31.76 | 78524.3 | 33.62 |  | 200.6 | 33.62
0 | 0 | 199.4 | 31.95 | 78997.5 | 33.79 |  | 199.4 | 33.79
0 | 0 | 198.2 | 32.21 | 79645 | 34.09 |  | 200 | 33.705
0 | 0 | 197 | 32.48 | 80311.4 | 34.34
0 | 0 | 195.8 | 32.71 | 80877.9 | 34.63
0 | 0 | 194.6 | 32.9 | 81329.8 | 34.82
0 | 0 | 193.4 | 33.18 | 82039.6 | 35.16
0 | 0 | 192.2 | 33.41 | 82589.6 | 35.32
0 | 0 | 191 | 33.66 | 83215.4 | 35.59
0 | 0 | 189.8 | 33.89 | 83780.9 | 35.89
0 | 0 | 188.5 | 34.14 | 84409 | 36.15
0 | 0 | 187.4 | 34.3 | 84798.5 | 36.34
0 | 0 | 186.2 | 34.7 | 85789.2 | 36.69
0 | 0 | 185 | 34.86 | 86188.9 | 36.97
0 | 0 | 183.8 | 35.21 | 87049.5 | 37.25
0 | 0 | 182.5 | 35.33 | 87356.6 | 37.4
0 | 0 | 181.4 | 35.68 | 88202.8 | 37.8
0 | 0 | 180.1 | 35.96 | 88899.2 | 38.14
0 | 0 | 178.9 | 36.27 | 89660 | 38.38
0 | 0 | 177.8 | 36.53 | 90326.8 | 38.69
0 | 0 | 176.5 | 36.84 | 91088.2 | 39.01
0 | 0 | 175.3 | 37 | 91472.9 | 39.23
0 | 0 | 174.1 | 37.26 | 92124.4 | 39.46
0 | 0 | 172.9 | 37.58 | 92917.3 | 39.79
0 | 0 | 171.7 | 37.84 | 93550.2 | 40.05
0 | 0 | 170.5 | 38.22 | 94495.1 | 40.49
0 | 0 | 169.3 | 38.42 | 94983.5 | 40.73
0 | 0 | 168.1 | 38.74 | 95778.3 | 41.05
0 | 0 | 166.9 | 39.04 | 96511.3 | 41.42
0 | 0 | 165.7 | 39.27 | 97086.2 | 41.55
0 | 0 | 164.5 | 39.55 | 97785.7 | 41.93
0 | 0 | 163.3 | 39.91 | 98663.1 | 42.26
0 | 0 | 162.1 | 40.12 | 99186.4 | 42.45
0 | 0 | 160.9 | 40.52 | 100177.2 | 42.85
0 | 0 | 159.7 | 40.75 | 100760.1 | 43.2
0 | 0 | 158.4 | 41.06 | 101518.2 | 43.47
0 | 0 | 157.3 | 41.41 | 102382.7 | 43.85
0 | 0 | 156 | 41.67 | 103032.9 | 44.17
0 | 0 | 154.9 | 41.91 | 103624.3 | 44.45
0 | 0 | 153.6 | 42.24 | 104437 | 44.74
0 | 0 | 152.5 | 42.56 | 105234.1 | 45.06
0 | 0 | 151.3 | 42.86 | 105970.3 | 45.44
0 | 0 | 150 | 43.19 | 106771.1 | 45.77
0 | 0 | 148.8 | 43.55 | 107681.4 | 46.14
0 | 0 | 147.6 | 43.88 | 108498.3 | 46.46
0 | 0 | 146.4 | 44.18 | 109234.4 | 46.72
0 | 0 | 145.2 | 44.6 | 110278.3 | 47.24
0 | 0 | 144 | 44.88 | 110969.3 | 47.58
0 | 0 | 142.8 | 45.2 | 111751.4 | 47.86
0 | 0 | 141.6 | 45.56 | 112642.6 | 48.25
0 | 0 | 140.4 | 45.94 | 113575.6 | 48.74
0 | 0 | 139.2 | 46.28 | 114424.2 | 49.09
0 | 0 | 138 | 46.54 | 115075.1 | 49.34 | R50 | z | D
0 | 0 | 136.8 | 46.9 | 115949.3 | 49.65 |  | 136.8 | 49.65
0 | 0 | 135.6 | 47.26 | 116845.4 | 50.09 |  | 135.6 | 50.09
0 | 0 | 134.4 | 47.66 | 117821.8 | 50.36 |  | 136.6717 | 50
0 | 0 | 133.2 | 47.97 | 118607.3 | 50.87
0 | 0 | 132 | 48.44 | 119761.2 | 51.34
0 | 0 | 130.8 | 48.71 | 120421 | 51.56
0 | 0 | 129.6 | 49.06 | 121297.3 | 51.91
0 | 0 | 128.4 | 49.45 | 122254.6 | 52.32
0 | 0 | 127.2 | 49.76 | 123014.7 | 52.78
0 | 0 | 125.9 | 50.19 | 124092.6 | 53.28
0 | 0 | 124.7 | 50.54 | 124954 | 53.56
0 | 0 | 123.5 | 50.89 | 125815.5 | 53.92
0 | 0 | 122.3 | 51.34 | 126940.6 | 54.33
0 | 0 | 121.2 | 51.62 | 127625.6 | 54.71
0 | 0 | 119.9 | 52.08 | 128747.7 | 55.16
0 | 0 | 118.7 | 52.51 | 129815.6 | 55.57
0 | 0 | 117.5 | 52.81 | 130554 | 55.91
0 | 0 | 116.3 | 53.26 | 131680.8 | 56.48
0 | 0 | 115.1 | 53.68 | 132715.6 | 56.85
0 | 0 | 113.9 | 54.05 | 133619.8 | 57.26
0 | 0 | 112.7 | 54.52 | 134791 | 57.87
0 | 0 | 111.5 | 54.81 | 135502.6 | 58.11
0 | 0 | 110.3 | 55.28 | 136676.6 | 58.6
0 | 0 | 109.1 | 55.79 | 137939.7 | 59.06
0 | 0 | 107.9 | 56.13 | 138779 | 59.38
0 | 0 | 106.7 | 56.54 | 139793.2 | 59.9
0 | 0 | 105.5 | 56.87 | 140607.8 | 60.33
0 | 0 | 104.3 | 57.29 | 141641 | 60.79
0 | 0 | 103.1 | 57.81 | 142935.6 | 61.16
0 | 0 | 101.9 | 58.13 | 143725.5 | 61.5 | D10 | z | D
0 | 0 | 100.7 | 58.63 | 144950.4 | 62.08 |  | 100.7 | 62.08
0 | 0 | 99.5 | 59.11 | 146151.6 | 62.57 |  | 99.5 | 62.57
0 | 0 | 98.3 | 59.59 | 147338 | 63.23 |  | 100 | 62.3658
0 | 0 | 97.1 | 59.95 | 148224.9 | 63.54
0 | 0 | 95.8 | 60.24 | 148932.2 | 63.84
0 | 0 | 94.7 | 60.76 | 150212.6 | 64.4
0 | 0 | 93.5 | 61.37 | 151727.4 | 65.02
0 | 0 | 92.2 | 61.65 | 152428.8 | 65.33
0 | 0 | 91.1 | 62.21 | 153798 | 65.84
0 | 0 | 89.8 | 62.69 | 154989.4 | 66.34
0 | 0 | 88.7 | 63.09 | 155981 | 66.89
0 | 0 | 87.4 | 63.63 | 157310.3 | 67.47
0 | 0 | 86.2 | 64.1 | 158484.5 | 67.99
0 | 0 | 85 | 64.48 | 159422.1 | 68.5
0 | 0 | 83.8 | 64.93 | 160520.1 | 68.91
0 | 0 | 82.6 | 65.5 | 161943.6 | 69.38
0 | 0 | 81.4 | 65.9 | 162916.5 | 69.79
0 | 0 | 80.2 | 66.56 | 164565.6 | 70.51
0 | 0 | 79 | 66.92 | 165461.9 | 70.79
0 | 0 | 77.8 | 67.39 | 166603.5 | 71.42
0 | 0 | 76.6 | 67.94 | 167962.3 | 71.98
0 | 0 | 75.4 | 68.43 | 169180.2 | 72.51
0 | 0 | 74.2 | 68.76 | 169996 | 72.95
0 | 0 | 73 | 69.25 | 171200.3 | 73.54
0 | 0 | 71.8 | 69.79 | 172545.8 | 74.05
0 | 0 | 70.6 | 70.39 | 174028.2 | 74.59
0 | 0 | 69.4 | 71.01 | 175567.3 | 75.16
0 | 0 | 68.1 | 71.39 | 176512.5 | 75.7
0 | 0 | 66.9 | 71.93 | 177839.8 | 76.22
0 | 0 | 65.7 | 72.27 | 178670.8 | 76.75
0 | 0 | 64.5 | 72.94 | 180327.2 | 77.42
0 | 0 | 63.3 | 73.56 | 181863.7 | 77.77
0 | 0 | 62.1 | 74.16 | 183341 | 78.55
0 | 0 | 60.9 | 74.48 | 184130.3 | 79.13
0 | 0 | 59.7 | 75.04 | 185531.8 | 79.58
0 | 0 | 58.5 | 75.65 | 187036 | 80.07
0 | 0 | 57.3 | 76.06 | 188057.6 | 80.69
0 | 0 | 56.1 | 76.66 | 189526.1 | 81.36
0 | 0 | 54.9 | 77.24 | 190966 | 81.79
0 | 0 | 53.6 | 77.8 | 192343 | 82.4
0 | 0 | 52.5 | 78.32 | 193622.3 | 82.89
0 | 0 | 51.2 | 78.97 | 195238.3 | 83.61
0 | 0 | 50 | 79.37 | 196238.4 | 84.23
0 | 0 | 48.8 | 80.03 | 197862.5 | 84.68
0 | 0 | 47.6 | 80.43 | 198863.4 | 85.39
0 | 0 | 46.4 | 81.27 | 200934.8 | 86.08
0 | 0 | 45.2 | 81.66 | 201887.7 | 86.52
0 | 0 | 44 | 82.11 | 202993.7 | 87.09
0 | 0 | 42.8 | 82.68 | 204406.9 | 87.72
0 | 0 | 41.6 | 83.24 | 205796.3 | 88.26
0 | 0 | 40.4 | 83.9 | 207432.8 | 88.71
0 | 0 | 39.2 | 84.41 | 208698.1 | 89.26
0 | 0 | 38 | 84.99 | 210136.8 | 90.05
0 | 0 | 36.7 | 85.49 | 211359.1 | 90.52
0 | 0 | 35.5 | 85.96 | 212513.6 | 91.09
0 | 0 | 34.3 | 86.61 | 214136.3 | 91.86
0 | 0 | 33.1 | 87.18 | 215548.5 | 92.44
0 | 0 | 31.9 | 87.66 | 216732.5 | 93.16
0 | 0 | 30.7 | 88.4 | 218557.7 | 93.63
0 | 0 | 29.5 | 88.81 | 219558.1 | 94.2
0 | 0 | 28.3 | 89.47 | 221199 | 95.03
0 | 0 | 27.1 | 90.07 | 222689.6 | 95.34
0 | 0 | 25.8 | 90.53 | 223812.3 | 95.96
0 | 0 | 24.7 | 91.08 | 225177.7 | 96.53
0 | 0 | 23.4 | 91.64 | 226555.2 | 97.03
0 | 0 | 22.2 | 92.09 | 227686.1 | 97.56
0 | 0 | 21 | 92.62 | 228992 | 97.96
0 | 0 | 19.8 | 92.94 | 229789.1 | 98.56
0 | 0 | 18.6 | 93.51 | 231184.1 | 99.26
0 | 0 | 17.4 | 93.84 | 231994.4 | 99.35
0 | 0 | 16.2 | 94.3 | 233153.1 | 99.85
0 | 0 | 15 | 94.53 | 233720.7 | 100
0 | 0 | 13.8 | 94.18 | 232851.6 | 99.82
0 | 0 | 12.6 | 93.77 | 231837.6 | 99.55
0 | 0 | 11.4 | 93.24 | 230520.1 | 98.97
0 | 0 | 10.2 | 92.32 | 228235.1 | 97.73
0 | 0 | 9 | 90.6 | 223984.7 | 95.93
0 | 0 | 7.8 | 87.97 | 217493.2 | 93.13
0 | 0 | 6.6 | 84.21 | 208207.4 | 89.14
0 | 0 | 5.3 | 78.78 | 194776 | 83.67
0 | 0 | 4.1 | 71.18 | 175970.3 | 75.46
0 | 0 | 2.9 | 60.49 | 149542.8 | 64.11
0 | 0 | 1.7 | 51.04 | 126190.1 | 54.16
0 | 0 | 0.5 | 46.29 | 114433.7 | 49.11
0 | 0 | -0.5 | 43.11 | 106574.1 | 45.77
Measurement time: | 12/08/2022 19:04:06
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 24.63 | 60901.6 | 22.75
0 | 0 | 309.1 | 23.53 | 58183.6 | 22.6
0 | 0 | 307.6 | 23.71 | 58631.3 | 22.79
0 | 0 | 306.6 | 23.83 | 58906.6 | 22.87
0 | 0 | 305.3 | 24.01 | 59350.3 | 23.03
0 | 0 | 304.1 | 24.17 | 59756.9 | 23.19
0 | 0 | 302.9 | 24.39 | 60312 | 23.41
0 | 0 | 301.7 | 24.48 | 60520.1 | 23.51
0 | 0 | 300.5 | 24.62 | 60857.4 | 23.63
0 | 0 | 299.3 | 24.78 | 61276.3 | 23.8
0 | 0 | 298.1 | 24.94 | 61656.9 | 23.91
0 | 0 | 296.9 | 25.14 | 62149.1 | 24.14
0 | 0 | 295.7 | 25.28 | 62489.2 | 24.25
0 | 0 | 294.5 | 25.44 | 62885.4 | 24.38
0 | 0 | 293.3 | 25.62 | 63336.9 | 24.62
0 | 0 | 292.1 | 25.79 | 63755.7 | 24.77
0 | 0 | 290.9 | 25.95 | 64163.2 | 24.9
0 | 0 | 289.7 | 26.06 | 64422.8 | 25.01
0 | 0 | 288.5 | 26.19 | 64753.6 | 25.15
0 | 0 | 287.3 | 26.41 | 65298.7 | 25.34
0 | 0 | 286.1 | 26.58 | 65710.8 | 25.53
0 | 0 | 284.9 | 26.83 | 66326.3 | 25.76
0 | 0 | 283.7 | 26.98 | 66714 | 25.92
0 | 0 | 282.5 | 27.17 | 67185.2 | 26.09
0 | 0 | 281.3 | 27.27 | 67408.9 | 26.22
0 | 0 | 280.1 | 27.49 | 67960.4 | 26.4
0 | 0 | 278.9 | 27.71 | 68510 | 26.66
0 | 0 | 277.7 | 27.87 | 68914.3 | 26.77
0 | 0 | 276.5 | 28.04 | 69329.5 | 26.9
0 | 0 | 275.3 | 28.21 | 69740.7 | 27.09
0 | 0 | 274.1 | 28.36 | 70126.3 | 27.24
0 | 0 | 272.8 | 28.52 | 70504.2 | 27.4
0 | 0 | 271.7 | 28.73 | 71028 | 27.62
0 | 0 | 270.4 | 28.91 | 71470.6 | 27.74
0 | 0 | 269.2 | 29.07 | 71863.7 | 27.92
0 | 0 | 268.1 | 29.34 | 72533.4 | 28.21
0 | 0 | 266.8 | 29.44 | 72797 | 28.29
0 | 0 | 265.6 | 29.68 | 73390.6 | 28.48
0 | 0 | 264.4 | 29.82 | 73737.4 | 28.69
0 | 0 | 263.2 | 30.05 | 74290.1 | 28.85
0 | 0 | 262 | 30.27 | 74844.8 | 29.11
0 | 0 | 260.8 | 30.46 | 75311 | 29.23
0 | 0 | 259.6 | 30.61 | 75678.7 | 29.41
0 | 0 | 258.4 | 30.77 | 76086.4 | 29.59
0 | 0 | 257.2 | 30.95 | 76524.6 | 29.75
0 | 0 | 256 | 31.24 | 77225.4 | 30.05
0 | 0 | 254.8 | 31.38 | 77579.2 | 30.19
0 | 0 | 253.6 | 31.6 | 78125.4 | 30.37
0 | 0 | 252.4 | 31.71 | 78389.1 | 30.49
0 | 0 | 251.2 | 31.99 | 79081.5 | 30.73
0 | 0 | 250 | 32.2 | 79620.4 | 30.94
0 | 0 | 248.8 | 32.48 | 80299.4 | 31.19
0 | 0 | 247.6 | 32.73 | 80910.1 | 31.4
0 | 0 | 246.4 | 32.87 | 81278.1 | 31.58
0 | 0 | 245.2 | 33.11 | 81853.5 | 31.81
0 | 0 | 244 | 33.28 | 82287.2 | 32.02
0 | 0 | 242.8 | 33.54 | 82930.3 | 32.24
0 | 0 | 241.6 | 33.72 | 83374.7 | 32.42
0 | 0 | 240.3 | 33.91 | 83838.5 | 32.62
0 | 0 | 239.1 | 34.12 | 84357.1 | 32.8
0 | 0 | 238 | 34.4 | 85050.4 | 33.08
0 | 0 | 236.7 | 34.59 | 85526.3 | 33.26
0 | 0 | 235.5 | 34.86 | 86191.7 | 33.48
0 | 0 | 234.3 | 35.1 | 86770.3 | 33.69
0 | 0 | 233.1 | 35.23 | 87102.9 | 33.81
0 | 0 | 231.9 | 35.53 | 87840 | 34.13
0 | 0 | 230.7 | 35.82 | 88566.5 | 34.4
0 | 0 | 229.5 | 36.01 | 89022.8 | 34.62
0 | 0 | 228.3 | 36.17 | 89423.2 | 34.73
0 | 0 | 227.1 | 36.43 | 90079 | 35.01
0 | 0 | 225.9 | 36.63 | 90571.7 | 35.24
0 | 0 | 224.7 | 36.93 | 91314.3 | 35.54
0 | 0 | 223.5 | 37.22 | 92018.5 | 35.78
0 | 0 | 222.3 | 37.39 | 92444 | 35.95
0 | 0 | 221.1 | 37.64 | 93065.2 | 36.17
0 | 0 | 219.9 | 37.96 | 93852 | 36.46
0 | 0 | 218.6 | 38.23 | 94520.6 | 36.79
0 | 0 | 217.5 | 38.42 | 94988.6 | 36.97
0 | 0 | 216.3 | 38.64 | 95528.8 | 37.14
0 | 0 | 215.1 | 38.86 | 96079.8 | 37.36
0 | 0 | 213.9 | 39.21 | 96931.7 | 37.68
0 | 0 | 212.6 | 39.41 | 97439.7 | 37.87
0 | 0 | 211.4 | 39.6 | 97908.2 | 38.09
0 | 0 | 210.2 | 39.91 | 98666 | 38.32
0 | 0 | 209 | 40.02 | 98939.8 | 38.56
0 | 0 | 207.8 | 40.42 | 99923.6 | 38.88
0 | 0 | 206.6 | 40.53 | 100211.1 | 38.99
0 | 0 | 205.4 | 40.84 | 100982.5 | 39.23
0 | 0 | 204.2 | 41.01 | 101396.3 | 39.44
0 | 0 | 203 | 41.45 | 102466.7 | 39.82
0 | 0 | 201.8 | 41.67 | 103019.1 | 40.1 | D20 | z | D
0 | 0 | 200.6 | 41.99 | 103820.5 | 40.38 |  | 200.6 | 40.38
0 | 0 | 199.4 | 42.26 | 104474.6 | 40.64 |  | 199.4 | 40.64
0 | 0 | 198.2 | 42.47 | 105001.3 | 40.87 |  | 200 | 40.51
0 | 0 | 197 | 42.84 | 105904.2 | 41.21
0 | 0 | 195.8 | 43.01 | 106325.2 | 41.39
0 | 0 | 194.6 | 43.33 | 107125.8 | 41.72
0 | 0 | 193.4 | 43.63 | 107878.2 | 41.95
0 | 0 | 192.2 | 43.86 | 108425.8 | 42.2
0 | 0 | 191 | 44.07 | 108961 | 42.39
0 | 0 | 189.8 | 44.4 | 109760.9 | 42.74
0 | 0 | 188.6 | 44.79 | 110741.9 | 43
0 | 0 | 187.4 | 45.13 | 111565.7 | 43.4
0 | 0 | 186.2 | 45.32 | 112036.4 | 43.64
0 | 0 | 185 | 45.66 | 112898.1 | 43.89
0 | 0 | 183.8 | 45.88 | 113434.2 | 44.13
0 | 0 | 182.5 | 46.26 | 114379.7 | 44.57
0 | 0 | 181.4 | 46.47 | 114890.2 | 44.67
0 | 0 | 180.1 | 46.78 | 115645.5 | 45.01
0 | 0 | 178.9 | 47.16 | 116592.6 | 45.29
0 | 0 | 177.8 | 47.47 | 117364 | 45.67
0 | 0 | 176.5 | 47.68 | 117886.3 | 45.84
0 | 0 | 175.3 | 48.04 | 118773 | 46.18
0 | 0 | 174.1 | 48.25 | 119293.9 | 46.43
0 | 0 | 172.9 | 48.62 | 120209 | 46.77
0 | 0 | 171.7 | 48.87 | 120828.7 | 47.07
0 | 0 | 170.5 | 49.23 | 121704.7 | 47.36
0 | 0 | 169.3 | 49.49 | 122349.6 | 47.59
0 | 0 | 168.1 | 49.8 | 123126.6 | 47.91
0 | 0 | 166.9 | 50.24 | 124200.7 | 48.21
0 | 0 | 165.7 | 50.57 | 125026.6 | 48.63
0 | 0 | 164.5 | 50.73 | 125423.5 | 48.84
0 | 0 | 163.3 | 51.07 | 126268.6 | 49.03
0 | 0 | 162.1 | 51.37 | 127015.3 | 49.42 | R50 | z | D
0 | 0 | 160.9 | 51.81 | 128087.2 | 49.77 |  | 160.9 | 49.77
0 | 0 | 159.7 | 52.13 | 128872 | 50.21 |  | 159.7 | 50.21
0 | 0 | 158.5 | 52.47 | 129729.1 | 50.37 |  | 160.8157 | 50
0 | 0 | 157.3 | 52.72 | 130341.7 | 50.74
0 | 0 | 156.1 | 53.02 | 131077.7 | 50.98
0 | 0 | 154.9 | 53.51 | 132295.6 | 51.51
0 | 0 | 153.7 | 53.78 | 132953.6 | 51.71
0 | 0 | 152.5 | 54.13 | 133835.3 | 52.16
0 | 0 | 151.3 | 54.52 | 134793.6 | 52.37
0 | 0 | 150.1 | 54.91 | 135764 | 52.76
0 | 0 | 148.9 | 55.22 | 136534.9 | 53.08
0 | 0 | 147.7 | 55.6 | 137457.3 | 53.5
0 | 0 | 146.4 | 55.84 | 138058.4 | 53.78
0 | 0 | 145.3 | 56.2 | 138942.1 | 54.1
0 | 0 | 144 | 56.6 | 139940.2 | 54.44
0 | 0 | 142.9 | 56.85 | 140544.1 | 54.72
0 | 0 | 141.7 | 57.29 | 141645.5 | 55.04
0 | 0 | 140.4 | 57.81 | 142924.8 | 55.43
0 | 0 | 139.3 | 58.03 | 143473.2 | 55.75
0 | 0 | 138 | 58.35 | 144271.1 | 56.08
0 | 0 | 136.8 | 58.69 | 145112.7 | 56.44
0 | 0 | 135.6 | 59.12 | 146171.4 | 56.92
0 | 0 | 134.4 | 59.59 | 147322.8 | 57.34
0 | 0 | 133.2 | 59.85 | 147960.3 | 57.48
0 | 0 | 132 | 60.34 | 149174.7 | 58.05
0 | 0 | 130.8 | 60.73 | 150143 | 58.44
0 | 0 | 129.6 | 61.08 | 151017.5 | 58.7
0 | 0 | 128.4 | 61.49 | 152028 | 59.03
0 | 0 | 127.2 | 61.83 | 152855.7 | 59.5
0 | 0 | 126 | 62.07 | 153454.1 | 59.76
0 | 0 | 124.8 | 62.57 | 154703.6 | 60.23
0 | 0 | 123.6 | 63.01 | 155772.2 | 60.55
0 | 0 | 122.4 | 63.17 | 156178.9 | 60.8
0 | 0 | 121.2 | 63.71 | 157520.7 | 61.3
0 | 0 | 120 | 64.18 | 158683.1 | 61.71
0 | 0 | 118.8 | 64.47 | 159396.4 | 61.95
0 | 0 | 117.6 | 64.95 | 160589.6 | 62.41
0 | 0 | 116.4 | 65.26 | 161337.5 | 62.76
0 | 0 | 115.2 | 65.57 | 162108.7 | 63.1
0 | 0 | 114 | 65.91 | 162961.2 | 63.44
0 | 0 | 112.8 | 66.41 | 164200.5 | 63.88
0 | 0 | 111.6 | 66.75 | 165035.9 | 64.24
0 | 0 | 110.3 | 67.16 | 166050.3 | 64.66
0 | 0 | 109.2 | 67.61 | 167164.3 | 64.97
0 | 0 | 107.9 | 67.95 | 167992.8 | 65.35
0 | 0 | 106.7 | 68.46 | 169255.2 | 65.82
0 | 0 | 105.5 | 69.05 | 170727.7 | 66.31
0 | 0 | 104.3 | 69.44 | 171684.2 | 66.73
0 | 0 | 103.2 | 69.84 | 172663.4 | 67.21
0 | 0 | 101.9 | 70.26 | 173714.4 | 67.61 | D10 | z | D
0 | 0 | 100.7 | 70.69 | 174781.1 | 67.82 |  | 100.7 | 67.82
0 | 0 | 99.5 | 71.1 | 175785.7 | 68.36 |  | 99.5 | 68.36
0 | 0 | 98.3 | 71.47 | 176697.9 | 68.72 |  | 100 | 68.135
0 | 0 | 97.1 | 71.78 | 177464.8 | 69.05
0 | 0 | 95.9 | 72.29 | 178724.1 | 69.6
0 | 0 | 94.7 | 72.71 | 179762.2 | 69.96
0 | 0 | 93.5 | 73.28 | 181179.3 | 70.47
0 | 0 | 92.3 | 73.68 | 182168.3 | 70.87
0 | 0 | 91.1 | 74 | 182963.7 | 71.22
0 | 0 | 89.9 | 74.49 | 184177.2 | 71.71
0 | 0 | 88.7 | 75.13 | 185744.6 | 72.2
0 | 0 | 87.5 | 75.67 | 187078.1 | 72.62
0 | 0 | 86.3 | 76.14 | 188239.8 | 73.17
0 | 0 | 85.1 | 76.52 | 189175.7 | 73.62
0 | 0 | 83.8 | 76.92 | 190171.6 | 74.02
0 | 0 | 82.6 | 77.37 | 191291.8 | 74.5
0 | 0 | 81.4 | 77.68 | 192043.2 | 74.69
0 | 0 | 80.2 | 78.22 | 193384.5 | 75.18
0 | 0 | 79 | 78.69 | 194548 | 75.85
0 | 0 | 77.8 | 79.2 | 195813.8 | 76.18
0 | 0 | 76.6 | 79.75 | 197162.6 | 76.63
0 | 0 | 75.4 | 80.11 | 198059.4 | 77.02
0 | 0 | 74.2 | 80.72 | 199566.2 | 77.59
0 | 0 | 73 | 81.05 | 200377.9 | 78.07
0 | 0 | 71.8 | 81.59 | 201713.4 | 78.5
0 | 0 | 70.6 | 82.02 | 202782.2 | 78.93
0 | 0 | 69.4 | 82.3 | 203483.6 | 79.21
0 | 0 | 68.1 | 82.7 | 204456.5 | 79.46
0 | 0 | 66.9 | 83.14 | 205560.6 | 79.96
0 | 0 | 65.7 | 83.61 | 206711.9 | 80.6
0 | 0 | 64.6 | 83.9 | 207424.1 | 80.94
0 | 0 | 63.3 | 84.55 | 209031.7 | 81.5
0 | 0 | 62.1 | 85.31 | 210921.2 | 81.92
0 | 0 | 60.9 | 85.88 | 212315.9 | 82.52
0 | 0 | 59.7 | 86.3 | 213354.3 | 82.94
0 | 0 | 58.5 | 86.74 | 214442.3 | 83.5
0 | 0 | 57.3 | 87.17 | 215509 | 83.92
0 | 0 | 56.1 | 87.76 | 216971.1 | 84.45
0 | 0 | 54.9 | 88.37 | 218475.2 | 84.94
0 | 0 | 53.6 | 88.7 | 219296.6 | 85.49
0 | 0 | 52.5 | 89.22 | 220590.3 | 85.84
0 | 0 | 51.3 | 89.73 | 221837.4 | 86.35
0 | 0 | 50 | 90.13 | 222838.1 | 86.62
0 | 0 | 48.8 | 90.73 | 224307.1 | 87.19
0 | 0 | 47.6 | 91.27 | 225659.7 | 87.6
0 | 0 | 46.4 | 91.62 | 226518.7 | 88.11
0 | 0 | 45.2 | 91.89 | 227187.7 | 88.4
0 | 0 | 44 | 92.68 | 229139.3 | 88.98
0 | 0 | 42.8 | 93.14 | 230270.5 | 89.63
0 | 0 | 41.6 | 93.54 | 231252.6 | 90.1
0 | 0 | 40.4 | 94.09 | 232618.2 | 90.42
0 | 0 | 39.2 | 94.43 | 233469.2 | 90.85
0 | 0 | 38 | 95.07 | 235039.3 | 91.35
0 | 0 | 36.8 | 95.49 | 236090.4 | 91.9
0 | 0 | 35.5 | 96.17 | 237775.9 | 92.45
0 | 0 | 34.3 | 96.72 | 239118.6 | 92.86
0 | 0 | 33.1 | 97.03 | 239889.7 | 93.45
0 | 0 | 31.9 | 97.7 | 241536.5 | 94.13
0 | 0 | 30.7 | 98.16 | 242691.5 | 94.5
0 | 0 | 29.5 | 98.56 | 243678.2 | 94.84
0 | 0 | 28.3 | 99.04 | 244856 | 95.26
0 | 0 | 27.1 | 99.6 | 246254.4 | 95.85
0 | 0 | 25.9 | 100.21 | 247752.2 | 96.29
0 | 0 | 24.7 | 100.64 | 248809.3 | 96.9
0 | 0 | 23.5 | 101.17 | 250121.4 | 97.22
0 | 0 | 22.3 | 101.43 | 250759.8 | 97.81
0 | 0 | 21.1 | 102.12 | 252476.5 | 98.1
0 | 0 | 19.8 | 102.61 | 253683.9 | 98.54
0 | 0 | 18.7 | 102.88 | 254349.3 | 99.09
0 | 0 | 17.4 | 102.98 | 254602.8 | 99.16
0 | 0 | 16.2 | 103.32 | 255440.2 | 99.58
0 | 0 | 15 | 103.55 | 256004.8 | 99.77
0 | 0 | 13.8 | 103.93 | 256959.7 | 100
0 | 0 | 12.6 | 103.82 | 256679.2 | 99.86
0 | 0 | 11.4 | 103.69 | 256360.9 | 99.66
0 | 0 | 10.2 | 102.8 | 254163.4 | 99.15
0 | 0 | 9 | 101.6 | 251179 | 97.85
0 | 0 | 7.8 | 99.73 | 246571.7 | 95.94
0 | 0 | 6.6 | 96.77 | 239249.5 | 92.88
0 | 0 | 5.4 | 91.69 | 226690.8 | 88.25
0 | 0 | 4.2 | 85.23 | 210719.2 | 82.04
0 | 0 | 3 | 74.97 | 185360.2 | 72.17
0 | 0 | 1.8 | 65.58 | 162129.5 | 63.11
0 | 0 | 0.6 | 60.53 | 149659.8 | 58.25
0 | 0 | -0.5 | 57.11 | 141201.1 | 54.97
Measurement time: | 12/08/2022 19:07:57
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 200 x 200 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 23.97 | 59258.7 | 24.07
0 | 0 | 309.1 | 25.91 | 64060.4 | 24.28
0 | 0 | 307.6 | 26.2 | 64772.1 | 24.52
0 | 0 | 306.6 | 26.31 | 65057.1 | 24.62
0 | 0 | 305.3 | 26.51 | 65543.8 | 24.84
0 | 0 | 304.1 | 26.69 | 65996.7 | 25.01
0 | 0 | 302.9 | 26.83 | 66333.9 | 25.08
0 | 0 | 301.7 | 26.94 | 66594.4 | 25.25
0 | 0 | 300.5 | 27.11 | 67030.4 | 25.42
0 | 0 | 299.3 | 27.27 | 67408.7 | 25.54
0 | 0 | 298.1 | 27.48 | 67938.6 | 25.72
0 | 0 | 296.9 | 27.62 | 68289.4 | 25.86
0 | 0 | 295.7 | 27.82 | 68775.3 | 26.06
0 | 0 | 294.5 | 27.92 | 69029.3 | 26.13
0 | 0 | 293.3 | 28.18 | 69674.3 | 26.37
0 | 0 | 292.1 | 28.35 | 70086.6 | 26.49
0 | 0 | 290.9 | 28.45 | 70326.4 | 26.64
0 | 0 | 289.7 | 28.67 | 70872.6 | 26.82
0 | 0 | 288.5 | 28.88 | 71401.7 | 27.02
0 | 0 | 287.3 | 29.04 | 71797.8 | 27.21
0 | 0 | 286.1 | 29.15 | 72075.4 | 27.3
0 | 0 | 284.9 | 29.47 | 72859.3 | 27.58
0 | 0 | 283.7 | 29.6 | 73179.8 | 27.67
0 | 0 | 282.5 | 29.82 | 73727.8 | 27.95
0 | 0 | 281.3 | 30.01 | 74193.8 | 28.05
0 | 0 | 280.1 | 30.22 | 74714.8 | 28.3
0 | 0 | 278.8 | 30.34 | 75010.5 | 28.41
0 | 0 | 277.7 | 30.52 | 75451.3 | 28.58
0 | 0 | 276.4 | 30.7 | 75908.7 | 28.78
0 | 0 | 275.2 | 31.01 | 76656 | 29.06
0 | 0 | 274.1 | 31.08 | 76838.7 | 29.17
0 | 0 | 272.8 | 31.35 | 77511.1 | 29.38
0 | 0 | 271.6 | 31.5 | 77879 | 29.53
0 | 0 | 270.4 | 31.69 | 78351.4 | 29.7
0 | 0 | 269.2 | 31.88 | 78826.8 | 29.87
0 | 0 | 268 | 32.06 | 79274 | 29.99
0 | 0 | 266.8 | 32.27 | 79779.3 | 30.23
0 | 0 | 265.6 | 32.45 | 80235.1 | 30.4
0 | 0 | 264.4 | 32.59 | 80584 | 30.5
0 | 0 | 263.2 | 32.89 | 81305.4 | 30.74
0 | 0 | 262 | 33.16 | 81981.1 | 31.05
0 | 0 | 260.8 | 33.27 | 82248.7 | 31.15
0 | 0 | 259.6 | 33.39 | 82552.1 | 31.28
0 | 0 | 258.4 | 33.64 | 83181 | 31.55
0 | 0 | 257.2 | 33.87 | 83741.2 | 31.77
0 | 0 | 256 | 34.1 | 84296.1 | 31.95
0 | 0 | 254.8 | 34.27 | 84739.3 | 32.11
0 | 0 | 253.6 | 34.49 | 85269.4 | 32.28
0 | 0 | 252.4 | 34.7 | 85780.6 | 32.5
0 | 0 | 251.2 | 34.89 | 86252.7 | 32.68
0 | 0 | 250 | 35.05 | 86654.3 | 32.9
0 | 0 | 248.8 | 35.2 | 87032.8 | 32.98
0 | 0 | 247.6 | 35.55 | 87888.4 | 33.27
0 | 0 | 246.4 | 35.66 | 88154.4 | 33.36
0 | 0 | 245.2 | 35.95 | 88887.4 | 33.65
0 | 0 | 244 | 36.28 | 89698.7 | 33.95
0 | 0 | 242.7 | 36.33 | 89820.5 | 34.08
0 | 0 | 241.6 | 36.66 | 90640.7 | 34.33
0 | 0 | 240.3 | 36.92 | 91281 | 34.55
0 | 0 | 239.1 | 37.12 | 91763.4 | 34.74
0 | 0 | 237.9 | 37.41 | 92501.5 | 35.02
0 | 0 | 236.7 | 37.65 | 93083.1 | 35.2
0 | 0 | 235.5 | 37.83 | 93518.6 | 35.44
0 | 0 | 234.3 | 38.05 | 94081.8 | 35.67
0 | 0 | 233.1 | 38.33 | 94766.2 | 35.91
0 | 0 | 231.9 | 38.53 | 95253.5 | 36.07
0 | 0 | 230.7 | 38.82 | 95973.2 | 36.41
0 | 0 | 229.5 | 39.03 | 96496 | 36.51
0 | 0 | 228.3 | 39.23 | 96979.3 | 36.67
0 | 0 | 227.1 | 39.49 | 97630.3 | 36.98
0 | 0 | 225.9 | 39.72 | 98189.7 | 37.22
0 | 0 | 224.7 | 39.95 | 98781.7 | 37.46
0 | 0 | 223.4 | 40.27 | 99564 | 37.71
0 | 0 | 222.2 | 40.35 | 99771.5 | 37.79
0 | 0 | 221 | 40.66 | 100519.9 | 38.09
0 | 0 | 219.9 | 40.97 | 101303.7 | 38.3
0 | 0 | 218.6 | 41.17 | 101775.3 | 38.59
0 | 0 | 217.4 | 41.43 | 102424.9 | 38.82
0 | 0 | 216.2 | 41.68 | 103039.7 | 39
0 | 0 | 215 | 42.01 | 103868.3 | 39.34
0 | 0 | 213.8 | 42.25 | 104452 | 39.6
0 | 0 | 212.6 | 42.54 | 105167.2 | 39.8
0 | 0 | 211.4 | 42.7 | 105577.8 | 39.97
0 | 0 | 210.2 | 43 | 106322.8 | 40.24
0 | 0 | 209 | 43.24 | 106899.3 | 40.49
0 | 0 | 207.8 | 43.47 | 107482.5 | 40.78
0 | 0 | 206.6 | 43.87 | 108452.3 | 41.08
0 | 0 | 205.4 | 44.01 | 108811 | 41.16
0 | 0 | 204.2 | 44.25 | 109392.8 | 41.43
0 | 0 | 203 | 44.54 | 110106.5 | 41.72
0 | 0 | 201.8 | 44.83 | 110836.2 | 42.04 | D20 | z | D
0 | 0 | 200.6 | 45.12 | 111551.8 | 42.23 |  | 200.6 | 42.23
0 | 0 | 199.4 | 45.4 | 112243.7 | 42.47 |  | 199.4 | 42.47
0 | 0 | 198.2 | 45.7 | 112995.1 | 42.71 |  | 200 | 42.35
0 | 0 | 197 | 45.87 | 113408 | 42.95
0 | 0 | 195.8 | 46.34 | 114558.6 | 43.4
0 | 0 | 194.6 | 46.65 | 115330.8 | 43.63
0 | 0 | 193.4 | 46.93 | 116019.2 | 43.88
0 | 0 | 192.1 | 47.07 | 116366.6 | 44.1
0 | 0 | 190.9 | 47.36 | 117086.3 | 44.4
0 | 0 | 189.8 | 47.7 | 117943.4 | 44.7
0 | 0 | 188.5 | 48 | 118668.5 | 44.96
0 | 0 | 187.3 | 48.36 | 119553.3 | 45.28
0 | 0 | 186.1 | 48.57 | 120080.6 | 45.51
0 | 0 | 184.9 | 48.85 | 120785.9 | 45.68
0 | 0 | 183.7 | 49.14 | 121490.2 | 46.05
0 | 0 | 182.5 | 49.57 | 122565.3 | 46.45
0 | 0 | 181.3 | 49.79 | 123104.1 | 46.56
0 | 0 | 180.1 | 50.03 | 123681.4 | 46.88
0 | 0 | 178.9 | 50.33 | 124426.5 | 47.13
0 | 0 | 177.7 | 50.72 | 125391.5 | 47.53
0 | 0 | 176.5 | 51 | 126087.5 | 47.73
0 | 0 | 175.3 | 51.29 | 126809.9 | 48.02
0 | 0 | 174.1 | 51.62 | 127632.2 | 48.37
0 | 0 | 172.9 | 51.91 | 128332.3 | 48.68
0 | 0 | 171.7 | 52.2 | 129062.6 | 48.91
0 | 0 | 170.4 | 52.53 | 129878 | 49.17
0 | 0 | 169.3 | 52.86 | 130681.8 | 49.49 | R50 | z | D
0 | 0 | 168.1 | 53.04 | 131141.4 | 49.66 |  | 168.1 | 49.66
0 | 0 | 166.9 | 53.52 | 132312.4 | 50.07 |  | 166.9 | 50.07
0 | 0 | 165.7 | 53.81 | 133045.8 | 50.46 |  | 167.9838 | 50
0 | 0 | 164.4 | 54.17 | 133930.8 | 50.77
0 | 0 | 163.2 | 54.47 | 134669 | 51.07
0 | 0 | 162 | 54.86 | 135633.4 | 51.37
0 | 0 | 160.8 | 55.15 | 136346.6 | 51.68
0 | 0 | 159.6 | 55.51 | 137251.1 | 52.03
0 | 0 | 158.4 | 55.88 | 138143.1 | 52.27
0 | 0 | 157.2 | 56.17 | 138881.3 | 52.63
0 | 0 | 156 | 56.41 | 139458.2 | 52.86
0 | 0 | 154.8 | 56.87 | 140605.1 | 53.25
0 | 0 | 153.6 | 57.19 | 141393.7 | 53.54
0 | 0 | 152.4 | 57.43 | 141977.9 | 53.83
0 | 0 | 151.2 | 57.95 | 143277.3 | 54.28
0 | 0 | 150 | 58.2 | 143896.1 | 54.57
0 | 0 | 148.8 | 58.53 | 144710.9 | 54.89
0 | 0 | 147.6 | 58.93 | 145686.7 | 55.26
0 | 0 | 146.4 | 59.29 | 146588.2 | 55.52
0 | 0 | 145.2 | 59.66 | 147499.2 | 55.94
0 | 0 | 144 | 59.99 | 148318.4 | 56.22
0 | 0 | 142.8 | 60.28 | 149023.6 | 56.52
0 | 0 | 141.6 | 60.71 | 150105.2 | 56.79
0 | 0 | 140.3 | 61.01 | 150850 | 57.17
0 | 0 | 139.2 | 61.35 | 151681.4 | 57.62
0 | 0 | 137.9 | 61.7 | 152547.3 | 57.93
0 | 0 | 136.8 | 62.11 | 153551.8 | 58.06
0 | 0 | 135.6 | 62.51 | 154558.5 | 58.59
0 | 0 | 134.3 | 62.91 | 155526.7 | 58.93
0 | 0 | 133.1 | 63.41 | 156760.2 | 59.34
0 | 0 | 131.9 | 63.72 | 157546.4 | 59.68
0 | 0 | 130.7 | 64.04 | 158318.7 | 60.01
0 | 0 | 129.5 | 64.46 | 159360.3 | 60.44
0 | 0 | 128.3 | 64.85 | 160343.8 | 60.77
0 | 0 | 127.1 | 65.19 | 161183.1 | 61.09
0 | 0 | 125.9 | 65.47 | 161860.6 | 61.34
0 | 0 | 124.7 | 66.01 | 163199.3 | 61.85
0 | 0 | 123.5 | 66.15 | 163539.1 | 62.07
0 | 0 | 122.3 | 66.7 | 164898.8 | 62.53
0 | 0 | 121.1 | 67.06 | 165796.8 | 62.85
0 | 0 | 119.9 | 67.37 | 166553.6 | 63.26
0 | 0 | 118.7 | 67.79 | 167602.3 | 63.64
0 | 0 | 117.5 | 68.17 | 168550 | 64.12
0 | 0 | 116.3 | 68.69 | 169823 | 64.54
0 | 0 | 115.1 | 68.98 | 170535.9 | 64.85
0 | 0 | 113.9 | 69.47 | 171754.1 | 65.2
0 | 0 | 112.7 | 70.11 | 173332.2 | 65.67
0 | 0 | 111.5 | 70.4 | 174043.8 | 65.77
0 | 0 | 110.2 | 70.91 | 175302.3 | 66.29
0 | 0 | 109.1 | 71.33 | 176359.3 | 66.65
0 | 0 | 107.8 | 71.45 | 176641.7 | 66.88
0 | 0 | 106.7 | 71.68 | 177223.6 | 67.18
0 | 0 | 105.5 | 72.31 | 178768.8 | 67.73
0 | 0 | 104.2 | 72.72 | 179780.3 | 68.11
0 | 0 | 103.1 | 73.1 | 180728.4 | 68.42
0 | 0 | 101.8 | 73.43 | 181553 | 68.85 | D10 | z | D
0 | 0 | 100.6 | 73.79 | 182432.8 | 69.03 |  | 100.6 | 69.03
0 | 0 | 99.5 | 74.31 | 183714.5 | 69.71 |  | 99.5 | 69.71
0 | 0 | 98.2 | 74.84 | 185029.9 | 70.08 |  | 100 | 69.4009
0 | 0 | 97 | 75.34 | 186260.4 | 70.54
0 | 0 | 95.8 | 75.61 | 186944.4 | 70.88
0 | 0 | 94.6 | 76.02 | 187957.8 | 71.32
0 | 0 | 93.4 | 76.74 | 189715.9 | 71.8
0 | 0 | 92.2 | 77 | 190379.4 | 72.19
0 | 0 | 91 | 77.37 | 191290.5 | 72.49
0 | 0 | 89.8 | 77.82 | 192402.7 | 72.93
0 | 0 | 88.6 | 78.26 | 193489.7 | 73.31
0 | 0 | 87.4 | 78.78 | 194767.3 | 73.86
0 | 0 | 86.2 | 79.07 | 195483.8 | 74.13
0 | 0 | 85 | 79.69 | 197026.5 | 74.66
0 | 0 | 83.8 | 80.07 | 197952.5 | 75.09
0 | 0 | 82.6 | 80.45 | 198905.2 | 75.56
0 | 0 | 81.4 | 81.03 | 200323.8 | 75.79
0 | 0 | 80.2 | 81.4 | 201257.5 | 76.23
0 | 0 | 79 | 81.74 | 202101.1 | 76.46
0 | 0 | 77.7 | 82.42 | 203779.9 | 77.01
0 | 0 | 76.5 | 82.6 | 204222.9 | 77.33
0 | 0 | 75.3 | 83.15 | 205574 | 77.86
0 | 0 | 74.1 | 83.63 | 206762.4 | 78.34
0 | 0 | 72.9 | 84.06 | 207835.3 | 78.79
0 | 0 | 71.7 | 84.47 | 208849.6 | 79.16
0 | 0 | 70.5 | 84.93 | 209981.4 | 79.59
0 | 0 | 69.3 | 85.68 | 211823.6 | 80.11
0 | 0 | 68.1 | 85.96 | 212517.5 | 80.55
0 | 0 | 66.9 | 86.25 | 213251.5 | 80.91
0 | 0 | 65.7 | 86.72 | 214412.5 | 81.29
0 | 0 | 64.5 | 87.3 | 215833.4 | 81.84
0 | 0 | 63.3 | 87.81 | 217107.9 | 82.16
0 | 0 | 62.1 | 88.33 | 218371.3 | 82.64
0 | 0 | 60.9 | 88.83 | 219625.1 | 83.05
0 | 0 | 59.6 | 89.08 | 220224.8 | 83.45
0 | 0 | 58.4 | 89.51 | 221306.2 | 83.79
0 | 0 | 57.2 | 90.25 | 223125.7 | 84.49
0 | 0 | 56 | 90.64 | 224086.3 | 84.82
0 | 0 | 54.8 | 91.11 | 225267.3 | 85.35
0 | 0 | 53.6 | 91.59 | 226454 | 85.73
0 | 0 | 52.4 | 92.07 | 227632.5 | 86.16
0 | 0 | 51.2 | 92.5 | 228703.5 | 86.67
0 | 0 | 50 | 93.04 | 230021.2 | 87.1
0 | 0 | 48.8 | 93.34 | 230779.7 | 87.46
0 | 0 | 47.5 | 94.07 | 232564.1 | 87.92
0 | 0 | 46.4 | 94.4 | 233383.2 | 88.45
0 | 0 | 45.1 | 95.17 | 235294.5 | 89.05
0 | 0 | 43.9 | 95.39 | 235849.9 | 89.22
0 | 0 | 42.7 | 95.89 | 237072.1 | 89.96
0 | 0 | 41.5 | 96.36 | 238223.9 | 90.42
0 | 0 | 40.3 | 96.87 | 239484.4 | 90.84
0 | 0 | 39.1 | 97.32 | 240602.5 | 91.2
0 | 0 | 37.9 | 97.98 | 242241.6 | 91.63
0 | 0 | 36.7 | 98.45 | 243401.6 | 92.1
0 | 0 | 35.5 | 98.55 | 243652.6 | 92.31
0 | 0 | 34.3 | 98.98 | 244719.3 | 92.82
0 | 0 | 33.1 | 99.5 | 245994.1 | 93.24
0 | 0 | 31.8 | 100.23 | 247803.9 | 93.85
0 | 0 | 30.7 | 100.71 | 248987.5 | 94.32
0 | 0 | 29.4 | 101.3 | 250444.9 | 94.83
0 | 0 | 28.2 | 101.63 | 251269.9 | 95.04
0 | 0 | 27 | 101.97 | 252108.5 | 95.58
0 | 0 | 25.8 | 102.64 | 253750.4 | 96.04
0 | 0 | 24.6 | 103.01 | 254677.5 | 96.59
0 | 0 | 23.4 | 103.55 | 256010.1 | 97.03
0 | 0 | 22.2 | 104.06 | 257279 | 97.64
0 | 0 | 21 | 104.63 | 258687.6 | 97.98
0 | 0 | 19.8 | 105.03 | 259678.9 | 98.27
0 | 0 | 18.5 | 105.33 | 260401.2 | 98.75
0 | 0 | 17.3 | 105.59 | 261044.4 | 98.96
0 | 0 | 16.1 | 105.97 | 262005.8 | 99.27
0 | 0 | 14.9 | 106.41 | 263081.8 | 99.7
0 | 0 | 13.7 | 106.69 | 263782.3 | 99.92
0 | 0 | 12.5 | 106.67 | 263735.5 | 100
0 | 0 | 11.3 | 106.29 | 262778.8 | 99.8
0 | 0 | 10.1 | 105.6 | 261070.6 | 98.94
0 | 0 | 8.9 | 104.6 | 258604.6 | 97.91
0 | 0 | 7.7 | 103.05 | 254785.6 | 96.53
0 | 0 | 6.5 | 100.01 | 247271.2 | 93.8
0 | 0 | 5.3 | 95.75 | 236721.5 | 89.56
0 | 0 | 4.1 | 89.42 | 221088 | 83.67
0 | 0 | 2.9 | 79.67 | 196961.5 | 74.58
0 | 0 | 1.7 | 71.34 | 176389.8 | 66.77
0 | 0 | 0.4 | 66.23 | 163736.6 | 62.19
0 | 0 | -0.5 | 63.02 | 155808.4 | 59.06