## 9_Profils taille de champ
Measurement time: | 12/08/2022 19:10:04
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 200 x 200 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-150 | 0 | 100 | 5.96 | 14730.1 | 7.72
-149.1 | 0 | 100 | 6.15 | 15207.3 | 8.12
-147.7 | 0 | 100 | 6.32 | 15617.5 | 8.34
-146.6 | 0 | 100 | 6.46 | 15983 | 8.55
-145.4 | 0 | 100 | 6.62 | 16366.2 | 8.76
-144.1 | 0 | 100 | 6.74 | 16664.6 | 8.9
-142.9 | 0 | 100 | 6.96 | 17200.6 | 9.19
-141.7 | 0 | 100 | 7.16 | 17703.2 | 9.46
-140.5 | 0 | 100 | 7.31 | 18074.4 | 9.66
-139.3 | 0 | 100 | 7.44 | 18382.9 | 9.83
-138.1 | 0 | 100 | 7.69 | 19002.5 | 10.16
-136.9 | 0 | 100 | 7.84 | 19373 | 10.35
-135.7 | 0 | 100 | 8.03 | 19858.9 | 10.61
-134.5 | 0 | 100 | 8.21 | 20308.5 | 10.84
-133.3 | 0 | 100 | 8.44 | 20855.5 | 11.15
-132.1 | 0 | 100 | 8.73 | 21574.5 | 11.53
-130.9 | 0 | 100 | 8.87 | 21927.1 | 11.73
-129.7 | 0 | 100 | 9.17 | 22675.5 | 12.12
-128.5 | 0 | 100 | 9.45 | 23352 | 12.52
-127.3 | 0 | 100 | 9.67 | 23897.5 | 12.78
-126.1 | 0 | 100 | 9.98 | 24666 | 13.16
-124.8 | 0 | 100 | 10.27 | 25397.2 | 13.59
-123.7 | 0 | 100 | 10.53 | 26025.9 | 13.9
-122.5 | 0 | 100 | 10.86 | 26860.4 | 14.36
-121.2 | 0 | 100 | 11.32 | 27982.6 | 14.99
-120.1 | 0 | 100 | 11.73 | 29005.6 | 15.5
-118.8 | 0 | 100 | 12.22 | 30204.2 | 16.16
-117.6 | 0 | 100 | 12.92 | 31936.4 | 17.09
-116.5 | 0 | 100 | 13.94 | 34460.5 | 18.45
-115.2 | 0 | 100 | 15.76 | 38966.8 | 20.81
-114 | 0 | 100 | 19.03 | 47036.5 | 25.11
-112.8 | 0 | 100 | 23.5 | 58095.3 | 31.04
-111.6 | 0 | 100 | 29.65 | 73295.6 | 39.14
-110.4 | 0 | 100 | 37.33 | 92296.9 | 49.32
-109.2 | 0 | 100 | 45.89 | 113446.8 | 60.59
-108 | 0 | 100 | 53.83 | 133091.3 | 71.15
-106.8 | 0 | 100 | 60.55 | 149695.6 | 80.04
-105.6 | 0 | 100 | 64.85 | 160324.9 | 85.73
-104.4 | 0 | 100 | 67.19 | 166119.4 | 88.74
-103.2 | 0 | 100 | 68.46 | 169245.7 | 90.49
-102 | 0 | 100 | 69.44 | 171679.2 | 91.93
-100.8 | 0 | 100 | 70.03 | 173143.2 | 92.77
-99.6 | 0 | 100 | 70.39 | 174027.3 | 93.01
-98.4 | 0 | 100 | 70.98 | 175480.7 | 93.76
-97.2 | 0 | 100 | 71.2 | 176029.8 | 94.17
-96 | 0 | 100 | 71.59 | 176987.7 | 94.76
-94.8 | 0 | 100 | 71.79 | 177494 | 95.02
-93.6 | 0 | 100 | 72.11 | 178292.9 | 95.34
-92.4 | 0 | 100 | 72.38 | 178945 | 95.73
-91.1 | 0 | 100 | 72.5 | 179257.6 | 95.95
-90 | 0 | 100 | 72.76 | 179886 | 96.14
-88.7 | 0 | 100 | 72.84 | 180095 | 96.16
-87.5 | 0 | 100 | 73.15 | 180853.8 | 96.66
-86.3 | 0 | 100 | 73.27 | 181137.1 | 97
-85.1 | 0 | 100 | 73.39 | 181438.3 | 97.05
-83.9 | 0 | 100 | 73.55 | 181843.8 | 97.27
-82.7 | 0 | 100 | 73.8 | 182465.2 | 97.66
-81.5 | 0 | 100 | 73.97 | 182875.6 | 97.66
-80.3 | 0 | 100 | 74.05 | 183076.4 | 97.97
-79.1 | 0 | 100 | 74.29 | 183676.7 | 98.24
-77.9 | 0 | 100 | 74.44 | 184034 | 98.28
-76.7 | 0 | 100 | 74.45 | 184056.4 | 98.46
-75.5 | 0 | 100 | 74.55 | 184309.7 | 98.51
-74.3 | 0 | 100 | 74.48 | 184144 | 98.54
-73.1 | 0 | 100 | 74.7 | 184693.4 | 98.82
-71.9 | 0 | 100 | 74.83 | 185001.9 | 99.04
-70.7 | 0 | 100 | 74.85 | 185054.2 | 99
-69.5 | 0 | 100 | 75.04 | 185519.5 | 99.31
-68.3 | 0 | 100 | 74.91 | 185196.7 | 99.13
-67.1 | 0 | 100 | 75.09 | 185646 | 99.29
-65.9 | 0 | 100 | 75.15 | 185795.6 | 99.33
-64.7 | 0 | 100 | 75.25 | 186053.1 | 99.49
-63.5 | 0 | 100 | 75.25 | 186049.2 | 99.52
-62.3 | 0 | 100 | 75.13 | 185742.1 | 99.42
-61 | 0 | 100 | 75.36 | 186325.8 | 99.78
-59.8 | 0 | 100 | 75.4 | 186404.5 | 99.76
-58.6 | 0 | 100 | 75.37 | 186346.3 | 99.78
-57.4 | 0 | 100 | 75.44 | 186523.5 | 99.77
-56.2 | 0 | 100 | 75.45 | 186535.5 | 99.72
-55 | 0 | 100 | 75.58 | 186849.4 | 99.92
-53.8 | 0 | 100 | 75.54 | 186772.8 | 100
-52.6 | 0 | 100 | 75.48 | 186621.8 | 99.89
-51.4 | 0 | 100 | 75.4 | 186419.3 | 99.75
-50.2 | 0 | 100 | 75.53 | 186728.4 | 99.78
-49 | 0 | 100 | 75.3 | 186163.1 | 99.71
-47.8 | 0 | 100 | 75.47 | 186580.4 | 99.89
-46.6 | 0 | 100 | 75.29 | 186152.9 | 99.58
-45.4 | 0 | 100 | 75.33 | 186250.9 | 99.71
-44.1 | 0 | 100 | 75.19 | 185889.2 | 99.6
-43 | 0 | 100 | 75.19 | 185900.9 | 99.55
-41.8 | 0 | 100 | 75.04 | 185524.2 | 99.4
-40.5 | 0 | 100 | 75.29 | 186155.3 | 99.56
-39.4 | 0 | 100 | 75.14 | 185782.5 | 99.41
-38.1 | 0 | 100 | 75.15 | 185800.5 | 99.28
-36.9 | 0 | 100 | 75.12 | 185726.1 | 99.47
-35.7 | 0 | 100 | 75.07 | 185593.4 | 99.42
-34.5 | 0 | 100 | 75.12 | 185721.7 | 99.33
-33.3 | 0 | 100 | 75.05 | 185556.5 | 99.42
-32.1 | 0 | 100 | 75.06 | 185566.9 | 99.22
-30.9 | 0 | 100 | 75.05 | 185544.5 | 99.35
-29.7 | 0 | 100 | 75.11 | 185691.7 | 99.37
-28.5 | 0 | 100 | 75.08 | 185625 | 99.45
-27.3 | 0 | 100 | 75.02 | 185473 | 99.27
-26.1 | 0 | 100 | 74.98 | 185376.9 | 99.3
-24.9 | 0 | 100 | 74.97 | 185358.8 | 99.16
-23.7 | 0 | 100 | 74.9 | 185172.9 | 99.08
-22.5 | 0 | 100 | 74.7 | 184684.6 | 98.91
-21.3 | 0 | 100 | 74.77 | 184863.3 | 99.06
-20.1 | 0 | 100 | 74.81 | 184967.8 | 99.1
-18.9 | 0 | 100 | 74.85 | 185054.9 | 98.87
-17.7 | 0 | 100 | 74.92 | 185233.5 | 99.11
-16.5 | 0 | 100 | 74.73 | 184759 | 98.88
-15.3 | 0 | 100 | 74.67 | 184612.5 | 98.97
-14 | 0 | 100 | 74.59 | 184410.5 | 98.63
-12.9 | 0 | 100 | 74.55 | 184314.4 | 98.73
-11.7 | 0 | 100 | 74.38 | 183893.6 | 98.43
-10.4 | 0 | 100 | 74.35 | 183826.1 | 98.49
-9.2 | 0 | 100 | 74.23 | 183529.8 | 98.26
-8 | 0 | 100 | 74.36 | 183855 | 98.37
-6.8 | 0 | 100 | 74.29 | 183664.1 | 98.36
-5.6 | 0 | 100 | 74.37 | 183862.9 | 98.27
-4.4 | 0 | 100 | 74.2 | 183439.3 | 98.23
-3.2 | 0 | 100 | 74.15 | 183330.6 | 98.17
-2 | 0 | 100 | 74.2 | 183449.2 | 98.27
-0.8 | 0 | 100 | 74.1 | 183207.4 | 97.98
0.4 | 0 | 100 | 74.2 | 183441.2 | 98.27
1.6 | 0 | 100 | 74.11 | 183215.8 | 98.03
2.8 | 0 | 100 | 74.05 | 183071 | 98.18
4 | 0 | 100 | 73.97 | 182888.1 | 97.98
5.2 | 0 | 100 | 73.85 | 182580.5 | 97.81
6.4 | 0 | 100 | 74.01 | 182981.3 | 98.03
7.6 | 0 | 100 | 74.15 | 183332.2 | 98.18
8.8 | 0 | 100 | 74.19 | 183432 | 98.31
10 | 0 | 100 | 74.31 | 183727.7 | 98.33
11.2 | 0 | 100 | 74.33 | 183773 | 98.34
12.4 | 0 | 100 | 74.37 | 183876.5 | 98.42
13.7 | 0 | 100 | 74.35 | 183823.4 | 98.48
14.8 | 0 | 100 | 74.28 | 183651 | 98.27
16 | 0 | 100 | 74.36 | 183843.1 | 98.44
17.2 | 0 | 100 | 74.32 | 183749.6 | 98.21
18.4 | 0 | 100 | 74.37 | 183856.9 | 98.38
19.6 | 0 | 100 | 74.42 | 183990.4 | 98.43
20.9 | 0 | 100 | 74.48 | 184149.9 | 98.6
22 | 0 | 100 | 74.55 | 184305.7 | 98.9
23.3 | 0 | 100 | 74.43 | 184024.6 | 98.67
24.5 | 0 | 100 | 74.49 | 184161.9 | 98.68
25.7 | 0 | 100 | 74.51 | 184206.7 | 98.72
26.9 | 0 | 100 | 74.46 | 184084.8 | 98.56
28.1 | 0 | 100 | 74.49 | 184174.3 | 98.64
29.3 | 0 | 100 | 74.37 | 183865.6 | 98.69
30.5 | 0 | 100 | 74.48 | 184140.8 | 98.77
31.7 | 0 | 100 | 74.5 | 184188.4 | 98.67
32.9 | 0 | 100 | 74.37 | 183876.4 | 98.45
34.1 | 0 | 100 | 74.47 | 184106.9 | 98.72
35.3 | 0 | 100 | 74.46 | 184095.9 | 98.91
36.5 | 0 | 100 | 74.45 | 184060.1 | 98.45
37.7 | 0 | 100 | 74.51 | 184226.6 | 98.72
38.9 | 0 | 100 | 74.35 | 183827.5 | 98.46
40.1 | 0 | 100 | 74.45 | 184075.9 | 98.64
41.3 | 0 | 100 | 74.35 | 183808.6 | 98.77
42.5 | 0 | 100 | 74.58 | 184391.4 | 98.71
43.7 | 0 | 100 | 74.52 | 184229.4 | 98.6
44.9 | 0 | 100 | 74.7 | 184695.9 | 98.96
46.1 | 0 | 100 | 74.48 | 184145 | 98.7
47.3 | 0 | 100 | 74.57 | 184358.3 | 98.73
48.5 | 0 | 100 | 74.61 | 184453.9 | 98.8
49.7 | 0 | 100 | 74.67 | 184613 | 98.87
50.9 | 0 | 100 | 74.67 | 184599.4 | 98.95
52.1 | 0 | 100 | 74.83 | 185016.5 | 99.13
53.3 | 0 | 100 | 74.7 | 184687.8 | 98.91
54.5 | 0 | 100 | 74.63 | 184500.5 | 98.85
55.7 | 0 | 100 | 74.65 | 184567.8 | 98.85
56.9 | 0 | 100 | 74.57 | 184368.1 | 98.65
58.1 | 0 | 100 | 74.35 | 183816.7 | 98.56
59.4 | 0 | 100 | 74.48 | 184151.8 | 98.7
60.5 | 0 | 100 | 74.38 | 183894.6 | 98.39
61.8 | 0 | 100 | 74.38 | 183891 | 98.35
63 | 0 | 100 | 74.41 | 183972.9 | 98.36
64.2 | 0 | 100 | 74.3 | 183706.6 | 98.44
65.4 | 0 | 100 | 74.17 | 183375.1 | 98.16
66.6 | 0 | 100 | 74.11 | 183217.1 | 98.14
67.8 | 0 | 100 | 74.11 | 183217.1 | 98.21
69 | 0 | 100 | 74.03 | 183018.4 | 97.94
70.2 | 0 | 100 | 73.92 | 182752.7 | 97.86
71.4 | 0 | 100 | 73.95 | 182835.6 | 97.83
72.6 | 0 | 100 | 73.84 | 182564.7 | 97.78
73.8 | 0 | 100 | 73.65 | 182083.6 | 97.49
75 | 0 | 100 | 73.63 | 182033.6 | 97.43
76.2 | 0 | 100 | 73.42 | 181525.7 | 97.14
77.4 | 0 | 100 | 73.2 | 180979.7 | 96.96
78.6 | 0 | 100 | 73.12 | 180784.9 | 96.89
79.8 | 0 | 100 | 73.22 | 181024.6 | 96.82
81 | 0 | 100 | 72.99 | 180452.8 | 96.53
82.3 | 0 | 100 | 72.73 | 179805.6 | 96.28
83.5 | 0 | 100 | 72.7 | 179732.2 | 96.25
84.6 | 0 | 100 | 72.44 | 179098.5 | 96
85.9 | 0 | 100 | 72.23 | 178585.7 | 95.65
87.1 | 0 | 100 | 72.2 | 178505.2 | 95.6
88.3 | 0 | 100 | 71.86 | 177658.5 | 95.23
89.5 | 0 | 100 | 71.65 | 177138.3 | 95.04
90.7 | 0 | 100 | 71.58 | 176974.1 | 94.67
91.9 | 0 | 100 | 71.37 | 176443.4 | 94.42
93.1 | 0 | 100 | 70.79 | 175005.8 | 93.86
94.3 | 0 | 100 | 70.58 | 174506.2 | 93.46
95.5 | 0 | 100 | 70.47 | 174227.2 | 93.29
96.7 | 0 | 100 | 70.02 | 173111.5 | 92.72
97.9 | 0 | 100 | 69.67 | 172239.8 | 92.19
99.1 | 0 | 100 | 69.11 | 170873.1 | 91.58
100.3 | 0 | 100 | 68.65 | 169723.7 | 90.81
101.5 | 0 | 100 | 67.88 | 167823.6 | 90.09
102.7 | 0 | 100 | 66.85 | 165282.9 | 88.66
103.9 | 0 | 100 | 64.95 | 160582.4 | 85.88
105.1 | 0 | 100 | 61.54 | 152158.1 | 81.48
106.3 | 0 | 100 | 56.96 | 140815.8 | 75.38
107.5 | 0 | 100 | 50.73 | 125431.7 | 67.23
108.7 | 0 | 100 | 43.1 | 106558 | 57.06
110 | 0 | 100 | 34.87 | 86202.8 | 46.15
111.2 | 0 | 100 | 27.19 | 67234.7 | 36
112.4 | 0 | 100 | 21.1 | 52178 | 27.91
113.5 | 0 | 100 | 17.09 | 42264.4 | 22.65
114.7 | 0 | 100 | 14.77 | 36512.4 | 19.56
116 | 0 | 100 | 13.52 | 33416 | 17.91
117.2 | 0 | 100 | 12.61 | 31181.8 | 16.7
118.4 | 0 | 100 | 12.03 | 29741.9 | 15.93
119.6 | 0 | 100 | 11.53 | 28506.2 | 15.27
120.8 | 0 | 100 | 11.12 | 27483.5 | 14.7
122 | 0 | 100 | 10.74 | 26557.2 | 14.22
123.2 | 0 | 100 | 10.41 | 25746.9 | 13.81
124.4 | 0 | 100 | 10.09 | 24934.8 | 13.38
125.6 | 0 | 100 | 9.82 | 24285.4 | 13.04
126.8 | 0 | 100 | 9.6 | 23735 | 12.73
128 | 0 | 100 | 9.36 | 23137.8 | 12.41
129.2 | 0 | 100 | 9.06 | 22410.3 | 12.02
130.4 | 0 | 100 | 8.85 | 21868.1 | 11.7
131.6 | 0 | 100 | 8.62 | 21312.1 | 11.41
132.8 | 0 | 100 | 8.41 | 20787.3 | 11.13
134 | 0 | 100 | 8.15 | 20161.3 | 10.79
135.2 | 0 | 100 | 7.97 | 19693.8 | 10.55
136.5 | 0 | 100 | 7.8 | 19277.1 | 10.33
137.7 | 0 | 100 | 7.61 | 18802.6 | 10.08
138.8 | 0 | 100 | 7.4 | 18305.2 | 9.81
140.1 | 0 | 100 | 7.22 | 17861.5 | 9.57
141.3 | 0 | 100 | 7.11 | 17567.8 | 9.41
142.5 | 0 | 100 | 6.94 | 17150.2 | 9.18
143.7 | 0 | 100 | 6.8 | 16803.6 | 9
144.9 | 0 | 100 | 6.66 | 16463.7 | 8.82
146.1 | 0 | 100 | 6.46 | 15971.2 | 8.55
147.3 | 0 | 100 | 6.34 | 15672.6 | 8.39
148.5 | 0 | 100 | 6.24 | 15426.9 | 8.27
150 | 0 | 100 | 6.11 | 15098.8 | 8.09
Measurement time: | 12/08/2022 19:12:58
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
120 | 0 | 100 | 4.7 | 11622.1 | 6.63
119 | 0 | 100 | 4.84 | 11967.9 | 6.75
117.5 | 0 | 100 | 5.04 | 12467 | 7.02
116.5 | 0 | 100 | 5.19 | 12823.8 | 7.22
115.2 | 0 | 100 | 5.3 | 13099.6 | 7.38
114 | 0 | 100 | 5.44 | 13443.9 | 7.59
112.8 | 0 | 100 | 5.59 | 13812.4 | 7.78
111.6 | 0 | 100 | 5.74 | 14202.9 | 8
110.4 | 0 | 100 | 5.88 | 14538.9 | 8.18
109.2 | 0 | 100 | 6.08 | 15024.5 | 8.47
108 | 0 | 100 | 6.25 | 15452.5 | 8.73
106.8 | 0 | 100 | 6.4 | 15812 | 8.9
105.6 | 0 | 100 | 6.54 | 16163.9 | 9.11
104.4 | 0 | 100 | 6.76 | 16720 | 9.43
103.2 | 0 | 100 | 6.96 | 17216.7 | 9.71
102 | 0 | 100 | 7.18 | 17754 | 10
100.8 | 0 | 100 | 7.4 | 18293.8 | 10.31
99.6 | 0 | 100 | 7.65 | 18921.4 | 10.68
98.4 | 0 | 100 | 7.94 | 19618.5 | 11.06
97.2 | 0 | 100 | 8.18 | 20220 | 11.41
96 | 0 | 100 | 8.45 | 20879.7 | 11.79
94.8 | 0 | 100 | 8.82 | 21798.7 | 12.28
93.6 | 0 | 100 | 9.19 | 22720.3 | 12.8
92.4 | 0 | 100 | 9.58 | 23687.5 | 13.32
91.1 | 0 | 100 | 10.07 | 24901.5 | 14.04
90 | 0 | 100 | 10.65 | 26330.2 | 14.84
88.8 | 0 | 100 | 11.39 | 28155.9 | 15.88
87.5 | 0 | 100 | 12.73 | 31478.7 | 17.76
86.4 | 0 | 100 | 14.86 | 36734 | 20.74
85.1 | 0 | 100 | 18.91 | 46763.3 | 26.4
83.9 | 0 | 100 | 25 | 61798.1 | 34.86
82.8 | 0 | 100 | 32.44 | 80195 | 45.24
81.5 | 0 | 100 | 40.72 | 100662.3 | 56.76
80.4 | 0 | 100 | 48.22 | 119206.2 | 67.14
79.1 | 0 | 100 | 54.55 | 134874.5 | 76.14
77.9 | 0 | 100 | 59.35 | 146735.8 | 82.68
76.8 | 0 | 100 | 62.46 | 154411.2 | 87.12
75.5 | 0 | 100 | 64.24 | 158823.8 | 89.65
74.4 | 0 | 100 | 65.51 | 161961.4 | 91.39
73.1 | 0 | 100 | 66.02 | 163234.7 | 92.22
71.9 | 0 | 100 | 66.66 | 164798.1 | 92.84
70.7 | 0 | 100 | 67.19 | 166104.9 | 93.76
69.5 | 0 | 100 | 67.32 | 166442.3 | 93.84
68.3 | 0 | 100 | 67.87 | 167807 | 94.51
67.1 | 0 | 100 | 68.05 | 168239.1 | 94.92
65.9 | 0 | 100 | 68.21 | 168639.9 | 95.13
64.7 | 0 | 100 | 68.51 | 169375.9 | 95.52
63.5 | 0 | 100 | 68.63 | 169680.1 | 95.87
62.3 | 0 | 100 | 68.85 | 170211.8 | 95.89
61.1 | 0 | 100 | 69.12 | 170879.9 | 96.36
59.9 | 0 | 100 | 69.23 | 171163.5 | 96.4
58.7 | 0 | 100 | 69.48 | 171783.5 | 96.82
57.5 | 0 | 100 | 69.52 | 171879.4 | 97.02
56.3 | 0 | 100 | 69.7 | 172313.7 | 97.35
55.1 | 0 | 100 | 69.93 | 172896.9 | 97.48
53.9 | 0 | 100 | 69.86 | 172720.8 | 97.53
52.7 | 0 | 100 | 70.06 | 173206.7 | 97.76
51.5 | 0 | 100 | 70.09 | 173278.2 | 97.76
50.3 | 0 | 100 | 70.14 | 173416.2 | 97.69
49.1 | 0 | 100 | 70.3 | 173799.4 | 98.04
47.9 | 0 | 100 | 70.35 | 173922.1 | 98.02
46.7 | 0 | 100 | 70.35 | 173942 | 98.2
45.4 | 0 | 100 | 70.56 | 174445.7 | 98.42
44.3 | 0 | 100 | 70.51 | 174335.2 | 98.18
43 | 0 | 100 | 70.37 | 173987.5 | 98.06
41.9 | 0 | 100 | 70.4 | 174057.4 | 98.2
40.7 | 0 | 100 | 70.45 | 174174.8 | 98.45
39.4 | 0 | 100 | 70.63 | 174627.5 | 98.53
38.3 | 0 | 100 | 70.5 | 174299.4 | 98.29
37 | 0 | 100 | 70.56 | 174443.9 | 98.43
35.8 | 0 | 100 | 70.77 | 174969.8 | 98.74
34.6 | 0 | 100 | 70.75 | 174922 | 98.71
33.4 | 0 | 100 | 70.68 | 174738 | 98.49
32.3 | 0 | 100 | 70.86 | 175182.6 | 98.86
31 | 0 | 100 | 70.84 | 175139.2 | 98.84
29.8 | 0 | 100 | 70.9 | 175288.8 | 99.09
28.6 | 0 | 100 | 70.99 | 175512.3 | 99.21
27.4 | 0 | 100 | 71 | 175541.7 | 98.96
26.2 | 0 | 100 | 71.09 | 175759.4 | 99.13
25 | 0 | 100 | 70.97 | 175474.1 | 99.03
23.8 | 0 | 100 | 71.12 | 175842.2 | 99.06
22.6 | 0 | 100 | 71.14 | 175878.6 | 99.23
21.4 | 0 | 100 | 71.08 | 175726.7 | 99.13
20.2 | 0 | 100 | 71.17 | 175951.7 | 99.29
19 | 0 | 100 | 71.02 | 175584 | 99.02
17.8 | 0 | 100 | 71.07 | 175701.3 | 99.05
16.6 | 0 | 100 | 70.99 | 175516.5 | 98.99
15.3 | 0 | 100 | 71.02 | 175586.7 | 99.16
14.2 | 0 | 100 | 70.97 | 175467.2 | 98.97
13 | 0 | 100 | 71.09 | 175770 | 99.14
11.8 | 0 | 100 | 71.04 | 175628.9 | 99.15
10.6 | 0 | 100 | 71.06 | 175688.9 | 99.11
9.3 | 0 | 100 | 70.77 | 174966.5 | 98.81
8.2 | 0 | 100 | 70.86 | 175178.5 | 98.99
6.9 | 0 | 100 | 71.01 | 175561.4 | 99.12
5.8 | 0 | 100 | 70.92 | 175341.6 | 98.82
4.5 | 0 | 100 | 71.03 | 175607 | 99.15
3.3 | 0 | 100 | 70.78 | 175005.1 | 98.97
2.2 | 0 | 100 | 71.01 | 175567.8 | 99.09
0.9 | 0 | 100 | 70.75 | 174915.8 | 98.74
-0.2 | 0 | 100 | 70.96 | 175426.2 | 98.89
-1.5 | 0 | 100 | 70.84 | 175149.7 | 98.9
-2.7 | 0 | 100 | 70.85 | 175178.3 | 98.92
-3.8 | 0 | 100 | 70.93 | 175365.9 | 99.27
-5.1 | 0 | 100 | 71.07 | 175716.6 | 99.3
-6.2 | 0 | 100 | 71.14 | 175871.4 | 99.28
-7.5 | 0 | 100 | 71.12 | 175845.2 | 99.2
-8.7 | 0 | 100 | 71.1 | 175795.8 | 99.31
-9.9 | 0 | 100 | 71.07 | 175704.6 | 99.18
-11.1 | 0 | 100 | 71.23 | 176097.9 | 99.36
-12.3 | 0 | 100 | 71.27 | 176197.4 | 99.45
-13.5 | 0 | 100 | 71.27 | 176208.3 | 99.53
-14.7 | 0 | 100 | 71.33 | 176360.5 | 99.66
-15.9 | 0 | 100 | 71.57 | 176958.5 | 100
-17.1 | 0 | 100 | 71.45 | 176642.8 | 99.9
-18.3 | 0 | 100 | 71.67 | 177193.8 | 99.93
-19.5 | 0 | 100 | 71.43 | 176602.4 | 99.83
-20.8 | 0 | 100 | 71.36 | 176432.3 | 99.73
-21.9 | 0 | 100 | 71.64 | 177122.6 | 99.96
-23.2 | 0 | 100 | 71.64 | 177120.9 | 99.96
-24.3 | 0 | 100 | 71.67 | 177201.6 | 99.95
-25.5 | 0 | 100 | 71.61 | 177037.3 | 99.93
-26.8 | 0 | 100 | 71.58 | 176963.1 | 99.95
-27.9 | 0 | 100 | 71.51 | 176802.1 | 99.74
-29.2 | 0 | 100 | 71.57 | 176942.2 | 99.94
-30.4 | 0 | 100 | 71.51 | 176789.4 | 99.76
-31.6 | 0 | 100 | 71.59 | 176996.4 | 99.79
-32.8 | 0 | 100 | 71.41 | 176557.8 | 99.68
-34 | 0 | 100 | 71.39 | 176504.4 | 99.57
-35.2 | 0 | 100 | 71.34 | 176379.1 | 99.48
-36.4 | 0 | 100 | 71.46 | 176686.3 | 99.68
-37.6 | 0 | 100 | 71.37 | 176456.6 | 99.68
-38.8 | 0 | 100 | 71.32 | 176324.2 | 99.57
-40 | 0 | 100 | 71.39 | 176490.8 | 99.63
-41.2 | 0 | 100 | 71.42 | 176570.8 | 99.67
-42.4 | 0 | 100 | 71.32 | 176316.5 | 99.46
-43.6 | 0 | 100 | 71.28 | 176217.2 | 99.7
-44.8 | 0 | 100 | 71.31 | 176292.1 | 99.49
-46 | 0 | 100 | 71.27 | 176207 | 99.48
-47.3 | 0 | 100 | 71.17 | 175969.4 | 99.52
-48.4 | 0 | 100 | 71.26 | 176181.8 | 99.43
-49.6 | 0 | 100 | 71.08 | 175739.4 | 99.32
-50.9 | 0 | 100 | 71.16 | 175929.1 | 99.39
-52 | 0 | 100 | 71.01 | 175556.2 | 99.01
-53.3 | 0 | 100 | 70.92 | 175343.3 | 99.05
-54.4 | 0 | 100 | 70.89 | 175262.6 | 98.91
-55.6 | 0 | 100 | 70.72 | 174852.1 | 98.65
-56.9 | 0 | 100 | 70.73 | 174880.2 | 98.65
-58 | 0 | 100 | 70.36 | 173964.7 | 98.5
-59.3 | 0 | 100 | 70.23 | 173642 | 98.24
-60.5 | 0 | 100 | 70.24 | 173660.8 | 98.18
-61.7 | 0 | 100 | 70.16 | 173472 | 98.02
-62.9 | 0 | 100 | 70.03 | 173131.8 | 97.69
-64.1 | 0 | 100 | 69.81 | 172606.2 | 97.37
-65.3 | 0 | 100 | 69.55 | 171942 | 97
-66.5 | 0 | 100 | 69.44 | 171668.9 | 96.86
-67.7 | 0 | 100 | 69.03 | 170671.4 | 96.51
-68.9 | 0 | 100 | 68.96 | 170489.5 | 96.27
-70.1 | 0 | 100 | 68.53 | 169423.4 | 95.71
-71.3 | 0 | 100 | 68.21 | 168639.9 | 95.18
-72.5 | 0 | 100 | 67.96 | 168026.1 | 94.83
-73.7 | 0 | 100 | 67.29 | 166361.5 | 93.88
-75 | 0 | 100 | 66.68 | 164845.2 | 93.22
-76.1 | 0 | 100 | 65.68 | 162383.5 | 91.86
-77.4 | 0 | 100 | 64.26 | 158867.7 | 89.69
-78.6 | 0 | 100 | 61.46 | 151950.9 | 85.79
-79.8 | 0 | 100 | 56.48 | 139641.7 | 78.86
-81 | 0 | 100 | 49.52 | 122433.1 | 69.16
-82.2 | 0 | 100 | 41.51 | 102636.4 | 58
-83.4 | 0 | 100 | 33 | 81575.5 | 46.08
-84.6 | 0 | 100 | 25.88 | 63974.9 | 36.1
-85.8 | 0 | 100 | 20.3 | 50191.4 | 28.34
-87 | 0 | 100 | 16.08 | 39755.7 | 22.43
-88.2 | 0 | 100 | 13.5 | 33372.7 | 18.84
-89.4 | 0 | 100 | 11.88 | 29375.1 | 16.59
-90.6 | 0 | 100 | 10.89 | 26924.7 | 15.21
-91.8 | 0 | 100 | 10.32 | 25519 | 14.41
-93 | 0 | 100 | 9.77 | 24147.7 | 13.65
-94.2 | 0 | 100 | 9.37 | 23155.6 | 13.08
-95.4 | 0 | 100 | 9.06 | 22404.4 | 12.66
-96.6 | 0 | 100 | 8.72 | 21552.8 | 12.17
-97.8 | 0 | 100 | 8.34 | 20609.1 | 11.65
-99 | 0 | 100 | 8.06 | 19934.6 | 11.26
-100.2 | 0 | 100 | 7.92 | 19570.7 | 11.06
-101.5 | 0 | 100 | 7.63 | 18857.1 | 10.65
-102.6 | 0 | 100 | 7.37 | 18227.7 | 10.29
-103.8 | 0 | 100 | 7.16 | 17708.6 | 10
-105.1 | 0 | 100 | 6.99 | 17273.9 | 9.75
-106.2 | 0 | 100 | 6.71 | 16592.1 | 9.38
-107.5 | 0 | 100 | 6.55 | 16183.2 | 9.13
-108.6 | 0 | 100 | 6.38 | 15770.4 | 8.9
-109.9 | 0 | 100 | 6.16 | 15241.4 | 8.6
-111 | 0 | 100 | 6.01 | 14861.9 | 8.4
-112.2 | 0 | 100 | 5.83 | 14407.3 | 8.16
-113.5 | 0 | 100 | 5.67 | 14018.3 | 7.92
-114.6 | 0 | 100 | 5.47 | 13524.6 | 7.63
-115.9 | 0 | 100 | 5.39 | 13335.4 | 7.53
-117.1 | 0 | 100 | 5.25 | 12968.6 | 7.32
-118.3 | 0 | 100 | 5.05 | 12474.1 | 7.04
-120 | 0 | 100 | 4.88 | 12076.3 | 6.82
Measurement time: | 12/08/2022 19:15:48
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-80 | 0 | 100 | 0.51 | 1271.9 | 0.81
-79.1 | 0 | 100 | 0.61 | 1509.9 | 1.02
-77.6 | 0 | 100 | 0.64 | 1593.3 | 1.08
-76.5 | 0 | 100 | 0.66 | 1641.3 | 1.11
-75.3 | 0 | 100 | 0.74 | 1832.8 | 1.24
-74 | 0 | 100 | 0.77 | 1898.1 | 1.29
-72.8 | 0 | 100 | 0.78 | 1937.2 | 1.31
-71.6 | 0 | 100 | 0.89 | 2193.8 | 1.49
-70.4 | 0 | 100 | 0.9 | 2231.4 | 1.51
-69.2 | 0 | 100 | 0.95 | 2353.5 | 1.6
-68 | 0 | 100 | 1.01 | 2507.4 | 1.7
-66.8 | 0 | 100 | 1.04 | 2572.3 | 1.74
-65.6 | 0 | 100 | 1.11 | 2737.9 | 1.86
-64.4 | 0 | 100 | 1.14 | 2821.3 | 1.91
-63.2 | 0 | 100 | 1.24 | 3063.1 | 2.08
-62 | 0 | 100 | 1.31 | 3239.8 | 2.2
-60.8 | 0 | 100 | 1.41 | 3477.4 | 2.36
-59.6 | 0 | 100 | 1.49 | 3685.6 | 2.5
-58.4 | 0 | 100 | 1.53 | 3794 | 2.58
-57.2 | 0 | 100 | 1.66 | 4095.2 | 2.78
-56 | 0 | 100 | 1.74 | 4309.7 | 2.92
-54.8 | 0 | 100 | 1.86 | 4605.9 | 3.13
-53.6 | 0 | 100 | 1.95 | 4811.1 | 3.26
-52.4 | 0 | 100 | 2.02 | 5001.7 | 3.4
-51.2 | 0 | 100 | 2.12 | 5245.9 | 3.56
-50 | 0 | 100 | 2.24 | 5538.6 | 3.77
-48.8 | 0 | 100 | 2.37 | 5867.1 | 4
-47.5 | 0 | 100 | 2.54 | 6288.7 | 4.28
-46.3 | 0 | 100 | 2.7 | 6664.1 | 4.54
-45.2 | 0 | 100 | 2.8 | 6926.5 | 4.71
-43.9 | 0 | 100 | 2.98 | 7364.4 | 5.02
-42.7 | 0 | 100 | 3.18 | 7860.6 | 5.35
-41.5 | 0 | 100 | 3.33 | 8234.5 | 5.62
-40.3 | 0 | 100 | 3.57 | 8827.3 | 6.02
-39.1 | 0 | 100 | 3.82 | 9441 | 6.44
-37.9 | 0 | 100 | 4.17 | 10309 | 7.01
-36.7 | 0 | 100 | 4.61 | 11385.5 | 7.75
-35.5 | 0 | 100 | 5.21 | 12886.6 | 8.77
-34.3 | 0 | 100 | 6.2 | 15321.9 | 10.45
-33.1 | 0 | 100 | 7.92 | 19591.6 | 13.36
-31.9 | 0 | 100 | 10.84 | 26794.7 | 18.24
-30.7 | 0 | 100 | 15.25 | 37708.9 | 25.66
-29.5 | 0 | 100 | 21.09 | 52138.4 | 35.53
-28.3 | 0 | 100 | 28.38 | 70156.6 | 47.75
-27.1 | 0 | 100 | 36.27 | 89665.4 | 61
-25.9 | 0 | 100 | 43.79 | 108266.8 | 73.63
-24.7 | 0 | 100 | 49.9 | 123360.2 | 84.08
-23.5 | 0 | 100 | 53.65 | 132643.6 | 90.35
-22.3 | 0 | 100 | 55.79 | 137931 | 93.91
-21.1 | 0 | 100 | 56.97 | 140838.9 | 96.29
-19.9 | 0 | 100 | 57.58 | 142351.3 | 96.98
-18.7 | 0 | 100 | 57.96 | 143292.6 | 97.88
-17.5 | 0 | 100 | 58.25 | 144019.2 | 98.05
-16.3 | 0 | 100 | 58.46 | 144545.2 | 98.7
-15.1 | 0 | 100 | 58.55 | 144753.5 | 98.77
-13.9 | 0 | 100 | 58.84 | 145466.5 | 99.1
-12.7 | 0 | 100 | 58.83 | 145446.2 | 99.18
-11.5 | 0 | 100 | 58.92 | 145665.1 | 99.42
-10.2 | 0 | 100 | 58.96 | 145760.3 | 99.54
-9.1 | 0 | 100 | 58.99 | 145836.4 | 99.38
-7.8 | 0 | 100 | 58.91 | 145640.3 | 99.39
-6.6 | 0 | 100 | 58.91 | 145643.3 | 99.6
-5.5 | 0 | 100 | 58.89 | 145592.3 | 99.55
-4.2 | 0 | 100 | 58.85 | 145502.9 | 99.59
-3 | 0 | 100 | 58.91 | 145654 | 99.65
-1.8 | 0 | 100 | 58.87 | 145538.1 | 99.66
-0.6 | 0 | 100 | 59.03 | 145951.4 | 100
0.6 | 0 | 100 | 58.83 | 145455.3 | 99.57
1.8 | 0 | 100 | 58.76 | 145269.5 | 99.45
3 | 0 | 100 | 58.87 | 145555.7 | 99.62
4.2 | 0 | 100 | 58.76 | 145283.8 | 99.54
5.4 | 0 | 100 | 58.7 | 145122.9 | 99.25
6.6 | 0 | 100 | 58.57 | 144816.3 | 99.38
7.8 | 0 | 100 | 58.55 | 144746.4 | 99.03
9 | 0 | 100 | 58.6 | 144889.1 | 99.48
10.2 | 0 | 100 | 58.49 | 144602.2 | 99.21
11.4 | 0 | 100 | 58.35 | 144262.9 | 98.98
12.6 | 0 | 100 | 58.28 | 144081.4 | 98.98
13.8 | 0 | 100 | 58.16 | 143795.2 | 98.61
15 | 0 | 100 | 57.86 | 143048 | 97.82
16.2 | 0 | 100 | 57.7 | 142657.9 | 97.63
17.4 | 0 | 100 | 57.51 | 142182.3 | 97.46
18.7 | 0 | 100 | 57.08 | 141123.9 | 96.75
19.8 | 0 | 100 | 56.45 | 139561.1 | 95.74
21.1 | 0 | 100 | 55.46 | 137108 | 94.01
22.3 | 0 | 100 | 53.69 | 132748.5 | 91.1
23.5 | 0 | 100 | 50.54 | 124950.4 | 85.71
24.7 | 0 | 100 | 46.2 | 114221.7 | 78.41
25.9 | 0 | 100 | 40.23 | 99461.6 | 68.1
27.1 | 0 | 100 | 32.99 | 81550.8 | 56.04
28.3 | 0 | 100 | 25.45 | 62914.2 | 43.21
29.5 | 0 | 100 | 18.25 | 45124.1 | 31
30.7 | 0 | 100 | 12.5 | 30905.4 | 21.24
31.9 | 0 | 100 | 8.57 | 21192.4 | 14.57
33.1 | 0 | 100 | 6.6 | 16316.7 | 11.22
34.3 | 0 | 100 | 5.48 | 13551 | 9.32
35.5 | 0 | 100 | 4.76 | 11765 | 8.09
36.7 | 0 | 100 | 4.24 | 10491.6 | 7.21
37.9 | 0 | 100 | 3.78 | 9350.7 | 6.42
39.1 | 0 | 100 | 3.53 | 8718.1 | 6
40.3 | 0 | 100 | 3.32 | 8214.5 | 5.66
41.5 | 0 | 100 | 3.12 | 7708 | 5.32
42.7 | 0 | 100 | 2.98 | 7357.7 | 5.07
43.9 | 0 | 100 | 2.79 | 6889.9 | 4.74
45.1 | 0 | 100 | 2.59 | 6394.8 | 4.41
46.3 | 0 | 100 | 2.45 | 6064.3 | 4.17
47.5 | 0 | 100 | 2.32 | 5742 | 3.96
48.7 | 0 | 100 | 2.17 | 5354.8 | 3.69
49.9 | 0 | 100 | 2.07 | 5108.5 | 3.52
51.1 | 0 | 100 | 1.94 | 4798.5 | 3.31
52.3 | 0 | 100 | 1.89 | 4665.6 | 3.22
53.5 | 0 | 100 | 1.76 | 4352 | 2.99
54.7 | 0 | 100 | 1.66 | 4096.5 | 2.82
55.9 | 0 | 100 | 1.62 | 4002.1 | 2.77
57.1 | 0 | 100 | 1.49 | 3688.1 | 2.54
58.4 | 0 | 100 | 1.42 | 3515.8 | 2.42
59.5 | 0 | 100 | 1.36 | 3354 | 2.31
60.7 | 0 | 100 | 1.3 | 3221.8 | 2.22
62 | 0 | 100 | 1.26 | 3118.8 | 2.15
63.2 | 0 | 100 | 1.18 | 2910 | 2.01
64.4 | 0 | 100 | 1.12 | 2773.8 | 1.91
65.6 | 0 | 100 | 1.05 | 2587 | 1.78
66.8 | 0 | 100 | 1.02 | 2530.4 | 1.74
68 | 0 | 100 | 0.94 | 2333.1 | 1.61
69.2 | 0 | 100 | 0.94 | 2325.9 | 1.61
70.4 | 0 | 100 | 0.86 | 2117.9 | 1.46
71.6 | 0 | 100 | 0.79 | 1955.3 | 1.35
72.8 | 0 | 100 | 0.76 | 1877.9 | 1.3
74 | 0 | 100 | 0.73 | 1795.5 | 1.24
75.2 | 0 | 100 | 0.71 | 1762.4 | 1.22
76.4 | 0 | 100 | 0.66 | 1630 | 1.13
77.6 | 0 | 100 | 0.64 | 1574.1 | 1.09
78.8 | 0 | 100 | 0.61 | 1519.4 | 1.05
80 | 0 | 100 | 0.54 | 1333.4 | 0.92
Measurement time: | 12/08/2022 19:17:58
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
100 | 0 | 100 | 2.51 | 6210.2 | 3.73
99.1 | 0 | 100 | 2.55 | 6304.8 | 3.81
97.6 | 0 | 100 | 2.68 | 6637.1 | 4.02
96.5 | 0 | 100 | 2.75 | 6805.1 | 4.11
95.4 | 0 | 100 | 2.86 | 7081.2 | 4.28
94.1 | 0 | 100 | 2.96 | 7318.4 | 4.42
92.9 | 0 | 100 | 3.07 | 7581.4 | 4.59
91.8 | 0 | 100 | 3.18 | 7861.4 | 4.75
90.5 | 0 | 100 | 3.27 | 8094.6 | 4.89
89.3 | 0 | 100 | 3.41 | 8430.9 | 5.1
88.1 | 0 | 100 | 3.5 | 8645.7 | 5.24
86.9 | 0 | 100 | 3.62 | 8951 | 5.41
85.7 | 0 | 100 | 3.77 | 9322.4 | 5.64
84.5 | 0 | 100 | 3.84 | 9485.1 | 5.74
83.3 | 0 | 100 | 4.03 | 9958.4 | 6.03
82.1 | 0 | 100 | 4.14 | 10237.4 | 6.19
80.9 | 0 | 100 | 4.34 | 10738.1 | 6.5
79.7 | 0 | 100 | 4.5 | 11121.5 | 6.73
78.5 | 0 | 100 | 4.67 | 11551.2 | 6.98
77.3 | 0 | 100 | 4.84 | 11971.7 | 7.23
76.1 | 0 | 100 | 5.03 | 12432.4 | 7.52
74.9 | 0 | 100 | 5.2 | 12853.3 | 7.77
73.7 | 0 | 100 | 5.35 | 13229.7 | 8
72.5 | 0 | 100 | 5.6 | 13842.4 | 8.38
71.3 | 0 | 100 | 5.81 | 14354.3 | 8.68
70 | 0 | 100 | 6.07 | 15015.5 | 9.09
68.9 | 0 | 100 | 6.31 | 15588.8 | 9.44
67.7 | 0 | 100 | 6.57 | 16239.5 | 9.83
66.4 | 0 | 100 | 6.95 | 17176.3 | 10.4
65.3 | 0 | 100 | 7.32 | 18087.8 | 10.96
64 | 0 | 100 | 7.81 | 19307.3 | 11.66
62.8 | 0 | 100 | 8.36 | 20670.3 | 12.51
61.6 | 0 | 100 | 9.18 | 22706.6 | 13.74
60.4 | 0 | 100 | 10.59 | 26174.3 | 15.84
59.3 | 0 | 100 | 12.9 | 31882.5 | 19.29
58 | 0 | 100 | 17.23 | 42593.5 | 25.78
56.8 | 0 | 100 | 23.55 | 58233.5 | 35.22
55.6 | 0 | 100 | 30.94 | 76503.9 | 46.3
54.4 | 0 | 100 | 39.29 | 97131.6 | 58.74
53.2 | 0 | 100 | 46.27 | 114403 | 69.36
52 | 0 | 100 | 52.16 | 128957.1 | 78.27
50.8 | 0 | 100 | 56.65 | 140060.7 | 84.93
49.6 | 0 | 100 | 59.29 | 146586.3 | 88.75
48.4 | 0 | 100 | 61.01 | 150848.4 | 91.39
47.2 | 0 | 100 | 62.08 | 153491.5 | 92.97
46 | 0 | 100 | 62.65 | 154904.8 | 93.73
44.8 | 0 | 100 | 62.97 | 155687 | 94.29
43.6 | 0 | 100 | 63.33 | 156563.7 | 94.81
42.3 | 0 | 100 | 63.42 | 156805.4 | 95.03
41.2 | 0 | 100 | 63.86 | 157895.5 | 95.58
39.9 | 0 | 100 | 64.08 | 158417.4 | 95.99
38.8 | 0 | 100 | 64.14 | 158577.7 | 96.17
37.5 | 0 | 100 | 64.39 | 159192.2 | 96.46
36.3 | 0 | 100 | 64.52 | 159503.7 | 96.56
35.1 | 0 | 100 | 64.74 | 160059.3 | 96.92
33.9 | 0 | 100 | 64.92 | 160515.6 | 97.33
32.8 | 0 | 100 | 65.08 | 160892.7 | 97.38
31.5 | 0 | 100 | 65.24 | 161304.7 | 97.58
30.3 | 0 | 100 | 65.37 | 161608.7 | 97.91
29.1 | 0 | 100 | 65.35 | 161564.6 | 97.95
27.9 | 0 | 100 | 65.69 | 162410.6 | 98.26
26.7 | 0 | 100 | 65.73 | 162516.4 | 98.5
25.5 | 0 | 100 | 65.8 | 162687.6 | 98.68
24.3 | 0 | 100 | 65.91 | 162940.9 | 98.65
23.1 | 0 | 100 | 66.02 | 163232.7 | 98.81
21.9 | 0 | 100 | 66.12 | 163467 | 98.9
20.7 | 0 | 100 | 66.04 | 163266.4 | 98.97
19.5 | 0 | 100 | 66.19 | 163644.1 | 99.29
18.3 | 0 | 100 | 66.11 | 163444.4 | 99.43
17.1 | 0 | 100 | 66.22 | 163711.1 | 99.18
15.9 | 0 | 100 | 66.26 | 163807.3 | 99.4
14.7 | 0 | 100 | 66.34 | 164005.3 | 99.09
13.5 | 0 | 100 | 66.49 | 164388.3 | 99.44
12.3 | 0 | 100 | 66.49 | 164390.1 | 99.4
11.1 | 0 | 100 | 66.32 | 163958.7 | 99.15
9.9 | 0 | 100 | 66.26 | 163817 | 99.28
8.7 | 0 | 100 | 66.27 | 163851.3 | 99.25
7.5 | 0 | 100 | 66.2 | 163680.1 | 99.11
6.3 | 0 | 100 | 66.19 | 163650.2 | 99.22
5.1 | 0 | 100 | 66.03 | 163240.2 | 99.01
3.8 | 0 | 100 | 66.13 | 163499.9 | 99.25
2.7 | 0 | 100 | 66.25 | 163786.2 | 99.24
1.5 | 0 | 100 | 66.35 | 164032.7 | 99.22
0.3 | 0 | 100 | 66.37 | 164091.1 | 99.51
-0.9 | 0 | 100 | 66.27 | 163854.5 | 99.16
-2.1 | 0 | 100 | 66.32 | 163967.4 | 99.26
-3.3 | 0 | 100 | 66.57 | 164587.5 | 99.68
-4.5 | 0 | 100 | 66.44 | 164265.1 | 99.49
-5.8 | 0 | 100 | 66.54 | 164498.2 | 99.6
-6.9 | 0 | 100 | 66.56 | 164547.5 | 99.58
-8.2 | 0 | 100 | 66.5 | 164409 | 99.55
-9.4 | 0 | 100 | 66.72 | 164953.2 | 99.9
-10.6 | 0 | 100 | 66.56 | 164569 | 99.95
-11.8 | 0 | 100 | 66.66 | 164801.9 | 99.78
-13 | 0 | 100 | 66.78 | 165095.9 | 99.96
-14.2 | 0 | 100 | 66.61 | 164691.4 | 99.78
-15.4 | 0 | 100 | 66.66 | 164817.4 | 99.96
-16.6 | 0 | 100 | 66.78 | 165098.2 | 99.93
-17.8 | 0 | 100 | 66.65 | 164786 | 99.83
-19 | 0 | 100 | 66.73 | 164982.2 | 99.96
-20.2 | 0 | 100 | 66.62 | 164705.1 | 99.68
-21.4 | 0 | 100 | 66.77 | 165075.7 | 100
-22.6 | 0 | 100 | 66.67 | 164826.2 | 99.64
-23.8 | 0 | 100 | 66.42 | 164214.3 | 99.52
-25 | 0 | 100 | 66.54 | 164504.2 | 99.77
-26.2 | 0 | 100 | 66.37 | 164084.4 | 99.66
-27.4 | 0 | 100 | 66.34 | 164024 | 99.28
-28.6 | 0 | 100 | 66.27 | 163846.7 | 99.34
-29.8 | 0 | 100 | 66.04 | 163273.7 | 99
-31 | 0 | 100 | 65.94 | 163015.9 | 98.91
-32.2 | 0 | 100 | 65.99 | 163153.1 | 98.88
-33.4 | 0 | 100 | 65.75 | 162562.2 | 98.46
-34.6 | 0 | 100 | 65.74 | 162526.9 | 98.66
-35.9 | 0 | 100 | 65.47 | 161854.8 | 98.02
-37 | 0 | 100 | 65.33 | 161506.7 | 97.81
-38.2 | 0 | 100 | 65.21 | 161218.9 | 97.75
-39.4 | 0 | 100 | 64.98 | 160656.5 | 97.32
-40.6 | 0 | 100 | 64.95 | 160573.2 | 97.32
-41.8 | 0 | 100 | 64.56 | 159606.8 | 96.8
-43 | 0 | 100 | 64.5 | 159456 | 96.78
-44.3 | 0 | 100 | 64.35 | 159099 | 96.47
-45.4 | 0 | 100 | 63.94 | 158084.1 | 95.8
-46.6 | 0 | 100 | 63.53 | 157068.8 | 95.37
-47.9 | 0 | 100 | 62.9 | 155519 | 94.16
-49 | 0 | 100 | 62.08 | 153484.6 | 93.08
-50.3 | 0 | 100 | 60.56 | 149724.3 | 90.73
-51.5 | 0 | 100 | 57.57 | 142337.7 | 86.31
-52.7 | 0 | 100 | 52.55 | 129927.8 | 78.68
-53.9 | 0 | 100 | 45.17 | 111673.4 | 67.79
-55.1 | 0 | 100 | 37.27 | 92149.6 | 55.87
-56.3 | 0 | 100 | 29.23 | 72258 | 43.72
-57.5 | 0 | 100 | 22.39 | 55346.9 | 33.56
-58.7 | 0 | 100 | 17 | 42037.2 | 25.47
-59.9 | 0 | 100 | 13.1 | 32380.9 | 19.63
-61.1 | 0 | 100 | 10.77 | 26635.1 | 16.12
-62.3 | 0 | 100 | 9.29 | 22957.6 | 13.93
-63.5 | 0 | 100 | 8.49 | 21000.8 | 12.73
-64.7 | 0 | 100 | 7.91 | 19568.3 | 11.86
-65.9 | 0 | 100 | 7.38 | 18254.1 | 11.06
-67.1 | 0 | 100 | 7 | 17304.5 | 10.5
-68.3 | 0 | 100 | 6.72 | 16619 | 10.07
-69.5 | 0 | 100 | 6.38 | 15785.8 | 9.57
-70.7 | 0 | 100 | 6.12 | 15127.1 | 9.16
-71.9 | 0 | 100 | 5.9 | 14592.4 | 8.83
-73.2 | 0 | 100 | 5.69 | 14056.6 | 8.52
-74.4 | 0 | 100 | 5.47 | 13527 | 8.19
-75.6 | 0 | 100 | 5.26 | 13008.5 | 7.89
-76.8 | 0 | 100 | 5.08 | 12548.2 | 7.6
-77.9 | 0 | 100 | 4.9 | 12126.7 | 7.35
-79.2 | 0 | 100 | 4.66 | 11514.1 | 6.98
-80.4 | 0 | 100 | 4.53 | 11211.3 | 6.79
-81.5 | 0 | 100 | 4.4 | 10880.7 | 6.59
-82.8 | 0 | 100 | 4.27 | 10560.3 | 6.4
-84 | 0 | 100 | 4.09 | 10112.2 | 6.13
-85.2 | 0 | 100 | 3.91 | 9658.6 | 5.85
-86.4 | 0 | 100 | 3.85 | 9529.4 | 5.77
-87.6 | 0 | 100 | 3.65 | 9017.3 | 5.47
-88.8 | 0 | 100 | 3.55 | 8787.9 | 5.33
-90 | 0 | 100 | 3.42 | 8448.9 | 5.12
-91.2 | 0 | 100 | 3.34 | 8246.3 | 5
-92.4 | 0 | 100 | 3.22 | 7949.2 | 4.82
-93.6 | 0 | 100 | 3.17 | 7829.9 | 4.74
-94.8 | 0 | 100 | 3.02 | 7476.9 | 4.53
-96 | 0 | 100 | 2.96 | 7318.1 | 4.44
-97.2 | 0 | 100 | 2.83 | 6993.8 | 4.24
-98.4 | 0 | 100 | 2.75 | 6797.1 | 4.12
-100 | 0 | 100 | 2.62 | 6475.8 | 3.92