## Feuil1
Clinac 2
Measurement time: | 12/08/2022 18:00:16
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 19.24 | 47561.7 | 20.48 | 0.2772 | 34.0955
0 | 0 | 309.1 | 20.74 | 51265.8 | 20.6 | 0.27844 | 34.2482
0 | 0 | 307.6 | 20.83 | 51504.6 | 20.67 | 0.27875 | 34.2859
0 | 0 | 306.6 | 21.02 | 51974.8 | 20.89 | 0.28129 | 34.5978
0 | 0 | 305.3 | 21.15 | 52292.6 | 21.01 | 0.28234 | 34.7273
0 | 0 | 304.1 | 21.35 | 52782.9 | 21.19 | 0.28423 | 34.9605
0 | 0 | 302.9 | 21.46 | 53066 | 21.33 | 0.28559 | 35.1267
0 | 0 | 301.7 | 21.62 | 53451.6 | 21.46 | 0.2868 | 35.2758
0 | 0 | 300.5 | 21.73 | 53713.5 | 21.55 | 0.28747 | 35.3584
0 | 0 | 299.3 | 21.83 | 53983.6 | 21.68 | 0.28867 | 35.5061
0 | 0 | 298.1 | 22.08 | 54597.4 | 21.97 | 0.29199 | 35.9146
0 | 0 | 296.9 | 22.18 | 54836.5 | 22.02 | 0.29211 | 35.9298
0 | 0 | 295.7 | 22.34 | 55238 | 22.21 | 0.29409 | 36.1728
0 | 0 | 294.5 | 22.52 | 55683.7 | 22.35 | 0.2954 | 36.3334
0 | 0 | 293.3 | 22.64 | 55983.6 | 22.46 | 0.2963 | 36.4446
0 | 0 | 292.1 | 22.78 | 56332 | 22.67 | 0.29852 | 36.7171
0 | 0 | 290.9 | 22.95 | 56744.8 | 22.82 | 0.29993 | 36.8914
0 | 0 | 289.7 | 23.07 | 57049 | 22.94 | 0.30095 | 37.0165
0 | 0 | 288.5 | 23.29 | 57574 | 23.13 | 0.30288 | 37.2537
0 | 0 | 287.3 | 23.45 | 57968 | 23.31 | 0.30467 | 37.4737
0 | 0 | 286.1 | 23.53 | 58174.1 | 23.42 | 0.30553 | 37.5804
0 | 0 | 284.9 | 23.7 | 58600 | 23.54 | 0.30653 | 37.7025
0 | 0 | 283.7 | 23.93 | 59151.6 | 23.76 | 0.30881 | 37.9838
0 | 0 | 282.5 | 24.01 | 59365.5 | 23.88 | 0.30979 | 38.1043
0 | 0 | 281.2 | 24.25 | 59960.5 | 24.07 | 0.31163 | 38.3296
0 | 0 | 280.1 | 24.39 | 60289 | 24.23 | 0.31316 | 38.5182
0 | 0 | 278.8 | 24.46 | 60477.1 | 24.31 | 0.31355 | 38.5669
0 | 0 | 277.7 | 24.68 | 61005.5 | 24.5 | 0.31546 | 38.8015
0 | 0 | 276.4 | 24.86 | 61469.1 | 24.72 | 0.31765 | 39.0703
0 | 0 | 275.2 | 25.02 | 61854 | 24.83 | 0.31846 | 39.1704
0 | 0 | 274 | 25.22 | 62360 | 25.07 | 0.32093 | 39.4746
0 | 0 | 272.8 | 25.32 | 62610.7 | 25.19 | 0.32186 | 39.5889
0 | 0 | 271.6 | 25.54 | 63143.9 | 25.37 | 0.32355 | 39.7966
0 | 0 | 270.4 | 25.66 | 63445.9 | 25.55 | 0.32523 | 40.0034
0 | 0 | 269.2 | 25.84 | 63885.9 | 25.73 | 0.32691 | 40.2091
0 | 0 | 268 | 26.05 | 64400.7 | 25.89 | 0.32832 | 40.3827
0 | 0 | 266.8 | 26.3 | 65024.5 | 26.08 | 0.3301 | 40.6021
0 | 0 | 265.6 | 26.42 | 65327.7 | 26.21 | 0.33112 | 40.7272
0 | 0 | 264.4 | 26.53 | 65584.6 | 26.36 | 0.33238 | 40.8827
0 | 0 | 263.2 | 26.75 | 66146.4 | 26.61 | 0.3349 | 41.1921
0 | 0 | 262 | 26.97 | 66670.8 | 26.79 | 0.33652 | 41.392
0 | 0 | 260.8 | 27.05 | 66881.8 | 26.88 | 0.33701 | 41.4521
0 | 0 | 259.6 | 27.34 | 67602.7 | 27.15 | 0.33975 | 41.7888
0 | 0 | 258.4 | 27.47 | 67924.8 | 27.26 | 0.34048 | 41.8782
0 | 0 | 257.2 | 27.65 | 68368.3 | 27.45 | 0.3422 | 42.0897
0 | 0 | 256 | 27.85 | 68849.4 | 27.69 | 0.34453 | 42.3767
0 | 0 | 254.8 | 28.04 | 69317.7 | 27.86 | 0.34598 | 42.5554
0 | 0 | 253.6 | 28.19 | 69703.5 | 28.04 | 0.34755 | 42.7485
0 | 0 | 252.4 | 28.45 | 70348.8 | 28.25 | 0.34948 | 42.9862
0 | 0 | 251.2 | 28.64 | 70815.2 | 28.46 | 0.35141 | 43.2228
0 | 0 | 250 | 28.82 | 71244.8 | 28.62 | 0.35271 | 43.3825
0 | 0 | 248.8 | 29.06 | 71852.1 | 28.86 | 0.35498 | 43.6623
0 | 0 | 247.6 | 29.19 | 72175.6 | 29.03 | 0.35639 | 43.8352
0 | 0 | 246.4 | 29.43 | 72757.2 | 29.3 | 0.35901 | 44.1578
0 | 0 | 245.1 | 29.65 | 73312.1 | 29.45 | 0.36009 | 44.2913
0 | 0 | 244 | 29.9 | 73917.3 | 29.71 | 0.36263 | 44.6034
0 | 0 | 242.7 | 30.03 | 74247.9 | 29.82 | 0.36321 | 44.675
0 | 0 | 241.6 | 30.24 | 74771.8 | 30.06 | 0.36549 | 44.9549
0 | 0 | 240.3 | 30.32 | 74956.5 | 30.08 | 0.36497 | 44.8907
0 | 0 | 239.1 | 30.57 | 75581.7 | 30.35 | 0.36753 | 45.206
0 | 0 | 237.9 | 30.83 | 76219.8 | 30.63 | 0.3702 | 45.5347
0 | 0 | 236.7 | 31.01 | 76674 | 30.79 | 0.37142 | 45.6839
0 | 0 | 235.5 | 31.17 | 77051.9 | 30.99 | 0.3731 | 45.8915
0 | 0 | 234.3 | 31.43 | 77707.8 | 31.23 | 0.37526 | 46.1571
0 | 0 | 233.1 | 31.52 | 77936.1 | 31.35 | 0.37597 | 46.2444
0 | 0 | 231.9 | 31.82 | 78671.4 | 31.59 | 0.37811 | 46.5077
0 | 0 | 230.7 | 32.12 | 79421.2 | 31.9 | 0.38108 | 46.8727
0 | 0 | 229.5 | 32.33 | 79929.1 | 32.1 | 0.38272 | 47.0746
0 | 0 | 228.3 | 32.5 | 80362.2 | 32.32 | 0.38459 | 47.3048
0 | 0 | 227.1 | 32.75 | 80967.7 | 32.54 | 0.38646 | 47.5338
0 | 0 | 225.9 | 32.97 | 81509.4 | 32.75 | 0.38819 | 47.747
0 | 0 | 224.7 | 33.21 | 82105.1 | 32.97 | 0.39003 | 47.9737
0 | 0 | 223.5 | 33.45 | 82702.7 | 33.24 | 0.39246 | 48.2718
0 | 0 | 222.2 | 33.64 | 83165.5 | 33.38 | 0.39327 | 48.3722
0 | 0 | 221 | 33.91 | 83846.6 | 33.69 | 0.39615 | 48.7256
0 | 0 | 219.9 | 34.1 | 84313.7 | 33.85 | 0.39731 | 48.8688
0 | 0 | 218.6 | 34.31 | 84814.5 | 34.1 | 0.39939 | 49.1249
0 | 0 | 217.5 | 34.65 | 85665 | 34.4 | 0.40218 | 49.4676
0 | 0 | 216.2 | 34.78 | 85995 | 34.58 | 0.40342 | 49.6203
0 | 0 | 215 | 35.02 | 86580.8 | 34.76 | 0.40472 | 49.7803
0 | 0 | 213.8 | 35.24 | 87132.5 | 34.98 | 0.40648 | 49.9964
0 | 0 | 212.6 | 35.45 | 87656.8 | 35.26 | 0.40892 | 50.297
0 | 0 | 211.4 | 35.79 | 88497.1 | 35.59 | 0.41193 | 50.6673
0 | 0 | 210.2 | 35.96 | 88911.9 | 35.76 | 0.41308 | 50.8085
0 | 0 | 209 | 36.22 | 89556.4 | 35.98 | 0.4148 | 51.0198
0 | 0 | 207.8 | 36.44 | 90090.1 | 36.21 | 0.41662 | 51.244
0 | 0 | 206.6 | 36.83 | 91054.7 | 36.49 | 0.41901 | 51.5377
0 | 0 | 205.4 | 37 | 91482 | 36.81 | 0.42184 | 51.8863
0 | 0 | 204.2 | 37.19 | 91958.6 | 36.95 | 0.42261 | 51.98
0 | 0 | 203 | 37.47 | 92627.4 | 37.21 | 0.42473 | 52.2415
0 | 0 | 201.8 | 37.77 | 93378.5 | 37.5 | 0.42719 | 52.5437
0 | 0 | 200.6 | 38.03 | 94030.3 | 37.81 | 0.42986 | 52.8723
0 | 0 | 199.4 | 38.27 | 94618.1 | 38.03 | 0.4315 | 53.0737
0 | 0 | 198.2 | 38.55 | 95320.8 | 38.28 | 0.43346 | 53.3157
0 | 0 | 197 | 38.79 | 95913.9 | 38.59 | 0.4361 | 53.6399
0 | 0 | 195.8 | 39.1 | 96666.5 | 38.8 | 0.43759 | 53.8237
0 | 0 | 194.6 | 39.25 | 97033 | 38.94 | 0.43829 | 53.9096
0 | 0 | 193.4 | 39.64 | 98008.2 | 39.35 | 0.44202 | 54.3678
0 | 0 | 192.2 | 39.93 | 98723.7 | 39.68 | 0.44483 | 54.7135
0 | 0 | 191 | 40.15 | 99276.6 | 39.92 | 0.44662 | 54.9337
0 | 0 | 189.8 | 40.52 | 100176.2 | 40.22 | 0.44907 | 55.2351
0 | 0 | 188.5 | 40.73 | 100696.1 | 40.48 | 0.45098 | 55.4707
0 | 0 | 187.4 | 41.03 | 101429 | 40.76 | 0.45326 | 55.751
0 | 0 | 186.1 | 41.28 | 102046.6 | 40.98 | 0.45471 | 55.9293
0 | 0 | 184.9 | 41.5 | 102609.4 | 41.24 | 0.45667 | 56.1703
0 | 0 | 183.7 | 41.87 | 103520.2 | 41.53 | 0.45895 | 56.4508
0 | 0 | 182.5 | 42.14 | 104174.2 | 41.87 | 0.46177 | 56.7976
0 | 0 | 181.3 | 42.38 | 104788.9 | 42.1 | 0.46337 | 56.9938
0 | 0 | 180.1 | 42.66 | 105475.2 | 42.38 | 0.4655 | 57.2563
0 | 0 | 178.9 | 42.95 | 106175.7 | 42.69 | 0.46795 | 57.5579
0 | 0 | 177.7 | 43.3 | 107043.9 | 43.02 | 0.47061 | 57.8848
0 | 0 | 176.5 | 43.56 | 107694.7 | 43.3 | 0.47271 | 58.1429
0 | 0 | 175.3 | 43.95 | 108666.8 | 43.64 | 0.47545 | 58.48
0 | 0 | 174.1 | 44.18 | 109239.1 | 43.86 | 0.47687 | 58.6548
0 | 0 | 172.9 | 44.42 | 109830.9 | 44.13 | 0.47883 | 58.8953
0 | 0 | 171.7 | 44.82 | 110804.1 | 44.5 | 0.48185 | 59.2676
0 | 0 | 170.5 | 45.09 | 111469.7 | 44.85 | 0.48465 | 59.6115
0 | 0 | 169.3 | 45.38 | 112194.3 | 45.07 | 0.48603 | 59.7811
0 | 0 | 168.1 | 45.79 | 113207.6 | 45.41 | 0.48869 | 60.1086
0 | 0 | 166.9 | 46.02 | 113773.1 | 45.71 | 0.49091 | 60.3814
0 | 0 | 165.7 | 46.26 | 114381.3 | 45.98 | 0.49279 | 60.6132
0 | 0 | 164.4 | 46.66 | 115368.9 | 46.4 | 0.49619 | 61.0305
0 | 0 | 163.3 | 46.97 | 116120.2 | 46.7 | 0.49845 | 61.3091
0 | 0 | 162 | 47.22 | 116749.8 | 46.99 | 0.50043 | 61.552
0 | 0 | 160.9 | 47.61 | 117715.4 | 47.29 | 0.50267 | 61.8278
0 | 0 | 159.6 | 47.99 | 118644.5 | 47.77 | 0.50663 | 62.3156
0 | 0 | 158.4 | 48.28 | 119363.4 | 47.94 | 0.50739 | 62.4079
0 | 0 | 157.2 | 48.55 | 120029.6 | 48.28 | 0.50993 | 62.7204
0 | 0 | 156 | 48.92 | 120939.2 | 48.58 | 0.51203 | 62.9793
0 | 0 | 154.8 | 49.25 | 121772.5 | 48.95 | 0.51486 | 63.3273
0 | 0 | 153.6 | 49.62 | 122682.4 | 49.35 | 0.51799 | 63.7122
0 | 0 | 152.4 | 50.03 | 123703.9 | 49.6 | 0.51953 | 63.9018
0 | 0 | 151.2 | 50.22 | 124168.6 | 49.91 | 0.52169 | 64.1673
0 | 0 | 150 | 50.46 | 124759.1 | 50.18 | 0.52342 | 64.38
0 | 0 | 148.8 | 50.87 | 125765.6 | 50.56 | 0.52628 | 64.7323
0 | 0 | 147.6 | 51.25 | 126705.6 | 50.86 | 0.5283 | 64.9804
0 | 0 | 146.4 | 51.59 | 127539 | 51.15 | 0.5302 | 65.2143
0 | 0 | 145.2 | 51.97 | 128493.7 | 51.62 | 0.53395 | 65.6758
0 | 0 | 144 | 52.33 | 129378.5 | 52.01 | 0.53686 | 66.0334
0 | 0 | 142.8 | 52.72 | 130336.2 | 52.47 | 0.54047 | 66.4778
0 | 0 | 141.6 | 53.01 | 131055.7 | 52.71 | 0.54181 | 66.6416
0 | 0 | 140.4 | 53.4 | 132017.8 | 52.97 | 0.54333 | 66.8296
0 | 0 | 139.2 | 53.66 | 132656.7 | 53.35 | 0.54608 | 67.1675
0 | 0 | 138 | 54.05 | 133641 | 53.65 | 0.548 | 67.403
0 | 0 | 136.8 | 54.43 | 134578.9 | 54.14 | 0.55184 | 67.8752
0 | 0 | 135.6 | 54.75 | 135369.7 | 54.45 | 0.55382 | 68.1198
0 | 0 | 134.4 | 55.14 | 136318.4 | 54.76 | 0.5558 | 68.3629
0 | 0 | 133.2 | 55.49 | 137188.5 | 55.1 | 0.55807 | 68.6419
0 | 0 | 131.9 | 55.85 | 138072 | 55.48 | 0.56063 | 68.9568
0 | 0 | 130.8 | 56.25 | 139075.6 | 55.89 | 0.56368 | 69.3315
0 | 0 | 129.6 | 56.71 | 140213 | 56.4 | 0.56761 | 69.8157
0 | 0 | 128.3 | 57.02 | 140961.3 | 56.57 | 0.56801 | 69.8651
0 | 0 | 127.1 | 57.39 | 141899.8 | 56.96 | 0.57071 | 70.1972
0 | 0 | 126 | 57.87 | 143076.8 | 57.42 | 0.5742 | 70.626
0 | 0 | 124.7 | 58.23 | 143964.5 | 57.84 | 0.57707 | 70.9784
0 | 0 | 123.5 | 58.57 | 144802.3 | 58.19 | 0.57932 | 71.2556
0 | 0 | 122.3 | 59.13 | 146193.7 | 58.73 | 0.58345 | 71.7634
0 | 0 | 121.1 | 59.48 | 147066 | 59.09 | 0.58577 | 72.0489
0 | 0 | 119.9 | 59.85 | 147964.3 | 59.48 | 0.58837 | 72.3693
0 | 0 | 118.7 | 60.24 | 148934.6 | 59.89 | 0.59116 | 72.712
0 | 0 | 117.5 | 60.6 | 149836.2 | 60.29 | 0.59383 | 73.0407
0 | 0 | 116.3 | 60.88 | 150522.4 | 60.55 | 0.59511 | 73.1983
0 | 0 | 115.1 | 61.27 | 151475.4 | 60.89 | 0.59717 | 73.4511
0 | 0 | 113.9 | 61.79 | 152758.1 | 61.33 | 0.60019 | 73.8227
0 | 0 | 112.7 | 62.09 | 153506.7 | 61.68 | 0.60232 | 74.0842
0 | 0 | 111.5 | 62.54 | 154614.8 | 62.13 | 0.6054 | 74.4638
0 | 0 | 110.3 | 63.01 | 155792.1 | 62.56 | 0.60828 | 74.8173
0 | 0 | 109.1 | 63.51 | 157019.6 | 63 | 0.61123 | 75.1808
0 | 0 | 107.9 | 64.02 | 158289.7 | 63.56 | 0.61533 | 75.685
0 | 0 | 106.7 | 64.51 | 159496.6 | 63.84 | 0.6167 | 75.8538
0 | 0 | 105.5 | 64.81 | 160233.8 | 64.36 | 0.62038 | 76.3059
0 | 0 | 104.2 | 65.18 | 161158.2 | 64.88 | 0.62392 | 76.7416
0 | 0 | 103.1 | 65.66 | 162345.9 | 65.19 | 0.62565 | 76.9548
0 | 0 | 101.9 | 66.12 | 163470.3 | 65.63 | 0.62851 | 77.3057
0 | 0 | 100.7 | 66.56 | 164553.7 | 66.23 | 0.63287 | 77.8426
0 | 0 | 99.5 | 66.8 | 165151.1 | 66.43 | 0.6334 | 77.9075
0 | 0 | 98.3 | 67.23 | 166216.7 | 66.84 | 0.63592 | 78.2174
0 | 0 | 97.1 | 67.74 | 167469.8 | 67.24 | 0.63833 | 78.5136
0 | 0 | 95.8 | 68.06 | 168266.5 | 67.73 | 0.64146 | 78.8984
0 | 0 | 94.7 | 68.39 | 169076.6 | 68.15 | 0.64414 | 79.2284
0 | 0 | 93.5 | 68.97 | 170521.4 | 68.6 | 0.64697 | 79.5768
0 | 0 | 92.2 | 69.61 | 172091.4 | 68.98 | 0.64901 | 79.8275
0 | 0 | 91.1 | 69.94 | 172923.4 | 69.54 | 0.65296 | 80.3135
0 | 0 | 89.8 | 70.5 | 174294.8 | 70.01 | 0.65581 | 80.6638
0 | 0 | 88.6 | 71.05 | 175650.2 | 70.44 | 0.65838 | 80.9806
0 | 0 | 87.4 | 71.54 | 176880.2 | 71.06 | 0.66272 | 81.5133
0 | 0 | 86.2 | 71.94 | 177867.6 | 71.55 | 0.66581 | 81.8944
0 | 0 | 85 | 72.45 | 179111.9 | 71.91 | 0.66769 | 82.1247
0 | 0 | 83.8 | 72.91 | 180252.2 | 72.35 | 0.67029 | 82.4445
0 | 0 | 82.6 | 73.37 | 181404 | 72.9 | 0.67389 | 82.8874
0 | 0 | 81.4 | 73.82 | 182515.6 | 73.36 | 0.67664 | 83.2256
0 | 0 | 80.2 | 74.32 | 183742 | 73.79 | 0.67909 | 83.5277
0 | 0 | 79 | 74.68 | 184629.3 | 74.18 | 0.68117 | 83.7827
0 | 0 | 77.8 | 75.26 | 186064.7 | 74.7 | 0.68442 | 84.1825
0 | 0 | 76.6 | 75.69 | 187144.5 | 75.09 | 0.68646 | 84.4337
0 | 0 | 75.4 | 75.94 | 187756.6 | 75.41 | 0.68785 | 84.6046
0 | 0 | 74.1 | 76.42 | 188938 | 75.96 | 0.69119 | 85.0157
0 | 0 | 73 | 76.9 | 190125.5 | 76.55 | 0.69513 | 85.5007
0 | 0 | 71.7 | 77.54 | 191710 | 77.05 | 0.69798 | 85.8507
0 | 0 | 70.5 | 78.1 | 193090.9 | 77.55 | 0.70094 | 86.2144
0 | 0 | 69.3 | 78.72 | 194624 | 77.98 | 0.70324 | 86.4982
0 | 0 | 68.1 | 79.04 | 195410 | 78.39 | 0.70535 | 86.7579
0 | 0 | 66.9 | 79.59 | 196786.3 | 79.02 | 0.70943 | 87.2588
0 | 0 | 65.7 | 80.07 | 197953.7 | 79.48 | 0.71195 | 87.5694
0 | 0 | 64.5 | 80.58 | 199215.5 | 79.89 | 0.71401 | 87.8231
0 | 0 | 63.3 | 81.05 | 200396.4 | 80.58 | 0.71856 | 88.382
0 | 0 | 62 | 81.7 | 201999.9 | 81.11 | 0.72152 | 88.7459
0 | 0 | 60.9 | 82.02 | 202779.4 | 81.45 | 0.72304 | 88.9334
0 | 0 | 59.6 | 82.47 | 203895.1 | 82.01 | 0.72623 | 89.3255
0 | 0 | 58.5 | 83.17 | 205630.1 | 82.51 | 0.72914 | 89.6836
0 | 0 | 57.2 | 83.78 | 207122.8 | 83.14 | 0.7329 | 90.1466
0 | 0 | 56 | 84.05 | 207801.8 | 83.42 | 0.7337 | 90.2449
0 | 0 | 54.8 | 84.53 | 208993.7 | 83.93 | 0.73651 | 90.5904
0 | 0 | 53.6 | 84.95 | 210038.5 | 84.43 | 0.73922 | 90.9229
0 | 0 | 52.4 | 85.63 | 211702.6 | 85.05 | 0.74295 | 91.382
0 | 0 | 51.2 | 86.21 | 213142.3 | 85.4 | 0.74431 | 91.5489
0 | 0 | 50 | 86.73 | 214435.1 | 86.13 | 0.74896 | 92.1208
0 | 0 | 48.8 | 87.06 | 215249.2 | 86.55 | 0.75089 | 92.3586
0 | 0 | 47.6 | 87.57 | 216492.8 | 86.98 | 0.75289 | 92.6052
0 | 0 | 46.4 | 88.17 | 217993.2 | 87.38 | 0.75462 | 92.818
0 | 0 | 45.2 | 88.51 | 218825 | 87.97 | 0.75798 | 93.2305
0 | 0 | 44 | 89.08 | 220237.2 | 88.49 | 0.76071 | 93.5664
0 | 0 | 42.8 | 89.42 | 221077.5 | 88.88 | 0.76231 | 93.7629
0 | 0 | 41.6 | 89.98 | 222465.5 | 89.46 | 0.76552 | 94.1577
0 | 0 | 40.4 | 90.66 | 224139.6 | 89.96 | 0.76802 | 94.4659
0 | 0 | 39.1 | 91.34 | 225820.5 | 90.54 | 0.77104 | 94.8375
0 | 0 | 38 | 91.63 | 226542.2 | 90.99 | 0.77324 | 95.1071
0 | 0 | 36.7 | 92.09 | 227684.5 | 91.53 | 0.77588 | 95.4321
0 | 0 | 35.5 | 92.62 | 228982 | 91.93 | 0.77746 | 95.6274
0 | 0 | 34.3 | 93.2 | 230429.1 | 92.41 | 0.77971 | 95.904
0 | 0 | 33.1 | 93.86 | 232066.6 | 93 | 0.78287 | 96.2925
0 | 0 | 31.9 | 94.45 | 233507.6 | 93.77 | 0.78752 | 96.8643
0 | 0 | 30.7 | 94.91 | 234651 | 94.29 | 0.79005 | 97.1751
0 | 0 | 29.5 | 95.33 | 235691.5 | 94.68 | 0.79147 | 97.35
0 | 0 | 28.3 | 95.98 | 237296.9 | 95.29 | 0.79471 | 97.7489
0 | 0 | 27.1 | 96.25 | 237958.1 | 95.63 | 0.79569 | 97.8688
0 | 0 | 25.9 | 96.81 | 239336 | 96.15 | 0.79815 | 98.1712
0 | 0 | 24.7 | 97.32 | 240618.6 | 96.6 | 0.80001 | 98.4001
0 | 0 | 23.5 | 98.01 | 242304.1 | 97.24 | 0.80342 | 98.8201
0 | 0 | 22.3 | 98.33 | 243106.8 | 97.82 | 0.80632 | 99.1766
0 | 0 | 21.1 | 99 | 244773.2 | 98.39 | 0.80912 | 99.5204
0 | 0 | 19.8 | 99.28 | 245457.3 | 98.79 | 0.81034 | 99.6708
0 | 0 | 18.6 | 99.83 | 246816.3 | 99.35 | 0.81301 | 100
0 | 0 | 17.4 | 99.99 | 247202.3 | 99.27 | 0.81045 | 99.6842
0 | 0 | 16.2 | 100.74 | 249072.3 | 99.76 | 0.81253 | 99.9401
0 | 0 | 15 | 100.58 | 248663 | 99.79 | 0.81085 | 99.7342
0 | 0 | 13.8 | 100.78 | 249172.3 | 99.99 | 0.81056 | 99.6979
0 | 0 | 12.6 | 100.54 | 248559 | 100 | 0.80872 | 99.4719
0 | 0 | 11.4 | 100.04 | 247344.5 | 99.33 | 0.8014 | 98.5714
0 | 0 | 10.2 | 99.02 | 244807.6 | 98.17 | 0.79016 | 97.1893
0 | 0 | 9 | 97.79 | 241773.5 | 97.12 | 0.77986 | 95.9215
0 | 0 | 7.8 | 95.33 | 235679.2 | 94.65 | 0.75822 | 93.2597
0 | 0 | 6.6 | 91.96 | 227360.3 | 91.27 | 0.7294 | 89.7153
0 | 0 | 5.4 | 86.89 | 214830.9 | 86.3 | 0.68804 | 84.6279
0 | 0 | 4.2 | 79.72 | 197103.6 | 79.19 | 0.62985 | 77.4704
0 | 0 | 3 | 69.42 | 171637.6 | 68.95 | 0.54709 | 67.2916
0 | 0 | 1.7 | 59.81 | 147860.3 | 59.41 | 0.47017 | 57.8309
0 | 0 | 0.6 | 54.3 | 134255.9 | 53.92 | 0.42579 | 52.3716
0 | 0 | -0.5 | 50.98 | 126046.2 | 50.65 | 0.39909 | 49.0874
Measurement time: | 12/08/2022 18:08:37
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 33.1 | 81833.9 | 33.6 | 0.32548 | 53.7855
0 | 0 | 309.1 | 34.77 | 85965.8 | 33.86 | 0.32755 | 54.1273
0 | 0 | 307.6 | 35.14 | 86867.2 | 34.19 | 0.32998 | 54.5296
0 | 0 | 306.6 | 35.13 | 86861.2 | 34.24 | 0.32996 | 54.5259
0 | 0 | 305.3 | 35.24 | 87125.6 | 34.33 | 0.33017 | 54.5605
0 | 0 | 304.2 | 35.38 | 87477.2 | 34.52 | 0.33144 | 54.77
0 | 0 | 302.9 | 35.77 | 88431.4 | 34.67 | 0.33222 | 54.8984
0 | 0 | 301.7 | 35.69 | 88236.2 | 34.83 | 0.33313 | 55.0502
0 | 0 | 300.5 | 35.81 | 88538.9 | 35.03 | 0.33443 | 55.2643
0 | 0 | 299.3 | 36.29 | 89714.5 | 35.26 | 0.336 | 55.5245
0 | 0 | 298.2 | 36.3 | 89741.2 | 35.37 | 0.33648 | 55.6035
0 | 0 | 296.9 | 36.37 | 89909.5 | 35.49 | 0.33695 | 55.6804
0 | 0 | 295.7 | 36.65 | 90610.7 | 35.73 | 0.3386 | 55.9533
0 | 0 | 294.5 | 36.81 | 91002.6 | 35.89 | 0.33949 | 56.0998
0 | 0 | 293.3 | 37.02 | 91532 | 36.11 | 0.34093 | 56.3391
0 | 0 | 292.1 | 37.24 | 92073.8 | 36.22 | 0.34134 | 56.4059
0 | 0 | 290.9 | 37.22 | 92019.1 | 36.38 | 0.34221 | 56.5499
0 | 0 | 289.7 | 37.58 | 92909.5 | 36.55 | 0.34317 | 56.7085
0 | 0 | 288.5 | 37.62 | 93011.1 | 36.73 | 0.34422 | 56.8818
0 | 0 | 287.3 | 37.97 | 93882.8 | 36.92 | 0.34535 | 57.0696
0 | 0 | 286.1 | 37.97 | 93884.1 | 37.04 | 0.34583 | 57.1484
0 | 0 | 284.9 | 38.3 | 94692.1 | 37.31 | 0.3477 | 57.4576
0 | 0 | 283.7 | 38.39 | 94907.2 | 37.48 | 0.34863 | 57.6117
0 | 0 | 282.5 | 38.67 | 95602.5 | 37.49 | 0.34808 | 57.5193
0 | 0 | 281.3 | 38.79 | 95903.7 | 37.79 | 0.35021 | 57.8712
0 | 0 | 280.1 | 39.1 | 96663.6 | 38.05 | 0.35195 | 58.1602
0 | 0 | 278.9 | 39.2 | 96917.8 | 38.26 | 0.35323 | 58.3716
0 | 0 | 277.7 | 39.33 | 97229.6 | 38.41 | 0.35395 | 58.4906
0 | 0 | 276.5 | 39.57 | 97820.6 | 38.54 | 0.35448 | 58.5783
0 | 0 | 275.2 | 39.79 | 98383.3 | 38.67 | 0.35496 | 58.6563
0 | 0 | 274.1 | 39.92 | 98698.7 | 38.95 | 0.35691 | 58.9791
0 | 0 | 272.8 | 40.23 | 99458.6 | 39.24 | 0.35883 | 59.297
0 | 0 | 271.6 | 40.47 | 100065.5 | 39.41 | 0.35971 | 59.4417
0 | 0 | 270.4 | 40.53 | 100198 | 39.53 | 0.36012 | 59.5102
0 | 0 | 269.2 | 40.75 | 100740.1 | 39.78 | 0.36172 | 59.7735
0 | 0 | 268 | 40.83 | 100949.8 | 39.95 | 0.36258 | 59.9155
0 | 0 | 266.8 | 41.2 | 101872.6 | 40.18 | 0.36397 | 60.1464
0 | 0 | 265.6 | 41.44 | 102466.2 | 40.27 | 0.3641 | 60.167
0 | 0 | 264.4 | 41.48 | 102555.3 | 40.55 | 0.36593 | 60.4705
0 | 0 | 263.2 | 41.73 | 103173.7 | 40.79 | 0.3674 | 60.713
0 | 0 | 262 | 41.8 | 103346.9 | 40.88 | 0.36751 | 60.7314
0 | 0 | 260.8 | 42.34 | 104690.9 | 41.15 | 0.36924 | 61.0163
0 | 0 | 259.6 | 42.28 | 104524.2 | 41.23 | 0.36925 | 61.0186
0 | 0 | 258.4 | 42.52 | 105128.6 | 41.43 | 0.37034 | 61.1978
0 | 0 | 257.2 | 42.7 | 105580.4 | 41.63 | 0.37141 | 61.376
0 | 0 | 256 | 42.91 | 106100.2 | 41.91 | 0.3732 | 61.671
0 | 0 | 254.8 | 43.13 | 106644.1 | 42.11 | 0.37426 | 61.8469
0 | 0 | 253.6 | 43.35 | 107175.1 | 42.24 | 0.3747 | 61.9192
0 | 0 | 252.4 | 43.66 | 107930.9 | 42.49 | 0.3762 | 62.1665
0 | 0 | 251.2 | 43.79 | 108254.8 | 42.69 | 0.37725 | 62.3395
0 | 0 | 250 | 44.01 | 108796.4 | 42.94 | 0.37873 | 62.5844
0 | 0 | 248.8 | 44.22 | 109320.4 | 43.14 | 0.37976 | 62.7552
0 | 0 | 247.6 | 44.34 | 109623.4 | 43.41 | 0.3814 | 63.0266
0 | 0 | 246.4 | 44.62 | 110311.2 | 43.65 | 0.38277 | 63.2532
0 | 0 | 245.2 | 45.02 | 111302.9 | 43.86 | 0.38388 | 63.4352
0 | 0 | 244 | 45.16 | 111655.1 | 44.01 | 0.38445 | 63.5296
0 | 0 | 242.7 | 45.19 | 111717 | 44.19 | 0.38521 | 63.6561
0 | 0 | 241.6 | 45.55 | 112607.3 | 44.5 | 0.38723 | 63.9893
0 | 0 | 240.4 | 45.78 | 113175.7 | 44.69 | 0.38813 | 64.1383
0 | 0 | 239.1 | 46.12 | 114034.2 | 44.98 | 0.38983 | 64.4193
0 | 0 | 238 | 46.3 | 114477.2 | 45.1 | 0.39018 | 64.4765
0 | 0 | 236.7 | 46.45 | 114836.5 | 45.43 | 0.39221 | 64.812
0 | 0 | 235.5 | 46.76 | 115595.3 | 45.64 | 0.39326 | 64.9853
0 | 0 | 234.3 | 46.96 | 116109.8 | 45.78 | 0.3937 | 65.058
0 | 0 | 233.1 | 47.15 | 116568.6 | 45.89 | 0.39388 | 65.0876
0 | 0 | 231.9 | 47.3 | 116937.3 | 46.25 | 0.39619 | 65.4706
0 | 0 | 230.7 | 47.51 | 117469.4 | 46.53 | 0.39782 | 65.7387
0 | 0 | 229.5 | 47.89 | 118410.4 | 46.69 | 0.39841 | 65.8362
0 | 0 | 228.3 | 48.21 | 119182.8 | 47.01 | 0.40035 | 66.1581
0 | 0 | 227.1 | 48.32 | 119459.3 | 47.15 | 0.40076 | 66.2255
0 | 0 | 225.9 | 48.46 | 119810.1 | 47.28 | 0.40108 | 66.2783
0 | 0 | 224.7 | 48.7 | 120409.2 | 47.73 | 0.40411 | 66.7782
0 | 0 | 223.5 | 49.19 | 121612.5 | 48.03 | 0.40585 | 67.0663
0 | 0 | 222.3 | 49.26 | 121778.5 | 48.13 | 0.4059 | 67.0742
0 | 0 | 221.1 | 49.43 | 122205.7 | 48.38 | 0.4072 | 67.2902
0 | 0 | 219.9 | 49.69 | 122855.5 | 48.69 | 0.40901 | 67.5884
0 | 0 | 218.7 | 49.99 | 123599.2 | 48.83 | 0.40938 | 67.6494
0 | 0 | 217.5 | 50.5 | 124851.1 | 49.18 | 0.4115 | 68.0002
0 | 0 | 216.3 | 50.43 | 124683.7 | 49.35 | 0.41211 | 68.1008
0 | 0 | 215.1 | 50.86 | 125741.9 | 49.67 | 0.41396 | 68.4072
0 | 0 | 213.9 | 51.04 | 126201.1 | 49.76 | 0.41389 | 68.3959
0 | 0 | 212.6 | 51.42 | 127121 | 50.07 | 0.41558 | 68.6746
0 | 0 | 211.5 | 51.59 | 127555.2 | 50.47 | 0.41814 | 69.0977
0 | 0 | 210.2 | 51.87 | 128241.4 | 50.55 | 0.41791 | 69.0588
0 | 0 | 209 | 51.9 | 128327.1 | 50.86 | 0.41964 | 69.3446
0 | 0 | 207.9 | 52.28 | 129256.6 | 50.93 | 0.41945 | 69.3137
0 | 0 | 206.6 | 52.55 | 129922.6 | 51.35 | 0.422 | 69.735
0 | 0 | 205.4 | 52.72 | 130336.3 | 51.58 | 0.42305 | 69.9081
0 | 0 | 204.2 | 52.84 | 130634.9 | 51.9 | 0.42482 | 70.2018
0 | 0 | 203 | 53.25 | 131647.5 | 52.15 | 0.42602 | 70.3995
0 | 0 | 201.8 | 53.68 | 132719 | 52.34 | 0.42672 | 70.5151
0 | 0 | 200.6 | 53.82 | 133070.9 | 52.63 | 0.42823 | 70.7642
0 | 0 | 199.4 | 54.28 | 134208.1 | 52.99 | 0.43029 | 71.1059
0 | 0 | 198.2 | 54.39 | 134472.9 | 53.13 | 0.43057 | 71.1512
0 | 0 | 197 | 54.53 | 134812.3 | 53.31 | 0.43116 | 71.2493
0 | 0 | 195.8 | 54.82 | 135523.3 | 53.4 | 0.43102 | 71.2266
0 | 0 | 194.6 | 55.28 | 136675.6 | 53.86 | 0.43387 | 71.696
0 | 0 | 193.4 | 55.44 | 137059.9 | 54.27 | 0.43629 | 72.0967
0 | 0 | 192.2 | 55.94 | 138313.7 | 54.42 | 0.43662 | 72.1507
0 | 0 | 191 | 55.89 | 138173 | 54.63 | 0.43742 | 72.2834
0 | 0 | 189.8 | 56.37 | 139370.8 | 54.85 | 0.4383 | 72.4283
0 | 0 | 188.5 | 56.54 | 139782.9 | 55.28 | 0.44077 | 72.8367
0 | 0 | 187.4 | 56.76 | 140334.3 | 55.51 | 0.44178 | 73.0044
0 | 0 | 186.2 | 57.11 | 141202.1 | 55.66 | 0.44208 | 73.0538
0 | 0 | 185 | 57.36 | 141825.2 | 55.97 | 0.44365 | 73.3121
0 | 0 | 183.8 | 57.51 | 142187.8 | 56.27 | 0.44512 | 73.5559
0 | 0 | 182.6 | 57.95 | 143262.1 | 56.66 | 0.4473 | 73.9156
0 | 0 | 181.4 | 58.36 | 144276.7 | 57 | 0.44907 | 74.2083
0 | 0 | 180.1 | 58.67 | 145054.6 | 57.29 | 0.45036 | 74.4218
0 | 0 | 178.9 | 58.77 | 145298.2 | 57.48 | 0.45094 | 74.5168
0 | 0 | 177.8 | 59.17 | 146297.5 | 57.81 | 0.45268 | 74.8049
0 | 0 | 176.5 | 59.38 | 146817.3 | 58.12 | 0.4541 | 75.0401
0 | 0 | 175.3 | 59.84 | 147933.2 | 58.26 | 0.45427 | 75.0674
0 | 0 | 174.1 | 60.06 | 148496.6 | 58.65 | 0.45638 | 75.4157
0 | 0 | 172.9 | 60.06 | 148491.9 | 58.88 | 0.45723 | 75.5568
0 | 0 | 171.7 | 60.57 | 149746.6 | 59.32 | 0.4597 | 75.9657
0 | 0 | 170.5 | 60.97 | 150738.4 | 59.52 | 0.46031 | 76.0658
0 | 0 | 169.3 | 61.38 | 151744.6 | 59.9 | 0.4623 | 76.3946
0 | 0 | 168.1 | 61.46 | 151941.3 | 60.12 | 0.46304 | 76.5178
0 | 0 | 166.9 | 61.85 | 152921.9 | 60.38 | 0.46409 | 76.6909
0 | 0 | 165.7 | 61.84 | 152889.6 | 60.61 | 0.4649 | 76.8248
0 | 0 | 164.5 | 62.51 | 154552.7 | 60.86 | 0.46586 | 76.983
0 | 0 | 163.3 | 62.51 | 154555.2 | 61.28 | 0.46811 | 77.3546
0 | 0 | 162.1 | 62.99 | 155721.9 | 61.58 | 0.46943 | 77.573
0 | 0 | 160.9 | 63.37 | 156677.2 | 61.78 | 0.46998 | 77.6643
0 | 0 | 159.7 | 63.84 | 157839.8 | 62.2 | 0.4722 | 78.0307
0 | 0 | 158.5 | 63.96 | 158125.9 | 62.42 | 0.47289 | 78.1447
0 | 0 | 157.3 | 64.15 | 158597.7 | 62.75 | 0.47441 | 78.3952
0 | 0 | 156.1 | 64.65 | 159834.5 | 63.14 | 0.47636 | 78.7189
0 | 0 | 154.9 | 64.74 | 160059.3 | 63.38 | 0.47718 | 78.8542
0 | 0 | 153.6 | 65.28 | 161406 | 63.55 | 0.47739 | 78.8878
0 | 0 | 152.4 | 65.49 | 161908.2 | 63.89 | 0.47894 | 79.1449
0 | 0 | 151.3 | 65.75 | 162559.3 | 64.34 | 0.4814 | 79.5503
0 | 0 | 150 | 66.03 | 163240.5 | 64.61 | 0.48232 | 79.7038
0 | 0 | 148.8 | 66.62 | 164701.8 | 65.13 | 0.48519 | 80.1777
0 | 0 | 147.6 | 66.91 | 165421.5 | 65.34 | 0.48574 | 80.2683
0 | 0 | 146.4 | 67.3 | 166391.9 | 65.69 | 0.48732 | 80.5296
0 | 0 | 145.2 | 67.48 | 166823.8 | 66.01 | 0.48867 | 80.7525
0 | 0 | 144 | 67.74 | 167478.5 | 66.09 | 0.48824 | 80.681
0 | 0 | 142.8 | 68.22 | 168672.6 | 66.45 | 0.48987 | 80.9504
0 | 0 | 141.6 | 68.57 | 169526.7 | 66.87 | 0.49193 | 81.2911
0 | 0 | 140.4 | 68.72 | 169893.8 | 67.4 | 0.49479 | 81.7632
0 | 0 | 139.2 | 69.14 | 170932.3 | 67.53 | 0.4947 | 81.7486
0 | 0 | 138 | 69.66 | 172223.8 | 67.83 | 0.49585 | 81.9389
0 | 0 | 136.8 | 69.83 | 172652.2 | 68.15 | 0.49714 | 82.1519
0 | 0 | 135.6 | 70.25 | 173688.4 | 68.49 | 0.49857 | 82.3876
0 | 0 | 134.4 | 70.56 | 174440.5 | 69.08 | 0.5018 | 82.9218
0 | 0 | 133.2 | 71.13 | 175854.2 | 69.38 | 0.50291 | 83.1058
0 | 0 | 132 | 71.28 | 176236.4 | 69.61 | 0.50351 | 83.2048
0 | 0 | 130.8 | 71.66 | 177173.7 | 70.09 | 0.50591 | 83.601
0 | 0 | 129.6 | 71.98 | 177967.1 | 70.42 | 0.50721 | 83.8164
0 | 0 | 128.4 | 72.48 | 179197.9 | 70.78 | 0.50872 | 84.066
0 | 0 | 127.2 | 73.11 | 180757.5 | 71.13 | 0.51015 | 84.3021
0 | 0 | 126 | 72.71 | 179764.5 | 71.32 | 0.51043 | 84.3474
0 | 0 | 124.8 | 73.56 | 181862.9 | 71.68 | 0.51191 | 84.5926
0 | 0 | 123.6 | 73.67 | 182138.4 | 72.07 | 0.5136 | 84.8715
0 | 0 | 122.3 | 74.12 | 183260.5 | 72.59 | 0.51611 | 85.2861
0 | 0 | 121.1 | 74.59 | 184417.5 | 72.88 | 0.51706 | 85.4438
0 | 0 | 119.9 | 74.86 | 185087.5 | 73.09 | 0.51744 | 85.5067
0 | 0 | 118.7 | 75.32 | 186226.4 | 73.51 | 0.5193 | 85.8139
0 | 0 | 117.6 | 75.32 | 186214.2 | 73.74 | 0.5199 | 85.9131
0 | 0 | 116.3 | 76.08 | 188090.1 | 74.31 | 0.5227 | 86.3759
0 | 0 | 115.1 | 76.29 | 188617.7 | 74.61 | 0.52368 | 86.5383
0 | 0 | 113.9 | 76.78 | 189822.8 | 74.99 | 0.52522 | 86.792
0 | 0 | 112.7 | 77.11 | 190648.5 | 75.32 | 0.52639 | 86.9862
0 | 0 | 111.5 | 77.32 | 191166.2 | 75.6 | 0.52721 | 87.1213
0 | 0 | 110.3 | 77.99 | 192828.2 | 75.93 | 0.52837 | 87.3128
0 | 0 | 109.1 | 78.05 | 192974.6 | 76.55 | 0.53153 | 87.8355
0 | 0 | 107.9 | 78.49 | 194044.7 | 76.75 | 0.53177 | 87.8746
0 | 0 | 106.7 | 79.09 | 195531.7 | 77.14 | 0.53331 | 88.1299
0 | 0 | 105.5 | 79.45 | 196422.9 | 77.31 | 0.53333 | 88.1327
0 | 0 | 104.3 | 79.64 | 196899.2 | 77.78 | 0.53541 | 88.4761
0 | 0 | 103.1 | 79.85 | 197410 | 78.43 | 0.53871 | 89.0217
0 | 0 | 101.9 | 80.69 | 199491.6 | 78.52 | 0.53816 | 88.93
0 | 0 | 100.7 | 80.88 | 199964.6 | 79.08 | 0.54081 | 89.3693
0 | 0 | 99.5 | 81.22 | 200802.8 | 79.39 | 0.54175 | 89.5241
0 | 0 | 98.3 | 81.66 | 201899.7 | 79.8 | 0.54336 | 89.7901
0 | 0 | 97.1 | 82.06 | 202874.7 | 80.12 | 0.54435 | 89.9533
0 | 0 | 95.9 | 82.34 | 203578.6 | 80.47 | 0.54553 | 90.1487
0 | 0 | 94.7 | 82.86 | 204863.3 | 81.08 | 0.54846 | 90.6333
0 | 0 | 93.5 | 83.28 | 205906.9 | 81.37 | 0.54922 | 90.7581
0 | 0 | 92.2 | 83.68 | 206889.1 | 82.04 | 0.55243 | 91.288
0 | 0 | 91.1 | 84.36 | 208557.7 | 82.21 | 0.55246 | 91.293
0 | 0 | 89.9 | 84.53 | 208986.4 | 82.73 | 0.55473 | 91.6685
0 | 0 | 88.7 | 84.97 | 210064.3 | 82.89 | 0.55458 | 91.6436
0 | 0 | 87.5 | 85.23 | 210707.1 | 83.32 | 0.55623 | 91.9161
0 | 0 | 86.2 | 85.86 | 212269.2 | 83.93 | 0.55896 | 92.3678
0 | 0 | 85.1 | 86.28 | 213309.7 | 84.21 | 0.55969 | 92.4883
0 | 0 | 83.8 | 86.51 | 213875.1 | 84.55 | 0.5606 | 92.6394
0 | 0 | 82.6 | 87.07 | 215264.3 | 85.1 | 0.563 | 93.0356
0 | 0 | 81.4 | 87.27 | 215761.8 | 85.35 | 0.5634 | 93.1022
0 | 0 | 80.2 | 87.72 | 216866.4 | 85.79 | 0.56505 | 93.3746
0 | 0 | 79 | 88.27 | 218232.7 | 86.04 | 0.56544 | 93.4387
0 | 0 | 77.8 | 88.51 | 218832.3 | 86.46 | 0.56694 | 93.6861
0 | 0 | 76.6 | 88.76 | 219437.7 | 86.88 | 0.56842 | 93.9317
0 | 0 | 75.4 | 89.51 | 221307.9 | 87.31 | 0.56996 | 94.1863
0 | 0 | 74.2 | 89.81 | 222037 | 87.8 | 0.57188 | 94.5036
0 | 0 | 73 | 90.55 | 223876.1 | 88.32 | 0.57399 | 94.8511
0 | 0 | 71.8 | 90.45 | 223614.5 | 88.58 | 0.57439 | 94.9176
0 | 0 | 70.5 | 91.08 | 225173 | 89.17 | 0.57681 | 95.3182
0 | 0 | 69.3 | 91.66 | 226607.8 | 89.55 | 0.57797 | 95.5099
0 | 0 | 68.1 | 91.97 | 227391.5 | 89.73 | 0.57784 | 95.4872
0 | 0 | 66.9 | 92.47 | 228615.1 | 90.37 | 0.58065 | 95.9523
0 | 0 | 65.7 | 92.99 | 229906.7 | 91.01 | 0.58345 | 96.4146
0 | 0 | 64.5 | 93.38 | 230873.5 | 91.29 | 0.58393 | 96.4935
0 | 0 | 63.3 | 93.72 | 231716.6 | 91.52 | 0.58408 | 96.5187
0 | 0 | 62.1 | 94.17 | 232830 | 91.95 | 0.5855 | 96.7534
0 | 0 | 60.9 | 94.7 | 234124.4 | 92.46 | 0.58742 | 97.0703
0 | 0 | 59.7 | 94.88 | 234581.7 | 92.85 | 0.58856 | 97.2594
0 | 0 | 58.5 | 95.4 | 235855.4 | 93.28 | 0.58995 | 97.4886
0 | 0 | 57.3 | 96.07 | 237512.2 | 93.54 | 0.59025 | 97.5388
0 | 0 | 56.1 | 96.47 | 238510.3 | 94.11 | 0.5925 | 97.9106
0 | 0 | 54.9 | 96.46 | 238493.9 | 94.7 | 0.59486 | 98.3006
0 | 0 | 53.6 | 97.28 | 240511.1 | 94.85 | 0.59434 | 98.2138
0 | 0 | 52.4 | 97.52 | 241096.7 | 95.35 | 0.59611 | 98.5068
0 | 0 | 51.2 | 97.99 | 242256.7 | 95.91 | 0.59824 | 98.8595
0 | 0 | 50 | 98.06 | 242449.4 | 96.14 | 0.59831 | 98.8704
0 | 0 | 48.8 | 98.69 | 243986.8 | 96.36 | 0.59831 | 98.8703
0 | 0 | 47.6 | 99.07 | 244945.8 | 96.78 | 0.59954 | 99.0741
0 | 0 | 46.4 | 99.32 | 245552.9 | 96.96 | 0.59928 | 99.0311
0 | 0 | 45.2 | 99.71 | 246528.1 | 97.43 | 0.60081 | 99.2831
0 | 0 | 44 | 100.15 | 247596.8 | 97.82 | 0.60183 | 99.4517
0 | 0 | 42.8 | 100.36 | 248131.4 | 98.12 | 0.60229 | 99.5275
0 | 0 | 41.6 | 100.88 | 249405.3 | 98.3 | 0.602 | 99.4808
0 | 0 | 40.4 | 101.24 | 250306.6 | 98.7 | 0.60306 | 99.6556
0 | 0 | 39.2 | 100.95 | 249574.9 | 99.27 | 0.60515 | 100
0 | 0 | 38 | 101.5 | 250941.2 | 99.27 | 0.60375 | 99.7692
0 | 0 | 36.8 | 102.1 | 252438.1 | 99.52 | 0.60387 | 99.7893
0 | 0 | 35.5 | 101.93 | 252018.7 | 99.72 | 0.60357 | 99.7393
0 | 0 | 34.3 | 102.32 | 252970.9 | 99.96 | 0.60362 | 99.7477
0 | 0 | 33.1 | 102.06 | 252335.1 | 100 | 0.60246 | 99.5562
0 | 0 | 31.9 | 102.23 | 252756.5 | 99.91 | 0.60052 | 99.2357
0 | 0 | 30.7 | 102.3 | 252930.6 | 99.7 | 0.59787 | 98.7969
0 | 0 | 29.5 | 102.27 | 252848.2 | 99.79 | 0.59701 | 98.656
0 | 0 | 28.3 | 101.94 | 252031.8 | 99.74 | 0.59532 | 98.3768
0 | 0 | 27.1 | 101.62 | 251250.9 | 99.46 | 0.59227 | 97.8718
0 | 0 | 25.9 | 101.66 | 251350.4 | 99.28 | 0.58981 | 97.4665
0 | 0 | 24.7 | 100.87 | 249397.5 | 98.66 | 0.58476 | 96.6314
0 | 0 | 23.5 | 100.65 | 248854.2 | 98.45 | 0.58215 | 96.2
0 | 0 | 22.3 | 99.78 | 246694.8 | 97.51 | 0.57524 | 95.0582
0 | 0 | 21 | 98.69 | 243992 | 96.83 | 0.56978 | 94.1554
0 | 0 | 19.8 | 98.07 | 242464.2 | 95.77 | 0.56222 | 92.9059
0 | 0 | 18.7 | 97 | 239813.6 | 94.61 | 0.55421 | 91.5827
0 | 0 | 17.4 | 95.7 | 236611 | 93.43 | 0.5459 | 90.2098
0 | 0 | 16.3 | 94 | 232402.7 | 91.95 | 0.53609 | 88.5889
0 | 0 | 15 | 92.08 | 227646.9 | 90 | 0.52338 | 86.4885
0 | 0 | 13.8 | 90.06 | 222648.8 | 87.87 | 0.50979 | 84.242
0 | 0 | 12.6 | 87.56 | 216481.9 | 85.42 | 0.4944 | 81.6994
0 | 0 | 11.4 | 84.36 | 208571.4 | 82.84 | 0.47833 | 79.0441
0 | 0 | 10.2 | 81.48 | 201459.3 | 79.5 | 0.45796 | 75.6773
0 | 0 | 9 | 77.7 | 192105.4 | 75.86 | 0.43595 | 72.0409
0 | 0 | 7.8 | 72.85 | 180115.4 | 71.02 | 0.40717 | 67.2842
0 | 0 | 6.6 | 67.39 | 166609.8 | 65.82 | 0.37646 | 62.2093
0 | 0 | 5.4 | 60.8 | 150329.2 | 59.23 | 0.33796 | 55.8474
0 | 0 | 4.2 | 53.13 | 131355.7 | 51.76 | 0.29463 | 48.6876
0 | 0 | 3 | 44.31 | 109548.9 | 43.46 | 0.24679 | 40.7826
0 | 0 | 1.8 | 37.36 | 92377 | 36.44 | 0.20644 | 34.1133
0 | 0 | 0.6 | 33.46 | 82728.8 | 32.77 | 0.1852 | 30.6042
0 | 0 | -0.5 | 31.45 | 77754.3 | 30.64 | 0.17278 | 28.5521