## Feuil1
Measurement time: | 02/09/2022 21:00:33 |  | 0,3 cm/s |  |  | Vitesse (cm/s) | Symétrie (%) | Homogénéité (%) | Pénombre G (cm) | Pénombre D (cm)
Radiation device: | Clinac3 Accelerator |  |  |  |  | 0.3 | 4.16 | 5.69 | 1.29 | 1.27
Energy: | 12 MeV |  |  |  |  | 1 | 3.25 | 5.78 | 1.29 | 1.28
Controller: | CCU |  |  |  |  | 2 | 3.99 | 5.82 | 1.28 | 1.27
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-151.4 | 0 | 28 | 2 | 3947.6 | 2.32
-150.7 | 0 | 28 | 1.28 | 2537.7 | 1.29
-149.3 | 0 | 28 | 1.28 | 2536.5 | 1.26
-148.2 | 0 | 28 | 1.27 | 2514.4 | 1.32
-146.9 | 0 | 28 | 1.26 | 2490.6 | 1.25
-145.7 | 0 | 28 | 1.25 | 2470.2 | 1.26
-144.5 | 0 | 28 | 1.24 | 2452.9 | 1.24
-143.3 | 0 | 28 | 1.26 | 2493 | 1.27
-142.2 | 0 | 28 | 1.23 | 2435.5 | 1.22
-140.9 | 0 | 28 | 1.24 | 2442.1 | 1.22
-139.7 | 0 | 28 | 1.23 | 2433.4 | 1.22
-138.5 | 0 | 28 | 1.28 | 2536.1 | 1.27
-137.3 | 0 | 28 | 1.24 | 2447.5 | 1.21
-136.1 | 0 | 28 | 1.27 | 2517.4 | 1.27
-134.9 | 0 | 28 | 1.26 | 2491.4 | 1.25
-133.7 | 0 | 28 | 1.26 | 2480.5 | 1.23
-132.5 | 0 | 28 | 1.28 | 2520.3 | 1.31
-131.3 | 0 | 28 | 1.28 | 2529.4 | 1.27
-130.1 | 0 | 28 | 1.29 | 2542.6 | 1.29
-128.9 | 0 | 28 | 1.27 | 2517.7 | 1.27
-127.7 | 0 | 28 | 1.32 | 2604.5 | 1.32
-126.5 | 0 | 28 | 1.31 | 2586.2 | 1.31
-125.3 | 0 | 28 | 1.3 | 2568.5 | 1.29
-124.1 | 0 | 28 | 1.28 | 2523.9 | 1.27
-122.9 | 0 | 28 | 1.31 | 2593.2 | 1.31
-121.7 | 0 | 28 | 1.26 | 2485.7 | 1.22
-120.5 | 0 | 28 | 1.3 | 2559.6 | 1.28
-119.3 | 0 | 28 | 1.34 | 2652.1 | 1.4
-118.1 | 0 | 28 | 1.32 | 2600.3 | 1.28
-116.9 | 0 | 28 | 1.26 | 2478.7 | 1.26
-115.7 | 0 | 28 | 1.3 | 2560.9 | 1.32
-114.5 | 0 | 28 | 1.26 | 2496.4 | 1.24
-113.3 | 0 | 28 | 1.29 | 2553.5 | 1.32
-112.1 | 0 | 28 | 1.24 | 2455 | 1.24
-110.9 | 0 | 28 | 1.26 | 2481.5 | 1.25
-109.7 | 0 | 28 | 1.25 | 2466 | 1.26
-108.5 | 0 | 28 | 1.27 | 2511.7 | 1.29
-107.3 | 0 | 28 | 1.23 | 2426.8 | 1.22
-106.1 | 0 | 28 | 1.26 | 2489.6 | 1.27
-104.9 | 0 | 28 | 1.24 | 2449.3 | 1.24
-103.7 | 0 | 28 | 1.17 | 2312.1 | 1.16
-102.5 | 0 | 28 | 1.22 | 2402.6 | 1.23
-101.3 | 0 | 28 | 1.18 | 2333.9 | 1.18
-100.1 | 0 | 28 | 1.17 | 2315.9 | 1.16
-98.9 | 0 | 28 | 1.12 | 2217.6 | 1.13
-97.7 | 0 | 28 | 1.1 | 2169.9 | 1.07
-96.5 | 0 | 28 | 1.11 | 2190.4 | 1.08
-95.3 | 0 | 28 | 1.1 | 2166.4 | 1.1
-94.1 | 0 | 28 | 1.05 | 2069.9 | 1.04
-92.9 | 0 | 28 | 1.03 | 2040 | 1.02
-91.6 | 0 | 28 | 0.98 | 1936 | 0.99
-90.5 | 0 | 28 | 0.9 | 1783.4 | 0.91
-89.3 | 0 | 28 | 0.8 | 1572.1 | 0.78
-88 | 0 | 28 | 0.74 | 1463.2 | 0.74
-86.9 | 0 | 28 | 0.69 | 1357.1 | 0.69
-85.6 | 0 | 28 | 0.47 | 925.3 | 0.45
-84.5 | 0 | 28 | 0.44 | 874.9 | 0.45
-83.2 | 0 | 28 | 0.32 | 639.8 | 0.32
-82 | 0 | 28 | 0.2 | 395.4 | 0.2
-80.8 | 0 | 28 | 0.03 | 63.6 | 0.03
-79.6 | 0 | 28 | 0.2 | 387 | 0.19
-78.4 | 0 | 28 | 0.46 | 901.6 | 0.44
-77.2 | 0 | 28 | 0.62 | 1217.3 | 0.62
-76 | 0 | 28 | 0.83 | 1639.2 | 0.86
-74.8 | 0 | 28 | 1.47 | 2893.6 | 1.39
-73.6 | 0 | 28 | 1.62 | 3190.9 | 1.67
-72.4 | 0 | 28 | 2.28 | 4494.1 | 2.24
-71.2 | 0 | 28 | 2.8 | 5523.5 | 2.78
-70 | 0 | 28 | 3.38 | 6676.6 | 3.45
-68.8 | 0 | 28 | 4.39 | 8669.3 | 4.31
-67.6 | 0 | 28 | 5.79 | 11440.4 | 5.45
-66.4 | 0 | 28 | 6.29 | 12419.3 | 6.45
-65.2 | 0 | 28 | 8.04 | 15887.7 | 8
-64 | 0 | 28 | 10.24 | 20222.7 | 9.88
-62.8 | 0 | 28 | 11.93 | 23571.9 | 11.99
-61.6 | 0 | 28 | 14.7 | 29025.1 | 14.78
-60.4 | 0 | 28 | 17.96 | 35474.8 | 17.96
-59.2 | 0 | 28 | 22.02 | 43497.3 | 22.02
-58 | 0 | 28 | 25.09 | 49547.4 | 26.32
-56.8 | 0 | 28 | 32.93 | 65035.8 | 31.68
-55.6 | 0 | 28 | 35.98 | 71071.3 | 37.6
-54.4 | 0 | 28 | 44.2 | 87294.8 | 44
-53.2 | 0 | 28 | 53.74 | 106141.9 | 51.05
-52 | 0 | 28 | 55.7 | 110011.8 | 57.45
-50.8 | 0 | 28 | 61.67 | 121810 | 63.66
-49.6 | 0 | 28 | 73.57 | 145308.9 | 69.86
-48.4 | 0 | 28 | 71.95 | 142108.4 | 74.64
-47.2 | 0 | 28 | 79.56 | 157131.4 | 79.31
-46 | 0 | 28 | 87.1 | 172030.6 | 83.52
-44.8 | 0 | 28 | 82.44 | 162832.8 | 85.97
-43.6 | 0 | 28 | 88.76 | 175306.8 | 88.6
-42.4 | 0 | 28 | 95.22 | 188069.7 | 90.97
-41.2 | 0 | 28 | 89.67 | 177110 | 92.03
-39.9 | 0 | 28 | 94.1 | 185864.6 | 93.44
-38.8 | 0 | 28 | 99.84 | 197193.2 | 94.71
-37.6 | 0 | 28 | 93.17 | 184016.8 | 95.24
-36.3 | 0 | 28 | 97.72 | 192999.5 | 96.2
-35.2 | 0 | 28 | 97.06 | 191705.7 | 96.7
-33.9 | 0 | 28 | 95.26 | 188151.2 | 97.06
-32.7 | 0 | 28 | 97.23 | 192033.1 | 97.41
-31.5 | 0 | 28 | 100.25 | 197998.1 | 97.9
-30.3 | 0 | 28 | 97.12 | 191826.3 | 98.01
-29.1 | 0 | 28 | 96.49 | 190583.3 | 98.23
-27.9 | 0 | 28 | 102.42 | 202284.4 | 98.7
-26.7 | 0 | 28 | 96.81 | 191216.5 | 98.82
-25.5 | 0 | 28 | 99.75 | 197015.7 | 98.96
-24.3 | 0 | 28 | 102.78 | 203006 | 99.29
-23.1 | 0 | 28 | 97.72 | 193010.7 | 99.31
-21.9 | 0 | 28 | 96.86 | 191307.4 | 99.3
-20.7 | 0 | 28 | 102.83 | 203108.3 | 99.55
-19.5 | 0 | 28 | 95.38 | 188390.8 | 99.37
-18.3 | 0 | 28 | 101.47 | 200405.3 | 99.65
-17.1 | 0 | 28 | 98.16 | 193869.4 | 99.72
-15.9 | 0 | 28 | 98.06 | 193667.6 | 99.54
-14.7 | 0 | 28 | 101.41 | 200300.5 | 99.77
-13.5 | 0 | 28 | 102.57 | 202594.6 | 99.97
-12.3 | 0 | 28 | 96.45 | 190497.8 | 99.63
-11.1 | 0 | 28 | 100.22 | 197950.7 | 99.7
-9.9 | 0 | 28 | 103.24 | 203916.1 | 100.2
-8.7 | 0 | 28 | 100.04 | 197593 | 100.2
-7.5 | 0 | 28 | 100.1 | 197715.6 | 99.99
-6.3 | 0 | 28 | 101.8 | 201070.4 | 100.27
-5.1 | 0 | 28 | 99.59 | 196704.3 | 100.11
-3.9 | 0 | 28 | 100.51 | 198520.3 | 100.08
-2.7 | 0 | 28 | 102.38 | 202203.6 | 100.19
-1.5 | 0 | 28 | 100.22 | 197943.8 | 100.07
-0.3 | 0 | 28 | 102.32 | 202100.1 | 100.3
0.9 | 0 | 28 | 97.65 | 192873.9 | 100.08
2.2 | 0 | 28 | 101.23 | 199933.7 | 100.33
3.3 | 0 | 28 | 96.98 | 191540 | 100.02
4.5 | 0 | 28 | 101.31 | 200095.1 | 100.19
5.7 | 0 | 28 | 99.66 | 196828.8 | 100.14
6.9 | 0 | 28 | 99.71 | 196930.3 | 100.07
8.2 | 0 | 28 | 100.54 | 198574.4 | 100.17
9.3 | 0 | 28 | 100.47 | 198444.7 | 100.02
10.5 | 0 | 28 | 101.37 | 200206.3 | 100.12
11.7 | 0 | 28 | 101.67 | 200801.7 | 100.01
12.9 | 0 | 28 | 96.79 | 191170.7 | 99.66
14.2 | 0 | 28 | 101.15 | 199787.1 | 99.73
15.3 | 0 | 28 | 101.58 | 200635 | 99.97
16.6 | 0 | 28 | 101.2 | 199881.2 | 99.46
17.7 | 0 | 28 | 94.99 | 187621.4 | 99.22
19 | 0 | 28 | 99.9 | 197302.3 | 99.76
20.2 | 0 | 28 | 98.13 | 193812.1 | 99.44
21.3 | 0 | 28 | 97.21 | 192007.4 | 99.2
22.6 | 0 | 28 | 101.92 | 201302.1 | 99.29
23.8 | 0 | 28 | 97.53 | 192635.1 | 98.97
25 | 0 | 28 | 98.63 | 194812 | 98.97
26.2 | 0 | 28 | 100.32 | 198131.4 | 98.66
27.4 | 0 | 28 | 98.83 | 195196.3 | 98.35
28.6 | 0 | 28 | 99.28 | 196077.4 | 98.08
29.8 | 0 | 28 | 99 | 195537.3 | 97.57
31 | 0 | 28 | 94.7 | 187037.2 | 97.28
32.2 | 0 | 28 | 97.82 | 193204.4 | 96.9
33.4 | 0 | 28 | 97.37 | 192309 | 96.23
34.6 | 0 | 28 | 95.85 | 189318.8 | 95.52
35.8 | 0 | 28 | 94.79 | 187220.6 | 94.5
37 | 0 | 28 | 93.63 | 184937.4 | 93.56
38.2 | 0 | 28 | 92.07 | 181843.6 | 92.34
39.4 | 0 | 28 | 90.74 | 179218.9 | 90.84
40.6 | 0 | 28 | 89.29 | 176347.4 | 88.66
41.8 | 0 | 28 | 86.53 | 170913.3 | 86.23
43 | 0 | 28 | 84.5 | 166897.5 | 83.3
44.2 | 0 | 28 | 81.65 | 161268.6 | 79.83
45.4 | 0 | 28 | 72.78 | 143747.1 | 75.13
46.6 | 0 | 28 | 71.46 | 141144.9 | 70.39
47.8 | 0 | 28 | 65.55 | 129475.1 | 64.78
49 | 0 | 28 | 57.96 | 114468.7 | 58.16
50.2 | 0 | 28 | 51.77 | 102243.1 | 51.65
51.4 | 0 | 28 | 44.55 | 87987.8 | 45.01
52.6 | 0 | 28 | 38.92 | 76861 | 38.61
53.8 | 0 | 28 | 32.18 | 63552.9 | 32.48
55 | 0 | 28 | 27.5 | 54324.4 | 26.96
56.2 | 0 | 28 | 21.95 | 43346 | 22.02
57.4 | 0 | 28 | 17.66 | 34886.8 | 17.92
58.6 | 0 | 28 | 14.9 | 29430.4 | 14.61
59.8 | 0 | 28 | 12.03 | 23752 | 11.91
61 | 0 | 28 | 9.62 | 18999.2 | 9.58
62.2 | 0 | 28 | 7.75 | 15301.6 | 7.79
63.4 | 0 | 28 | 6.52 | 12871.9 | 6.38
64.6 | 0 | 28 | 5.06 | 10002.3 | 5.15
65.8 | 0 | 28 | 3.98 | 7863.7 | 4.15
67.1 | 0 | 28 | 3.48 | 6878.7 | 3.44
68.2 | 0 | 28 | 2.76 | 5460.2 | 2.74
69.5 | 0 | 28 | 2.25 | 4442 | 2.22
70.7 | 0 | 28 | 1.77 | 3486.6 | 1.73
71.9 | 0 | 28 | 1.2 | 2376.8 | 1.24
73.1 | 0 | 28 | 1.03 | 2038.9 | 1.02
74.3 | 0 | 28 | 0.73 | 1447.7 | 0.73
75.5 | 0 | 28 | 0.49 | 970.1 | 0.49
76.7 | 0 | 28 | 0.23 | 445.2 | 0.23
77.9 | 0 | 28 | 0.11 | 215.9 | 0.11
79.1 | 0 | 28 | 0.07 | 129.3 | 0.07
80.3 | 0 | 28 | 0.24 | 470.5 | 0.24
81.5 | 0 | 28 | 0.36 | 719.7 | 0.36
82.7 | 0 | 28 | 0.46 | 899.3 | 0.45
83.9 | 0 | 28 | 0.55 | 1093.1 | 0.54
85.1 | 0 | 28 | 0.68 | 1349.7 | 0.68
86.3 | 0 | 28 | 0.75 | 1483.4 | 0.77
87.5 | 0 | 28 | 0.86 | 1698.4 | 0.85
88.7 | 0 | 28 | 0.97 | 1908.1 | 0.97
89.9 | 0 | 28 | 1.03 | 2033 | 1.04
91.1 | 0 | 28 | 1.1 | 2180.2 | 1.11
92.3 | 0 | 28 | 1.03 | 2039.3 | 1
93.5 | 0 | 28 | 1.13 | 2236.7 | 1.17
94.7 | 0 | 28 | 1.14 | 2250.3 | 1.12
95.9 | 0 | 28 | 1.17 | 2304.3 | 1.15
97.1 | 0 | 28 | 1.26 | 2484.6 | 1.29
98.3 | 0 | 28 | 1.2 | 2379.8 | 1.19
99.5 | 0 | 28 | 1.25 | 2477.6 | 1.24
100.7 | 0 | 28 | 1.31 | 2596.3 | 1.37
101.9 | 0 | 28 | 1.29 | 2549.9 | 1.29
103.1 | 0 | 28 | 1.28 | 2520.2 | 1.23
104.3 | 0 | 28 | 1.3 | 2561.5 | 1.31
105.5 | 0 | 28 | 1.35 | 2661.4 | 1.41
106.7 | 0 | 28 | 1.34 | 2640.9 | 1.33
107.9 | 0 | 28 | 1.31 | 2588.4 | 1.3
109.1 | 0 | 28 | 1.33 | 2629.6 | 1.32
110.3 | 0 | 28 | 1.34 | 2648.3 | 1.33
111.5 | 0 | 28 | 1.36 | 2688.5 | 1.4
112.7 | 0 | 28 | 1.33 | 2617.6 | 1.31
113.9 | 0 | 28 | 1.37 | 2712.4 | 1.37
115.1 | 0 | 28 | 1.35 | 2670.5 | 1.36
116.3 | 0 | 28 | 1.38 | 2722.5 | 1.37
117.5 | 0 | 28 | 1.37 | 2712.5 | 1.36
118.8 | 0 | 28 | 1.35 | 2670.3 | 1.34
119.9 | 0 | 28 | 1.37 | 2707.5 | 1.35
121.1 | 0 | 28 | 1.39 | 2740.3 | 1.37
122.3 | 0 | 28 | 1.37 | 2699.9 | 1.4
123.5 | 0 | 28 | 1.35 | 2675.7 | 1.36
124.7 | 0 | 28 | 1.4 | 2758.6 | 1.35
126 | 0 | 28 | 1.38 | 2727 | 1.42
127.1 | 0 | 28 | 1.37 | 2704.1 | 1.36
128.3 | 0 | 28 | 1.4 | 2768.4 | 1.4
129.5 | 0 | 28 | 1.4 | 2767 | 1.41
130.7 | 0 | 28 | 1.38 | 2719.2 | 1.37
132 | 0 | 28 | 1.44 | 2853.9 | 1.44
133.2 | 0 | 28 | 1.43 | 2820.8 | 1.41
134.4 | 0 | 28 | 1.46 | 2879.6 | 1.49
135.6 | 0 | 28 | 1.42 | 2795 | 1.41
136.8 | 0 | 28 | 1.44 | 2846.5 | 1.45
138 | 0 | 28 | 1.43 | 2815.1 | 1.44
139.2 | 0 | 28 | 1.45 | 2870.3 | 1.43
140.4 | 0 | 28 | 1.44 | 2841.8 | 1.43
141.6 | 0 | 28 | 1.44 | 2840.2 | 1.43
142.8 | 0 | 28 | 1.47 | 2912 | 1.48
144 | 0 | 28 | 1.44 | 2851.5 | 1.43
145.2 | 0 | 28 | 1.44 | 2838.4 | 1.43
146.4 | 0 | 28 | 1.45 | 2856.8 | 1.44
147.6 | 0 | 28 | 1.41 | 2787.9 | 1.38
148.8 | 0 | 28 | 1.38 | 2726.1 | 1.44
150 | 0 | 28 | 1.4 | 2759 | 1.39
151.4 | 0 | 28 | 1.46 | 2888.1 | 1.47
Measurement time: | 02/09/2022 21:04:31 |  | 1 cm/s
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
151.4 | 0 | 28 | 0.69 | 1372.4 | 0.55
150.4 | 0 | 28 | 1.51 | 2990.8 | 1.52
148.6 | 0 | 28 | 1.51 | 2989.7 | 1.5
147.2 | 0 | 28 | 1.4 | 2764.8 | 1.18
146 | 0 | 28 | 1.52 | 3006.5 | 1.81
144.9 | 0 | 28 | 1.44 | 2850.3 | 1.42
143.9 | 0 | 28 | 1.47 | 2907.6 | 1.41
142.6 | 0 | 28 | 1.48 | 2916.6 | 1.54
141.4 | 0 | 28 | 1.45 | 2872.2 | 1.27
140.1 | 0 | 28 | 1.52 | 2993.8 | 1.73
138.9 | 0 | 28 | 1.46 | 2885.1 | 1.41
137.7 | 0 | 28 | 1.47 | 2897.8 | 1.42
136.5 | 0 | 28 | 1.49 | 2952 | 1.58
135.3 | 0 | 28 | 1.42 | 2797.5 | 1.22
134.1 | 0 | 28 | 1.57 | 3104.3 | 1.76
132.9 | 0 | 28 | 1.45 | 2857.1 | 1.39
131.7 | 0 | 28 | 1.43 | 2821.4 | 1.39
130.5 | 0 | 28 | 1.44 | 2845.6 | 1.49
129.3 | 0 | 28 | 1.38 | 2719.7 | 1.17
128.1 | 0 | 28 | 1.46 | 2889.5 | 1.66
126.9 | 0 | 28 | 1.41 | 2786.5 | 1.36
125.7 | 0 | 28 | 1.45 | 2865.5 | 1.41
124.5 | 0 | 28 | 1.46 | 2887.6 | 1.59
123.3 | 0 | 28 | 1.43 | 2819.2 | 1.36
122.1 | 0 | 28 | 1.44 | 2839.3 | 1.61
120.9 | 0 | 28 | 1.26 | 2494.8 | 1.2
119.7 | 0 | 28 | 1.43 | 2829.7 | 1.44
118.5 | 0 | 28 | 1.42 | 2808.8 | 1.47
117.2 | 0 | 28 | 1.37 | 2701.9 | 1.18
116.1 | 0 | 28 | 1.38 | 2716.5 | 1.55
114.8 | 0 | 28 | 1.34 | 2650 | 1.29
113.6 | 0 | 28 | 1.38 | 2719.9 | 1.39
112.4 | 0 | 28 | 1.37 | 2699.8 | 1.42
111.2 | 0 | 28 | 1.25 | 2474.3 | 1.09
110 | 0 | 28 | 1.35 | 2674.6 | 1.49
108.8 | 0 | 28 | 1.45 | 2854.6 | 1.41
107.6 | 0 | 28 | 1.36 | 2689.3 | 1.46
106.4 | 0 | 28 | 1.35 | 2674 | 1.36
105.1 | 0 | 28 | 1.18 | 2329.8 | 1.09
104 | 0 | 28 | 1.31 | 2596.1 | 1.37
102.7 | 0 | 28 | 1.32 | 2611.1 | 1.34
101.5 | 0 | 28 | 1.32 | 2604.3 | 1.37
100.3 | 0 | 28 | 1.28 | 2535.5 | 1.28
99.1 | 0 | 28 | 1.22 | 2409.9 | 1.1
97.9 | 0 | 28 | 1.27 | 2506.8 | 1.3
96.6 | 0 | 28 | 1.12 | 2212.4 | 1.14
95.5 | 0 | 28 | 1.21 | 2386.3 | 1.26
94.3 | 0 | 28 | 1.19 | 2357.2 | 1.15
93 | 0 | 28 | 1.1 | 2172.1 | 1.02
91.8 | 0 | 28 | 1.05 | 2070.7 | 1.06
90.6 | 0 | 28 | 1.13 | 2223.3 | 1.2
89.4 | 0 | 28 | 1 | 1984.3 | 1.01
88.2 | 0 | 28 | 0.92 | 1815.6 | 0.88
87 | 0 | 28 | 0.74 | 1455.1 | 0.68
85.8 | 0 | 28 | 0.7 | 1376.8 | 0.7
84.5 | 0 | 28 | 0.63 | 1237.5 | 0.66
83.4 | 0 | 28 | 0.42 | 823 | 0.41
82.1 | 0 | 28 | 0.4 | 786.6 | 0.41
80.9 | 0 | 28 | 0.31 | 614.6 | 0.32
79.7 | 0 | 28 | 0.14 | 285.6 | 0.14
78.5 | 0 | 28 | 0.01 | 16.8 | 0.01
77.3 | 0 | 28 | 0.19 | 374.2 | 0.19
76.1 | 0 | 28 | 0.47 | 923 | 0.45
74.9 | 0 | 28 | 0.64 | 1273.5 | 0.61
73.7 | 0 | 28 | 0.94 | 1866.4 | 0.96
72.5 | 0 | 28 | 1.05 | 2066.3 | 1.1
71.3 | 0 | 28 | 1.58 | 3124.4 | 1.59
70.1 | 0 | 28 | 2.17 | 4293.9 | 2.11
68.8 | 0 | 28 | 2.71 | 5353.6 | 2.59
67.6 | 0 | 28 | 2.7 | 5336.5 | 2.97
66.4 | 0 | 28 | 3.81 | 7533.5 | 3.9
65.2 | 0 | 28 | 5.68 | 11228.4 | 5.08
64 | 0 | 28 | 5.44 | 10737.7 | 5.84
62.8 | 0 | 28 | 7.17 | 14164.2 | 7.34
61.6 | 0 | 28 | 9.14 | 18049 | 8.96
60.3 | 0 | 28 | 11.4 | 22510.7 | 11.09
59.2 | 0 | 28 | 13.18 | 26032.6 | 13.58
58 | 0 | 28 | 16.12 | 31844.9 | 16.64
56.7 | 0 | 28 | 21.01 | 41488.7 | 20.83
55.6 | 0 | 28 | 24.13 | 47653 | 25.32
54.3 | 0 | 28 | 31.74 | 62698.2 | 31
53.1 | 0 | 28 | 37.87 | 74795.4 | 36.88
51.9 | 0 | 28 | 45.95 | 90754.8 | 43.5
50.7 | 0 | 28 | 51 | 100734.3 | 49.95
49.5 | 0 | 28 | 54.16 | 106978.1 | 56.65
48.3 | 0 | 28 | 63.78 | 125970.6 | 63.12
47.1 | 0 | 28 | 72.41 | 143007.2 | 69.16
45.9 | 0 | 28 | 66.67 | 131679.2 | 73.59
44.6 | 0 | 28 | 74.79 | 147718 | 78.75
43.5 | 0 | 28 | 80.41 | 158823.1 | 82.34
42.2 | 0 | 28 | 88.98 | 175742.6 | 85.84
41 | 0 | 28 | 91.78 | 181276.4 | 88.6
39.8 | 0 | 28 | 96.38 | 190353.8 | 90.87
38.6 | 0 | 28 | 91.24 | 180209.4 | 92.02
37.5 | 0 | 28 | 88.43 | 174652.5 | 93.21
36.2 | 0 | 28 | 96.66 | 190904.2 | 94.81
35.1 | 0 | 28 | 99.06 | 195658.4 | 95.45
33.8 | 0 | 28 | 98.41 | 194363.9 | 96.05
32.6 | 0 | 28 | 90.09 | 177944 | 96.26
31.4 | 0 | 28 | 96.4 | 190395.5 | 97.2
30.1 | 0 | 28 | 98.39 | 194325 | 97.66
29 | 0 | 28 | 99.26 | 196051.9 | 98.2
27.7 | 0 | 28 | 101.52 | 200516 | 98.71
26.6 | 0 | 28 | 100.86 | 199209 | 98.9
25.3 | 0 | 28 | 100.53 | 198556.1 | 99.06
24.1 | 0 | 28 | 95.47 | 188559.1 | 98.96
22.9 | 0 | 28 | 92.33 | 182367.9 | 98.66
21.7 | 0 | 28 | 99.29 | 196106.8 | 99.59
20.5 | 0 | 28 | 91.76 | 181225.8 | 99.27
19.3 | 0 | 28 | 112.97 | 223117.9 | 99.97
18.1 | 0 | 28 | 95.62 | 188852.2 | 99.53
16.9 | 0 | 28 | 96.43 | 190462.8 | 99.78
15.7 | 0 | 28 | 103.23 | 203898.1 | 100.32
14.5 | 0 | 28 | 91.89 | 181500.3 | 99.55
13.3 | 0 | 28 | 99.61 | 196732 | 100.09
12 | 0 | 28 | 103.57 | 204564.8 | 100.19
10.9 | 0 | 28 | 104.19 | 205789.6 | 100.22
9.6 | 0 | 28 | 103.39 | 204210 | 99.93
8.4 | 0 | 28 | 88.92 | 175619.8 | 99.59
7.2 | 0 | 28 | 103.34 | 204110.7 | 99.94
6 | 0 | 28 | 102.89 | 203224.3 | 100.23
4.8 | 0 | 28 | 103.42 | 204264.3 | 100.47
3.6 | 0 | 28 | 106.24 | 209824.2 | 100.47
2.3 | 0 | 28 | 89.37 | 176522 | 99.8
1.1 | 0 | 28 | 103.7 | 204815.5 | 100.32
-0.1 | 0 | 28 | 104.82 | 207030.1 | 100.39
-1.2 | 0 | 28 | 90.87 | 179470.4 | 99.81
-2.5 | 0 | 28 | 104.08 | 205566.4 | 100.38
-3.7 | 0 | 28 | 90.76 | 179251.5 | 99.9
-4.9 | 0 | 28 | 102.91 | 203261.7 | 100.39
-6.1 | 0 | 28 | 104.24 | 205880.1 | 100.48
-7.3 | 0 | 28 | 102.24 | 201942 | 100.39
-8.5 | 0 | 28 | 108.35 | 213998 | 100.6
-9.7 | 0 | 28 | 89.16 | 176106.1 | 99.68
-10.9 | 0 | 28 | 101.67 | 200809.6 | 100.26
-12.1 | 0 | 28 | 102.75 | 202937 | 100.52
-13.3 | 0 | 28 | 99.33 | 196176.3 | 100.16
-14.6 | 0 | 28 | 108.47 | 214233.2 | 100.3
-15.8 | 0 | 28 | 87.37 | 172559.4 | 99.17
-17 | 0 | 28 | 102.48 | 202402 | 99.95
-18.2 | 0 | 28 | 103.31 | 204037.5 | 99.76
-19.4 | 0 | 28 | 95.73 | 189073.5 | 99.53
-20.6 | 0 | 28 | 111.05 | 219332.7 | 100.08
-21.8 | 0 | 28 | 87.47 | 172769.4 | 99.04
-23 | 0 | 28 | 102.07 | 201595.5 | 99.45
-24.3 | 0 | 28 | 102.55 | 202554.2 | 99.14
-25.4 | 0 | 28 | 93.11 | 183897.5 | 98.38
-26.7 | 0 | 28 | 114.25 | 225661.5 | 99.26
-27.9 | 0 | 28 | 85.68 | 169231.8 | 98.04
-29.1 | 0 | 28 | 100.87 | 199232.4 | 98.31
-30.3 | 0 | 28 | 100.99 | 199471.2 | 97.99
-31.4 | 0 | 28 | 92.86 | 183403.5 | 97.66
-32.7 | 0 | 28 | 113.36 | 223904.3 | 97.96
-33.9 | 0 | 28 | 85.53 | 168922.2 | 96.55
-35.1 | 0 | 28 | 99.83 | 197170.3 | 96.49
-36.3 | 0 | 28 | 99.33 | 196193.3 | 96.22
-37.4 | 0 | 28 | 90.36 | 178468 | 94.67
-38.7 | 0 | 28 | 110.79 | 218823.1 | 94.92
-39.9 | 0 | 28 | 83.5 | 164928.2 | 93.04
-41.1 | 0 | 28 | 97.88 | 193318.9 | 92.06
-42.4 | 0 | 28 | 79.69 | 157388.7 | 90.07
-43.5 | 0 | 28 | 91.85 | 181419.2 | 88.52
-44.8 | 0 | 28 | 88.68 | 175155.9 | 85.68
-45.9 | 0 | 28 | 71.21 | 140651.1 | 82.14
-47.1 | 0 | 28 | 81.69 | 161347 | 79.09
-48.3 | 0 | 28 | 75.44 | 149006.9 | 74.48
-49.5 | 0 | 28 | 65.54 | 129444.6 | 69.12
-50.8 | 0 | 28 | 73.97 | 146105.5 | 63.38
-51.9 | 0 | 28 | 49.14 | 97049.5 | 55.99
-53.2 | 0 | 28 | 51.3 | 101316.6 | 49.85
-54.4 | 0 | 28 | 43.62 | 86145.1 | 42.86
-55.5 | 0 | 28 | 37.62 | 74310.8 | 37.01
-56.8 | 0 | 28 | 31.11 | 61446.6 | 30.95
-58 | 0 | 28 | 23.99 | 47384.1 | 25.59
-59.2 | 0 | 28 | 21.22 | 41920.3 | 21.33
-60.4 | 0 | 28 | 17.62 | 34798.5 | 17.44
-61.6 | 0 | 28 | 14.7 | 29041.3 | 14.09
-62.9 | 0 | 28 | 11.35 | 22415.3 | 11.37
-64.1 | 0 | 28 | 8.56 | 16904.1 | 9.18
-65.3 | 0 | 28 | 8.25 | 16291.3 | 7.93
-66.5 | 0 | 28 | 6.07 | 11985.6 | 6.38
-67.7 | 0 | 28 | 5.14 | 10150.3 | 5.07
-68.9 | 0 | 28 | 4.88 | 9635 | 4.37
-70.1 | 0 | 28 | 3.1 | 6115.4 | 3.23
-71.3 | 0 | 28 | 2.74 | 5404 | 2.71
-72.5 | 0 | 28 | 2.01 | 3967.5 | 2.08
-73.7 | 0 | 28 | 1.71 | 3370.9 | 1.69
-75 | 0 | 28 | 1.53 | 3020.2 | 1.39
-76.1 | 0 | 28 | 0.83 | 1634.2 | 0.83
-77.4 | 0 | 28 | 0.54 | 1068.2 | 0.55
-78.6 | 0 | 28 | 0.23 | 459.6 | 0.24
-79.8 | 0 | 28 | 0.12 | 238.3 | 0.12
-81 | 0 | 28 | 0.15 | 305.1 | 0.16
-82.2 | 0 | 28 | 0.23 | 460.9 | 0.23
-83.4 | 0 | 28 | 0.48 | 943 | 0.51
-84.6 | 0 | 28 | 0.41 | 817.4 | 0.41
-85.9 | 0 | 28 | 0.54 | 1059.2 | 0.52
-87 | 0 | 28 | 0.61 | 1197.6 | 0.58
-88.2 | 0 | 28 | 0.86 | 1706.5 | 0.88
-89.5 | 0 | 28 | 0.89 | 1749.9 | 0.94
-90.7 | 0 | 28 | 0.87 | 1726.8 | 0.89
-91.9 | 0 | 28 | 0.85 | 1681.8 | 0.79
-93.1 | 0 | 28 | 1.04 | 2045.7 | 1.07
-94.3 | 0 | 28 | 1.02 | 2013.1 | 1.01
-95.5 | 0 | 28 | 1.01 | 1990.5 | 1
-96.7 | 0 | 28 | 1.07 | 2114.7 | 1.06
-97.9 | 0 | 28 | 1.15 | 2268.9 | 1.11
-99.2 | 0 | 28 | 1.15 | 2270.5 | 1.09
-100.3 | 0 | 28 | 1.23 | 2431.2 | 1.33
-101.6 | 0 | 28 | 1.2 | 2375.8 | 1.24
-102.8 | 0 | 28 | 1.22 | 2405.6 | 1.18
-104 | 0 | 28 | 1.22 | 2402.3 | 1.18
-105.2 | 0 | 28 | 1.25 | 2475.5 | 1.21
-106.4 | 0 | 28 | 1.28 | 2520.6 | 1.39
-107.6 | 0 | 28 | 1.19 | 2347.1 | 1.19
-108.8 | 0 | 28 | 1.19 | 2358 | 1.18
-110 | 0 | 28 | 1.26 | 2494.8 | 1.23
-111.3 | 0 | 28 | 1.24 | 2445.2 | 1.2
-112.4 | 0 | 28 | 1.38 | 2721.4 | 1.5
-113.6 | 0 | 28 | 1.3 | 2566.2 | 1.29
-114.9 | 0 | 28 | 1.34 | 2650.2 | 1.3
-116.1 | 0 | 28 | 1.29 | 2543.2 | 1.25
-117.3 | 0 | 28 | 1.3 | 2571.5 | 1.26
-118.5 | 0 | 28 | 1.35 | 2672 | 1.44
-119.7 | 0 | 28 | 1.3 | 2567.1 | 1.28
-120.9 | 0 | 28 | 1.3 | 2567.8 | 1.25
-122.1 | 0 | 28 | 1.23 | 2426.3 | 1.15
-123.3 | 0 | 28 | 1.28 | 2534.3 | 1.27
-124.5 | 0 | 28 | 1.29 | 2545.9 | 1.36
-125.8 | 0 | 28 | 1.25 | 2474.1 | 1.22
-126.9 | 0 | 28 | 1.26 | 2492 | 1.22
-128.1 | 0 | 28 | 1.27 | 2506.4 | 1.21
-129.4 | 0 | 28 | 1.28 | 2527.3 | 1.34
-130.6 | 0 | 28 | 1.29 | 2552.9 | 1.34
-131.8 | 0 | 28 | 1.25 | 2473.7 | 1.22
-132.9 | 0 | 28 | 1.28 | 2533 | 1.24
-134.2 | 0 | 28 | 1.35 | 2673.2 | 1.32
-135.4 | 0 | 28 | 1.39 | 2752 | 1.49
-136.6 | 0 | 28 | 1.29 | 2554 | 1.29
-137.9 | 0 | 28 | 1.31 | 2580.5 | 1.29
-139 | 0 | 28 | 1.35 | 2675.8 | 1.33
-140.3 | 0 | 28 | 1.34 | 2641.1 | 1.29
-141.5 | 0 | 28 | 1.34 | 2649.3 | 1.44
-142.6 | 0 | 28 | 1.29 | 2553.1 | 1.3
-143.9 | 0 | 28 | 1.24 | 2458.2 | 1.22
-145 | 0 | 28 | 1.28 | 2528 | 1.25
-146.3 | 0 | 28 | 1.26 | 2490.9 | 1.21
-147.5 | 0 | 28 | 1.32 | 2608 | 1.43
-148.7 | 0 | 28 | 1.34 | 2648.3 | 1.34
-149.9 | 0 | 28 | 1.26 | 2488.1 | 1.23
-151.4 | 0 | 28 | 1.31 | 2592.9 | 1.28
Measurement time: | 02/09/2022 21:07:10 |  | 2 cm/s
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
151.4 | 0 | 28 | 1.61 | 3178 | 2.13
150.2 | 0 | 28 | 1.47 | 2898.9 | 1.43
147.9 | 0 | 28 | 1.46 | 2886.2 | 1.59
146.9 | 0 | 28 | 1.4 | 2759.1 | 1.43
145.4 | 0 | 28 | 1.47 | 2897.8 | 1.26
144.1 | 0 | 28 | 1.45 | 2862.9 | 1.42
142.7 | 0 | 28 | 1.5 | 2972 | 1.8
141.4 | 0 | 28 | 1.49 | 2939.1 | 1.37
140.2 | 0 | 28 | 1.43 | 2815.5 | 1.59
139 | 0 | 28 | 1.41 | 2777.9 | 1.21
137.8 | 0 | 28 | 1.39 | 2751 | 1.61
136.7 | 0 | 28 | 1.39 | 2748.3 | 1.33
135.6 | 0 | 28 | 1.38 | 2722.9 | 1.31
134.3 | 0 | 28 | 1.46 | 2892.5 | 1.6
133.2 | 0 | 28 | 1.39 | 2748.1 | 1.21
132 | 0 | 28 | 1.34 | 2652.3 | 1.21
130.8 | 0 | 28 | 1.53 | 3013.3 | 1.9
129.5 | 0 | 28 | 1.43 | 2833.8 | 1.34
128.3 | 0 | 28 | 1.4 | 2756.4 | 1.5
127.1 | 0 | 28 | 1.43 | 2825 | 1.27
125.8 | 0 | 28 | 1.4 | 2756.1 | 1.46
124.6 | 0 | 28 | 1.34 | 2655.3 | 1.4
123.3 | 0 | 28 | 1.32 | 2610.3 | 1.16
122.2 | 0 | 28 | 1.41 | 2779.2 | 1.59
121 | 0 | 28 | 1.23 | 2438.6 | 1.02
119.7 | 0 | 28 | 1.46 | 2888.7 | 1.48
118.6 | 0 | 28 | 1.4 | 2758 | 1.68
117.3 | 0 | 28 | 1.34 | 2647.3 | 1.15
116.1 | 0 | 28 | 1.36 | 2690.9 | 1.49
114.8 | 0 | 28 | 1.4 | 2771 | 1.42
113.6 | 0 | 28 | 1.38 | 2727.5 | 1.25
112.4 | 0 | 28 | 1.37 | 2696.3 | 1.46
111.1 | 0 | 28 | 1.28 | 2537.4 | 1.11
110 | 0 | 28 | 1.31 | 2587 | 1.42
108.7 | 0 | 28 | 1.27 | 2517.1 | 1.09
107.5 | 0 | 28 | 1.24 | 2448 | 1.26
106.3 | 0 | 28 | 1.4 | 2761.8 | 1.61
105.1 | 0 | 28 | 1.16 | 2298.2 | 1.02
103.9 | 0 | 28 | 1.32 | 2610.2 | 1.42
102.7 | 0 | 28 | 1.3 | 2571.6 | 1.32
101.4 | 0 | 28 | 1.21 | 2381.2 | 1.1
100.3 | 0 | 28 | 1.31 | 2594.8 | 1.41
99 | 0 | 28 | 1.31 | 2578.8 | 1.49
97.8 | 0 | 28 | 1.22 | 2415.3 | 1.21
96.6 | 0 | 28 | 1.22 | 2404.4 | 1.12
95.4 | 0 | 28 | 1.22 | 2417.1 | 1.35
94.2 | 0 | 28 | 1.27 | 2516.9 | 1.44
92.9 | 0 | 28 | 1.17 | 2319.1 | 1.08
91.8 | 0 | 28 | 1.04 | 2048.9 | 1.03
90.5 | 0 | 28 | 1 | 1976.1 | 1.09
89.2 | 0 | 28 | 0.84 | 1653.5 | 0.72
88.1 | 0 | 28 | 0.99 | 1961.7 | 1.09
86.8 | 0 | 28 | 0.8 | 1585.4 | 0.82
85.6 | 0 | 28 | 0.67 | 1328.2 | 0.62
84.4 | 0 | 28 | 0.36 | 709.4 | 0.3
83.1 | 0 | 28 | 0.56 | 1100.5 | 0.61
82 | 0 | 28 | 0.57 | 1132.1 | 0.65
80.7 | 0 | 28 | 0.37 | 730 | 0.38
79.6 | 0 | 28 | 0.02 | 40.2 | 0.02
78.3 | 0 | 28 | 0.14 | 274 | 0.15
77.1 | 0 | 28 | 0.43 | 839.5 | 0.37
75.9 | 0 | 28 | 0.27 | 523.5 | 0.29
74.6 | 0 | 28 | 0.49 | 966.7 | 0.5
73.5 | 0 | 28 | 1.11 | 2196.3 | 1.03
72.2 | 0 | 28 | 1.9 | 3751.1 | 1.54
71 | 0 | 28 | 1.17 | 2301.7 | 1.37
69.8 | 0 | 28 | 1.71 | 3375.8 | 1.87
68.6 | 0 | 28 | 2.33 | 4595.8 | 2.4
67.4 | 0 | 28 | 3.61 | 7128.6 | 3.28
66.1 | 0 | 28 | 3.16 | 6238.7 | 3.62
64.9 | 0 | 28 | 5.26 | 10392.5 | 4.98
63.8 | 0 | 28 | 5.53 | 10914.7 | 5.8
62.5 | 0 | 28 | 6.51 | 12857.6 | 6.99
61.3 | 0 | 28 | 10.62 | 20971.5 | 9.29
60.1 | 0 | 28 | 12.74 | 25154.3 | 11.26
58.9 | 0 | 28 | 9.61 | 18980.1 | 12.9
57.6 | 0 | 28 | 18.41 | 36351.9 | 17.3
56.4 | 0 | 28 | 18.41 | 36353.8 | 20.49
55.3 | 0 | 28 | 29.24 | 57754.4 | 26.05
53.9 | 0 | 28 | 29.07 | 57409 | 31.67
52.8 | 0 | 28 | 35.57 | 70258.9 | 36.73
51.5 | 0 | 28 | 47.94 | 94688.4 | 44.34
50.3 | 0 | 28 | 46.95 | 92724.3 | 50.25
49.1 | 0 | 28 | 65.35 | 129074 | 57.8
47.9 | 0 | 28 | 69.33 | 136932.3 | 63.56
46.7 | 0 | 28 | 56.5 | 111599.9 | 69.06
45.5 | 0 | 28 | 81.15 | 160282.6 | 75.39
44.2 | 0 | 28 | 73.38 | 144932.8 | 78.92
43 | 0 | 28 | 90.38 | 178515.6 | 83.32
41.8 | 0 | 28 | 83.29 | 164513.6 | 85.73
40.6 | 0 | 28 | 85.39 | 168645.7 | 88.48
39.4 | 0 | 28 | 103.59 | 204603.2 | 91.14
38.1 | 0 | 28 | 81.06 | 160099.9 | 92.11
37 | 0 | 28 | 112.18 | 221571.2 | 94.34
35.7 | 0 | 28 | 93.64 | 184950.4 | 94.24
34.5 | 0 | 28 | 81.29 | 160549.4 | 94.96
33.2 | 0 | 28 | 111.6 | 220415.9 | 96.34
32 | 0 | 28 | 89.54 | 176842 | 96.61
30.8 | 0 | 28 | 95.53 | 188679.5 | 97.36
29.5 | 0 | 28 | 107.08 | 211493.2 | 98.29
28.3 | 0 | 28 | 92.24 | 182189 | 97.35
27.1 | 0 | 28 | 112.09 | 221391.4 | 99.08
25.9 | 0 | 28 | 75.73 | 149582 | 97.92
24.7 | 0 | 28 | 103.96 | 205325.4 | 98.91
23.5 | 0 | 28 | 92.63 | 182944.6 | 98.88
22.3 | 0 | 28 | 86.69 | 171211.2 | 98.62
21 | 0 | 28 | 112.97 | 223135.2 | 99.86
19.9 | 0 | 28 | 92.99 | 183666.8 | 99.4
18.7 | 0 | 28 | 97.48 | 192539.4 | 99.02
17.4 | 0 | 28 | 109.02 | 215329.9 | 100.13
16.3 | 0 | 28 | 94.25 | 186160.6 | 99.85
15 | 0 | 28 | 112.1 | 221408.3 | 100.77
13.8 | 0 | 28 | 95.5 | 188622.7 | 99.53
12.5 | 0 | 28 | 122.66 | 242267.7 | 100.79
11.3 | 0 | 28 | 92.95 | 183592.8 | 99.38
10.2 | 0 | 28 | 87.3 | 172433.5 | 99.71
8.9 | 0 | 28 | 102.97 | 203375 | 100.23
7.7 | 0 | 28 | 106.5 | 210342.4 | 100.46
6.5 | 0 | 28 | 91.71 | 181130 | 99.08
5.3 | 0 | 28 | 116.8 | 230700.4 | 100.69
4.1 | 0 | 28 | 92.42 | 182537.5 | 99.83
2.8 | 0 | 28 | 96.45 | 190501.9 | 100
1.6 | 0 | 28 | 85.2 | 168283.5 | 99.18
0.4 | 0 | 28 | 115.3 | 227719.1 | 100
-0.8 | 0 | 28 | 89.82 | 177411.6 | 99.62
-2 | 0 | 28 | 91.54 | 180807.2 | 99.91
-3.3 | 0 | 28 | 98.41 | 194378 | 99.83
-4.4 | 0 | 28 | 109.58 | 216428.7 | 100.64
-5.7 | 0 | 28 | 93.41 | 184488.8 | 100.08
-6.9 | 0 | 28 | 108.5 | 214305.5 | 100.72
-8.2 | 0 | 28 | 97.85 | 193270 | 99.37
-9.4 | 0 | 28 | 96.55 | 190686.5 | 99.91
-10.6 | 0 | 28 | 114.21 | 225581.4 | 100.61
-11.8 | 0 | 28 | 119.57 | 236170.9 | 100.7
-13 | 0 | 28 | 80.53 | 159053.4 | 99.21
-14.3 | 0 | 28 | 95.8 | 189208.3 | 99.89
-15.5 | 0 | 28 | 97.19 | 191963.6 | 100.09
-16.7 | 0 | 28 | 114.42 | 225984.5 | 100.58
-17.9 | 0 | 28 | 87.37 | 172555.9 | 99.26
-19.1 | 0 | 28 | 108.83 | 214949.3 | 100.08
-20.3 | 0 | 28 | 101.27 | 200015 | 99.56
-21.5 | 0 | 28 | 90.05 | 177852.5 | 98.96
-22.7 | 0 | 28 | 115.9 | 228919.9 | 99.85
-24 | 0 | 28 | 111.47 | 220155.8 | 99.88
-25.2 | 0 | 28 | 77.96 | 153978.7 | 98.15
-26.4 | 0 | 28 | 106.52 | 210381.7 | 98.99
-27.6 | 0 | 28 | 91.64 | 181006.1 | 98.41
-28.8 | 0 | 28 | 110.96 | 219151.8 | 98.88
-30.1 | 0 | 28 | 90.58 | 178912.2 | 98.01
-31.3 | 0 | 28 | 95.1 | 187837.6 | 97.75
-32.5 | 0 | 28 | 104.85 | 207081.6 | 97.33
-33.8 | 0 | 28 | 88.78 | 175339.1 | 97.01
-35 | 0 | 28 | 113.47 | 224119.8 | 97.3
-36.2 | 0 | 28 | 97.59 | 192747.6 | 96.41
-37.4 | 0 | 28 | 77.78 | 153629.2 | 94.23
-38.6 | 0 | 28 | 107.17 | 211667.8 | 95.37
-39.8 | 0 | 28 | 83.25 | 164419.6 | 93.16
-41 | 0 | 28 | 93.48 | 184622.9 | 92.89
-42.3 | 0 | 28 | 94.61 | 186859.7 | 91.02
-43.4 | 0 | 28 | 81.93 | 161815.7 | 89.04
-44.7 | 0 | 28 | 98.31 | 194176.7 | 87
-46 | 0 | 28 | 92.36 | 182410.6 | 84.17
-47.2 | 0 | 28 | 74.46 | 147067.8 | 79.18
-48.5 | 0 | 28 | 72.56 | 143322.2 | 75.47
-49.6 | 0 | 28 | 56.24 | 111078.4 | 69.34
-50.8 | 0 | 28 | 72.6 | 143393 | 64.43
-52 | 0 | 28 | 65.14 | 128654.7 | 58.62
-53.2 | 0 | 28 | 49.01 | 96805.9 | 50.93
-54.5 | 0 | 28 | 40.62 | 80231.2 | 44.65
-55.6 | 0 | 28 | 43.05 | 85026.4 | 38.02
-56.9 | 0 | 28 | 27.4 | 54124.5 | 31.24
-58.2 | 0 | 28 | 26.41 | 52157.8 | 26.78
-59.3 | 0 | 28 | 22.67 | 44769.2 | 21.44
-60.5 | 0 | 28 | 15.71 | 31031.3 | 17.84
-61.8 | 0 | 28 | 12.25 | 24200.6 | 14.23
-63 | 0 | 28 | 14.35 | 28346.2 | 12.09
-64.2 | 0 | 28 | 9.66 | 19077.3 | 9.82
-65.4 | 0 | 28 | 8.58 | 16948 | 7.94
-66.6 | 0 | 28 | 5.76 | 11379.7 | 6.31
-67.8 | 0 | 28 | 6.14 | 12127.3 | 5.51
-69 | 0 | 28 | 3.72 | 7342.8 | 4.1
-70.3 | 0 | 28 | 3.33 | 6575.4 | 3.45
-71.5 | 0 | 28 | 2.93 | 5793.4 | 2.69
-72.7 | 0 | 28 | 1.77 | 3486.5 | 1.93
-73.9 | 0 | 28 | 1.07 | 2121.7 | 1.24
-75.1 | 0 | 28 | 1.29 | 2547.8 | 1.26
-76.3 | 0 | 28 | 0.93 | 1829.3 | 0.88
-77.6 | 0 | 28 | 0.39 | 766.7 | 0.43
-78.8 | 0 | 28 | 0.59 | 1164.3 | 0.51
-80 | 0 | 28 | 0.12 | 234.6 | 0.13
-81.2 | 0 | 28 | 0.18 | 353 | 0.19
-82.4 | 0 | 28 | 0.16 | 323.1 | 0.15
-83.7 | 0 | 28 | 0.11 | 226.3 | 0.09
-84.9 | 0 | 28 | 0.68 | 1345.6 | 0.81
-86.1 | 0 | 28 | 0.65 | 1274.7 | 0.7
-87.3 | 0 | 28 | 0.72 | 1430.1 | 0.74
-88.5 | 0 | 28 | 0.64 | 1271.8 | 0.58
-89.8 | 0 | 28 | 0.98 | 1942.3 | 1.14
-91 | 0 | 28 | 0.86 | 1694.4 | 0.83
-92.3 | 0 | 28 | 1.01 | 2002.9 | 0.97
-93.5 | 0 | 28 | 1.14 | 2261.2 | 1.27
-94.7 | 0 | 28 | 0.9 | 1777.3 | 0.77
-95.9 | 0 | 28 | 1.01 | 1997.1 | 0.88
-97.1 | 0 | 28 | 1.25 | 2459.7 | 1.6
-98.3 | 0 | 28 | 1.14 | 2258.1 | 1.1
-99.6 | 0 | 28 | 1.25 | 2461.8 | 1.38
-100.7 | 0 | 28 | 1.09 | 2150.7 | 0.94
-102 | 0 | 28 | 1.25 | 2461.1 | 1.38
-103.2 | 0 | 28 | 1.22 | 2418.8 | 1.25
-104.4 | 0 | 28 | 1.22 | 2401.1 | 1.12
-105.7 | 0 | 28 | 1.27 | 2511.8 | 1.37
-106.8 | 0 | 28 | 1.27 | 2510.1 | 1.12
-108.1 | 0 | 28 | 1.23 | 2432.5 | 1.18
-109.3 | 0 | 28 | 1.31 | 2583 | 1.53
-110.6 | 0 | 28 | 1.12 | 2212.9 | 1.02
-111.8 | 0 | 28 | 1.3 | 2558.9 | 1.41
-113 | 0 | 28 | 1.19 | 2354.6 | 1.08
-114.2 | 0 | 28 | 1.24 | 2451.2 | 1.31
-115.4 | 0 | 28 | 1.43 | 2830.9 | 1.45
-116.6 | 0 | 28 | 1.29 | 2555.8 | 1.15
-117.8 | 0 | 28 | 1.41 | 2783.4 | 1.59
-119 | 0 | 28 | 1.19 | 2344 | 0.99
-120.3 | 0 | 28 | 1.4 | 2767.7 | 1.4
-121.5 | 0 | 28 | 1.32 | 2609.7 | 1.56
-122.8 | 0 | 28 | 1.36 | 2677.7 | 1.18
-124 | 0 | 28 | 1.31 | 2596.1 | 1.44
-125.1 | 0 | 28 | 1.33 | 2626 | 1.36
-126.4 | 0 | 28 | 1.29 | 2538.3 | 1.17
-127.6 | 0 | 28 | 1.31 | 2584.1 | 1.39
-128.8 | 0 | 28 | 1.28 | 2518.6 | 1.1
-130 | 0 | 28 | 1.4 | 2767.7 | 1.76
-131.2 | 0 | 28 | 1.32 | 2608.2 | 1.37
-132.4 | 0 | 28 | 1.37 | 2701 | 1.42
-133.6 | 0 | 28 | 1.39 | 2746.5 | 1.58
-134.8 | 0 | 28 | 1.25 | 2475.4 | 1.15
-136 | 0 | 28 | 1.36 | 2676.4 | 1.43
-137.2 | 0 | 28 | 1.29 | 2551.7 | 1.31
-138.5 | 0 | 28 | 1.21 | 2395.2 | 1.1
-139.7 | 0 | 28 | 1.46 | 2874.3 | 1.56
-140.9 | 0 | 28 | 1.31 | 2595.9 | 1.19
-142.1 | 0 | 28 | 1.36 | 2677.5 | 1.42
-143.3 | 0 | 28 | 1.25 | 2475.1 | 1.02
-144.6 | 0 | 28 | 1.33 | 2636.6 | 1.46
-145.7 | 0 | 28 | 1.4 | 2772.4 | 1.61
-147 | 0 | 28 | 1.33 | 2631.7 | 1.21
-148.2 | 0 | 28 | 1.34 | 2651.7 | 1.4
-149.4 | 0 | 28 | 1.45 | 2870.5 | 1.52
-150.5 | 0 | 28 | 1.19 | 2352.4 | 1.04
-151.4 | 0 | 28 | 1.35 | 2665 | 1.41