## 3_Profils profondeur de mesure_
Measurement time: | 07/09/2022 17:12:56 |  |  |  |  | Profondeur de mesure (cm) | Symétrie (%) | Homogénéité (%) | Pénombre G (cm) | Pénombre D (cm)
Radiation device: | Clinac3 Accelerator |  |  |  |  | 0 | 1.09 | 1.58 | 0.46 | 0.46
Energy: | 12 MeV |  |  |  |  | 2.8 | 2.23 | 5.87 | 1.36 | 1.35
Controller: | CCU |  |  |  |  | 10 | 3.47 | 11.71 | 2.15 | 2.13
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
150 | 0 | 0 | 0.32 | 634.9 | 0.71
149.2 | 0 | 0 | 0.7 | 1376.3 | 0.77
147.8 | 0 | 0 | 0.7 | 1374.1 | 0.76
146.6 | 0 | 0 | 0.66 | 1295.8 | 0.73
145.4 | 0 | 0 | 0.69 | 1350.4 | 0.74
144.2 | 0 | 0 | 0.67 | 1317 | 0.72
143 | 0 | 0 | 0.69 | 1357.9 | 0.77
141.8 | 0 | 0 | 0.73 | 1430.8 | 0.8
140.6 | 0 | 0 | 0.66 | 1285.8 | 0.71
139.4 | 0 | 0 | 0.71 | 1380.8 | 0.78
138.2 | 0 | 0 | 0.68 | 1324.6 | 0.76
137 | 0 | 0 | 0.69 | 1354.8 | 0.74
135.8 | 0 | 0 | 0.68 | 1336.7 | 0.76
134.6 | 0 | 0 | 0.68 | 1338.6 | 0.75
133.4 | 0 | 0 | 0.73 | 1421.2 | 0.78
132.2 | 0 | 0 | 0.7 | 1377.3 | 0.78
131 | 0 | 0 | 0.73 | 1432.9 | 0.78
129.8 | 0 | 0 | 0.71 | 1390.8 | 0.76
128.6 | 0 | 0 | 0.67 | 1321.1 | 0.75
127.4 | 0 | 0 | 0.7 | 1360.8 | 0.74
126.2 | 0 | 0 | 0.67 | 1302.5 | 0.74
125 | 0 | 0 | 0.66 | 1292 | 0.73
123.8 | 0 | 0 | 0.69 | 1350.1 | 0.76
122.6 | 0 | 0 | 0.67 | 1304.4 | 0.7
121.4 | 0 | 0 | 0.66 | 1288.8 | 0.73
120.2 | 0 | 0 | 0.67 | 1314.7 | 0.75
119 | 0 | 0 | 0.71 | 1392.2 | 0.76
117.8 | 0 | 0 | 0.66 | 1301.5 | 0.75
116.6 | 0 | 0 | 0.69 | 1352.8 | 0.76
115.4 | 0 | 0 | 0.73 | 1432.6 | 0.78
114.2 | 0 | 0 | 0.69 | 1346.8 | 0.77
113 | 0 | 0 | 0.71 | 1397.9 | 0.77
111.8 | 0 | 0 | 0.76 | 1485.9 | 0.8
110.6 | 0 | 0 | 0.7 | 1371.7 | 0.79
109.4 | 0 | 0 | 0.69 | 1341.2 | 0.78
108.2 | 0 | 0 | 0.78 | 1529.6 | 0.82
106.9 | 0 | 0 | 0.71 | 1382.1 | 0.8
105.8 | 0 | 0 | 0.79 | 1545.7 | 0.86
104.6 | 0 | 0 | 0.81 | 1579.6 | 0.84
103.4 | 0 | 0 | 0.77 | 1500.9 | 0.87
102.2 | 0 | 0 | 0.81 | 1578 | 0.89
101 | 0 | 0 | 0.88 | 1723.3 | 0.92
99.8 | 0 | 0 | 0.8 | 1567.4 | 0.91
98.6 | 0 | 0 | 0.86 | 1682.4 | 0.95
97.3 | 0 | 0 | 0.88 | 1716.3 | 0.94
96.2 | 0 | 0 | 0.96 | 1872.2 | 1.04
94.9 | 0 | 0 | 0.89 | 1750.3 | 1.04
93.7 | 0 | 0 | 1 | 1956.4 | 1.06
92.5 | 0 | 0 | 0.99 | 1941.5 | 1.13
91.3 | 0 | 0 | 1.11 | 2169.3 | 1.22
90.2 | 0 | 0 | 1.15 | 2252.8 | 1.21
88.9 | 0 | 0 | 1.11 | 2172 | 1.27
87.7 | 0 | 0 | 1.16 | 2276.8 | 1.26
86.5 | 0 | 0 | 1.27 | 2492.9 | 1.32
85.3 | 0 | 0 | 1.18 | 2300.7 | 1.33
84.1 | 0 | 0 | 1.33 | 2601.7 | 1.44
82.9 | 0 | 0 | 1.42 | 2779.2 | 1.5
81.7 | 0 | 0 | 1.38 | 2700.6 | 1.56
80.5 | 0 | 0 | 1.43 | 2797.3 | 1.57
79.3 | 0 | 0 | 1.55 | 3036 | 1.64
78.1 | 0 | 0 | 1.51 | 2948.8 | 1.72
76.9 | 0 | 0 | 1.58 | 3096.6 | 1.74
75.7 | 0 | 0 | 1.77 | 3460.4 | 1.86
74.5 | 0 | 0 | 1.75 | 3419.8 | 1.98
73.3 | 0 | 0 | 1.95 | 3820.3 | 2.11
72.1 | 0 | 0 | 2.09 | 4087.7 | 2.18
70.9 | 0 | 0 | 1.98 | 3879.4 | 2.27
69.7 | 0 | 0 | 2.19 | 4287.8 | 2.38
68.5 | 0 | 0 | 2.34 | 4583.1 | 2.43
67.3 | 0 | 0 | 2.28 | 4461.2 | 2.59
66.1 | 0 | 0 | 2.36 | 4625.8 | 2.7
64.9 | 0 | 0 | 2.68 | 5238.7 | 2.8
63.7 | 0 | 0 | 2.58 | 5050.5 | 2.95
62.5 | 0 | 0 | 2.84 | 5553.6 | 3.12
61.3 | 0 | 0 | 3.16 | 6179.2 | 3.32
60.1 | 0 | 0 | 3.01 | 5885.7 | 3.44
58.8 | 0 | 0 | 3.26 | 6371.3 | 3.62
57.7 | 0 | 0 | 3.69 | 7216 | 3.91
56.4 | 0 | 0 | 3.77 | 7376.1 | 4.25
55.2 | 0 | 0 | 4.74 | 9284.5 | 5.24
54.1 | 0 | 0 | 7.02 | 13733.6 | 7.51
52.8 | 0 | 0 | 12.08 | 23644.1 | 13.5
51.6 | 0 | 0 | 22.29 | 43620.2 | 24.34
50.5 | 0 | 0 | 37.54 | 73474.1 | 40.01
49.2 | 0 | 0 | 53.13 | 103994.9 | 59.04
48 | 0 | 0 | 69.62 | 136276.7 | 74.64
46.8 | 0 | 0 | 81.71 | 159939.1 | 86.35
45.6 | 0 | 0 | 81.69 | 159886.3 | 92.55
44.4 | 0 | 0 | 87.87 | 171986.7 | 95.3
43.2 | 0 | 0 | 90.84 | 177804 | 96.67
42 | 0 | 0 | 87.61 | 171476.3 | 97.39
40.8 | 0 | 0 | 89.28 | 174757.1 | 97.78
39.6 | 0 | 0 | 91.6 | 179293.6 | 97.99
38.4 | 0 | 0 | 89.41 | 175006.4 | 98.27
37.2 | 0 | 0 | 90.15 | 176450.4 | 98.41
36 | 0 | 0 | 92.42 | 180895.7 | 98.57
34.8 | 0 | 0 | 88.8 | 173818.6 | 98.81
33.6 | 0 | 0 | 91.21 | 178535.6 | 98.95
32.4 | 0 | 0 | 90.95 | 178013.3 | 98.88
31.2 | 0 | 0 | 89.29 | 174764.9 | 99.09
30 | 0 | 0 | 90.81 | 177735.8 | 99.15
28.7 | 0 | 0 | 90.24 | 176624.1 | 99.2
27.6 | 0 | 0 | 88.86 | 173927.3 | 99.43
26.4 | 0 | 0 | 93.87 | 183733.6 | 99.44
25.2 | 0 | 0 | 88.81 | 173823.6 | 99.21
24 | 0 | 0 | 93.05 | 182125.2 | 99.55
22.8 | 0 | 0 | 89.35 | 174895 | 99.71
21.5 | 0 | 0 | 89.62 | 175421.6 | 99.36
20.4 | 0 | 0 | 90.53 | 177188.3 | 99.8
19.2 | 0 | 0 | 90.59 | 177308.3 | 99.85
18 | 0 | 0 | 91.81 | 179702.7 | 99.51
16.7 | 0 | 0 | 90.84 | 177808.8 | 99.9
15.6 | 0 | 0 | 92.36 | 180781.4 | 99.88
14.3 | 0 | 0 | 92.04 | 180161.1 | 99.93
13.1 | 0 | 0 | 90.54 | 177208.2 | 99.92
11.9 | 0 | 0 | 91.83 | 179744.8 | 99.88
10.7 | 0 | 0 | 92.83 | 181705.7 | 99.53
9.5 | 0 | 0 | 92.05 | 180167.6 | 99.97
8.3 | 0 | 0 | 91.6 | 179297 | 99.99
7.1 | 0 | 0 | 92.65 | 181341.8 | 99.61
5.9 | 0 | 0 | 90.35 | 176836.2 | 99.97
4.7 | 0 | 0 | 93.06 | 182155.8 | 99.76
3.5 | 0 | 0 | 91.8 | 179688.7 | 99.76
2.3 | 0 | 0 | 89.88 | 175917.3 | 99.77
1.1 | 0 | 0 | 91.7 | 179490.2 | 99.96
-0.1 | 0 | 0 | 89.76 | 175683 | 99.72
-1.3 | 0 | 0 | 92.39 | 180834.3 | 99.84
-2.5 | 0 | 0 | 90.39 | 176928.2 | 100
-3.7 | 0 | 0 | 90.71 | 177554.7 | 99.65
-4.9 | 0 | 0 | 90.84 | 177797.9 | 99.87
-6.1 | 0 | 0 | 88.82 | 173853.1 | 99.85
-7.3 | 0 | 0 | 93.13 | 182293 | 99.54
-8.5 | 0 | 0 | 88.87 | 173942.2 | 99.75
-9.7 | 0 | 0 | 91.89 | 179867.3 | 99.79
-10.9 | 0 | 0 | 91.82 | 179727.9 | 99.58
-12.1 | 0 | 0 | 89.81 | 175794.6 | 99.71
-13.3 | 0 | 0 | 92.96 | 181953.2 | 99.69
-14.5 | 0 | 0 | 90.77 | 177667 | 99.56
-15.7 | 0 | 0 | 90.6 | 177324 | 99.6
-16.9 | 0 | 0 | 91.08 | 178267.6 | 99.52
-18.1 | 0 | 0 | 90.63 | 177390.4 | 99.44
-19.3 | 0 | 0 | 92.4 | 180847 | 99.43
-20.5 | 0 | 0 | 90.49 | 177125.1 | 99.31
-21.7 | 0 | 0 | 90.48 | 177106.5 | 99.04
-22.9 | 0 | 0 | 91.14 | 178381.7 | 99.26
-24.1 | 0 | 0 | 90.77 | 177668.7 | 99
-25.3 | 0 | 0 | 90.98 | 178082.4 | 98.77
-26.5 | 0 | 0 | 91.06 | 178227.7 | 98.91
-27.7 | 0 | 0 | 90.98 | 178083.7 | 98.78
-28.9 | 0 | 0 | 91.2 | 178508 | 98.62
-30.1 | 0 | 0 | 90.35 | 176850.3 | 98.62
-31.4 | 0 | 0 | 91.55 | 179184.7 | 98.42
-32.6 | 0 | 0 | 86.58 | 169469.7 | 98.26
-33.8 | 0 | 0 | 89.79 | 175747 | 98.11
-35 | 0 | 0 | 90.06 | 176277.9 | 98.05
-36.2 | 0 | 0 | 89.81 | 175776.9 | 97.48
-37.4 | 0 | 0 | 89.44 | 175059.1 | 97.63
-38.6 | 0 | 0 | 88.7 | 173616.5 | 97.37
-39.8 | 0 | 0 | 87.84 | 171927.3 | 96.89
-41 | 0 | 0 | 88.24 | 172714.2 | 96.9
-42.2 | 0 | 0 | 88.21 | 172649 | 96.51
-43.4 | 0 | 0 | 88.06 | 172354.4 | 96.19
-44.6 | 0 | 0 | 87.32 | 170920.9 | 95.83
-45.8 | 0 | 0 | 85.66 | 167664.1 | 94.58
-47 | 0 | 0 | 83.83 | 164084.7 | 91.55
-48.2 | 0 | 0 | 77.66 | 151998.2 | 85.24
-49.4 | 0 | 0 | 65.8 | 128792.9 | 72.78
-50.6 | 0 | 0 | 51.9 | 101592.2 | 56.78
-51.8 | 0 | 0 | 35.04 | 68592.7 | 38.17
-53 | 0 | 0 | 20.47 | 40058.9 | 22.33
-54.2 | 0 | 0 | 11.77 | 23028.9 | 12.72
-55.4 | 0 | 0 | 6.89 | 13477.7 | 7.45
-56.6 | 0 | 0 | 4.93 | 9655.3 | 5.34
-57.8 | 0 | 0 | 3.99 | 7807.4 | 4.35
-59 | 0 | 0 | 3.55 | 6949.6 | 3.9
-60.2 | 0 | 0 | 3.36 | 6569.9 | 3.64
-61.4 | 0 | 0 | 3.16 | 6185.9 | 3.47
-62.6 | 0 | 0 | 2.9 | 5682.5 | 3.21
-63.8 | 0 | 0 | 2.9 | 5678.9 | 3.08
-65 | 0 | 0 | 2.47 | 4832 | 2.86
-66.2 | 0 | 0 | 2.61 | 5099.2 | 2.76
-67.4 | 0 | 0 | 2.35 | 4609.4 | 2.59
-68.6 | 0 | 0 | 2.18 | 4270.6 | 2.51
-69.8 | 0 | 0 | 2.26 | 4414 | 2.4
-71.1 | 0 | 0 | 2.05 | 4010.8 | 2.32
-72.2 | 0 | 0 | 2.03 | 3980 | 2.21
-73.5 | 0 | 0 | 1.94 | 3792.1 | 2.12
-74.6 | 0 | 0 | 1.83 | 3577.2 | 1.99
-75.9 | 0 | 0 | 1.76 | 3451.4 | 1.92
-77.1 | 0 | 0 | 1.66 | 3250.2 | 1.81
-78.2 | 0 | 0 | 1.64 | 3215.3 | 1.78
-79.5 | 0 | 0 | 1.52 | 2976.9 | 1.66
-80.6 | 0 | 0 | 1.49 | 2923.2 | 1.62
-81.9 | 0 | 0 | 1.47 | 2869.9 | 1.6
-83 | 0 | 0 | 1.39 | 2716.8 | 1.5
-84.3 | 0 | 0 | 1.32 | 2587.1 | 1.44
-85.5 | 0 | 0 | 1.24 | 2418.6 | 1.34
-86.7 | 0 | 0 | 1.2 | 2355.8 | 1.3
-87.9 | 0 | 0 | 1.16 | 2275.9 | 1.27
-89.1 | 0 | 0 | 1.08 | 2117.9 | 1.17
-90.3 | 0 | 0 | 1.07 | 2103.5 | 1.15
-91.5 | 0 | 0 | 0.99 | 1941.4 | 1.09
-92.7 | 0 | 0 | 1.03 | 2013.4 | 1.11
-93.9 | 0 | 0 | 0.97 | 1903.1 | 1.05
-95.1 | 0 | 0 | 0.97 | 1890.8 | 1.05
-96.3 | 0 | 0 | 0.94 | 1847.8 | 1.02
-97.5 | 0 | 0 | 0.9 | 1753.7 | 0.98
-98.7 | 0 | 0 | 0.86 | 1678.8 | 0.96
-99.9 | 0 | 0 | 0.87 | 1695.8 | 0.95
-101.1 | 0 | 0 | 0.85 | 1660.6 | 0.92
-102.3 | 0 | 0 | 0.81 | 1579.1 | 0.9
-103.5 | 0 | 0 | 0.82 | 1599.2 | 0.9
-104.7 | 0 | 0 | 0.8 | 1570.5 | 0.86
-105.9 | 0 | 0 | 0.76 | 1487.8 | 0.85
-107.1 | 0 | 0 | 0.77 | 1500 | 0.84
-108.3 | 0 | 0 | 0.79 | 1549.8 | 0.85
-109.5 | 0 | 0 | 0.79 | 1555.4 | 0.9
-110.7 | 0 | 0 | 0.75 | 1467.2 | 0.82
-111.9 | 0 | 0 | 0.75 | 1462.9 | 0.81
-113.1 | 0 | 0 | 0.78 | 1517.1 | 0.84
-114.3 | 0 | 0 | 0.78 | 1527.8 | 0.84
-115.5 | 0 | 0 | 0.81 | 1579 | 0.88
-116.7 | 0 | 0 | 0.75 | 1458.8 | 0.84
-117.9 | 0 | 0 | 0.8 | 1560.4 | 0.87
-119.1 | 0 | 0 | 0.79 | 1551.3 | 0.86
-120.3 | 0 | 0 | 0.83 | 1615.1 | 0.89
-121.5 | 0 | 0 | 0.79 | 1545.8 | 0.87
-122.7 | 0 | 0 | 0.88 | 1731.9 | 0.95
-123.9 | 0 | 0 | 0.85 | 1670.1 | 0.93
-125.1 | 0 | 0 | 0.85 | 1668.6 | 0.93
-126.3 | 0 | 0 | 0.83 | 1625.2 | 0.9
-127.5 | 0 | 0 | 0.83 | 1634.2 | 0.91
-128.7 | 0 | 0 | 0.83 | 1622.6 | 0.9
-129.9 | 0 | 0 | 0.88 | 1714.6 | 0.95
-131.1 | 0 | 0 | 0.84 | 1643 | 0.92
-132.3 | 0 | 0 | 0.86 | 1686.4 | 0.95
-133.5 | 0 | 0 | 0.84 | 1652.6 | 0.92
-134.7 | 0 | 0 | 0.85 | 1666 | 0.92
-135.9 | 0 | 0 | 0.84 | 1651.7 | 0.93
-137.1 | 0 | 0 | 0.83 | 1624.1 | 0.9
-138.3 | 0 | 0 | 0.82 | 1596.5 | 0.89
-139.5 | 0 | 0 | 0.82 | 1596.1 | 0.91
-140.7 | 0 | 0 | 0.82 | 1597.9 | 0.9
-142 | 0 | 0 | 0.79 | 1544.2 | 0.86
-143.2 | 0 | 0 | 0.84 | 1653.8 | 0.93
-144.3 | 0 | 0 | 0.81 | 1581.2 | 0.9
-145.6 | 0 | 0 | 0.81 | 1586.8 | 0.88
-146.8 | 0 | 0 | 0.78 | 1531.3 | 0.87
-148 | 0 | 0 | 0.82 | 1602 | 0.9
-150 | 0 | 0 | 0.78 | 1525.7 | 0.86
Measurement time: | 07/09/2022 17:14:44
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-151.4 | 0 | 28 | 0.4 | 782.9 | 0.4
-150.7 | 0 | 28 | 0.27 | 535.8 | 0.26
-149.3 | 0 | 28 | 0.32 | 617 | 0.31
-148.1 | 0 | 28 | 0.28 | 539.6 | 0.28
-146.9 | 0 | 28 | 0.33 | 645.6 | 0.32
-145.7 | 0 | 28 | 0.33 | 641.4 | 0.31
-144.5 | 0 | 28 | 0.31 | 598.1 | 0.3
-143.3 | 0 | 28 | 0.32 | 622.7 | 0.32
-142.1 | 0 | 28 | 0.31 | 609.3 | 0.29
-140.9 | 0 | 28 | 0.29 | 573.4 | 0.3
-139.7 | 0 | 28 | 0.35 | 682.3 | 0.34
-138.5 | 0 | 28 | 0.37 | 721.1 | 0.34
-137.3 | 0 | 28 | 0.36 | 696.5 | 0.36
-136.1 | 0 | 28 | 0.34 | 666 | 0.33
-134.9 | 0 | 28 | 0.38 | 744.5 | 0.35
-133.7 | 0 | 28 | 0.3 | 594.8 | 0.31
-132.5 | 0 | 28 | 0.33 | 647.8 | 0.32
-131.3 | 0 | 28 | 0.33 | 641.3 | 0.3
-130.1 | 0 | 28 | 0.34 | 657.1 | 0.34
-128.9 | 0 | 28 | 0.35 | 692.3 | 0.34
-127.7 | 0 | 28 | 0.36 | 698.2 | 0.33
-126.5 | 0 | 28 | 0.35 | 686.9 | 0.35
-125.3 | 0 | 28 | 0.34 | 661.5 | 0.32
-124.1 | 0 | 28 | 0.34 | 675.1 | 0.32
-122.9 | 0 | 28 | 0.37 | 728.1 | 0.37
-121.7 | 0 | 28 | 0.35 | 682.5 | 0.33
-120.4 | 0 | 28 | 0.33 | 646.4 | 0.32
-119.3 | 0 | 28 | 0.31 | 599 | 0.3
-118.1 | 0 | 28 | 0.39 | 764.8 | 0.38
-116.9 | 0 | 28 | 0.32 | 621 | 0.31
-115.7 | 0 | 28 | 0.3 | 595 | 0.3
-114.5 | 0 | 28 | 0.35 | 680.7 | 0.33
-113.3 | 0 | 28 | 0.39 | 763.1 | 0.36
-112 | 0 | 28 | 0.34 | 655.8 | 0.33
-110.9 | 0 | 28 | 0.36 | 714.1 | 0.35
-109.7 | 0 | 28 | 0.39 | 766.1 | 0.36
-108.4 | 0 | 28 | 0.32 | 629.6 | 0.32
-107.2 | 0 | 28 | 0.37 | 715.9 | 0.35
-106 | 0 | 28 | 0.43 | 839.9 | 0.4
-104.8 | 0 | 28 | 0.42 | 821.5 | 0.42
-103.6 | 0 | 28 | 0.42 | 812.6 | 0.4
-102.4 | 0 | 28 | 0.46 | 893.2 | 0.42
-101.2 | 0 | 28 | 0.45 | 889.6 | 0.45
-100 | 0 | 28 | 0.48 | 933.2 | 0.46
-98.8 | 0 | 28 | 0.5 | 986.1 | 0.47
-97.6 | 0 | 28 | 0.5 | 983.4 | 0.49
-96.4 | 0 | 28 | 0.56 | 1090.8 | 0.54
-95.2 | 0 | 28 | 0.63 | 1242.4 | 0.6
-94 | 0 | 28 | 0.62 | 1216.4 | 0.61
-92.8 | 0 | 28 | 0.62 | 1221.4 | 0.6
-91.6 | 0 | 28 | 0.71 | 1397.3 | 0.67
-90.4 | 0 | 28 | 0.66 | 1298.4 | 0.65
-89.2 | 0 | 28 | 0.78 | 1531.2 | 0.75
-88 | 0 | 28 | 0.86 | 1680 | 0.81
-86.8 | 0 | 28 | 0.93 | 1822.2 | 0.92
-85.6 | 0 | 28 | 1.06 | 2073.9 | 1.02
-84.4 | 0 | 28 | 1.18 | 2302 | 1.12
-83.2 | 0 | 28 | 1.24 | 2423.7 | 1.22
-82 | 0 | 28 | 1.39 | 2712.6 | 1.35
-80.8 | 0 | 28 | 1.54 | 3009.4 | 1.49
-79.6 | 0 | 28 | 1.7 | 3324.7 | 1.67
-78.4 | 0 | 28 | 1.86 | 3646.2 | 1.82
-77.2 | 0 | 28 | 2.2 | 4310.1 | 2.11
-76 | 0 | 28 | 2.51 | 4903.4 | 2.44
-74.8 | 0 | 28 | 2.77 | 5430 | 2.7
-73.6 | 0 | 28 | 3.25 | 6358.9 | 3.16
-72.4 | 0 | 28 | 3.77 | 7376.1 | 3.66
-71.2 | 0 | 28 | 4.3 | 8422.1 | 4.22
-70 | 0 | 28 | 5.06 | 9913.7 | 4.9
-68.8 | 0 | 28 | 5.83 | 11414.3 | 5.68
-67.6 | 0 | 28 | 6.81 | 13328.6 | 6.65
-66.4 | 0 | 28 | 8.11 | 15869.3 | 7.82
-65.2 | 0 | 28 | 9.46 | 18518.9 | 9.16
-63.9 | 0 | 28 | 11.22 | 21963.2 | 10.84
-62.7 | 0 | 28 | 13.45 | 26332 | 12.87
-61.5 | 0 | 28 | 15.87 | 31069.7 | 15.26
-60.3 | 0 | 28 | 18.89 | 36979.5 | 18.19
-59.1 | 0 | 28 | 23.09 | 45187.9 | 21.74
-57.9 | 0 | 28 | 26.76 | 52373.7 | 25.94
-56.7 | 0 | 28 | 32.03 | 62699.5 | 30.69
-55.5 | 0 | 28 | 37.69 | 73771.2 | 36.11
-54.3 | 0 | 28 | 43.63 | 85405.9 | 42.11
-53.1 | 0 | 28 | 50.46 | 98766.5 | 48.23
-51.9 | 0 | 28 | 57.34 | 112222.8 | 54.52
-50.7 | 0 | 28 | 62.73 | 122784.2 | 60.84
-49.5 | 0 | 28 | 70.07 | 137151.8 | 66.77
-48.3 | 0 | 28 | 71.92 | 140771.6 | 71.95
-47.1 | 0 | 28 | 79.79 | 156169.8 | 76.75
-45.9 | 0 | 28 | 84.43 | 165264.5 | 80.67
-44.7 | 0 | 28 | 88.23 | 172703.5 | 84
-43.5 | 0 | 28 | 90.14 | 176433.7 | 86.83
-42.3 | 0 | 28 | 92.96 | 181954.9 | 89.06
-41.1 | 0 | 28 | 94.54 | 185039.1 | 90.7
-39.9 | 0 | 28 | 95.54 | 186995.6 | 92.12
-38.7 | 0 | 28 | 96.6 | 189078.9 | 93.39
-37.5 | 0 | 28 | 98.02 | 191864.7 | 94.43
-36.3 | 0 | 28 | 97.65 | 191134.4 | 95.11
-35.1 | 0 | 28 | 98.58 | 192959.9 | 95.88
-33.9 | 0 | 28 | 100.18 | 196090.2 | 96.32
-32.7 | 0 | 28 | 99.32 | 194408.6 | 96.62
-31.5 | 0 | 28 | 99.71 | 195164.3 | 97.26
-30.3 | 0 | 28 | 101.42 | 198518.9 | 97.47
-29.1 | 0 | 28 | 99.96 | 195652.8 | 97.92
-27.9 | 0 | 28 | 102.28 | 200198 | 98.1
-26.7 | 0 | 28 | 102 | 199644.4 | 98.35
-25.5 | 0 | 28 | 102.22 | 200080.3 | 98.62
-24.3 | 0 | 28 | 102.47 | 200556.5 | 98.63
-23.1 | 0 | 28 | 103.06 | 201728 | 98.7
-21.9 | 0 | 28 | 103.9 | 203372.9 | 99.11
-20.7 | 0 | 28 | 97.88 | 191579.5 | 98.99
-19.5 | 0 | 28 | 105.33 | 206167.8 | 99.34
-18.3 | 0 | 28 | 102.18 | 200004.5 | 99.35
-17.1 | 0 | 28 | 103.79 | 203140 | 99.48
-15.9 | 0 | 28 | 104.36 | 204265.3 | 99.46
-14.7 | 0 | 28 | 102.03 | 199711.8 | 99.33
-13.5 | 0 | 28 | 104.35 | 204253.4 | 99.53
-12.3 | 0 | 28 | 103.39 | 202361.2 | 99.64
-11.1 | 0 | 28 | 102.13 | 199900.6 | 99.53
-9.8 | 0 | 28 | 104.46 | 204466.5 | 99.67
-8.6 | 0 | 28 | 102.17 | 199977.2 | 99.62
-7.4 | 0 | 28 | 102.84 | 201283.9 | 99.69
-6.2 | 0 | 28 | 102.56 | 200745 | 99.73
-5 | 0 | 28 | 102.77 | 201149.1 | 99.75
-3.8 | 0 | 28 | 106.67 | 208783.7 | 99.85
-2.6 | 0 | 28 | 99.14 | 194044.5 | 99.95
-1.4 | 0 | 28 | 104.04 | 203630.3 | 100
-0.2 | 0 | 28 | 103.57 | 202715 | 99.87
1 | 0 | 28 | 103.99 | 203538.8 | 99.95
2.2 | 0 | 28 | 104.7 | 204927.9 | 99.93
3.4 | 0 | 28 | 104.59 | 204715.3 | 99.99
4.6 | 0 | 28 | 98.73 | 193237.9 | 99.75
5.8 | 0 | 28 | 104.79 | 205107.9 | 99.97
7 | 0 | 28 | 103.91 | 203393.9 | 99.87
8.2 | 0 | 28 | 104.07 | 203697.2 | 99.83
9.4 | 0 | 28 | 103.79 | 203147.3 | 99.89
10.6 | 0 | 28 | 103.78 | 203127.5 | 99.58
11.8 | 0 | 28 | 103.65 | 202884.8 | 99.74
13 | 0 | 28 | 103.51 | 202607 | 99.53
14.2 | 0 | 28 | 102.99 | 201591.1 | 99.38
15.4 | 0 | 28 | 102.45 | 200533.7 | 99.55
16.6 | 0 | 28 | 100.89 | 197476.3 | 99.36
17.8 | 0 | 28 | 105.47 | 206434.4 | 99.32
19 | 0 | 28 | 98.5 | 192786.1 | 99.52
20.2 | 0 | 28 | 102.33 | 200286.7 | 99.35
21.4 | 0 | 28 | 102.67 | 200948.2 | 99.06
22.6 | 0 | 28 | 102.7 | 201012.1 | 99.21
23.8 | 0 | 28 | 102.48 | 200585.7 | 99.04
25 | 0 | 28 | 102.36 | 200343.6 | 98.7
26.2 | 0 | 28 | 101.81 | 199268 | 98.6
27.4 | 0 | 28 | 102.62 | 200856.2 | 98.51
28.6 | 0 | 28 | 102.31 | 200254.5 | 98.11
29.8 | 0 | 28 | 102.13 | 199903.9 | 97.74
31 | 0 | 28 | 102.07 | 199780 | 97.28
32.2 | 0 | 28 | 101.65 | 198956.6 | 96.85
33.4 | 0 | 28 | 94.67 | 185308.3 | 96.44
34.6 | 0 | 28 | 99.62 | 194983.7 | 95.66
35.8 | 0 | 28 | 97.94 | 191704.1 | 94.87
37 | 0 | 28 | 98.54 | 192866.4 | 93.78
38.2 | 0 | 28 | 96.82 | 189514.8 | 92.76
39.4 | 0 | 28 | 94.4 | 184765.1 | 91.24
40.6 | 0 | 28 | 91.92 | 179922 | 89.55
41.8 | 0 | 28 | 92.48 | 181010.3 | 87.03
43 | 0 | 28 | 86.81 | 169920.1 | 84.5
44.2 | 0 | 28 | 82.87 | 162210.3 | 81.18
45.4 | 0 | 28 | 75.82 | 148399.5 | 77.39
46.7 | 0 | 28 | 78.55 | 153748.4 | 72.85
47.9 | 0 | 28 | 67.07 | 131268.4 | 67.48
49.1 | 0 | 28 | 64.65 | 126532.4 | 61.89
50.3 | 0 | 28 | 58.06 | 113644.1 | 55.72
51.5 | 0 | 28 | 51.59 | 100986.5 | 49.28
52.7 | 0 | 28 | 43.17 | 84489.7 | 43.03
53.9 | 0 | 28 | 38.53 | 75414.1 | 36.9
55.1 | 0 | 28 | 32.97 | 64523.8 | 31.33
56.3 | 0 | 28 | 25.9 | 50687.7 | 26.32
57.5 | 0 | 28 | 23.04 | 45100.7 | 21.93
58.7 | 0 | 28 | 19.08 | 37340.5 | 18.23
59.9 | 0 | 28 | 15.76 | 30852.3 | 15.13
61.1 | 0 | 28 | 13.29 | 26014.2 | 12.71
62.3 | 0 | 28 | 11.03 | 21591.6 | 10.63
63.5 | 0 | 28 | 9.34 | 18286.5 | 8.96
64.7 | 0 | 28 | 7.87 | 15394.5 | 7.59
65.9 | 0 | 28 | 6.75 | 13216.3 | 6.49
67.1 | 0 | 28 | 5.83 | 11417.2 | 5.59
68.3 | 0 | 28 | 4.96 | 9701.4 | 4.78
69.5 | 0 | 28 | 4.31 | 8443.9 | 4.13
70.7 | 0 | 28 | 3.72 | 7286.8 | 3.62
71.9 | 0 | 28 | 3.26 | 6372.1 | 3.14
73.1 | 0 | 28 | 2.91 | 5689.9 | 2.77
74.3 | 0 | 28 | 2.53 | 4945 | 2.46
75.5 | 0 | 28 | 2.26 | 4432.6 | 2.16
76.8 | 0 | 28 | 1.99 | 3893.2 | 1.94
77.9 | 0 | 28 | 1.85 | 3611.4 | 1.77
79.1 | 0 | 28 | 1.65 | 3233.2 | 1.59
80.4 | 0 | 28 | 1.45 | 2829.9 | 1.42
81.5 | 0 | 28 | 1.31 | 2567.3 | 1.26
82.8 | 0 | 28 | 1.19 | 2328.8 | 1.16
83.9 | 0 | 28 | 1.03 | 2010.5 | 1
85.2 | 0 | 28 | 0.97 | 1905.5 | 0.95
86.4 | 0 | 28 | 0.86 | 1687 | 0.84
87.6 | 0 | 28 | 0.74 | 1441.7 | 0.72
88.8 | 0 | 28 | 0.71 | 1388.3 | 0.69
90 | 0 | 28 | 0.63 | 1231.6 | 0.61
91.2 | 0 | 28 | 0.57 | 1116.3 | 0.55
92.4 | 0 | 28 | 0.52 | 1010.3 | 0.5
93.6 | 0 | 28 | 0.48 | 948.7 | 0.47
94.8 | 0 | 28 | 0.46 | 903 | 0.45
96 | 0 | 28 | 0.43 | 834.6 | 0.41
97.2 | 0 | 28 | 0.41 | 807.9 | 0.39
98.4 | 0 | 28 | 0.37 | 716.5 | 0.35
99.6 | 0 | 28 | 0.37 | 724.2 | 0.36
100.8 | 0 | 28 | 0.32 | 624.1 | 0.31
102 | 0 | 28 | 0.34 | 667.1 | 0.33
103.2 | 0 | 28 | 0.31 | 605.9 | 0.3
104.4 | 0 | 28 | 0.34 | 674.1 | 0.34
105.6 | 0 | 28 | 0.34 | 669.5 | 0.33
106.8 | 0 | 28 | 0.3 | 593.4 | 0.29
108 | 0 | 28 | 0.32 | 626.4 | 0.31
109.2 | 0 | 28 | 0.28 | 555.8 | 0.28
110.4 | 0 | 28 | 0.26 | 515.3 | 0.25
111.6 | 0 | 28 | 0.28 | 540.4 | 0.26
112.8 | 0 | 28 | 0.25 | 484.8 | 0.24
114 | 0 | 28 | 0.27 | 533 | 0.26
115.2 | 0 | 28 | 0.23 | 459.7 | 0.22
116.4 | 0 | 28 | 0.25 | 494 | 0.25
117.6 | 0 | 28 | 0.26 | 508.2 | 0.25
118.8 | 0 | 28 | 0.21 | 414.3 | 0.2
120 | 0 | 28 | 0.24 | 462.3 | 0.23
121.2 | 0 | 28 | 0.18 | 351.9 | 0.17
122.4 | 0 | 28 | 0.2 | 382.8 | 0.19
123.6 | 0 | 28 | 0.22 | 429.3 | 0.21
124.8 | 0 | 28 | 0.21 | 412.6 | 0.21
126 | 0 | 28 | 0.21 | 416 | 0.21
127.2 | 0 | 28 | 0.15 | 295 | 0.15
128.4 | 0 | 28 | 0.2 | 400.5 | 0.2
129.6 | 0 | 28 | 0.19 | 368.6 | 0.18
130.9 | 0 | 28 | 0.18 | 350.7 | 0.18
132 | 0 | 28 | 0.17 | 330.7 | 0.16
133.2 | 0 | 28 | 0.17 | 333.4 | 0.16
134.4 | 0 | 28 | 0.14 | 278.3 | 0.14
135.6 | 0 | 28 | 0.17 | 336.1 | 0.16
136.8 | 0 | 28 | 0.18 | 350.8 | 0.17
138 | 0 | 28 | 0.18 | 349.8 | 0.18
139.2 | 0 | 28 | 0.2 | 395.5 | 0.19
140.4 | 0 | 28 | 0.16 | 319.6 | 0.15
141.6 | 0 | 28 | 0.17 | 339.4 | 0.17
142.9 | 0 | 28 | 0.19 | 367.3 | 0.18
144.1 | 0 | 28 | 0.14 | 272.7 | 0.13
145.3 | 0 | 28 | 0.2 | 386.3 | 0.2
146.5 | 0 | 28 | 0.15 | 297.3 | 0.15
147.7 | 0 | 28 | 0.17 | 337.8 | 0.16
148.9 | 0 | 28 | 0.18 | 361.2 | 0.18
150.1 | 0 | 28 | 0.18 | 349.5 | 0.17
151.4 | 0 | 28 | 0.16 | 308.1 | 0.15
Measurement time: | 07/09/2022 17:16:32
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
152.5 | 0 | 50 | 0.16 | 319.2 | 0.25
151.8 | 0 | 50 | 0.1 | 192.6 | 0.2
150.4 | 0 | 50 | 0.08 | 162.1 | 0.17
149.2 | 0 | 50 | 0.05 | 103 | 0.11
148 | 0 | 50 | 0.06 | 120.4 | 0.13
146.8 | 0 | 50 | 0.07 | 136 | 0.14
145.6 | 0 | 50 | 0.06 | 112.2 | 0.12
144.4 | 0 | 50 | 0.06 | 118.5 | 0.13
143.2 | 0 | 50 | 0.09 | 184 | 0.19
142 | 0 | 50 | 0.11 | 209.7 | 0.23
140.8 | 0 | 50 | 0.12 | 236.4 | 0.25
139.6 | 0 | 50 | 0.12 | 231.6 | 0.23
138.4 | 0 | 50 | 0.1 | 196.8 | 0.21
137.2 | 0 | 50 | 0.1 | 194.2 | 0.2
136 | 0 | 50 | 0.11 | 215.9 | 0.22
134.8 | 0 | 50 | 0.14 | 270.7 | 0.29
133.6 | 0 | 50 | 0.12 | 227.8 | 0.24
132.4 | 0 | 50 | 0.1 | 195.7 | 0.19
131.2 | 0 | 50 | 0.09 | 171.7 | 0.19
130 | 0 | 50 | 0.14 | 272 | 0.28
128.8 | 0 | 50 | 0.12 | 237 | 0.24
127.5 | 0 | 50 | 0.11 | 222.2 | 0.24
126.4 | 0 | 50 | 0.16 | 314.8 | 0.33
125.1 | 0 | 50 | 0.12 | 226.9 | 0.24
124 | 0 | 50 | 0.13 | 262.1 | 0.29
122.8 | 0 | 50 | 0.17 | 336.2 | 0.35
121.5 | 0 | 50 | 0.14 | 277.5 | 0.29
120.4 | 0 | 50 | 0.15 | 291.3 | 0.31
119.1 | 0 | 50 | 0.14 | 282 | 0.29
117.9 | 0 | 50 | 0.17 | 325 | 0.32
116.7 | 0 | 50 | 0.12 | 225.4 | 0.24
115.5 | 0 | 50 | 0.1 | 191.5 | 0.2
114.4 | 0 | 50 | 0.2 | 388.8 | 0.39
113.1 | 0 | 50 | 0.12 | 242.4 | 0.26
111.9 | 0 | 50 | 0.17 | 325 | 0.34
110.7 | 0 | 50 | 0.17 | 332 | 0.34
109.5 | 0 | 50 | 0.17 | 337.1 | 0.36
108.4 | 0 | 50 | 0.18 | 356.4 | 0.37
107.1 | 0 | 50 | 0.23 | 454.9 | 0.46
105.9 | 0 | 50 | 0.21 | 414.2 | 0.45
104.7 | 0 | 50 | 0.18 | 347.8 | 0.37
103.5 | 0 | 50 | 0.21 | 417.4 | 0.42
102.3 | 0 | 50 | 0.23 | 449.1 | 0.49
101.1 | 0 | 50 | 0.22 | 426.5 | 0.45
99.9 | 0 | 50 | 0.22 | 434.4 | 0.45
98.7 | 0 | 50 | 0.23 | 442.1 | 0.48
97.5 | 0 | 50 | 0.22 | 427.2 | 0.45
96.3 | 0 | 50 | 0.27 | 522.6 | 0.54
95.1 | 0 | 50 | 0.3 | 583.2 | 0.62
93.9 | 0 | 50 | 0.29 | 566.3 | 0.59
92.7 | 0 | 50 | 0.33 | 637.7 | 0.66
91.5 | 0 | 50 | 0.33 | 654.5 | 0.7
90.3 | 0 | 50 | 0.38 | 734.6 | 0.77
89.1 | 0 | 50 | 0.43 | 842.4 | 0.87
87.9 | 0 | 50 | 0.46 | 895.2 | 0.95
86.7 | 0 | 50 | 0.5 | 980.2 | 1.01
85.5 | 0 | 50 | 0.54 | 1047.3 | 1.08
84.3 | 0 | 50 | 0.54 | 1063.5 | 1.13
83.1 | 0 | 50 | 0.61 | 1203.1 | 1.26
81.9 | 0 | 50 | 0.71 | 1397.2 | 1.43
80.7 | 0 | 50 | 0.78 | 1529.2 | 1.61
79.5 | 0 | 50 | 0.91 | 1790.9 | 1.88
78.3 | 0 | 50 | 1.11 | 2176.2 | 2.21
77.1 | 0 | 50 | 1.17 | 2282.8 | 2.47
75.9 | 0 | 50 | 1.46 | 2857.4 | 3
74.7 | 0 | 50 | 1.83 | 3584.4 | 3.67
73.5 | 0 | 50 | 2.1 | 4100.7 | 4.33
72.3 | 0 | 50 | 2.58 | 5058.7 | 5.19
71 | 0 | 50 | 3.04 | 5959 | 6.14
69.9 | 0 | 50 | 3.57 | 6989.8 | 7.37
68.7 | 0 | 50 | 4.31 | 8441.1 | 8.72
67.4 | 0 | 50 | 5.19 | 10157.9 | 10.48
66.3 | 0 | 50 | 5.93 | 11603.9 | 12.12
65 | 0 | 50 | 6.93 | 13556.9 | 14.17
63.8 | 0 | 50 | 8.17 | 15992 | 16.49
62.6 | 0 | 50 | 9.18 | 17966.8 | 18.8
61.4 | 0 | 50 | 10.43 | 20409.7 | 21.35
60.2 | 0 | 50 | 11.88 | 23249.2 | 24.15
59 | 0 | 50 | 13.42 | 26259.8 | 27.28
57.8 | 0 | 50 | 14.91 | 29184.5 | 30.48
56.6 | 0 | 50 | 16.44 | 32180.4 | 33.86
55.4 | 0 | 50 | 18.65 | 36503.5 | 37.65
54.2 | 0 | 50 | 19.87 | 38899.1 | 41.13
53 | 0 | 50 | 22.39 | 43830.1 | 45.08
51.8 | 0 | 50 | 24.1 | 47167.5 | 49.13
50.6 | 0 | 50 | 25.17 | 49268 | 52.68
49.4 | 0 | 50 | 28.2 | 55199.8 | 56.66
48.2 | 0 | 50 | 28.76 | 56289 | 60.35
47 | 0 | 50 | 31.39 | 61440.1 | 64.04
45.8 | 0 | 50 | 32.89 | 64370.3 | 67.59
44.6 | 0 | 50 | 34.87 | 68244.1 | 70.75
43.4 | 0 | 50 | 35.29 | 69079.8 | 74.11
42.2 | 0 | 50 | 36.14 | 70737.3 | 76.6
41 | 0 | 50 | 39.73 | 77763.5 | 79.55
39.8 | 0 | 50 | 39.5 | 77318.4 | 82.24
38.6 | 0 | 50 | 39.94 | 78181.3 | 84.5
37.4 | 0 | 50 | 43.57 | 85280.2 | 87
36.2 | 0 | 50 | 43.71 | 85550.3 | 88.7
34.9 | 0 | 50 | 42.05 | 82296.2 | 90.8
33.8 | 0 | 50 | 45.23 | 88529.8 | 91.87
32.6 | 0 | 50 | 43.54 | 85227.7 | 93.18
31.3 | 0 | 50 | 45.37 | 88800.8 | 93.69
30.2 | 0 | 50 | 47.14 | 92268.1 | 94.36
28.9 | 0 | 50 | 45.82 | 89678.5 | 95.58
27.8 | 0 | 50 | 45.43 | 88916.4 | 95.74
26.6 | 0 | 50 | 46.72 | 91437.2 | 96.5
25.3 | 0 | 50 | 47.95 | 93862.3 | 97.54
24.1 | 0 | 50 | 47.61 | 93186.3 | 97.68
22.9 | 0 | 50 | 48.2 | 94342.5 | 98.44
21.7 | 0 | 50 | 48.31 | 94550 | 98.85
20.5 | 0 | 50 | 48.21 | 94354.8 | 98.41
19.3 | 0 | 50 | 48.76 | 95442.5 | 98.8
18.1 | 0 | 50 | 48.1 | 94151 | 98.69
16.9 | 0 | 50 | 48.5 | 94929.9 | 98.85
15.7 | 0 | 50 | 48.17 | 94287.4 | 98.92
14.5 | 0 | 50 | 48.07 | 94081.3 | 98.64
13.3 | 0 | 50 | 48.96 | 95838.8 | 99.56
12.1 | 0 | 50 | 48.73 | 95373.2 | 99.36
10.9 | 0 | 50 | 48.75 | 95424 | 99.25
9.7 | 0 | 50 | 46.76 | 91530.7 | 99.56
8.5 | 0 | 50 | 48.95 | 95809.4 | 99.62
7.3 | 0 | 50 | 49.35 | 96597.4 | 99.42
6.1 | 0 | 50 | 49 | 95906.3 | 99.81
4.9 | 0 | 50 | 49.02 | 95949.9 | 99.53
3.7 | 0 | 50 | 49.18 | 96253.3 | 99.84
2.5 | 0 | 50 | 48.98 | 95869.2 | 99.68
1.3 | 0 | 50 | 49.07 | 96050.5 | 99.6
0.1 | 0 | 50 | 49.33 | 96552.8 | 100
-1.1 | 0 | 50 | 48.29 | 94523.4 | 99.06
-2.3 | 0 | 50 | 48.94 | 95800.3 | 99.53
-3.5 | 0 | 50 | 49.04 | 95987.3 | 99.69
-4.7 | 0 | 50 | 48.41 | 94758.1 | 99.3
-5.9 | 0 | 50 | 49.08 | 96066.7 | 99.39
-7.1 | 0 | 50 | 48.38 | 94691.5 | 99.43
-8.3 | 0 | 50 | 48.75 | 95423.3 | 99.39
-9.5 | 0 | 50 | 48.92 | 95754.4 | 99.71
-10.7 | 0 | 50 | 48.9 | 95717.5 | 99.63
-11.9 | 0 | 50 | 48.9 | 95707.1 | 99.35
-13.1 | 0 | 50 | 48.21 | 94364 | 99.43
-14.3 | 0 | 50 | 48.2 | 94345.3 | 99.14
-15.5 | 0 | 50 | 47.9 | 93753 | 99.21
-16.7 | 0 | 50 | 49.98 | 97818.7 | 99.12
-17.9 | 0 | 50 | 47.38 | 92738 | 99.02
-19.1 | 0 | 50 | 47.13 | 92249 | 99.35
-20.3 | 0 | 50 | 47.81 | 93576 | 98.77
-21.5 | 0 | 50 | 47.65 | 93275.6 | 98.53
-22.8 | 0 | 50 | 49.02 | 95953.8 | 98.65
-24 | 0 | 50 | 46.82 | 91644.3 | 98.03
-25.2 | 0 | 50 | 47.79 | 93543.1 | 97.73
-26.4 | 0 | 50 | 45.69 | 89433.6 | 97.37
-27.6 | 0 | 50 | 47.75 | 93463.6 | 96.96
-28.8 | 0 | 50 | 47.22 | 92423.5 | 96.4
-30 | 0 | 50 | 45.17 | 88419.1 | 95.58
-31.2 | 0 | 50 | 46.87 | 91731.5 | 94.8
-32.4 | 0 | 50 | 45.04 | 88159.4 | 94.12
-33.6 | 0 | 50 | 43.91 | 85950.5 | 92.72
-34.8 | 0 | 50 | 44.7 | 87498.4 | 91.51
-36 | 0 | 50 | 43.64 | 85422.5 | 90.33
-37.2 | 0 | 50 | 41.8 | 81816.4 | 88.45
-38.4 | 0 | 50 | 42.66 | 83503.2 | 86.85
-39.6 | 0 | 50 | 40.96 | 80170.4 | 84.93
-40.8 | 0 | 50 | 39.26 | 76849.2 | 82.58
-42 | 0 | 50 | 39.53 | 77381.2 | 80.33
-43.2 | 0 | 50 | 37.71 | 73804.8 | 77.45
-44.4 | 0 | 50 | 35.54 | 69553.7 | 74.94
-45.6 | 0 | 50 | 35.16 | 68825.4 | 71.64
-46.8 | 0 | 50 | 32.67 | 63946 | 68.51
-48 | 0 | 50 | 32.01 | 62662.7 | 65.63
-49.2 | 0 | 50 | 30.33 | 59360.7 | 61.91
-50.4 | 0 | 50 | 28.3 | 55391.6 | 58.26
-51.6 | 0 | 50 | 26.18 | 51245.2 | 54.54
-52.8 | 0 | 50 | 24.65 | 48239 | 50.45
-54 | 0 | 50 | 22.41 | 43858.8 | 46.65
-55.2 | 0 | 50 | 20.63 | 40377.5 | 43.03
-56.4 | 0 | 50 | 19.36 | 37894.5 | 39.27
-57.6 | 0 | 50 | 17.4 | 34051.1 | 35.67
-58.8 | 0 | 50 | 15.73 | 30787.6 | 32.12
-60 | 0 | 50 | 14.18 | 27759.2 | 28.82
-61.2 | 0 | 50 | 12.46 | 24388.3 | 25.63
-62.4 | 0 | 50 | 11.17 | 21860.5 | 22.75
-63.6 | 0 | 50 | 9.84 | 19252.6 | 19.99
-64.8 | 0 | 50 | 8.5 | 16631.5 | 17.45
-66 | 0 | 50 | 7.51 | 14689.9 | 15.16
-67.2 | 0 | 50 | 6.36 | 12444.9 | 13.09
-68.4 | 0 | 50 | 5.46 | 10681.1 | 11.06
-69.7 | 0 | 50 | 4.62 | 9034.1 | 9.36
-70.9 | 0 | 50 | 3.87 | 7572.1 | 7.95
-72 | 0 | 50 | 3.3 | 6451.4 | 6.67
-73.3 | 0 | 50 | 2.77 | 5429 | 5.65
-74.5 | 0 | 50 | 2.31 | 4512.4 | 4.7
-75.7 | 0 | 50 | 1.95 | 3811.5 | 3.95
-76.9 | 0 | 50 | 1.57 | 3082 | 3.23
-78 | 0 | 50 | 1.38 | 2696.9 | 2.8
-79.3 | 0 | 50 | 1.13 | 2220.4 | 2.32
-80.4 | 0 | 50 | 1 | 1955.2 | 2.04
-81.7 | 0 | 50 | 0.85 | 1656.9 | 1.72
-82.9 | 0 | 50 | 0.77 | 1503.6 | 1.57
-84.1 | 0 | 50 | 0.7 | 1378.3 | 1.44
-85.3 | 0 | 50 | 0.64 | 1248.1 | 1.29
-86.5 | 0 | 50 | 0.59 | 1152.3 | 1.2
-87.7 | 0 | 50 | 0.53 | 1044.2 | 1.09
-88.9 | 0 | 50 | 0.52 | 1016.3 | 1.05
-90.1 | 0 | 50 | 0.44 | 863.4 | 0.92
-91.3 | 0 | 50 | 0.47 | 927.5 | 0.95
-92.5 | 0 | 50 | 0.44 | 858.3 | 0.89
-93.7 | 0 | 50 | 0.38 | 734.5 | 0.77
-94.9 | 0 | 50 | 0.35 | 688.7 | 0.71
-96.1 | 0 | 50 | 0.33 | 636.8 | 0.65
-97.3 | 0 | 50 | 0.35 | 691 | 0.72
-98.5 | 0 | 50 | 0.3 | 593.6 | 0.62
-99.7 | 0 | 50 | 0.28 | 554.4 | 0.57
-100.9 | 0 | 50 | 0.3 | 584.1 | 0.61
-102.1 | 0 | 50 | 0.27 | 527.3 | 0.56
-103.3 | 0 | 50 | 0.3 | 588.8 | 0.6
-104.5 | 0 | 50 | 0.27 | 536.2 | 0.56
-105.7 | 0 | 50 | 0.23 | 457.2 | 0.5
-106.9 | 0 | 50 | 0.26 | 500.9 | 0.51
-108.1 | 0 | 50 | 0.24 | 462.5 | 0.49
-109.3 | 0 | 50 | 0.27 | 522.8 | 0.54
-110.5 | 0 | 50 | 0.24 | 469.1 | 0.48
-111.7 | 0 | 50 | 0.27 | 538 | 0.57
-112.9 | 0 | 50 | 0.19 | 373 | 0.39
-114.1 | 0 | 50 | 0.26 | 501.4 | 0.52
-115.3 | 0 | 50 | 0.23 | 448.5 | 0.48
-116.5 | 0 | 50 | 0.22 | 433.8 | 0.46
-117.7 | 0 | 50 | 0.2 | 399.7 | 0.4
-118.9 | 0 | 50 | 0.19 | 375.5 | 0.41
-120.1 | 0 | 50 | 0.17 | 337.8 | 0.36
-121.3 | 0 | 50 | 0.18 | 360.1 | 0.37
-122.5 | 0 | 50 | 0.17 | 342.1 | 0.37
-123.7 | 0 | 50 | 0.18 | 360.7 | 0.37
-124.9 | 0 | 50 | 0.2 | 401 | 0.43
-126.1 | 0 | 50 | 0.18 | 342.7 | 0.36
-127.3 | 0 | 50 | 0.18 | 354.6 | 0.39
-128.5 | 0 | 50 | 0.21 | 414.3 | 0.42
-129.7 | 0 | 50 | 0.19 | 362.9 | 0.38
-130.9 | 0 | 50 | 0.16 | 313.3 | 0.33
-132.1 | 0 | 50 | 0.16 | 321.6 | 0.33
-133.3 | 0 | 50 | 0.2 | 400.1 | 0.44
-134.5 | 0 | 50 | 0.19 | 378.1 | 0.4
-135.7 | 0 | 50 | 0.18 | 360.5 | 0.36
-137 | 0 | 50 | 0.18 | 359.6 | 0.39
-138.1 | 0 | 50 | 0.18 | 353.7 | 0.37
-139.4 | 0 | 50 | 0.19 | 379.9 | 0.38
-140.5 | 0 | 50 | 0.18 | 351.2 | 0.38
-141.8 | 0 | 50 | 0.18 | 344.1 | 0.36
-143 | 0 | 50 | 0.16 | 316.2 | 0.32
-144.1 | 0 | 50 | 0.12 | 240.4 | 0.26
-145.4 | 0 | 50 | 0.11 | 222.3 | 0.23
-146.6 | 0 | 50 | 0.15 | 300.4 | 0.3
-147.8 | 0 | 50 | 0.11 | 224.2 | 0.24
-149 | 0 | 50 | 0.14 | 273.5 | 0.28
-150.2 | 0 | 50 | 0.09 | 175.5 | 0.18
-151.4 | 0 | 50 | 0.11 | 219.5 | 0.24
-152.5 | 0 | 50 | 0.14 | 267.6 | 0.28