## Feuil1
Measurement time: | 02/09/2022 20:27:12 | [5 empty] | Taille de champ (cmxcm) | Symérie (%) | Homogénéité (%) | Pénombre G (cm) | Pénombre D (cm)
Radiation device: | Clinac3 Accelerator | [5 empty] | 5x5 | 3.01 | 10.66 | 1.22 | 1.2
Energy: | 12 MeV | [5 empty] | 10x10 | 4.09 | 5.59 | 1.27 | 1.28
Controller: | CCU | [5 empty] | 15x15 | 1.76 | 3.56 | 1.38 | 1.39
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | 20x20 | 0.6 | 1.17 | 1.34 | 1.35
Field detector: | CC 13 - F | [5 empty] | 25x25 | 0.52 | 1.15 | 1.28 | 1.28
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-125.7 | 0 | 28 | 1.37 | 2008.8 | 1.58 | 1.2244
-125 | 0 | 28 | 1.67 | 2446.6 | 1.6 | 1.2399
-123.6 | 0 | 28 | 1.65 | 2414.4 | 1.6 | 1.2399
-122.4 | 0 | 28 | 1.66 | 2433.9 | 1.64 | 1.2709
-121.2 | 0 | 28 | 1.64 | 2395.5 | 1.6 | 1.2399
-120 | 0 | 28 | 1.62 | 2367.3 | 1.56 | 1.2089
-118.8 | 0 | 28 | 1.65 | 2410.5 | 1.62 | 1.2554
-117.6 | 0 | 28 | 1.74 | 2547.4 | 1.73 | 1.3407
-116.4 | 0 | 28 | 1.64 | 2404.2 | 1.55 | 1.2012
-115.2 | 0 | 28 | 1.69 | 2482.8 | 1.69 | 1.3097
-114 | 0 | 28 | 1.66 | 2430.2 | 1.6 | 1.2399
-112.8 | 0 | 28 | 1.61 | 2362.5 | 1.51 | 1.1702
-111.6 | 0 | 28 | 1.67 | 2447.7 | 1.67 | 1.2942
-110.4 | 0 | 28 | 1.68 | 2453.8 | 1.6 | 1.2399
-109.2 | 0 | 28 | 1.58 | 2321.1 | 1.48 | 1.1469
-108 | 0 | 28 | 1.64 | 2408 | 1.63 | 1.2632
-106.8 | 0 | 28 | 1.7 | 2491.2 | 1.63 | 1.2632
-105.6 | 0 | 28 | 1.67 | 2440.9 | 1.56 | 1.2089
-104.4 | 0 | 28 | 1.66 | 2438 | 1.64 | 1.2709
-103.2 | 0 | 28 | 1.75 | 2563.9 | 1.68 | 1.3019
-102 | 0 | 28 | 1.75 | 2561.8 | 1.63 | 1.2632
-100.8 | 0 | 28 | 1.77 | 2596.8 | 1.77 | 1.3717
-99.6 | 0 | 28 | 1.73 | 2530.9 | 1.67 | 1.2942
-98.4 | 0 | 28 | 1.7 | 2489.6 | 1.57 | 1.2167
-97.2 | 0 | 28 | 1.71 | 2504.3 | 1.71 | 1.3252
-96 | 0 | 28 | 1.72 | 2521.6 | 1.66 | 1.2864
-94.7 | 0 | 28 | 1.7 | 2490.5 | 1.58 | 1.2244
-93.6 | 0 | 28 | 1.71 | 2503.3 | 1.71 | 1.3252
-92.3 | 0 | 28 | 1.67 | 2442.8 | 1.62 | 1.2554
-91.1 | 0 | 28 | 1.61 | 2364.5 | 1.5 | 1.1624
-89.9 | 0 | 28 | 1.7 | 2489.6 | 1.71 | 1.3252
-88.7 | 0 | 28 | 1.7 | 2491.7 | 1.66 | 1.2864
-87.5 | 0 | 28 | 1.63 | 2381.3 | 1.52 | 1.1779
-86.3 | 0 | 28 | 1.66 | 2425.2 | 1.61 | 1.2477
-85.1 | 0 | 28 | 1.64 | 2403.3 | 1.68 | 1.3019
-83.9 | 0 | 28 | 1.6 | 2338.8 | 1.5 | 1.1624
-82.7 | 0 | 28 | 1.64 | 2399.3 | 1.6 | 1.2399
-81.5 | 0 | 28 | 1.62 | 2369.4 | 1.59 | 1.2322
-80.3 | 0 | 28 | 1.59 | 2327 | 1.57 | 1.2167
-79.1 | 0 | 28 | 1.52 | 2231.6 | 1.49 | 1.1547
-77.9 | 0 | 28 | 1.54 | 2255.3 | 1.52 | 1.1779
-76.7 | 0 | 28 | 1.51 | 2213.5 | 1.47 | 1.1392
-75.5 | 0 | 28 | 1.45 | 2129.6 | 1.44 | 1.1159
-74.3 | 0 | 28 | 1.41 | 2066.9 | 1.38 | 1.0694
-73.1 | 0 | 28 | 1.3 | 1898 | 1.27 | 0.98419
-71.9 | 0 | 28 | 1.32 | 1929.3 | 1.31 | 1.0152
-70.7 | 0 | 28 | 1.19 | 1749.5 | 1.16 | 0.89895
-69.5 | 0 | 28 | 1.17 | 1718.2 | 1.15 | 0.8912
-68.3 | 0 | 28 | 1.08 | 1575.2 | 1.06 | 0.82145
-67.1 | 0 | 28 | 0.91 | 1340.1 | 0.88 | 0.68196
-65.9 | 0 | 28 | 0.9 | 1317.4 | 0.89 | 0.68971
-64.7 | 0 | 28 | 0.74 | 1083.9 | 0.72 | 0.55797
-63.5 | 0 | 28 | 0.69 | 1014.2 | 0.69 | 0.53472
-62.3 | 0 | 28 | 0.44 | 650.7 | 0.42 | 0.32548
-61.1 | 0 | 28 | 0.4 | 579 | 0.39 | 0.30223
-59.9 | 0 | 28 | 0.17 | 247.9 | 0.16 | 0.12399
-58.6 | 0 | 28 | 0.06 | 89.2 | 0.06 | 0.046497
-57.4 | 0 | 28 | 0.27 | 397.8 | 0.27 | 0.20924
-56.2 | 0 | 28 | 0.63 | 925.9 | 0.61 | 0.47272
-55 | 0 | 28 | 1.09 | 1594 | 1.03 | 0.7982
-53.8 | 0 | 28 | 1.46 | 2141.5 | 1.42 | 1.1004
-52.6 | 0 | 28 | 1.99 | 2908.9 | 1.91 | 1.4802
-51.4 | 0 | 28 | 2.67 | 3907.3 | 2.54 | 1.9684
-50.2 | 0 | 28 | 3.33 | 4879.5 | 3.23 | 2.5031
-49 | 0 | 28 | 4.24 | 6204.7 | 4.06 | 3.1463
-47.8 | 0 | 28 | 5.24 | 7680.4 | 4.99 | 3.867
-46.6 | 0 | 28 | 6.26 | 9163.8 | 6.15 | 4.766
-45.4 | 0 | 28 | 7.86 | 11506.9 | 7.6 | 5.8896
-44.2 | 0 | 28 | 9.58 | 14029.5 | 9.38 | 7.2691
-43 | 0 | 28 | 11.72 | 17168.6 | 11.51 | 8.9197
-41.8 | 0 | 28 | 14.69 | 21519.5 | 14.28 | 11.0663
-40.6 | 0 | 28 | 18.16 | 26600 | 17.58 | 13.6237
-39.4 | 0 | 28 | 22.25 | 32598.1 | 21.65 | 16.7777
-38.2 | 0 | 28 | 27.52 | 40308.3 | 26.52 | 20.5518
-37 | 0 | 28 | 34.04 | 49863.6 | 32.35 | 25.0697
-35.8 | 0 | 28 | 40.8 | 59771.4 | 39.13 | 30.3239
-34.6 | 0 | 28 | 49.2 | 72070.3 | 47 | 36.4228
-33.4 | 0 | 28 | 58.19 | 85244 | 55.36 | 42.9014
-32.2 | 0 | 28 | 66.58 | 97532.7 | 64.32 | 49.845
-31 | 0 | 28 | 76.3 | 111764.5 | 73.26 | 56.7731
-29.8 | 0 | 28 | 85.53 | 125282.2 | 81.88 | 63.4532
-28.6 | 0 | 28 | 92.47 | 135458 | 89.91 | 69.6761
-27.4 | 0 | 28 | 100.51 | 147222.8 | 96.91 | 75.1007
-26.2 | 0 | 28 | 107.22 | 157062.5 | 103.29 | 80.0449
-25 | 0 | 28 | 110.98 | 162565.6 | 108.32 | 83.943
-23.8 | 0 | 28 | 116.72 | 170974.6 | 112.48 | 87.1668
-22.6 | 0 | 28 | 120.09 | 175916.5 | 115.82 | 89.7551
-21.4 | 0 | 28 | 121.54 | 178027.9 | 118.45 | 91.7932
-20.2 | 0 | 28 | 122.99 | 180162.4 | 120.46 | 93.3509
-19 | 0 | 28 | 125.33 | 183580.4 | 122.11 | 94.6296
-17.7 | 0 | 28 | 127.01 | 186051.8 | 123.41 | 95.637
-16.5 | 0 | 28 | 127.77 | 187161.7 | 124.46 | 96.4507
-15.4 | 0 | 28 | 129.26 | 189348.8 | 125.55 | 97.2954
-14.1 | 0 | 28 | 129.29 | 189381.7 | 125.74 | 97.4427
-12.9 | 0 | 28 | 131.87 | 193162.6 | 126.5 | 98.0316
-11.7 | 0 | 28 | 132.04 | 193409.4 | 126.95 | 98.3803
-10.5 | 0 | 28 | 133.1 | 194974.1 | 127.28 | 98.6361
-9.3 | 0 | 28 | 133.25 | 195186.2 | 127.88 | 99.1011
-8.1 | 0 | 28 | 133.91 | 196155 | 128.08 | 99.256
-6.9 | 0 | 28 | 132.99 | 194812.1 | 128.24 | 99.38
-5.7 | 0 | 28 | 133.99 | 196265.2 | 128.63 | 99.6823
-4.5 | 0 | 28 | 134.48 | 196994.4 | 128.86 | 99.8605
-3.3 | 0 | 28 | 133.68 | 195819.5 | 128.82 | 99.8295
-2.1 | 0 | 28 | 132.64 | 194292.5 | 129.02 | 99.9845
-0.9 | 0 | 28 | 134.3 | 196721.9 | 129.01 | 99.9768
0.3 | 0 | 28 | 132.49 | 194073.7 | 128.99 | 99.9613
1.5 | 0 | 28 | 134.09 | 196420.4 | 129.04 | 100
2.7 | 0 | 28 | 131.7 | 192910 | 128.93 | 99.9148
3.9 | 0 | 28 | 133.27 | 195211.8 | 128.69 | 99.7288
5.1 | 0 | 28 | 137.14 | 200886.5 | 128.92 | 99.907
6.3 | 0 | 28 | 135.58 | 198606.6 | 128.57 | 99.6358
7.5 | 0 | 28 | 131.93 | 193250.5 | 128.22 | 99.3645
8.7 | 0 | 28 | 134.04 | 196339.9 | 127.76 | 99.0081
9.9 | 0 | 28 | 128.72 | 188545.9 | 127.42 | 98.7446
11.1 | 0 | 28 | 129.89 | 190262.6 | 126.68 | 98.1711
12.3 | 0 | 28 | 126.48 | 185266.5 | 126.16 | 97.7681
13.5 | 0 | 28 | 130.83 | 191648.8 | 125.53 | 97.2799
14.7 | 0 | 28 | 129.49 | 189675.5 | 124.72 | 96.6522
15.9 | 0 | 28 | 128.51 | 188237.1 | 123.43 | 95.6525
17.1 | 0 | 28 | 127.74 | 187121.5 | 122.02 | 94.5598
18.3 | 0 | 28 | 123.53 | 180954.5 | 120.43 | 93.3277
19.5 | 0 | 28 | 123.16 | 180412.7 | 118.48 | 91.8165
20.7 | 0 | 28 | 119.27 | 174711.3 | 115.93 | 89.8404
21.9 | 0 | 28 | 117.05 | 171461.4 | 112.46 | 87.1513
23.1 | 0 | 28 | 113.27 | 165920.5 | 108.84 | 84.3459
24.3 | 0 | 28 | 108.72 | 159260.7 | 103.81 | 80.4479
25.5 | 0 | 28 | 102.06 | 149502.7 | 98.17 | 76.0772
26.7 | 0 | 28 | 94.99 | 139141.5 | 91.28 | 70.7378
27.9 | 0 | 28 | 87.07 | 127543.6 | 83.53 | 64.7319
29.1 | 0 | 28 | 78.02 | 114286.6 | 74.89 | 58.0363
30.3 | 0 | 28 | 67.86 | 99397.4 | 65.85 | 51.0307
31.5 | 0 | 28 | 58.62 | 85865.2 | 56.81 | 44.0251
32.7 | 0 | 28 | 50.43 | 73871 | 47.98 | 37.1823
33.9 | 0 | 28 | 40.62 | 59497.2 | 39.61 | 30.6959
35.1 | 0 | 28 | 32.48 | 47572 | 32.63 | 25.2867
36.3 | 0 | 28 | 27.2 | 39845.2 | 26.42 | 20.4743
37.5 | 0 | 28 | 21.96 | 32171.1 | 21.3 | 16.5065
38.7 | 0 | 28 | 17.68 | 25904.9 | 17.18 | 13.3137
39.9 | 0 | 28 | 14.25 | 20878.7 | 13.84 | 10.7254
41.1 | 0 | 28 | 11.48 | 16812.3 | 11.17 | 8.6562
42.4 | 0 | 28 | 9.26 | 13561.9 | 8.97 | 6.9513
43.5 | 0 | 28 | 7.43 | 10888.6 | 7.16 | 5.5487
44.8 | 0 | 28 | 6.15 | 9005 | 5.93 | 4.5955
46 | 0 | 28 | 4.87 | 7132.1 | 4.69 | 3.6345
47.2 | 0 | 28 | 3.89 | 5698.8 | 3.73 | 2.8906
48.4 | 0 | 28 | 3.13 | 4585.4 | 3 | 2.3249
49.6 | 0 | 28 | 2.52 | 3694.8 | 2.42 | 1.8754
50.8 | 0 | 28 | 1.85 | 2713.5 | 1.78 | 1.3794
52 | 0 | 28 | 1.46 | 2139 | 1.4 | 1.0849
53.2 | 0 | 28 | 1.07 | 1566 | 1.03 | 0.7982
54.4 | 0 | 28 | 0.74 | 1085.6 | 0.71 | 0.55022
55.6 | 0 | 28 | 0.43 | 630.7 | 0.41 | 0.31773
56.8 | 0 | 28 | 0.15 | 222.8 | 0.15 | 0.11624
58 | 0 | 28 | 0.05 | 74.5 | 0.05 | 0.038748
59.2 | 0 | 28 | 0.28 | 411.3 | 0.27 | 0.20924
60.4 | 0 | 28 | 0.49 | 719.4 | 0.47 | 0.36423
61.6 | 0 | 28 | 0.63 | 922 | 0.61 | 0.47272
62.8 | 0 | 28 | 0.68 | 1001.4 | 0.66 | 0.51147
64 | 0 | 28 | 0.84 | 1233.8 | 0.81 | 0.62771
65.2 | 0 | 28 | 1.04 | 1518.4 | 1 | 0.77495
66.4 | 0 | 28 | 1.07 | 1567.6 | 1.04 | 0.80595
67.6 | 0 | 28 | 1.24 | 1810.6 | 1.2 | 0.92994
68.8 | 0 | 28 | 1.26 | 1844.4 | 1.2 | 0.92994
70 | 0 | 28 | 1.3 | 1902.1 | 1.26 | 0.97644
71.2 | 0 | 28 | 1.41 | 2068.2 | 1.41 | 1.0927
72.5 | 0 | 28 | 1.42 | 2082 | 1.35 | 1.0462
73.7 | 0 | 28 | 1.54 | 2258.6 | 1.52 | 1.1779
74.8 | 0 | 28 | 1.62 | 2375.6 | 1.55 | 1.2012
76.1 | 0 | 28 | 1.6 | 2347.4 | 1.56 | 1.2089
77.2 | 0 | 28 | 1.65 | 2412 | 1.59 | 1.2322
78.5 | 0 | 28 | 1.65 | 2421.4 | 1.59 | 1.2322
79.6 | 0 | 28 | 1.64 | 2399.3 | 1.57 | 1.2167
80.8 | 0 | 28 | 1.7 | 2487 | 1.62 | 1.2554
82.1 | 0 | 28 | 1.7 | 2488.2 | 1.62 | 1.2554
83.3 | 0 | 28 | 1.66 | 2438.5 | 1.6 | 1.2399
84.5 | 0 | 28 | 1.72 | 2515.5 | 1.65 | 1.2787
85.7 | 0 | 28 | 1.76 | 2579.3 | 1.7 | 1.3174
86.9 | 0 | 28 | 1.73 | 2537.1 | 1.66 | 1.2864
88.1 | 0 | 28 | 1.77 | 2585.8 | 1.7 | 1.3174
89.3 | 0 | 28 | 1.75 | 2559.3 | 1.69 | 1.3097
90.5 | 0 | 28 | 1.76 | 2582.8 | 1.69 | 1.3097
91.7 | 0 | 28 | 1.75 | 2563.5 | 1.69 | 1.3097
92.9 | 0 | 28 | 1.8 | 2630.2 | 1.73 | 1.3407
94.1 | 0 | 28 | 1.86 | 2717.7 | 1.81 | 1.4027
95.3 | 0 | 28 | 1.84 | 2692.8 | 1.78 | 1.3794
96.5 | 0 | 28 | 1.84 | 2699.5 | 1.8 | 1.3949
97.7 | 0 | 28 | 1.84 | 2692.7 | 1.79 | 1.3872
98.9 | 0 | 28 | 1.81 | 2654.6 | 1.77 | 1.3717
100.1 | 0 | 28 | 1.81 | 2651.8 | 1.81 | 1.4027
101.3 | 0 | 28 | 1.78 | 2601.8 | 1.67 | 1.2942
102.5 | 0 | 28 | 1.85 | 2709.7 | 1.8 | 1.3949
103.7 | 0 | 28 | 1.82 | 2666.7 | 1.8 | 1.3949
104.9 | 0 | 28 | 1.8 | 2642.4 | 1.73 | 1.3407
106.1 | 0 | 28 | 1.88 | 2754.5 | 1.86 | 1.4414
107.3 | 0 | 28 | 1.8 | 2631.7 | 1.75 | 1.3562
108.5 | 0 | 28 | 1.75 | 2569.1 | 1.69 | 1.3097
109.7 | 0 | 28 | 1.82 | 2664.1 | 1.78 | 1.3794
110.9 | 0 | 28 | 1.77 | 2589.6 | 1.7 | 1.3174
112.1 | 0 | 28 | 1.83 | 2683.1 | 1.76 | 1.3639
113.3 | 0 | 28 | 1.83 | 2682.5 | 1.76 | 1.3639
114.5 | 0 | 28 | 1.86 | 2719.3 | 1.79 | 1.3872
115.7 | 0 | 28 | 1.79 | 2629 | 1.73 | 1.3407
116.9 | 0 | 28 | 1.83 | 2683.1 | 1.78 | 1.3794
118.1 | 0 | 28 | 1.86 | 2728.6 | 1.78 | 1.3794
119.3 | 0 | 28 | 1.83 | 2677.6 | 1.75 | 1.3562
120.5 | 0 | 28 | 1.83 | 2683.8 | 1.76 | 1.3639
121.7 | 0 | 28 | 1.85 | 2705.6 | 1.76 | 1.3639
122.9 | 0 | 28 | 1.87 | 2735.5 | 1.78 | 1.3794
124.2 | 0 | 28 | 1.89 | 2762.3 | 1.83 | 1.4182
125.7 | 0 | 28 | 1.87 | 2740.4 | 1.81 | 1.4027
Measurement time: | 02/09/2022 20:44:47
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-151.4 | 0 | 28 | 1.76 | 2581.1 | 1.94 | 1.5706
-150.6 | 0 | 28 | 1.72 | 2526.5 | 1.58 | 1.2791
-149.3 | 0 | 28 | 1.71 | 2502 | 1.56 | 1.263
-148.1 | 0 | 28 | 1.73 | 2530.3 | 1.51 | 1.2225
-146.9 | 0 | 28 | 1.73 | 2530.4 | 1.61 | 1.3034
-145.7 | 0 | 28 | 1.71 | 2505.3 | 1.55 | 1.2549
-144.5 | 0 | 28 | 1.75 | 2566 | 1.51 | 1.2225
-143.3 | 0 | 28 | 1.83 | 2679.8 | 1.73 | 1.4006
-142.1 | 0 | 28 | 1.76 | 2571.7 | 1.65 | 1.3358
-140.9 | 0 | 28 | 1.7 | 2495.2 | 1.47 | 1.1901
-139.7 | 0 | 28 | 1.77 | 2588.1 | 1.66 | 1.3439
-138.5 | 0 | 28 | 1.77 | 2589.5 | 1.68 | 1.3601
-137.3 | 0 | 28 | 1.67 | 2450.5 | 1.44 | 1.1658
-136.1 | 0 | 28 | 1.72 | 2524.8 | 1.63 | 1.3196
-134.9 | 0 | 28 | 1.77 | 2591.1 | 1.67 | 1.352
-133.7 | 0 | 28 | 1.74 | 2547.9 | 1.51 | 1.2225
-132.5 | 0 | 28 | 1.74 | 2549.7 | 1.62 | 1.3115
-131.3 | 0 | 28 | 1.72 | 2519.9 | 1.56 | 1.263
-130.1 | 0 | 28 | 1.78 | 2606.6 | 1.54 | 1.2468
-128.9 | 0 | 28 | 1.77 | 2589.6 | 1.67 | 1.352
-127.7 | 0 | 28 | 1.79 | 2627.4 | 1.62 | 1.3115
-126.5 | 0 | 28 | 1.73 | 2535.8 | 1.51 | 1.2225
-125.3 | 0 | 28 | 1.74 | 2544.3 | 1.65 | 1.3358
-124.1 | 0 | 28 | 1.69 | 2482.2 | 1.55 | 1.2549
-122.9 | 0 | 28 | 1.73 | 2533.2 | 1.52 | 1.2306
-121.7 | 0 | 28 | 1.74 | 2551.5 | 1.66 | 1.3439
-120.5 | 0 | 28 | 1.73 | 2538.4 | 1.53 | 1.2387
-119.3 | 0 | 28 | 1.73 | 2540.7 | 1.61 | 1.3034
-118.1 | 0 | 28 | 1.71 | 2504.9 | 1.55 | 1.2549
-116.9 | 0 | 28 | 1.71 | 2509.6 | 1.63 | 1.3196
-115.7 | 0 | 28 | 1.69 | 2473.8 | 1.49 | 1.2063
-114.5 | 0 | 28 | 1.8 | 2631.3 | 1.67 | 1.352
-113.3 | 0 | 28 | 1.68 | 2465.4 | 1.58 | 1.2791
-112.1 | 0 | 28 | 1.67 | 2441 | 1.48 | 1.1982
-110.9 | 0 | 28 | 1.68 | 2456.4 | 1.54 | 1.2468
-109.6 | 0 | 28 | 1.71 | 2504.8 | 1.6 | 1.2953
-108.4 | 0 | 28 | 1.62 | 2373.4 | 1.45 | 1.1739
-107.2 | 0 | 28 | 1.6 | 2349.4 | 1.48 | 1.1982
-106 | 0 | 28 | 1.63 | 2381.6 | 1.52 | 1.2306
-104.9 | 0 | 28 | 1.58 | 2307.3 | 1.4 | 1.1334
-103.6 | 0 | 28 | 1.57 | 2299.4 | 1.45 | 1.1739
-102.4 | 0 | 28 | 1.6 | 2341.7 | 1.44 | 1.1658
-101.2 | 0 | 28 | 1.54 | 2249.7 | 1.36 | 1.101
-100 | 0 | 28 | 1.59 | 2329.3 | 1.46 | 1.182
-98.8 | 0 | 28 | 1.48 | 2166 | 1.33 | 1.0767
-97.6 | 0 | 28 | 1.42 | 2082.6 | 1.25 | 1.012
-96.4 | 0 | 28 | 1.42 | 2075.7 | 1.31 | 1.0606
-95.2 | 0 | 28 | 1.42 | 2079.8 | 1.29 | 1.0444
-94 | 0 | 28 | 1.42 | 2079.3 | 1.32 | 1.0687
-92.8 | 0 | 28 | 1.3 | 1906.7 | 1.19 | 0.96341
-91.6 | 0 | 28 | 1.22 | 1793.7 | 1.1 | 0.89054
-90.4 | 0 | 28 | 1.13 | 1651.3 | 1 | 0.80959
-89.2 | 0 | 28 | 1.07 | 1565.8 | 0.98 | 0.79339
-88 | 0 | 28 | 0.9 | 1319.6 | 0.81 | 0.65576
-86.8 | 0 | 28 | 0.8 | 1164.9 | 0.71 | 0.57481
-85.6 | 0 | 28 | 0.71 | 1046.1 | 0.65 | 0.52623
-84.4 | 0 | 28 | 0.67 | 978.5 | 0.63 | 0.51004
-83.2 | 0 | 28 | 0.44 | 650.2 | 0.4 | 0.32383
-82 | 0 | 28 | 0.24 | 345.7 | 0.21 | 0.17001
-80.8 | 0 | 28 | 0.01 | 15.2 | 0.01 | 0.0080959
-79.6 | 0 | 28 | 0.21 | 307.1 | 0.19 | 0.15382
-78.4 | 0 | 28 | 0.54 | 787.5 | 0.49 | 0.3967
-77.2 | 0 | 28 | 0.86 | 1255 | 0.77 | 0.62338
-76 | 0 | 28 | 1.27 | 1860.9 | 1.15 | 0.93102
-74.8 | 0 | 28 | 1.74 | 2554.7 | 1.59 | 1.2872
-73.6 | 0 | 28 | 2.33 | 3418.8 | 2.15 | 1.7406
-72.4 | 0 | 28 | 3.01 | 4415.3 | 2.74 | 2.2183
-71.2 | 0 | 28 | 3.69 | 5404.3 | 3.39 | 2.7445
-70 | 0 | 28 | 4.73 | 6928.3 | 4.34 | 3.5136
-68.8 | 0 | 28 | 5.76 | 8443.7 | 5.29 | 4.2827
-67.6 | 0 | 28 | 7.23 | 10583.4 | 6.61 | 5.3514
-66.4 | 0 | 28 | 8.86 | 12980.4 | 8.12 | 6.5738
-65.2 | 0 | 28 | 10.97 | 16075.8 | 9.97 | 8.0716
-63.9 | 0 | 28 | 13.35 | 19553 | 12.15 | 9.8365
-62.8 | 0 | 28 | 16.1 | 23577.8 | 14.68 | 11.8847
-61.5 | 0 | 28 | 20.37 | 29837.9 | 18.21 | 14.7426
-60.3 | 0 | 28 | 24.25 | 35515.2 | 22.12 | 17.908
-59.1 | 0 | 28 | 29.9 | 43803.2 | 27 | 21.8588
-57.9 | 0 | 28 | 36.28 | 53146.8 | 32.75 | 26.5139
-56.7 | 0 | 28 | 43.08 | 63106.3 | 39.29 | 31.8086
-55.5 | 0 | 28 | 52.41 | 76767.7 | 46.75 | 37.8481
-54.3 | 0 | 28 | 58 | 84962.2 | 54.44 | 44.0738
-53.1 | 0 | 28 | 69.25 | 101444.7 | 62.91 | 50.931
-51.9 | 0 | 28 | 79.02 | 115743.6 | 71.06 | 57.5291
-50.7 | 0 | 28 | 86.8 | 127141.9 | 78.95 | 63.9168
-49.5 | 0 | 28 | 95.39 | 139734.3 | 86.22 | 69.8025
-48.3 | 0 | 28 | 100.92 | 147834.9 | 92.47 | 74.8624
-47.1 | 0 | 28 | 107.55 | 157547.6 | 98.13 | 79.4446
-45.9 | 0 | 28 | 112.24 | 164407.3 | 102.62 | 83.0797
-44.7 | 0 | 28 | 115.96 | 169859.7 | 106.14 | 85.9294
-43.5 | 0 | 28 | 119.52 | 175068.4 | 109.36 | 88.5363
-42.3 | 0 | 28 | 121.97 | 178662.5 | 111.54 | 90.3012
-41.1 | 0 | 28 | 124.96 | 183044 | 113.63 | 91.9932
-39.9 | 0 | 28 | 128.95 | 188887.2 | 115.3 | 93.3452
-38.7 | 0 | 28 | 126.97 | 185986.1 | 116.45 | 94.2762
-37.5 | 0 | 28 | 131.39 | 192459.4 | 117.6 | 95.2073
-36.3 | 0 | 28 | 131.67 | 192877.3 | 118.29 | 95.7659
-35.1 | 0 | 28 | 130.81 | 191611.9 | 119.08 | 96.4054
-33.9 | 0 | 28 | 126.7 | 185596.5 | 119.18 | 96.4864
-32.7 | 0 | 28 | 132.69 | 194362.4 | 120.21 | 97.3203
-31.5 | 0 | 28 | 135.43 | 198380.1 | 120.61 | 97.6441
-30.3 | 0 | 28 | 131.52 | 192648.6 | 120.76 | 97.7655
-29.1 | 0 | 28 | 133.63 | 195742.4 | 121.27 | 98.1784
-27.9 | 0 | 28 | 133.33 | 195301.7 | 121.35 | 98.2432
-26.7 | 0 | 28 | 132.89 | 194654.7 | 121.62 | 98.4618
-25.5 | 0 | 28 | 136.53 | 199995.8 | 122.07 | 98.8261
-24.3 | 0 | 28 | 133.36 | 195342.4 | 122.05 | 98.8099
-23.1 | 0 | 28 | 130.51 | 191170.9 | 122.13 | 98.8747
-21.9 | 0 | 28 | 134.22 | 196604.4 | 122.41 | 99.1014
-20.7 | 0 | 28 | 134.82 | 197493.1 | 122.8 | 99.4171
-19.5 | 0 | 28 | 134.3 | 196732 | 122.79 | 99.409
-18.3 | 0 | 28 | 135.67 | 198726.8 | 122.92 | 99.5142
-17 | 0 | 28 | 135.52 | 198509.6 | 122.81 | 99.4252
-15.8 | 0 | 28 | 136.51 | 199966.6 | 122.86 | 99.4657
-14.6 | 0 | 28 | 129.76 | 190072 | 123.01 | 99.5871
-13.4 | 0 | 28 | 137.47 | 201362.9 | 123.06 | 99.6276
-12.2 | 0 | 28 | 136.58 | 200059.9 | 122.97 | 99.5547
-11 | 0 | 28 | 130.46 | 191098 | 123.08 | 99.6438
-9.8 | 0 | 28 | 137.42 | 201295 | 123.23 | 99.7652
-8.6 | 0 | 28 | 137.13 | 200874.9 | 123.19 | 99.7328
-7.4 | 0 | 28 | 136.26 | 199592 | 123.32 | 99.8381
-6.2 | 0 | 28 | 137.02 | 200711.1 | 123.44 | 99.9352
-5 | 0 | 28 | 137.34 | 201171.9 | 123.13 | 99.6843
-3.8 | 0 | 28 | 130.05 | 190500.8 | 123.17 | 99.7166
-2.6 | 0 | 28 | 137.63 | 201597.9 | 123.46 | 99.9514
-1.4 | 0 | 28 | 136.49 | 199935.9 | 123.21 | 99.749
-0.2 | 0 | 28 | 136.8 | 200386.7 | 123.36 | 99.8705
1 | 0 | 28 | 136.64 | 200154.9 | 123.25 | 99.7814
2.2 | 0 | 28 | 135.82 | 198954.9 | 123.36 | 99.8705
3.4 | 0 | 28 | 135.96 | 199159.9 | 123.52 | 100
4.6 | 0 | 28 | 135.68 | 198749.5 | 123.37 | 99.8786
5.8 | 0 | 28 | 135.01 | 197769.5 | 123.37 | 99.8786
7 | 0 | 28 | 135.8 | 198923.4 | 123.37 | 99.8786
8.2 | 0 | 28 | 134.75 | 197379.3 | 123.23 | 99.7652
9.4 | 0 | 28 | 134.72 | 197337.5 | 123.17 | 99.7166
10.6 | 0 | 28 | 134.49 | 196997.1 | 123.27 | 99.7976
11.8 | 0 | 28 | 139.41 | 204215.9 | 123.14 | 99.6924
13 | 0 | 28 | 133.72 | 195882.7 | 122.8 | 99.4171
14.2 | 0 | 28 | 127.98 | 187471.3 | 122.66 | 99.3038
15.4 | 0 | 28 | 139.75 | 204704.5 | 122.98 | 99.5628
16.6 | 0 | 28 | 133.45 | 195479.2 | 122.59 | 99.2471
17.8 | 0 | 28 | 127.96 | 187434.1 | 122.56 | 99.2228
19 | 0 | 28 | 139.58 | 204460 | 122.81 | 99.4252
20.2 | 0 | 28 | 132.73 | 194429.6 | 122.25 | 98.9718
21.4 | 0 | 28 | 131.15 | 192112.8 | 122.31 | 99.0204
22.6 | 0 | 28 | 133.38 | 195381 | 122.13 | 98.8747
23.8 | 0 | 28 | 132.62 | 194262.1 | 121.79 | 98.5994
25 | 0 | 28 | 128.1 | 187643.9 | 121.56 | 98.4132
26.2 | 0 | 28 | 136.81 | 200405.7 | 121.43 | 98.308
27.4 | 0 | 28 | 127.93 | 187396.9 | 120.77 | 97.7736
28.6 | 0 | 28 | 131.63 | 192811.3 | 120.81 | 97.806
29.8 | 0 | 28 | 135.84 | 198979.3 | 120.51 | 97.5631
31 | 0 | 28 | 130.28 | 190834.3 | 119.79 | 96.9802
32.2 | 0 | 28 | 125.7 | 184131.4 | 119.02 | 96.3569
33.4 | 0 | 28 | 133.01 | 194831.8 | 118.55 | 95.9764
34.6 | 0 | 28 | 129.35 | 189476.7 | 117.55 | 95.1668
35.8 | 0 | 28 | 123.04 | 180229.1 | 116.33 | 94.1791
37 | 0 | 28 | 128.87 | 188770.2 | 115.21 | 93.2723
38.2 | 0 | 28 | 120.54 | 176573.1 | 113.27 | 91.7017
39.4 | 0 | 28 | 120.95 | 177175.5 | 111.89 | 90.5845
40.6 | 0 | 28 | 123.67 | 181155.8 | 109.43 | 88.5929
41.8 | 0 | 28 | 112.19 | 164335.3 | 106.03 | 85.8403
43 | 0 | 28 | 112.45 | 164723.3 | 102.69 | 83.1363
44.2 | 0 | 28 | 107.69 | 157750.8 | 97.82 | 79.1937
45.4 | 0 | 28 | 102.23 | 149752.8 | 92.56 | 74.9352
46.7 | 0 | 28 | 94.94 | 139073.2 | 86.45 | 69.9887
47.9 | 0 | 28 | 87.18 | 127706.4 | 79.47 | 64.3378
49 | 0 | 28 | 79.03 | 115758.9 | 71.64 | 57.9987
50.3 | 0 | 28 | 70.06 | 102621.9 | 63.5 | 51.4087
51.5 | 0 | 28 | 61.22 | 89682 | 55.27 | 44.7458
52.7 | 0 | 28 | 51.94 | 76086.8 | 47 | 38.0505
53.9 | 0 | 28 | 44.1 | 64595 | 39.68 | 32.1244
55.1 | 0 | 28 | 36.25 | 53099.6 | 32.75 | 26.5139
56.3 | 0 | 28 | 29.9 | 43795.1 | 26.91 | 21.7859
57.5 | 0 | 28 | 24.34 | 35657.7 | 21.93 | 17.7542
58.7 | 0 | 28 | 19.78 | 28970.7 | 17.84 | 14.443
59.9 | 0 | 28 | 16.01 | 23455.8 | 14.45 | 11.6985
61.1 | 0 | 28 | 12.25 | 17948 | 11.57 | 9.3669
62.3 | 0 | 28 | 10.65 | 15602.3 | 9.59 | 7.7639
63.5 | 0 | 28 | 8.69 | 12735.1 | 7.79 | 6.3067
64.7 | 0 | 28 | 6.54 | 9573.4 | 6.22 | 5.0356
65.9 | 0 | 28 | 5.73 | 8387 | 5.15 | 4.1694
67.1 | 0 | 28 | 4.58 | 6711.7 | 4.14 | 3.3517
68.3 | 0 | 28 | 3.7 | 5420.7 | 3.35 | 2.7121
69.5 | 0 | 28 | 3 | 4394.6 | 2.71 | 2.194
70.7 | 0 | 28 | 2.32 | 3398 | 2.11 | 1.7082
71.9 | 0 | 28 | 1.85 | 2708.1 | 1.68 | 1.3601
73.1 | 0 | 28 | 1.34 | 1956.5 | 1.22 | 0.98769
74.3 | 0 | 28 | 0.95 | 1385.7 | 0.86 | 0.69624
75.5 | 0 | 28 | 0.71 | 1040.4 | 0.65 | 0.52623
76.7 | 0 | 28 | 0.36 | 521.2 | 0.33 | 0.26716
77.9 | 0 | 28 | 0.02 | 33.1 | 0.02 | 0.016192
79.1 | 0 | 28 | 0.11 | 158.2 | 0.1 | 0.080959
80.3 | 0 | 28 | 0.32 | 475.5 | 0.29 | 0.23478
81.5 | 0 | 28 | 0.53 | 778 | 0.49 | 0.3967
82.7 | 0 | 28 | 0.66 | 968.7 | 0.6 | 0.48575
84 | 0 | 28 | 0.77 | 1121.3 | 0.69 | 0.55861
85.2 | 0 | 28 | 1.01 | 1483.7 | 0.93 | 0.75291
86.3 | 0 | 28 | 1.08 | 1578.4 | 0.98 | 0.79339
87.6 | 0 | 28 | 1.12 | 1642.9 | 1 | 0.80959
88.8 | 0 | 28 | 1.25 | 1825.8 | 1.15 | 0.93102
90 | 0 | 28 | 1.36 | 1993.8 | 1.24 | 1.0039
91.2 | 0 | 28 | 1.4 | 2047.8 | 1.24 | 1.0039
92.4 | 0 | 28 | 1.5 | 2197.4 | 1.39 | 1.1253
93.6 | 0 | 28 | 1.51 | 2215.4 | 1.35 | 1.0929
94.8 | 0 | 28 | 1.59 | 2323.3 | 1.41 | 1.1415
96 | 0 | 28 | 1.58 | 2309 | 1.45 | 1.1739
97.2 | 0 | 28 | 1.67 | 2445.3 | 1.49 | 1.2063
98.4 | 0 | 28 | 1.68 | 2454.9 | 1.49 | 1.2063
99.6 | 0 | 28 | 1.7 | 2486.5 | 1.55 | 1.2549
100.8 | 0 | 28 | 1.72 | 2526.8 | 1.64 | 1.3277
102 | 0 | 28 | 1.72 | 2516.9 | 1.51 | 1.2225
103.2 | 0 | 28 | 1.74 | 2552.4 | 1.61 | 1.3034
104.4 | 0 | 28 | 1.78 | 2601.1 | 1.68 | 1.3601
105.6 | 0 | 28 | 1.75 | 2556.5 | 1.53 | 1.2387
106.8 | 0 | 28 | 1.8 | 2636 | 1.68 | 1.3601
108 | 0 | 28 | 1.79 | 2628.3 | 1.63 | 1.3196
109.2 | 0 | 28 | 1.77 | 2592.7 | 1.55 | 1.2549
110.4 | 0 | 28 | 1.81 | 2647.1 | 1.71 | 1.3844
111.6 | 0 | 28 | 1.84 | 2689.2 | 1.67 | 1.352
112.8 | 0 | 28 | 1.84 | 2696.7 | 1.63 | 1.3196
114 | 0 | 28 | 1.82 | 2659.2 | 1.73 | 1.4006
115.2 | 0 | 28 | 1.86 | 2719.4 | 1.7 | 1.3763
116.4 | 0 | 28 | 1.79 | 2623.8 | 1.57 | 1.271
117.6 | 0 | 28 | 1.9 | 2778.8 | 1.8 | 1.4573
118.8 | 0 | 28 | 1.88 | 2750.3 | 1.71 | 1.3844
120 | 0 | 28 | 1.83 | 2683.6 | 1.6 | 1.2953
121.3 | 0 | 28 | 1.85 | 2713.4 | 1.77 | 1.433
122.4 | 0 | 28 | 1.85 | 2703.6 | 1.68 | 1.3601
123.6 | 0 | 28 | 1.84 | 2693.1 | 1.61 | 1.3034
124.9 | 0 | 28 | 1.97 | 2887.6 | 1.89 | 1.5301
126 | 0 | 28 | 1.89 | 2773.6 | 1.71 | 1.3844
127.2 | 0 | 28 | 1.89 | 2762.5 | 1.64 | 1.3277
128.4 | 0 | 28 | 1.9 | 2779.7 | 1.78 | 1.4411
129.7 | 0 | 28 | 1.91 | 2802.2 | 1.72 | 1.3925
130.8 | 0 | 28 | 1.93 | 2833.1 | 1.76 | 1.4249
132.1 | 0 | 28 | 1.92 | 2808.4 | 1.8 | 1.4573
133.3 | 0 | 28 | 1.96 | 2877.7 | 1.77 | 1.433
134.5 | 0 | 28 | 1.96 | 2864.8 | 1.7 | 1.3763
135.7 | 0 | 28 | 1.96 | 2865 | 1.84 | 1.4896
136.9 | 0 | 28 | 1.94 | 2849 | 1.75 | 1.4168
138.1 | 0 | 28 | 1.86 | 2721.4 | 1.62 | 1.3115
139.3 | 0 | 28 | 1.96 | 2877.7 | 1.84 | 1.4896
140.5 | 0 | 28 | 1.96 | 2867.1 | 1.85 | 1.4977
141.7 | 0 | 28 | 1.91 | 2802.6 | 1.65 | 1.3358
142.9 | 0 | 28 | 2.01 | 2942.1 | 1.91 | 1.5463
144.1 | 0 | 28 | 1.91 | 2799.4 | 1.74 | 1.4087
145.3 | 0 | 28 | 1.92 | 2817.3 | 1.68 | 1.3601
146.5 | 0 | 28 | 2.01 | 2937.9 | 1.92 | 1.5544
147.7 | 0 | 28 | 1.95 | 2861.5 | 1.78 | 1.4411
148.9 | 0 | 28 | 2.02 | 2958.4 | 1.78 | 1.4411
150.1 | 0 | 28 | 1.97 | 2880.6 | 1.81 | 1.4653
151.4 | 0 | 28 | 1.99 | 2922.2 | 1.84 | 1.4896
Measurement time: | 02/09/2022 20:15:50
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 250 x 250 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
228.5 | 0 | 28 | 1.78 | 2610.7 | 2.5 | 1.8612
227.8 | 0 | 28 | 1.58 | 2318.6 | 1.63 | 1.2135
226.3 | 0 | 28 | 1.63 | 2382.4 | 1.66 | 1.2359
225.2 | 0 | 28 | 1.57 | 2297.9 | 1.66 | 1.2359
224 | 0 | 28 | 1.6 | 2344.7 | 1.68 | 1.2507
222.8 | 0 | 28 | 1.55 | 2273.4 | 1.6 | 1.1912
221.6 | 0 | 28 | 1.53 | 2240.1 | 1.61 | 1.1986
220.4 | 0 | 28 | 1.53 | 2244.6 | 1.6 | 1.1912
219.2 | 0 | 28 | 1.56 | 2287.6 | 1.63 | 1.2135
218 | 0 | 28 | 1.57 | 2302.5 | 1.65 | 1.2284
216.8 | 0 | 28 | 1.56 | 2281 | 1.65 | 1.2284
215.6 | 0 | 28 | 1.55 | 2277 | 1.64 | 1.221
214.4 | 0 | 28 | 1.52 | 2222.9 | 1.6 | 1.1912
213.2 | 0 | 28 | 1.6 | 2341.5 | 1.69 | 1.2582
212 | 0 | 28 | 1.56 | 2290.4 | 1.66 | 1.2359
210.8 | 0 | 28 | 1.49 | 2187.5 | 1.56 | 1.1614
209.6 | 0 | 28 | 1.58 | 2312 | 1.67 | 1.2433
208.4 | 0 | 28 | 1.49 | 2186.5 | 1.56 | 1.1614
207.2 | 0 | 28 | 1.49 | 2187.9 | 1.56 | 1.1614
206 | 0 | 28 | 1.54 | 2260.2 | 1.65 | 1.2284
204.8 | 0 | 28 | 1.52 | 2225.2 | 1.56 | 1.1614
203.6 | 0 | 28 | 1.59 | 2330.7 | 1.69 | 1.2582
202.3 | 0 | 28 | 1.49 | 2189.4 | 1.53 | 1.1391
201.2 | 0 | 28 | 1.46 | 2145.5 | 1.52 | 1.1316
200 | 0 | 28 | 1.52 | 2219.5 | 1.61 | 1.1986
198.7 | 0 | 28 | 1.43 | 2089.3 | 1.44 | 1.0721
197.6 | 0 | 28 | 1.47 | 2156.1 | 1.56 | 1.1614
196.4 | 0 | 28 | 1.42 | 2079.5 | 1.47 | 1.0944
195.2 | 0 | 28 | 1.37 | 2007.9 | 1.43 | 1.0646
194 | 0 | 28 | 1.47 | 2158.9 | 1.58 | 1.1763
192.7 | 0 | 28 | 1.41 | 2069.4 | 1.47 | 1.0944
191.6 | 0 | 28 | 1.52 | 2222.8 | 1.59 | 1.1837
190.3 | 0 | 28 | 1.41 | 2070.5 | 1.54 | 1.1465
189.1 | 0 | 28 | 1.37 | 2012.7 | 1.42 | 1.0572
187.9 | 0 | 28 | 1.42 | 2086.4 | 1.54 | 1.1465
186.7 | 0 | 28 | 1.39 | 2033.7 | 1.43 | 1.0646
185.6 | 0 | 28 | 1.42 | 2081.1 | 1.51 | 1.1242
184.3 | 0 | 28 | 1.41 | 2066.8 | 1.47 | 1.0944
183.1 | 0 | 28 | 1.37 | 2001 | 1.43 | 1.0646
181.9 | 0 | 28 | 1.36 | 1993.3 | 1.43 | 1.0646
180.7 | 0 | 28 | 1.35 | 1977.7 | 1.4 | 1.0423
179.5 | 0 | 28 | 1.35 | 1983.6 | 1.42 | 1.0572
178.3 | 0 | 28 | 1.33 | 1951.2 | 1.39 | 1.0348
177.1 | 0 | 28 | 1.31 | 1920.2 | 1.37 | 1.02
175.9 | 0 | 28 | 1.31 | 1920.6 | 1.37 | 1.02
174.7 | 0 | 28 | 1.32 | 1940.4 | 1.38 | 1.0274
173.5 | 0 | 28 | 1.24 | 1823.3 | 1.26 | 0.93806
172.3 | 0 | 28 | 1.23 | 1797.8 | 1.34 | 0.99762
171.1 | 0 | 28 | 1.11 | 1621.3 | 1.13 | 0.84127
169.9 | 0 | 28 | 1.14 | 1676.5 | 1.21 | 0.90083
168.7 | 0 | 28 | 1.09 | 1598.4 | 1.18 | 0.8785
167.5 | 0 | 28 | 1 | 1465.7 | 1.06 | 0.78916
166.3 | 0 | 28 | 0.97 | 1427.6 | 1.01 | 0.75194
165.1 | 0 | 28 | 0.87 | 1277.2 | 0.95 | 0.70727
163.9 | 0 | 28 | 0.76 | 1119.3 | 0.78 | 0.5807
162.7 | 0 | 28 | 0.72 | 1052 | 0.77 | 0.57326
161.5 | 0 | 28 | 0.53 | 769.2 | 0.57 | 0.42436
160.3 | 0 | 28 | 0.29 | 424.5 | 0.29 | 0.2159
159.1 | 0 | 28 | 0.16 | 235.5 | 0.17 | 0.12656
157.9 | 0 | 28 | 0.03 | 38.6 | 0.03 | 0.022335
156.7 | 0 | 28 | 0.22 | 319.4 | 0.23 | 0.17123
155.5 | 0 | 28 | 0.53 | 777 | 0.55 | 0.40947
154.3 | 0 | 28 | 0.82 | 1204.3 | 0.86 | 0.64026
153.1 | 0 | 28 | 1.19 | 1739.6 | 1.24 | 0.92317
151.9 | 0 | 28 | 1.67 | 2450.9 | 1.73 | 1.288
150.6 | 0 | 28 | 2.12 | 3109.5 | 2.23 | 1.6602
149.5 | 0 | 28 | 2.71 | 3971.1 | 2.82 | 2.0995
148.2 | 0 | 28 | 3.37 | 4932 | 3.55 | 2.6429
147.1 | 0 | 28 | 4.09 | 5987.7 | 4.29 | 3.1939
145.8 | 0 | 28 | 5.3 | 7763.7 | 5.37 | 3.9979
144.6 | 0 | 28 | 6.14 | 8986.7 | 6.47 | 4.8169
143.4 | 0 | 28 | 7 | 10249.6 | 7.64 | 5.6879
142.2 | 0 | 28 | 9.24 | 13540 | 9.44 | 7.028
141 | 0 | 28 | 10.8 | 15817.5 | 11.39 | 8.4797
139.8 | 0 | 28 | 12.63 | 18503.6 | 13.67 | 10.1772
138.6 | 0 | 28 | 16.76 | 24550.3 | 16.91 | 12.5893
137.4 | 0 | 28 | 18.55 | 27166.4 | 20.17 | 15.0164
136.2 | 0 | 28 | 23.46 | 34366.9 | 24.48 | 18.2251
135 | 0 | 28 | 28.65 | 41962.9 | 29.82 | 22.2007
133.8 | 0 | 28 | 34.72 | 50865.2 | 35.93 | 26.7496
132.6 | 0 | 28 | 41.49 | 60771 | 43.25 | 32.1992
131.4 | 0 | 28 | 49.5 | 72510.7 | 51.05 | 38.0063
130.2 | 0 | 28 | 57.07 | 83596.6 | 59.78 | 44.5057
129 | 0 | 28 | 66.14 | 96876.9 | 68.78 | 51.2061
127.8 | 0 | 28 | 76.14 | 111536.2 | 77.71 | 57.8544
126.6 | 0 | 28 | 82.56 | 120930.5 | 86.71 | 64.5548
125.4 | 0 | 28 | 85.38 | 125060.3 | 93.99 | 69.9747
124.2 | 0 | 28 | 100.03 | 146530.8 | 101.41 | 75.4988
123 | 0 | 28 | 97.96 | 143498.9 | 106.77 | 79.4893
121.8 | 0 | 28 | 106.46 | 155948.1 | 111.57 | 83.0628
120.6 | 0 | 28 | 114.3 | 167427.1 | 115.95 | 86.3237
119.4 | 0 | 28 | 110.42 | 161742.7 | 118.58 | 88.2817
118.2 | 0 | 28 | 115.42 | 169066.4 | 120.97 | 90.061
117 | 0 | 28 | 119.58 | 175166.4 | 123.28 | 91.7808
115.8 | 0 | 28 | 119.94 | 175691.7 | 124.48 | 92.6742
114.6 | 0 | 28 | 121.04 | 177296.6 | 125.69 | 93.575
113.4 | 0 | 28 | 122.37 | 179246.6 | 127.24 | 94.729
112.2 | 0 | 28 | 120.89 | 177080.9 | 127.75 | 95.1087
111 | 0 | 28 | 123.85 | 181412.5 | 128.64 | 95.7713
109.8 | 0 | 28 | 125.28 | 183516.7 | 129.47 | 96.3892
108.6 | 0 | 28 | 123.72 | 181231.8 | 129.71 | 96.5679
107.4 | 0 | 28 | 125.1 | 183250 | 130.19 | 96.9253
106.2 | 0 | 28 | 124.85 | 182881 | 130.54 | 97.1858
105 | 0 | 28 | 124.56 | 182454.6 | 130.81 | 97.3868
103.8 | 0 | 28 | 123.65 | 181131 | 131.09 | 97.5953
102.6 | 0 | 28 | 126.42 | 185183.5 | 131.54 | 97.9303
101.4 | 0 | 28 | 125 | 183102.7 | 131.49 | 97.8931
100.2 | 0 | 28 | 124.15 | 181863.1 | 131.85 | 98.1611
99 | 0 | 28 | 126.78 | 185713.4 | 131.93 | 98.2207
97.8 | 0 | 28 | 128.25 | 187865.4 | 132.39 | 98.5631
96.5 | 0 | 28 | 126.24 | 184922.6 | 132.5 | 98.645
95.4 | 0 | 28 | 128.08 | 187609.3 | 132.31 | 98.5036
94.1 | 0 | 28 | 125.93 | 184465.3 | 133.04 | 99.0471
93 | 0 | 28 | 126.75 | 185661.8 | 132.69 | 98.7865
91.8 | 0 | 28 | 126.97 | 185986.6 | 132.66 | 98.7641
90.5 | 0 | 28 | 124.8 | 182808.5 | 132.73 | 98.8163
89.3 | 0 | 28 | 125.41 | 183702.3 | 132.78 | 98.8535
88.1 | 0 | 28 | 125.95 | 184496 | 132.87 | 98.9205
86.9 | 0 | 28 | 126.86 | 185829.6 | 132.67 | 98.7716
85.7 | 0 | 28 | 126.37 | 185107.1 | 132.94 | 98.9726
84.5 | 0 | 28 | 126.48 | 185270.8 | 132.99 | 99.0098
83.3 | 0 | 28 | 127.69 | 187044.4 | 133.11 | 99.0992
82.1 | 0 | 28 | 126.82 | 185766.1 | 132.88 | 98.9279
80.9 | 0 | 28 | 128.66 | 188466.5 | 133.27 | 99.2183
79.7 | 0 | 28 | 127.58 | 186874.4 | 133.16 | 99.1364
78.5 | 0 | 28 | 128.19 | 187770.9 | 133.02 | 99.0322
77.3 | 0 | 28 | 128.57 | 188337.4 | 133.17 | 99.1438
76.1 | 0 | 28 | 127.56 | 186849.7 | 133.23 | 99.1885
74.9 | 0 | 28 | 128.1 | 187637.5 | 133.23 | 99.1885
73.7 | 0 | 28 | 128.47 | 188190.9 | 133.29 | 99.2332
72.5 | 0 | 28 | 127.91 | 187371.8 | 133.36 | 99.2853
71.3 | 0 | 28 | 124.67 | 182616.8 | 133.22 | 99.1811
70.1 | 0 | 28 | 130.3 | 190862.8 | 133.18 | 99.1513
68.9 | 0 | 28 | 126.96 | 185974.4 | 133.37 | 99.2927
67.7 | 0 | 28 | 125.21 | 183409.3 | 133.12 | 99.1066
66.5 | 0 | 28 | 125.86 | 184366.4 | 133.1 | 99.0917
65.3 | 0 | 28 | 127.02 | 186064.5 | 133.46 | 99.3597
64.1 | 0 | 28 | 125.19 | 183382.1 | 133.33 | 99.263
62.9 | 0 | 28 | 126.8 | 185740.3 | 133.16 | 99.1364
61.7 | 0 | 28 | 126.82 | 185764.3 | 132.96 | 98.9875
60.5 | 0 | 28 | 121.48 | 177948.8 | 132.79 | 98.8609
59.3 | 0 | 28 | 131.06 | 191984.5 | 133.31 | 99.2481
58 | 0 | 28 | 125.69 | 184119.9 | 132.77 | 98.846
56.8 | 0 | 28 | 121.41 | 177848.9 | 132.67 | 98.7716
55.6 | 0 | 28 | 133.21 | 195132.9 | 133.53 | 99.4119
54.4 | 0 | 28 | 121.52 | 177997.5 | 132.87 | 98.9205
53.2 | 0 | 28 | 127.76 | 187144.9 | 133.26 | 99.2108
52 | 0 | 28 | 131.55 | 192698 | 133.61 | 99.4714
50.8 | 0 | 28 | 122.91 | 180046.6 | 133.12 | 99.1066
49.6 | 0 | 28 | 127.1 | 186183.3 | 133.36 | 99.2853
48.4 | 0 | 28 | 132.52 | 194118.1 | 133.51 | 99.397
47.2 | 0 | 28 | 123.36 | 180701.2 | 133.22 | 99.1811
46 | 0 | 28 | 127.42 | 186641.5 | 133.41 | 99.3225
44.8 | 0 | 28 | 133.4 | 195412.8 | 133.44 | 99.3448
43.6 | 0 | 28 | 122.16 | 178946.7 | 133.06 | 99.0619
42.4 | 0 | 28 | 127.53 | 186806.1 | 133.53 | 99.4119
41.2 | 0 | 28 | 132.94 | 194739.8 | 133.71 | 99.5459
40 | 0 | 28 | 122.84 | 179938.6 | 133.12 | 99.1066
38.8 | 0 | 28 | 127.22 | 186356.8 | 133.49 | 99.3821
37.6 | 0 | 28 | 133.52 | 195578.8 | 133.51 | 99.397
36.4 | 0 | 28 | 121.8 | 178414.5 | 133.19 | 99.1587
35.2 | 0 | 28 | 128.42 | 188118.4 | 133.42 | 99.33
34 | 0 | 28 | 131.78 | 193033.3 | 133.42 | 99.33
32.8 | 0 | 28 | 123.07 | 180280.7 | 132.98 | 99.0024
31.6 | 0 | 28 | 126.96 | 185966.3 | 133.05 | 99.0545
30.4 | 0 | 28 | 133.04 | 194880.7 | 133.73 | 99.5608
29.2 | 0 | 28 | 125.12 | 183276.4 | 133.09 | 99.0843
28 | 0 | 28 | 126.79 | 185726.3 | 132.96 | 98.9875
26.8 | 0 | 28 | 125.89 | 184403.9 | 133.39 | 99.3076
25.6 | 0 | 28 | 125.5 | 183839.2 | 133.27 | 99.2183
24.4 | 0 | 28 | 130.84 | 191650.7 | 133.53 | 99.4119
23.2 | 0 | 28 | 124.5 | 182363.6 | 133.41 | 99.3225
22 | 0 | 28 | 126.9 | 185889.6 | 133.47 | 99.3672
20.8 | 0 | 28 | 126.02 | 184602.2 | 133.26 | 99.2108
19.6 | 0 | 28 | 126.27 | 184956.3 | 133.57 | 99.4416
18.4 | 0 | 28 | 126.07 | 184673.9 | 133.52 | 99.4044
17.2 | 0 | 28 | 123.5 | 180912.1 | 133.37 | 99.2927
16 | 0 | 28 | 127.26 | 186417.9 | 133.31 | 99.2481
14.8 | 0 | 28 | 125.58 | 183947.8 | 133.42 | 99.33
13.6 | 0 | 28 | 125.59 | 183959.3 | 133.5 | 99.3895
12.4 | 0 | 28 | 127.9 | 187344.8 | 133.42 | 99.33
11.2 | 0 | 28 | 124.85 | 182882.1 | 133.59 | 99.4565
9.9 | 0 | 28 | 125 | 183104.5 | 133.69 | 99.531
8.8 | 0 | 28 | 130.42 | 191036.4 | 133.52 | 99.4044
7.6 | 0 | 28 | 124.97 | 183057.2 | 133.83 | 99.6352
6.3 | 0 | 28 | 126.94 | 185946.5 | 133.83 | 99.6352
5.2 | 0 | 28 | 132.2 | 193646 | 134.32 | 100
3.9 | 0 | 28 | 126.42 | 185189.5 | 133.84 | 99.6426
2.7 | 0 | 28 | 128.46 | 188172.3 | 133.7 | 99.5384
1.5 | 0 | 28 | 131.33 | 192372.8 | 133.89 | 99.6799
0.3 | 0 | 28 | 125.54 | 183891.8 | 133.66 | 99.5086
-0.9 | 0 | 28 | 129.72 | 190010.8 | 133.74 | 99.5682
-2.1 | 0 | 28 | 131.77 | 193019.3 | 134 | 99.7618
-3.2 | 0 | 28 | 124 | 181640.9 | 133.7 | 99.5384
-4.5 | 0 | 28 | 128.92 | 188842.8 | 133.96 | 99.732
-5.7 | 0 | 28 | 127.85 | 187278.6 | 134.04 | 99.7915
-6.9 | 0 | 28 | 125.52 | 183863.7 | 133.73 | 99.5608
-8.1 | 0 | 28 | 124.6 | 182523.6 | 133.84 | 99.6426
-9.3 | 0 | 28 | 129.13 | 189147.5 | 134.14 | 99.866
-10.5 | 0 | 28 | 127.09 | 186160.3 | 133.7 | 99.5384
-11.7 | 0 | 28 | 127.26 | 186416.9 | 133.82 | 99.6278
-12.9 | 0 | 28 | 127.4 | 186622.7 | 133.65 | 99.5012
-14.1 | 0 | 28 | 130.48 | 191134.1 | 134.07 | 99.8139
-15.3 | 0 | 28 | 124.61 | 182527.6 | 133.85 | 99.6501
-16.5 | 0 | 28 | 130.24 | 190770.9 | 133.79 | 99.6054
-17.7 | 0 | 28 | 125.03 | 183152.5 | 133.69 | 99.531
-18.9 | 0 | 28 | 128.69 | 188502.5 | 133.86 | 99.6575
-20.1 | 0 | 28 | 131.68 | 192891.2 | 133.85 | 99.6501
-21.3 | 0 | 28 | 125.37 | 183640.9 | 133.75 | 99.5756
-22.5 | 0 | 28 | 129.47 | 189643.3 | 133.88 | 99.6724
-23.7 | 0 | 28 | 131.1 | 192035.3 | 133.98 | 99.7469
-24.9 | 0 | 28 | 126.86 | 185825.7 | 133.44 | 99.3448
-26.1 | 0 | 28 | 128.02 | 187523.3 | 133.71 | 99.5459
-27.3 | 0 | 28 | 131.68 | 192891.7 | 133.89 | 99.6799
-28.5 | 0 | 28 | 126.87 | 185843.5 | 133.42 | 99.33
-29.7 | 0 | 28 | 127.75 | 187128.7 | 133.28 | 99.2257
-30.9 | 0 | 28 | 130.29 | 190849.8 | 133.89 | 99.6799
-32.1 | 0 | 28 | 125.65 | 184060.3 | 133.5 | 99.3895
-33.3 | 0 | 28 | 128.47 | 188181.8 | 133.37 | 99.2927
-34.5 | 0 | 28 | 129.47 | 189655.1 | 133.74 | 99.5682
-35.7 | 0 | 28 | 125.66 | 184062.3 | 133.29 | 99.2332
-36.9 | 0 | 28 | 128.42 | 188111.8 | 133.64 | 99.4937
-38.1 | 0 | 28 | 128.82 | 188704 | 133.5 | 99.3895
-39.3 | 0 | 28 | 126.24 | 184923.8 | 133.34 | 99.2704
-40.5 | 0 | 28 | 125.16 | 183336.4 | 133.43 | 99.3374
-41.7 | 0 | 28 | 128.25 | 187864.6 | 133.23 | 99.1885
-42.9 | 0 | 28 | 126.72 | 185617.7 | 133.55 | 99.4267
-44.2 | 0 | 28 | 126.58 | 185412.6 | 133.44 | 99.3448
-45.3 | 0 | 28 | 126.25 | 184928.4 | 133.16 | 99.1364
-46.5 | 0 | 28 | 126.26 | 184941.8 | 133.38 | 99.3002
-47.7 | 0 | 28 | 126.27 | 184965 | 133.28 | 99.2257
-48.9 | 0 | 28 | 127.28 | 186444.4 | 133.33 | 99.263
-50.2 | 0 | 28 | 126.68 | 185561.6 | 133.48 | 99.3746
-51.3 | 0 | 28 | 125.66 | 184072.3 | 133.42 | 99.33
-52.6 | 0 | 28 | 126.85 | 185809 | 133.04 | 99.0471
-53.7 | 0 | 28 | 127.21 | 186345 | 133.01 | 99.0247
-54.9 | 0 | 28 | 126.42 | 185178.5 | 133.3 | 99.2406
-56.2 | 0 | 28 | 127.59 | 186901 | 133.23 | 99.1885
-57.4 | 0 | 28 | 125.9 | 184413.7 | 133.22 | 99.1811
-58.6 | 0 | 28 | 126.13 | 184763 | 133 | 99.0173
-59.8 | 0 | 28 | 128 | 187500.7 | 133.4 | 99.3151
-61 | 0 | 28 | 128.22 | 187814.7 | 133.43 | 99.3374
-62.2 | 0 | 28 | 127.11 | 186195.6 | 133.33 | 99.263
-63.4 | 0 | 28 | 129.89 | 190262.3 | 133.66 | 99.5086
-64.6 | 0 | 28 | 127.28 | 186445.5 | 133.48 | 99.3746
-65.8 | 0 | 28 | 126.76 | 185673.5 | 133.48 | 99.3746
-67 | 0 | 28 | 129.55 | 189763.4 | 133.63 | 99.4863
-68.2 | 0 | 28 | 127.18 | 186292.4 | 133.46 | 99.3597
-69.4 | 0 | 28 | 128.81 | 188678 | 133.47 | 99.3672
-70.6 | 0 | 28 | 128.78 | 188645.8 | 133.58 | 99.4491
-71.8 | 0 | 28 | 127.54 | 186819.4 | 133.52 | 99.4044
-73 | 0 | 28 | 126.13 | 184762.6 | 133.14 | 99.1215
-74.2 | 0 | 28 | 130.28 | 190831.7 | 133.43 | 99.3374
-75.4 | 0 | 28 | 126.65 | 185522.2 | 133.45 | 99.3523
-76.6 | 0 | 28 | 129.39 | 189529.3 | 133.84 | 99.6426
-77.8 | 0 | 28 | 128.87 | 188771.9 | 133.55 | 99.4267
-79 | 0 | 28 | 126.39 | 185133.8 | 133.52 | 99.4044
-80.2 | 0 | 28 | 129.57 | 189796.2 | 133.43 | 99.3374
-81.4 | 0 | 28 | 127.12 | 186214.1 | 133.43 | 99.3374
-82.6 | 0 | 28 | 128.68 | 188485.7 | 133.26 | 99.2108
-83.8 | 0 | 28 | 129.07 | 189069.4 | 133.28 | 99.2257
-85 | 0 | 28 | 126.31 | 185019 | 133.38 | 99.3002
-86.2 | 0 | 28 | 129.88 | 190255.6 | 133.06 | 99.0619
-87.4 | 0 | 28 | 127.04 | 186096.3 | 133.06 | 99.0619
-88.6 | 0 | 28 | 128 | 187500.2 | 133.23 | 99.1885
-89.8 | 0 | 28 | 128.01 | 187509.1 | 133.01 | 99.0247
-91 | 0 | 28 | 126.45 | 185232.4 | 132.97 | 98.9949
-92.2 | 0 | 28 | 129.79 | 190120.3 | 133.13 | 99.1141
-93.4 | 0 | 28 | 125.86 | 184364.5 | 132.93 | 98.9652
-94.6 | 0 | 28 | 127.52 | 186790.9 | 132.74 | 98.8237
-95.8 | 0 | 28 | 127.47 | 186722.4 | 132.81 | 98.8758
-97 | 0 | 28 | 126.54 | 185353.5 | 132.57 | 98.6971
-98.2 | 0 | 28 | 127.11 | 186186.2 | 132.66 | 98.7641
-99.4 | 0 | 28 | 127.2 | 186327.3 | 132.31 | 98.5036
-100.6 | 0 | 28 | 126.42 | 185178.7 | 132.06 | 98.3175
-101.9 | 0 | 28 | 126.76 | 185674.2 | 131.95 | 98.2356
-103 | 0 | 28 | 126.62 | 185480.1 | 131.71 | 98.0569
-104.3 | 0 | 28 | 126.12 | 184741.7 | 131.66 | 98.0197
-105.4 | 0 | 28 | 125.56 | 183929.6 | 131.48 | 97.8856
-106.7 | 0 | 28 | 126.4 | 185159.1 | 131.2 | 97.6772
-107.9 | 0 | 28 | 126.6 | 185442.4 | 130.89 | 97.4464
-109.1 | 0 | 28 | 124.79 | 182800.4 | 130.38 | 97.0667
-110.3 | 0 | 28 | 125.75 | 184194.2 | 129.98 | 96.7689
-111.5 | 0 | 28 | 123.49 | 180889.3 | 129.35 | 96.2999
-112.7 | 0 | 28 | 123.72 | 181232.7 | 128.97 | 96.017
-113.9 | 0 | 28 | 123.01 | 180184.5 | 128.47 | 95.6447
-115.1 | 0 | 28 | 121.3 | 177688.7 | 127.23 | 94.7216
-116.3 | 0 | 28 | 121.54 | 178036.9 | 126.52 | 94.193
-117.5 | 0 | 28 | 118.67 | 173825.9 | 124.83 | 92.9348
-118.7 | 0 | 28 | 122.58 | 179554.1 | 123.7 | 92.0935
-119.9 | 0 | 28 | 115.43 | 169085.1 | 120.84 | 89.9643
-121.1 | 0 | 28 | 108.72 | 159250.6 | 118.22 | 88.0137
-122.3 | 0 | 28 | 112.51 | 164801.5 | 114.98 | 85.6015
-123.5 | 0 | 28 | 105.11 | 153971 | 110.18 | 82.028
-124.7 | 0 | 28 | 94.4 | 138271.9 | 104.3 | 77.6504
-125.9 | 0 | 28 | 96.4 | 141207.4 | 98.32 | 73.1983
-127.1 | 0 | 28 | 86.11 | 126142.8 | 90.54 | 67.4062
-128.3 | 0 | 28 | 73.9 | 108245.1 | 81.92 | 60.9887
-129.5 | 0 | 28 | 72.19 | 105739.2 | 73.59 | 54.7871
-130.7 | 0 | 28 | 60 | 87884.5 | 64.34 | 47.9005
-131.9 | 0 | 28 | 50.57 | 74074.4 | 55.65 | 41.4309
-133.1 | 0 | 28 | 47.05 | 68926.6 | 47.68 | 35.4973
-134.3 | 0 | 28 | 37.18 | 54456.9 | 39.63 | 29.5042
-135.5 | 0 | 28 | 30.55 | 44746.6 | 33.34 | 24.8213
-136.7 | 0 | 28 | 27.58 | 40397.7 | 27.8 | 20.6968
-137.9 | 0 | 28 | 21.37 | 31305.2 | 22.79 | 16.9669
-139.1 | 0 | 28 | 17.54 | 25690.7 | 19.03 | 14.1677
-140.3 | 0 | 28 | 15.66 | 22933.7 | 15.87 | 11.8151
-141.5 | 0 | 28 | 11.93 | 17470.9 | 12.96 | 9.6486
-142.7 | 0 | 28 | 10.41 | 15245.7 | 10.9 | 8.1149
-143.9 | 0 | 28 | 8.58 | 12571.4 | 8.96 | 6.6706
-145.1 | 0 | 28 | 7.08 | 10377.3 | 7.38 | 5.4943
-146.3 | 0 | 28 | 5.76 | 8443.3 | 6.08 | 4.5265
-147.5 | 0 | 28 | 4.81 | 7046.8 | 5.01 | 3.7299
-148.7 | 0 | 28 | 3.85 | 5637.8 | 4.03 | 3.0003
-149.9 | 0 | 28 | 3.03 | 4439.8 | 3.16 | 2.3526
-151.1 | 0 | 28 | 2.39 | 3505 | 2.48 | 1.8463
-152.4 | 0 | 28 | 1.76 | 2576.1 | 1.85 | 1.3773
-153.5 | 0 | 28 | 1.34 | 1968.2 | 1.39 | 1.0348
-154.8 | 0 | 28 | 0.9 | 1313.9 | 0.94 | 0.69982
-156 | 0 | 28 | 0.61 | 896.4 | 0.64 | 0.47647
-157.2 | 0 | 28 | 0.4 | 592.9 | 0.42 | 0.31269
-158.4 | 0 | 28 | 0.13 | 189.8 | 0.14 | 0.10423
-159.6 | 0 | 28 | 0.07 | 109 | 0.08 | 0.059559
-160.8 | 0 | 28 | 0.28 | 411.6 | 0.29 | 0.2159
-162 | 0 | 28 | 0.41 | 604.5 | 0.43 | 0.32013
-163.2 | 0 | 28 | 0.53 | 770.9 | 0.55 | 0.40947
-164.4 | 0 | 28 | 0.59 | 861.4 | 0.61 | 0.45414
-165.6 | 0 | 28 | 0.67 | 984.6 | 0.7 | 0.52114
-166.8 | 0 | 28 | 0.72 | 1050.6 | 0.74 | 0.55092
-168 | 0 | 28 | 0.77 | 1133.5 | 0.8 | 0.59559
-169.2 | 0 | 28 | 0.86 | 1262.3 | 0.9 | 0.67004
-170.4 | 0 | 28 | 0.89 | 1306 | 0.93 | 0.69238
-171.6 | 0 | 28 | 0.88 | 1285.2 | 0.91 | 0.67749
-172.8 | 0 | 28 | 0.95 | 1389 | 0.99 | 0.73705
-174 | 0 | 28 | 0.95 | 1394.1 | 0.99 | 0.73705
-175.2 | 0 | 28 | 1.02 | 1490.9 | 1.06 | 0.78916
-176.4 | 0 | 28 | 0.95 | 1393.4 | 1 | 0.74449
-177.6 | 0 | 28 | 0.99 | 1448.4 | 1.03 | 0.76683
-178.8 | 0 | 28 | 1.01 | 1479.5 | 1.07 | 0.79661
-180 | 0 | 28 | 1.02 | 1496.4 | 1.05 | 0.78172
-181.2 | 0 | 28 | 1.02 | 1488.3 | 1.08 | 0.80405
-182.4 | 0 | 28 | 1.04 | 1529.2 | 1.13 | 0.84127
-183.6 | 0 | 28 | 1.01 | 1483.2 | 1.05 | 0.78172
-184.8 | 0 | 28 | 1.05 | 1542.5 | 1.11 | 0.82638
-186 | 0 | 28 | 1.06 | 1546.6 | 1.12 | 0.83383
-187.2 | 0 | 28 | 1.04 | 1527.2 | 1.1 | 0.81894
-188.4 | 0 | 28 | 1.09 | 1590.8 | 1.15 | 0.85616
-189.6 | 0 | 28 | 1.05 | 1539.6 | 1.12 | 0.83383
-190.8 | 0 | 28 | 1.1 | 1611.9 | 1.15 | 0.85616
-192 | 0 | 28 | 1.12 | 1641.9 | 1.2 | 0.89339
-193.2 | 0 | 28 | 1.14 | 1668 | 1.21 | 0.90083
-194.4 | 0 | 28 | 1.12 | 1639.2 | 1.18 | 0.8785
-195.6 | 0 | 28 | 1.16 | 1696.6 | 1.22 | 0.90828
-196.8 | 0 | 28 | 1.15 | 1679.4 | 1.22 | 0.90828
-198 | 0 | 28 | 1.17 | 1708.6 | 1.21 | 0.90083
-199.2 | 0 | 28 | 1.14 | 1675.4 | 1.21 | 0.90083
-200.4 | 0 | 28 | 1.15 | 1684.6 | 1.21 | 0.90083
-201.6 | 0 | 28 | 1.17 | 1713.6 | 1.21 | 0.90083
-202.8 | 0 | 28 | 1.23 | 1802.2 | 1.3 | 0.96784
-204 | 0 | 28 | 1.19 | 1737.4 | 1.24 | 0.92317
-205.2 | 0 | 28 | 1.18 | 1724.6 | 1.23 | 0.91572
-206.4 | 0 | 28 | 1.2 | 1761.4 | 1.26 | 0.93806
-207.6 | 0 | 28 | 1.2 | 1760.5 | 1.25 | 0.93061
-208.8 | 0 | 28 | 1.2 | 1763.6 | 1.24 | 0.92317
-210 | 0 | 28 | 1.27 | 1860.8 | 1.34 | 0.99762
-211.2 | 0 | 28 | 1.24 | 1811.6 | 1.29 | 0.96039
-212.4 | 0 | 28 | 1.27 | 1854.9 | 1.33 | 0.99017
-213.6 | 0 | 28 | 1.26 | 1849.8 | 1.31 | 0.97528
-214.8 | 0 | 28 | 1.26 | 1847.3 | 1.32 | 0.98273
-216 | 0 | 28 | 1.3 | 1902 | 1.35 | 1.0051
-217.3 | 0 | 28 | 1.29 | 1884.7 | 1.34 | 0.99762
-218.5 | 0 | 28 | 1.33 | 1948.3 | 1.39 | 1.0348
-219.7 | 0 | 28 | 1.39 | 2030.4 | 1.44 | 1.0721
-220.9 | 0 | 28 | 1.31 | 1915.9 | 1.36 | 1.0125
-222.1 | 0 | 28 | 1.36 | 1989.5 | 1.42 | 1.0572
-223.3 | 0 | 28 | 1.35 | 1974.6 | 1.4 | 1.0423
-224.5 | 0 | 28 | 1.37 | 2011.5 | 1.43 | 1.0646
-225.7 | 0 | 28 | 1.37 | 2007.6 | 1.44 | 1.0721
-226.9 | 0 | 28 | 1.37 | 2002.2 | 1.41 | 1.0497
-228.5 | 0 | 28 | 1.35 | 1970.5 | 1.4 | 1.0423
Measurement time: | 07/09/2022 17:05:55
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 200 x 200 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-202.8 | 0 | 28 | 0.2 | 383.3 | 0.53
-202.1 | 0 | 28 | 0.5 | 980 | 0.49
-200.8 | 0 | 28 | 0.45 | 884.5 | 0.46
-199.6 | 0 | 28 | 0.47 | 919.6 | 0.47
-198.4 | 0 | 28 | 0.54 | 1064 | 0.52
-197.2 | 0 | 28 | 0.47 | 921.5 | 0.49
-196 | 0 | 28 | 0.55 | 1077.7 | 0.55
-194.8 | 0 | 28 | 0.56 | 1104.3 | 0.55
-193.6 | 0 | 28 | 0.51 | 990.8 | 0.51
-192.4 | 0 | 28 | 0.51 | 1002.3 | 0.53
-191.2 | 0 | 28 | 0.56 | 1091.9 | 0.54
-190 | 0 | 28 | 0.51 | 1006 | 0.53
-188.8 | 0 | 28 | 0.55 | 1080.1 | 0.55
-187.5 | 0 | 28 | 0.58 | 1139.1 | 0.55
-186.3 | 0 | 28 | 0.49 | 962 | 0.51
-185.2 | 0 | 28 | 0.53 | 1038.1 | 0.53
-184 | 0 | 28 | 0.55 | 1083.4 | 0.53
-182.7 | 0 | 28 | 0.5 | 969.3 | 0.52
-181.6 | 0 | 28 | 0.58 | 1129.9 | 0.57
-180.3 | 0 | 28 | 0.56 | 1098.7 | 0.54
-179.1 | 0 | 28 | 0.52 | 1015.3 | 0.54
-178 | 0 | 28 | 0.53 | 1040.8 | 0.52
-176.7 | 0 | 28 | 0.56 | 1098.8 | 0.57
-175.5 | 0 | 28 | 0.54 | 1052.8 | 0.54
-174.3 | 0 | 28 | 0.5 | 982.6 | 0.52
-173.1 | 0 | 28 | 0.58 | 1136.4 | 0.57
-171.9 | 0 | 28 | 0.52 | 1021.5 | 0.53
-170.7 | 0 | 28 | 0.53 | 1032.2 | 0.52
-169.5 | 0 | 28 | 0.56 | 1101.4 | 0.55
-168.3 | 0 | 28 | 0.51 | 1002.2 | 0.52
-167.1 | 0 | 28 | 0.56 | 1101.4 | 0.56
-165.9 | 0 | 28 | 0.58 | 1139.7 | 0.57
-164.7 | 0 | 28 | 0.59 | 1153.1 | 0.59
-163.5 | 0 | 28 | 0.59 | 1159.6 | 0.59
-162.3 | 0 | 28 | 0.63 | 1238 | 0.63
-161.1 | 0 | 28 | 0.61 | 1192.7 | 0.62
-159.9 | 0 | 28 | 0.64 | 1261.4 | 0.64
-158.7 | 0 | 28 | 0.67 | 1302.8 | 0.67
-157.5 | 0 | 28 | 0.69 | 1344.4 | 0.7
-156.3 | 0 | 28 | 0.67 | 1303 | 0.68
-155.1 | 0 | 28 | 0.68 | 1327 | 0.68
-153.9 | 0 | 28 | 0.7 | 1379.8 | 0.71
-152.7 | 0 | 28 | 0.71 | 1392.7 | 0.72
-151.5 | 0 | 28 | 0.75 | 1462.8 | 0.74
-150.3 | 0 | 28 | 0.75 | 1460 | 0.75
-149.1 | 0 | 28 | 0.74 | 1455.5 | 0.74
-147.9 | 0 | 28 | 0.77 | 1515.5 | 0.77
-146.7 | 0 | 28 | 0.77 | 1500.2 | 0.77
-145.5 | 0 | 28 | 0.78 | 1526.5 | 0.78
-144.3 | 0 | 28 | 0.85 | 1658.6 | 0.83
-143.1 | 0 | 28 | 0.82 | 1614.1 | 0.82
-141.9 | 0 | 28 | 0.86 | 1687.9 | 0.86
-140.7 | 0 | 28 | 0.96 | 1886.4 | 0.95
-139.5 | 0 | 28 | 0.96 | 1886.8 | 0.97
-138.3 | 0 | 28 | 1.06 | 2071.7 | 1.05
-137.1 | 0 | 28 | 1.16 | 2268.6 | 1.14
-135.9 | 0 | 28 | 1.16 | 2272 | 1.16
-134.7 | 0 | 28 | 1.33 | 2599.6 | 1.32
-133.5 | 0 | 28 | 1.42 | 2774.9 | 1.4
-132.2 | 0 | 28 | 1.53 | 3001.1 | 1.54
-131.1 | 0 | 28 | 1.76 | 3450 | 1.74
-129.9 | 0 | 28 | 1.98 | 3874.8 | 1.96
-128.6 | 0 | 28 | 2.27 | 4448.4 | 2.27
-127.4 | 0 | 28 | 2.53 | 4952.9 | 2.55
-126.2 | 0 | 28 | 2.97 | 5805.7 | 2.91
-125.1 | 0 | 28 | 3.3 | 6453.1 | 3.34
-123.8 | 0 | 28 | 3.87 | 7580.3 | 3.85
-122.7 | 0 | 28 | 4.42 | 8657.1 | 4.44
-121.5 | 0 | 28 | 5.11 | 9992.4 | 5.13
-120.2 | 0 | 28 | 5.86 | 11472.3 | 5.99
-119.1 | 0 | 28 | 6.99 | 13688.8 | 6.93
-117.8 | 0 | 28 | 8.13 | 15911.5 | 8.15
-116.6 | 0 | 28 | 9.38 | 18354.2 | 9.49
-115.4 | 0 | 28 | 11 | 21533.6 | 11.15
-114.2 | 0 | 28 | 13.18 | 25789.4 | 13.19
-113 | 0 | 28 | 15.5 | 30339.6 | 15.49
-111.8 | 0 | 28 | 18.73 | 36668 | 18.4
-110.6 | 0 | 28 | 21.65 | 42371.8 | 21.76
-109.4 | 0 | 28 | 25.79 | 50481 | 25.66
-108.2 | 0 | 28 | 30.75 | 60180.5 | 30.48
-107 | 0 | 28 | 35.62 | 69710 | 35.75
-105.8 | 0 | 28 | 41.98 | 82172.9 | 41.57
-104.6 | 0 | 28 | 47.98 | 93906 | 47.66
-103.4 | 0 | 28 | 54.37 | 106427.1 | 54.27
-102.2 | 0 | 28 | 60.78 | 118956.1 | 60.69
-101 | 0 | 28 | 66.4 | 129974.8 | 66.56
-99.8 | 0 | 28 | 72.27 | 141463.8 | 72.33
-98.6 | 0 | 28 | 76.43 | 149598.3 | 77.35
-97.4 | 0 | 28 | 81.14 | 158809.3 | 81.58
-96.2 | 0 | 28 | 85.14 | 166652.6 | 85.51
-95 | 0 | 28 | 88.09 | 172423.3 | 88.33
-93.8 | 0 | 28 | 90.05 | 176261.5 | 90.61
-92.6 | 0 | 28 | 90.63 | 177389 | 92.59
-91.4 | 0 | 28 | 94.79 | 185539.9 | 94.07
-90.2 | 0 | 28 | 97.14 | 190133.5 | 95.25
-89 | 0 | 28 | 95.5 | 186930 | 96.2
-87.8 | 0 | 28 | 99.18 | 194124.1 | 96.96
-86.6 | 0 | 28 | 96.42 | 188722.9 | 97.5
-85.4 | 0 | 28 | 98.2 | 192210.8 | 97.94
-84.2 | 0 | 28 | 98.9 | 193581.2 | 98.36
-83 | 0 | 28 | 97.89 | 191604.9 | 98.48
-81.8 | 0 | 28 | 98.83 | 193446 | 98.84
-80.6 | 0 | 28 | 98.19 | 192195.9 | 98.98
-79.4 | 0 | 28 | 101.78 | 199219.5 | 99.16
-78.2 | 0 | 28 | 98.46 | 192726 | 99.57
-77 | 0 | 28 | 101.41 | 198497.3 | 99.46
-75.8 | 0 | 28 | 98.02 | 191856.4 | 99.73
-74.6 | 0 | 28 | 101.02 | 197734.4 | 99.71
-73.4 | 0 | 28 | 102.46 | 200555.5 | 99.62
-72.2 | 0 | 28 | 99.2 | 194159.6 | 99.8
-71 | 0 | 28 | 100.03 | 195789.8 | 99.92
-69.7 | 0 | 28 | 100.11 | 195937.3 | 99.97
-68.6 | 0 | 28 | 100.73 | 197167.2 | 100.02
-67.3 | 0 | 28 | 100.07 | 195864.7 | 100.05
-66.1 | 0 | 28 | 100.44 | 196596.4 | 99.96
-64.9 | 0 | 28 | 100.27 | 196267.8 | 99.83
-63.7 | 0 | 28 | 101.2 | 198071.4 | 100.15
-62.5 | 0 | 28 | 100.74 | 197178.4 | 100.09
-61.3 | 0 | 28 | 100.36 | 196427.4 | 100.08
-60.1 | 0 | 28 | 101.3 | 198282.6 | 100.27
-58.9 | 0 | 28 | 101.14 | 197967.4 | 100.31
-57.7 | 0 | 28 | 100.98 | 197658.3 | 100.11
-56.5 | 0 | 28 | 100.94 | 197562.3 | 100.23
-55.3 | 0 | 28 | 100.94 | 197567.2 | 100.28
-54.1 | 0 | 28 | 100.2 | 196130.3 | 100.32
-52.9 | 0 | 28 | 100.23 | 196178.7 | 100.24
-51.7 | 0 | 28 | 99.14 | 194044 | 100.05
-50.5 | 0 | 28 | 98.78 | 193335.4 | 100.27
-49.3 | 0 | 28 | 104.38 | 204304.4 | 100.19
-48.1 | 0 | 28 | 99.33 | 194418.5 | 100.09
-46.9 | 0 | 28 | 98.68 | 193151.1 | 100
-45.7 | 0 | 28 | 98.23 | 192270.7 | 99.99
-44.5 | 0 | 28 | 99.03 | 193842.1 | 100.05
-43.3 | 0 | 28 | 95.31 | 186559.9 | 99.98
-42.1 | 0 | 28 | 100.02 | 195761.1 | 100.21
-40.9 | 0 | 28 | 100.21 | 196138.8 | 100.18
-39.7 | 0 | 28 | 100.7 | 197100.1 | 99.9
-38.5 | 0 | 28 | 100.27 | 196253.9 | 100.16
-37.3 | 0 | 28 | 100.56 | 196825.3 | 100.15
-36.1 | 0 | 28 | 99.91 | 195547.1 | 99.87
-34.9 | 0 | 28 | 101.44 | 198555 | 100.07
-33.7 | 0 | 28 | 100.43 | 196576.2 | 100.2
-32.5 | 0 | 28 | 101.63 | 198931.5 | 99.94
-31.3 | 0 | 28 | 101.22 | 198119.3 | 100.23
-30.1 | 0 | 28 | 100.81 | 197321.3 | 100.2
-28.9 | 0 | 28 | 100.94 | 197570.7 | 100.18
-27.7 | 0 | 28 | 100.78 | 197263.2 | 100.07
-26.5 | 0 | 28 | 100.82 | 197333.1 | 100.12
-25.3 | 0 | 28 | 100.58 | 196875.9 | 100.22
-24.1 | 0 | 28 | 100.81 | 197312.1 | 100.17
-22.9 | 0 | 28 | 100.52 | 196741.5 | 100.21
-21.7 | 0 | 28 | 100.02 | 195768.9 | 100.12
-20.5 | 0 | 28 | 101.31 | 198289.2 | 100.22
-19.3 | 0 | 28 | 99.85 | 195438 | 100.13
-18.1 | 0 | 28 | 101.13 | 197942.1 | 100.06
-16.9 | 0 | 28 | 100.48 | 196662.8 | 100.3
-15.6 | 0 | 28 | 100.27 | 196256.2 | 100.09
-14.4 | 0 | 28 | 100.17 | 196069.6 | 100.01
-13.2 | 0 | 28 | 99.91 | 195564.3 | 100.29
-12 | 0 | 28 | 99.22 | 194204.2 | 100.23
-10.8 | 0 | 28 | 99.78 | 195298.2 | 99.82
-9.6 | 0 | 28 | 100.22 | 196169.7 | 100.26
-8.4 | 0 | 28 | 99.45 | 194658.5 | 100.31
-7.2 | 0 | 28 | 97.96 | 191741.4 | 100.02
-6 | 0 | 28 | 100.46 | 196622.2 | 100.18
-4.8 | 0 | 28 | 99.26 | 194289.1 | 100.13
-3.6 | 0 | 28 | 98.65 | 193089.9 | 100.05
-2.4 | 0 | 28 | 100.84 | 197385.1 | 100.16
-1.2 | 0 | 28 | 99.57 | 194881.6 | 100.15
0 | 0 | 28 | 100.49 | 196691.6 | 100.3
1.2 | 0 | 28 | 101.23 | 198147.7 | 100.2
2.4 | 0 | 28 | 100.57 | 196851.5 | 100.1
3.6 | 0 | 28 | 101.03 | 197740.6 | 100.26
4.8 | 0 | 28 | 100.8 | 197300.5 | 100.2
6 | 0 | 28 | 101.27 | 198208 | 100.11
7.2 | 0 | 28 | 99.96 | 195661.3 | 100.04
8.4 | 0 | 28 | 100.84 | 197381 | 100.29
9.6 | 0 | 28 | 100.54 | 196797.3 | 99.99
10.8 | 0 | 28 | 100.27 | 196268 | 99.96
12 | 0 | 28 | 100.47 | 196652.9 | 100.12
13.2 | 0 | 28 | 100.39 | 196487.7 | 100.06
14.4 | 0 | 28 | 100.98 | 197652.6 | 100.03
15.6 | 0 | 28 | 100.53 | 196761.8 | 100.15
16.8 | 0 | 28 | 99.49 | 194733.9 | 100.08
18 | 0 | 28 | 99.28 | 194316 | 99.93
19.2 | 0 | 28 | 99.24 | 194235.4 | 100.05
20.4 | 0 | 28 | 98.8 | 193376.9 | 100.09
21.6 | 0 | 28 | 98.17 | 192145.6 | 99.72
22.8 | 0 | 28 | 99 | 193778.9 | 99.97
24 | 0 | 28 | 98.01 | 191845.2 | 100.05
25.2 | 0 | 28 | 98.83 | 193443.6 | 99.9
26.5 | 0 | 28 | 99.82 | 195386.4 | 100
27.6 | 0 | 28 | 99.65 | 195048.8 | 99.96
28.8 | 0 | 28 | 98.54 | 192872.3 | 99.9
30 | 0 | 28 | 99.66 | 195059.3 | 100.05
31.2 | 0 | 28 | 98.88 | 193544.1 | 99.82
32.5 | 0 | 28 | 98.83 | 193439.9 | 100.03
33.6 | 0 | 28 | 100.8 | 197288.9 | 99.99
34.9 | 0 | 28 | 99.83 | 195399.1 | 99.78
36 | 0 | 28 | 100.41 | 196536.6 | 99.84
37.3 | 0 | 28 | 100.89 | 197472.2 | 100.16
38.5 | 0 | 28 | 99.95 | 195631.4 | 99.81
39.7 | 0 | 28 | 99.77 | 195284.3 | 99.71
40.9 | 0 | 28 | 100.47 | 196653.8 | 99.93
42.1 | 0 | 28 | 99.97 | 195679.8 | 99.93
43.3 | 0 | 28 | 101.16 | 197999.7 | 99.83
44.5 | 0 | 28 | 100.12 | 195964.8 | 99.72
45.7 | 0 | 28 | 99.43 | 194623.9 | 99.86
46.9 | 0 | 28 | 98.4 | 192602.4 | 99.65
48.1 | 0 | 28 | 99.81 | 195355.1 | 99.87
49.3 | 0 | 28 | 99.42 | 194586.6 | 99.92
50.5 | 0 | 28 | 98.67 | 193127.1 | 99.84
51.7 | 0 | 28 | 100.14 | 196007.2 | 99.74
52.9 | 0 | 28 | 98.73 | 193239.9 | 99.93
54.1 | 0 | 28 | 99.06 | 193893.8 | 99.76
55.3 | 0 | 28 | 100.41 | 196529.3 | 99.76
56.5 | 0 | 28 | 98.46 | 192710.6 | 99.87
57.7 | 0 | 28 | 97.87 | 191568.3 | 99.74
58.9 | 0 | 28 | 103.04 | 201681.4 | 100
60.1 | 0 | 28 | 97.33 | 190500.5 | 99.79
61.3 | 0 | 28 | 100.68 | 197069.1 | 99.76
62.5 | 0 | 28 | 104.51 | 204559 | 99.91
63.7 | 0 | 28 | 96.51 | 188910.1 | 100.03
64.9 | 0 | 28 | 100.35 | 196423.3 | 99.65
66.1 | 0 | 28 | 104.04 | 203636.8 | 99.85
67.3 | 0 | 28 | 95.44 | 186814.8 | 99.71
68.5 | 0 | 28 | 100.31 | 196344.2 | 99.68
69.7 | 0 | 28 | 103.66 | 202886.2 | 99.69
70.9 | 0 | 28 | 96.33 | 188548.1 | 99.72
72.1 | 0 | 28 | 99.76 | 195259 | 99.5
73.3 | 0 | 28 | 103.64 | 202856 | 99.57
74.5 | 0 | 28 | 95.79 | 187500.3 | 99.39
75.7 | 0 | 28 | 98.92 | 193625.7 | 99.14
76.9 | 0 | 28 | 102.93 | 201470.5 | 99.22
78.1 | 0 | 28 | 95.63 | 187184.1 | 99.08
79.3 | 0 | 28 | 98.7 | 193179.5 | 98.79
80.5 | 0 | 28 | 101.96 | 199559.6 | 98.83
81.8 | 0 | 28 | 98.29 | 192375.4 | 98.39
82.9 | 0 | 28 | 97.44 | 190730.4 | 98.04
84.1 | 0 | 28 | 96.9 | 189659.4 | 97.92
85.4 | 0 | 28 | 96.21 | 188315.5 | 97.56
86.5 | 0 | 28 | 91.65 | 179383.5 | 97.05
87.7 | 0 | 28 | 100.57 | 196851.8 | 96.44
89 | 0 | 28 | 94.9 | 185752.3 | 95.7
90.2 | 0 | 28 | 90.02 | 176196.2 | 94.75
91.4 | 0 | 28 | 97.26 | 190368.9 | 93.7
92.6 | 0 | 28 | 91.51 | 179121.4 | 91.98
93.8 | 0 | 28 | 89.39 | 174967.4 | 90.01
95 | 0 | 28 | 87.49 | 171242.9 | 87.74
96.2 | 0 | 28 | 83.87 | 164163 | 84.79
97.4 | 0 | 28 | 81.15 | 158829.2 | 81.16
98.6 | 0 | 28 | 76.37 | 149475.7 | 76.96
99.8 | 0 | 28 | 71.61 | 140170.5 | 71.95
101 | 0 | 28 | 63.41 | 124109.9 | 66.3
102.2 | 0 | 28 | 63.27 | 123838.5 | 60.54
103.4 | 0 | 28 | 51.85 | 101490.2 | 54.09
104.6 | 0 | 28 | 48.06 | 94071.5 | 47.82
105.8 | 0 | 28 | 42.57 | 83319 | 41.58
107 | 0 | 28 | 34.45 | 67435.2 | 35.59
108.2 | 0 | 28 | 30.45 | 59591.4 | 30.43
109.4 | 0 | 28 | 26.36 | 51585.2 | 25.73
110.6 | 0 | 28 | 21.36 | 41807.2 | 21.56
111.8 | 0 | 28 | 18.42 | 36060.9 | 18.27
113 | 0 | 28 | 15.81 | 30942.3 | 15.4
114.2 | 0 | 28 | 12.74 | 24937.6 | 12.95
115.4 | 0 | 28 | 11.07 | 21661 | 10.94
116.6 | 0 | 28 | 9.5 | 18585.3 | 9.31
117.8 | 0 | 28 | 7.91 | 15477.5 | 7.95
119 | 0 | 28 | 6.84 | 13384.1 | 6.84
120.2 | 0 | 28 | 6.09 | 11924.4 | 5.93
121.4 | 0 | 28 | 5.08 | 9933.8 | 5.13
122.6 | 0 | 28 | 4.44 | 8694.5 | 4.44
123.8 | 0 | 28 | 3.98 | 7783.5 | 3.89
125 | 0 | 28 | 3.34 | 6547 | 3.39
126.2 | 0 | 28 | 3.02 | 5910.7 | 3.01
127.4 | 0 | 28 | 2.73 | 5345.3 | 2.67
128.6 | 0 | 28 | 2.29 | 4473.5 | 2.33
129.8 | 0 | 28 | 2.12 | 4142.1 | 2.12
131.1 | 0 | 28 | 1.92 | 3748.5 | 1.85
132.2 | 0 | 28 | 1.66 | 3258.7 | 1.71
133.4 | 0 | 28 | 1.51 | 2949.8 | 1.5
134.6 | 0 | 28 | 1.45 | 2838.4 | 1.41
135.8 | 0 | 28 | 1.24 | 2434.3 | 1.28
137 | 0 | 28 | 1.12 | 2185 | 1.11
138.2 | 0 | 28 | 1.08 | 2117.9 | 1.06
139.5 | 0 | 28 | 0.93 | 1827.1 | 0.96
140.6 | 0 | 28 | 0.89 | 1749.6 | 0.89
141.9 | 0 | 28 | 0.85 | 1657.2 | 0.83
143.1 | 0 | 28 | 0.75 | 1476.8 | 0.77
144.3 | 0 | 28 | 0.78 | 1527 | 0.78
145.5 | 0 | 28 | 0.73 | 1430.2 | 0.72
146.7 | 0 | 28 | 0.65 | 1273.1 | 0.67
147.9 | 0 | 28 | 0.7 | 1362 | 0.69
149.1 | 0 | 28 | 0.65 | 1280.6 | 0.65
150.3 | 0 | 28 | 0.59 | 1155.9 | 0.6
151.5 | 0 | 28 | 0.56 | 1096.7 | 0.56
152.7 | 0 | 28 | 0.57 | 1125.4 | 0.57
153.9 | 0 | 28 | 0.57 | 1121.4 | 0.58
155.1 | 0 | 28 | 0.52 | 1024 | 0.53
156.3 | 0 | 28 | 0.54 | 1048.2 | 0.52
157.5 | 0 | 28 | 0.5 | 979.8 | 0.5
158.7 | 0 | 28 | 0.47 | 925.5 | 0.48
159.9 | 0 | 28 | 0.53 | 1028.4 | 0.52
161.1 | 0 | 28 | 0.45 | 878.6 | 0.45
162.3 | 0 | 28 | 0.45 | 876.8 | 0.45
163.5 | 0 | 28 | 0.49 | 963.4 | 0.49
164.7 | 0 | 28 | 0.43 | 844.4 | 0.43
165.9 | 0 | 28 | 0.48 | 937.1 | 0.48
167.1 | 0 | 28 | 0.47 | 911 | 0.47
168.3 | 0 | 28 | 0.38 | 735.1 | 0.38
169.5 | 0 | 28 | 0.45 | 879.2 | 0.45
170.7 | 0 | 28 | 0.39 | 765.7 | 0.39
171.9 | 0 | 28 | 0.41 | 798 | 0.41
173.1 | 0 | 28 | 0.37 | 726.5 | 0.37
174.3 | 0 | 28 | 0.41 | 803.8 | 0.4
175.5 | 0 | 28 | 0.36 | 701.5 | 0.36
176.7 | 0 | 28 | 0.38 | 743.4 | 0.38
177.9 | 0 | 28 | 0.36 | 700.7 | 0.35
179.1 | 0 | 28 | 0.36 | 707.6 | 0.36
180.3 | 0 | 28 | 0.38 | 739.5 | 0.37
181.5 | 0 | 28 | 0.34 | 672.9 | 0.34
182.7 | 0 | 28 | 0.34 | 657.6 | 0.34
183.9 | 0 | 28 | 0.34 | 667.9 | 0.34
185.2 | 0 | 28 | 0.35 | 691 | 0.35
186.3 | 0 | 28 | 0.35 | 682.5 | 0.35
187.5 | 0 | 28 | 0.34 | 663.9 | 0.34
188.8 | 0 | 28 | 0.38 | 734.3 | 0.37
189.9 | 0 | 28 | 0.36 | 709.6 | 0.37
191.2 | 0 | 28 | 0.34 | 669.2 | 0.34
192.4 | 0 | 28 | 0.36 | 702.6 | 0.35
193.6 | 0 | 28 | 0.34 | 673.5 | 0.35
194.8 | 0 | 28 | 0.34 | 658.3 | 0.34
196 | 0 | 28 | 0.35 | 685.8 | 0.35
197.2 | 0 | 28 | 0.34 | 663.1 | 0.33
198.4 | 0 | 28 | 0.34 | 657.8 | 0.34
199.6 | 0 | 28 | 0.35 | 686.4 | 0.35
200.8 | 0 | 28 | 0.35 | 690.9 | 0.36
202.8 | 0 | 28 | 0.32 | 629.9 | 0.32
Measurement time: | 07/09/2022 16:34:35
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-200 | 0 | 28 | 0.25 | 508.3 | 0.21
-199.3 | 0 | 28 | 0.28 | 566.7 | 0.28
-197.9 | 0 | 28 | 0.3 | 589.7 | 0.3
-196.7 | 0 | 28 | 0.3 | 590.4 | 0.3
-195.5 | 0 | 28 | 0.26 | 523.9 | 0.26
-194.3 | 0 | 28 | 0.3 | 592.5 | 0.3
-193.1 | 0 | 28 | 0.31 | 613.9 | 0.3
-191.9 | 0 | 28 | 0.31 | 623 | 0.31
-190.7 | 0 | 28 | 0.32 | 649.1 | 0.32
-189.5 | 0 | 28 | 0.3 | 598.9 | 0.3
-188.3 | 0 | 28 | 0.31 | 616.5 | 0.3
-187.1 | 0 | 28 | 0.27 | 537.5 | 0.28
-185.9 | 0 | 28 | 0.34 | 674.2 | 0.34
-184.7 | 0 | 28 | 0.28 | 568.3 | 0.29
-183.5 | 0 | 28 | 0.33 | 649.8 | 0.32
-182.3 | 0 | 28 | 0.32 | 636.3 | 0.32
-181.1 | 0 | 28 | 0.32 | 648.5 | 0.32
-179.9 | 0 | 28 | 0.37 | 729.2 | 0.36
-178.7 | 0 | 28 | 0.34 | 681.3 | 0.35
-177.5 | 0 | 28 | 0.34 | 682.3 | 0.34
-176.3 | 0 | 28 | 0.38 | 767.3 | 0.39
-175.1 | 0 | 28 | 0.39 | 769.7 | 0.39
-173.9 | 0 | 28 | 0.35 | 699.1 | 0.35
-172.7 | 0 | 28 | 0.4 | 790.7 | 0.4
-171.5 | 0 | 28 | 0.39 | 784.2 | 0.38
-170.3 | 0 | 28 | 0.36 | 724.9 | 0.36
-169.1 | 0 | 28 | 0.34 | 685.3 | 0.34
-167.9 | 0 | 28 | 0.35 | 695.2 | 0.34
-166.7 | 0 | 28 | 0.35 | 701.2 | 0.36
-165.5 | 0 | 28 | 0.36 | 720.1 | 0.36
-164.3 | 0 | 28 | 0.34 | 685.4 | 0.34
-163.1 | 0 | 28 | 0.39 | 773.2 | 0.39
-161.9 | 0 | 28 | 0.37 | 732.2 | 0.38
-160.7 | 0 | 28 | 0.43 | 859.9 | 0.42
-159.5 | 0 | 28 | 0.37 | 730.2 | 0.37
-158.3 | 0 | 28 | 0.36 | 711.4 | 0.35
-157.1 | 0 | 28 | 0.37 | 748.1 | 0.36
-155.9 | 0 | 28 | 0.35 | 709 | 0.36
-154.7 | 0 | 28 | 0.38 | 755 | 0.38
-153.5 | 0 | 28 | 0.39 | 780.9 | 0.38
-152.2 | 0 | 28 | 0.36 | 718.5 | 0.37
-151 | 0 | 28 | 0.38 | 758.2 | 0.38
-149.8 | 0 | 28 | 0.4 | 808.7 | 0.39
-148.6 | 0 | 28 | 0.38 | 763.9 | 0.39
-147.4 | 0 | 28 | 0.4 | 791.4 | 0.39
-146.2 | 0 | 28 | 0.39 | 777.9 | 0.37
-145 | 0 | 28 | 0.35 | 692.4 | 0.36
-143.8 | 0 | 28 | 0.38 | 766.5 | 0.39
-142.6 | 0 | 28 | 0.41 | 824 | 0.4
-141.4 | 0 | 28 | 0.42 | 833.3 | 0.42
-140.2 | 0 | 28 | 0.39 | 770.5 | 0.4
-139 | 0 | 28 | 0.42 | 839.8 | 0.4
-137.8 | 0 | 28 | 0.39 | 781.2 | 0.4
-136.6 | 0 | 28 | 0.42 | 845.9 | 0.42
-135.4 | 0 | 28 | 0.45 | 903.2 | 0.43
-134.2 | 0 | 28 | 0.39 | 781.6 | 0.4
-133 | 0 | 28 | 0.45 | 900.9 | 0.45
-131.8 | 0 | 28 | 0.42 | 846.6 | 0.41
-130.6 | 0 | 28 | 0.43 | 857.2 | 0.44
-129.4 | 0 | 28 | 0.42 | 847.7 | 0.44
-128.2 | 0 | 28 | 0.48 | 967.6 | 0.47
-127 | 0 | 28 | 0.44 | 880.7 | 0.46
-125.8 | 0 | 28 | 0.5 | 998 | 0.5
-124.6 | 0 | 28 | 0.56 | 1126.5 | 0.54
-123.4 | 0 | 28 | 0.48 | 959.9 | 0.49
-122.2 | 0 | 28 | 0.48 | 958.4 | 0.47
-121 | 0 | 28 | 0.54 | 1084.4 | 0.52
-119.8 | 0 | 28 | 0.53 | 1066.9 | 0.56
-118.6 | 0 | 28 | 0.58 | 1161.4 | 0.58
-117.4 | 0 | 28 | 0.65 | 1288.6 | 0.63
-116.2 | 0 | 28 | 0.65 | 1306.4 | 0.68
-115 | 0 | 28 | 0.71 | 1416.5 | 0.7
-113.8 | 0 | 28 | 0.79 | 1580.9 | 0.77
-112.6 | 0 | 28 | 0.81 | 1616.8 | 0.82
-111.4 | 0 | 28 | 0.95 | 1893.3 | 0.93
-110.2 | 0 | 28 | 1.04 | 2087.1 | 1.01
-109 | 0 | 28 | 1.08 | 2154 | 1.11
-107.8 | 0 | 28 | 1.3 | 2587.3 | 1.28
-106.6 | 0 | 28 | 1.45 | 2899.3 | 1.46
-105.3 | 0 | 28 | 1.59 | 3179.9 | 1.64
-104.2 | 0 | 28 | 1.88 | 3759.1 | 1.86
-102.9 | 0 | 28 | 2.23 | 4457.6 | 2.13
-101.8 | 0 | 28 | 2.4 | 4803 | 2.46
-100.5 | 0 | 28 | 2.79 | 5572.5 | 2.75
-99.3 | 0 | 28 | 3.38 | 6744.4 | 3.24
-98.1 | 0 | 28 | 3.63 | 7253 | 3.72
-96.9 | 0 | 28 | 4.27 | 8528 | 4.3
-95.7 | 0 | 28 | 4.98 | 9958.2 | 4.99
-94.5 | 0 | 28 | 5.77 | 11518.4 | 5.86
-93.3 | 0 | 28 | 7.01 | 14005.7 | 6.8
-92.1 | 0 | 28 | 8.16 | 16302.6 | 7.96
-90.9 | 0 | 28 | 9.29 | 18553.1 | 9.42
-89.7 | 0 | 28 | 11.15 | 22269.1 | 11.06
-88.5 | 0 | 28 | 13.52 | 27016.5 | 13.13
-87.3 | 0 | 28 | 15.57 | 31098 | 15.67
-86.1 | 0 | 28 | 18.48 | 36910.1 | 18.57
-84.9 | 0 | 28 | 22.46 | 44869 | 22.24
-83.7 | 0 | 28 | 26.07 | 52076.3 | 26.28
-82.5 | 0 | 28 | 31.26 | 62442.3 | 31.1
-81.3 | 0 | 28 | 37.37 | 74667.2 | 36.6
-80.1 | 0 | 28 | 42.6 | 85104.2 | 42.53
-78.9 | 0 | 28 | 48.91 | 97712.4 | 48.78
-77.7 | 0 | 28 | 56.15 | 112176.4 | 55.05
-76.5 | 0 | 28 | 60.38 | 120633.6 | 61.14
-75.3 | 0 | 28 | 66.12 | 132095.8 | 66.93
-74.1 | 0 | 28 | 73.41 | 146655.7 | 72.04
-72.9 | 0 | 28 | 75.96 | 151751.3 | 76.71
-71.7 | 0 | 28 | 80.38 | 160575.8 | 80.5
-70.5 | 0 | 28 | 84.68 | 169180.9 | 83.84
-69.3 | 0 | 28 | 86.02 | 171853.8 | 86.53
-68.1 | 0 | 28 | 88.61 | 177025.8 | 88.43
-66.9 | 0 | 28 | 90.33 | 180465 | 90.51
-65.7 | 0 | 28 | 90.97 | 181735.6 | 91.88
-64.5 | 0 | 28 | 92.31 | 184427.9 | 93.1
-63.3 | 0 | 28 | 94.18 | 188163.4 | 93.88
-62.1 | 0 | 28 | 94.31 | 188415.3 | 94.71
-60.9 | 0 | 28 | 95.26 | 190302.4 | 95.28
-59.7 | 0 | 28 | 97.81 | 195408.3 | 95.82
-58.4 | 0 | 28 | 96.43 | 192646.7 | 96.53
-57.3 | 0 | 28 | 96.6 | 192988 | 96.71
-56.1 | 0 | 28 | 97.27 | 194329.4 | 97.19
-54.8 | 0 | 28 | 97.06 | 193916 | 97.47
-53.6 | 0 | 28 | 98.58 | 196935.7 | 97.87
-52.4 | 0 | 28 | 99.21 | 198204.7 | 98.02
-51.2 | 0 | 28 | 98.96 | 197697.9 | 98.28
-50 | 0 | 28 | 99.31 | 198399.8 | 98.55
-48.8 | 0 | 28 | 100.14 | 200055 | 98.7
-47.6 | 0 | 28 | 98.43 | 196636 | 98.58
-46.4 | 0 | 28 | 97.69 | 195158.4 | 98.99
-45.2 | 0 | 28 | 99.73 | 199247.5 | 99.01
-44 | 0 | 28 | 98.56 | 196914 | 99.3
-42.8 | 0 | 28 | 100.2 | 200180.7 | 99.04
-41.6 | 0 | 28 | 99.77 | 199318.7 | 99.33
-40.4 | 0 | 28 | 100.18 | 200148 | 99.54
-39.2 | 0 | 28 | 100.68 | 201132.2 | 99.48
-38 | 0 | 28 | 99.42 | 198625 | 99.54
-36.8 | 0 | 28 | 98.1 | 195985.1 | 99.75
-35.6 | 0 | 28 | 98.69 | 197169.8 | 99.66
-34.4 | 0 | 28 | 99.55 | 198875.1 | 99.81
-33.2 | 0 | 28 | 99.85 | 199476 | 99.76
-32 | 0 | 28 | 99.44 | 198654.8 | 99.93
-30.8 | 0 | 28 | 101.47 | 202719.6 | 99.75
-29.6 | 0 | 28 | 99.64 | 199058.1 | 100.01
-28.4 | 0 | 28 | 100.9 | 201585.3 | 99.92
-27.2 | 0 | 28 | 99.49 | 198766 | 100.06
-26 | 0 | 28 | 100.04 | 199853.8 | 99.92
-24.8 | 0 | 28 | 100.43 | 200646.7 | 100
-23.6 | 0 | 28 | 100.55 | 200872.5 | 100.13
-22.4 | 0 | 28 | 102.03 | 203843.7 | 99.87
-21.2 | 0 | 28 | 100.53 | 200830.6 | 100.23
-20 | 0 | 28 | 102.02 | 203824.6 | 100.09
-18.8 | 0 | 28 | 98.29 | 196362.4 | 100.1
-17.6 | 0 | 28 | 99 | 197774.5 | 99.94
-16.4 | 0 | 28 | 100.56 | 200894.3 | 99.92
-15.2 | 0 | 28 | 101.6 | 202971.6 | 100.03
-14 | 0 | 28 | 100.38 | 200534.5 | 99.96
-12.8 | 0 | 28 | 97.35 | 194496 | 100.06
-11.6 | 0 | 28 | 100.4 | 200579.5 | 100.13
-10.4 | 0 | 28 | 99.2 | 198176.4 | 99.8
-9.2 | 0 | 28 | 100.87 | 201514.6 | 100.06
-8 | 0 | 28 | 99.32 | 198416.8 | 100.17
-6.7 | 0 | 28 | 96.56 | 192910.7 | 100.17
-5.5 | 0 | 28 | 100.34 | 200468.4 | 100.09
-4.3 | 0 | 28 | 100.45 | 200682 | 100.16
-3.1 | 0 | 28 | 101.24 | 202255.3 | 100.21
-1.9 | 0 | 28 | 96.19 | 192173.7 | 100.02
-0.7 | 0 | 28 | 101.15 | 202086.2 | 100.22
0.5 | 0 | 28 | 99.8 | 199384.5 | 100.24
1.7 | 0 | 28 | 100.38 | 200532.4 | 100.19
2.9 | 0 | 28 | 100.47 | 200726.6 | 100.13
4.1 | 0 | 28 | 100.75 | 201271.3 | 100.26
5.3 | 0 | 28 | 100.81 | 201402 | 100.13
6.5 | 0 | 28 | 101.63 | 203044.8 | 99.99
7.7 | 0 | 28 | 101.4 | 202581 | 100.24
8.9 | 0 | 28 | 100 | 199774 | 100.04
10.1 | 0 | 28 | 103.54 | 206850.5 | 100.16
11.3 | 0 | 28 | 95.63 | 191056.2 | 100.06
12.5 | 0 | 28 | 101.02 | 201822.9 | 100
13.7 | 0 | 28 | 100.96 | 201700.5 | 100.13
14.9 | 0 | 28 | 100.74 | 201264.1 | 99.87
16.1 | 0 | 28 | 96.73 | 193253.8 | 99.94
17.3 | 0 | 28 | 99.28 | 198334.8 | 100.01
18.5 | 0 | 28 | 102.16 | 204098.3 | 99.81
19.7 | 0 | 28 | 95.25 | 190298.4 | 99.83
20.9 | 0 | 28 | 100.27 | 200320.2 | 100
22.1 | 0 | 28 | 101.32 | 202419.7 | 100.05
23.3 | 0 | 28 | 100.4 | 200574.9 | 99.81
24.5 | 0 | 28 | 101.04 | 201860.9 | 100.04
25.7 | 0 | 28 | 100.77 | 201327.8 | 99.85
26.9 | 0 | 28 | 100.06 | 199902.1 | 99.74
28.1 | 0 | 28 | 99.04 | 197869.4 | 99.84
29.3 | 0 | 28 | 100 | 199776.3 | 99.82
30.5 | 0 | 28 | 99.78 | 199348.1 | 99.67
31.7 | 0 | 28 | 100.11 | 199995.5 | 99.44
32.9 | 0 | 28 | 100.75 | 201280.4 | 99.63
34.1 | 0 | 28 | 101.28 | 202344.4 | 99.66
35.3 | 0 | 28 | 99.57 | 198920.7 | 99.3
36.5 | 0 | 28 | 98.89 | 197565.3 | 99.38
37.7 | 0 | 28 | 98.1 | 195994.9 | 99.23
38.9 | 0 | 28 | 102.08 | 203944.9 | 99.16
40.1 | 0 | 28 | 97.31 | 194400.7 | 99.07
41.4 | 0 | 28 | 95.28 | 190343.9 | 98.82
42.5 | 0 | 28 | 100.6 | 200983 | 98.96
43.8 | 0 | 28 | 99.44 | 198660.2 | 98.57
45 | 0 | 28 | 98.36 | 196508 | 98.45
46.2 | 0 | 28 | 98.33 | 196436.3 | 98.53
47.4 | 0 | 28 | 97.62 | 195032 | 98.19
48.6 | 0 | 28 | 96.25 | 192297.5 | 98.02
49.8 | 0 | 28 | 98.69 | 197165.6 | 97.85
51 | 0 | 28 | 98.36 | 196495.8 | 97.44
52.2 | 0 | 28 | 98.56 | 196895.1 | 97.3
53.4 | 0 | 28 | 94.48 | 188750.2 | 97.07
54.6 | 0 | 28 | 95.93 | 191646.6 | 96.66
55.8 | 0 | 28 | 96.59 | 192972.1 | 96.33
57 | 0 | 28 | 95.78 | 191346.3 | 95.82
58.2 | 0 | 28 | 96.15 | 192096.2 | 95.47
59.4 | 0 | 28 | 94.86 | 189516.4 | 94.72
60.6 | 0 | 28 | 95.94 | 191663.2 | 93.95
61.8 | 0 | 28 | 94.36 | 188504 | 93.11
63 | 0 | 28 | 92.55 | 184907 | 92.11
64.2 | 0 | 28 | 90.41 | 180625.8 | 91.01
65.4 | 0 | 28 | 88.66 | 177122.4 | 89.41
66.6 | 0 | 28 | 88.54 | 176880.3 | 87.56
67.8 | 0 | 28 | 85.91 | 171629.6 | 85.64
69 | 0 | 28 | 83.34 | 166489.5 | 82.69
70.2 | 0 | 28 | 80.33 | 160494.4 | 79.39
71.4 | 0 | 28 | 76.35 | 152532 | 75.76
72.6 | 0 | 28 | 71.44 | 142715.8 | 71.05
73.9 | 0 | 28 | 65.28 | 130422.8 | 65.81
75 | 0 | 28 | 59.52 | 118912.4 | 60.24
76.2 | 0 | 28 | 53.84 | 107553.1 | 54.08
77.4 | 0 | 28 | 47.49 | 94869.2 | 47.81
78.6 | 0 | 28 | 41.97 | 83848.7 | 41.72
79.9 | 0 | 28 | 35.98 | 71887.5 | 35.63
81.1 | 0 | 28 | 30.44 | 60817 | 30.28
82.3 | 0 | 28 | 25.67 | 51281.6 | 25.31
83.5 | 0 | 28 | 21.29 | 42530.5 | 21.26
84.7 | 0 | 28 | 17.49 | 34940.5 | 17.69
85.9 | 0 | 28 | 14.71 | 29389.3 | 14.66
87.1 | 0 | 28 | 12.26 | 24497.3 | 12.31
88.3 | 0 | 28 | 10.35 | 20673.5 | 10.25
89.5 | 0 | 28 | 8.82 | 17620.9 | 8.67
90.7 | 0 | 28 | 7.43 | 14834.9 | 7.36
91.9 | 0 | 28 | 6.03 | 12044.8 | 6.27
93.1 | 0 | 28 | 5.33 | 10649.3 | 5.31
94.3 | 0 | 28 | 4.54 | 9067.3 | 4.66
95.5 | 0 | 28 | 4.07 | 8131.9 | 4.02
96.7 | 0 | 28 | 3.57 | 7130.6 | 3.48
97.9 | 0 | 28 | 3.01 | 6010.9 | 3.05
99.1 | 0 | 28 | 2.68 | 5359.6 | 2.68
100.3 | 0 | 28 | 2.43 | 4857.4 | 2.38
101.5 | 0 | 28 | 2.04 | 4079.5 | 2.07
102.7 | 0 | 28 | 1.84 | 3685.6 | 1.86
103.9 | 0 | 28 | 1.7 | 3387.1 | 1.69
105.1 | 0 | 28 | 1.42 | 2838.9 | 1.47
106.3 | 0 | 28 | 1.32 | 2632.5 | 1.31
107.5 | 0 | 28 | 1.24 | 2473 | 1.19
108.7 | 0 | 28 | 1.04 | 2069.2 | 1.06
109.9 | 0 | 28 | 0.95 | 1892.2 | 0.95
111.1 | 0 | 28 | 0.9 | 1800.5 | 0.87
112.3 | 0 | 28 | 0.73 | 1466.9 | 0.76
113.5 | 0 | 28 | 0.7 | 1394.6 | 0.69
114.7 | 0 | 28 | 0.61 | 1214.1 | 0.61
115.9 | 0 | 28 | 0.57 | 1146.3 | 0.59
117.1 | 0 | 28 | 0.54 | 1073.3 | 0.54
118.3 | 0 | 28 | 0.52 | 1047.7 | 0.5
119.6 | 0 | 28 | 0.41 | 827.8 | 0.43
120.7 | 0 | 28 | 0.43 | 862.8 | 0.43
122 | 0 | 28 | 0.45 | 905.4 | 0.43
123.2 | 0 | 28 | 0.43 | 850.9 | 0.44
124.4 | 0 | 28 | 0.43 | 868.8 | 0.42
125.6 | 0 | 28 | 0.36 | 724 | 0.37
126.8 | 0 | 28 | 0.35 | 699.7 | 0.36
128 | 0 | 28 | 0.38 | 764 | 0.38
129.2 | 0 | 28 | 0.32 | 641.2 | 0.31
130.4 | 0 | 28 | 0.29 | 581.4 | 0.3
131.6 | 0 | 28 | 0.31 | 610.7 | 0.3
132.8 | 0 | 28 | 0.31 | 626.1 | 0.3
134 | 0 | 28 | 0.25 | 500.5 | 0.26
135.2 | 0 | 28 | 0.28 | 556.9 | 0.28
136.4 | 0 | 28 | 0.31 | 622.8 | 0.3
137.6 | 0 | 28 | 0.27 | 531.3 | 0.28
138.8 | 0 | 28 | 0.28 | 554 | 0.27
140 | 0 | 28 | 0.25 | 491.1 | 0.24
141.2 | 0 | 28 | 0.22 | 443.3 | 0.23
142.4 | 0 | 28 | 0.3 | 589.4 | 0.29
143.6 | 0 | 28 | 0.3 | 597.8 | 0.28
144.8 | 0 | 28 | 0.23 | 460.9 | 0.24
146 | 0 | 28 | 0.25 | 506.3 | 0.25
147.2 | 0 | 28 | 0.27 | 548.8 | 0.27
148.4 | 0 | 28 | 0.22 | 442 | 0.22
149.6 | 0 | 28 | 0.25 | 494.3 | 0.25
150.8 | 0 | 28 | 0.25 | 505.4 | 0.26
152 | 0 | 28 | 0.24 | 470.7 | 0.24
153.2 | 0 | 28 | 0.2 | 396.2 | 0.21
154.4 | 0 | 28 | 0.27 | 547.3 | 0.26
155.6 | 0 | 28 | 0.23 | 451.2 | 0.23
156.8 | 0 | 28 | 0.19 | 382.9 | 0.19
158 | 0 | 28 | 0.23 | 454.7 | 0.22
159.2 | 0 | 28 | 0.22 | 435.1 | 0.22
160.4 | 0 | 28 | 0.2 | 395 | 0.2
161.6 | 0 | 28 | 0.22 | 435.2 | 0.21
162.8 | 0 | 28 | 0.17 | 344 | 0.18
164 | 0 | 28 | 0.19 | 373.2 | 0.19
165.2 | 0 | 28 | 0.22 | 443.9 | 0.21
166.4 | 0 | 28 | 0.22 | 436.7 | 0.22
167.7 | 0 | 28 | 0.18 | 361.8 | 0.18
168.8 | 0 | 28 | 0.17 | 341.3 | 0.17
170 | 0 | 28 | 0.15 | 302.8 | 0.15
171.2 | 0 | 28 | 0.15 | 304.3 | 0.15
172.4 | 0 | 28 | 0.16 | 316.3 | 0.16
173.6 | 0 | 28 | 0.13 | 269.3 | 0.14
174.8 | 0 | 28 | 0.15 | 296.5 | 0.15
176 | 0 | 28 | 0.15 | 308.4 | 0.15
177.2 | 0 | 28 | 0.17 | 343.9 | 0.17
178.4 | 0 | 28 | 0.17 | 337.9 | 0.17
179.6 | 0 | 28 | 0.19 | 372.4 | 0.18
180.8 | 0 | 28 | 0.15 | 300.6 | 0.15
182 | 0 | 28 | 0.2 | 390.1 | 0.2
183.2 | 0 | 28 | 0.16 | 318.1 | 0.16
184.5 | 0 | 28 | 0.18 | 355.4 | 0.18
185.6 | 0 | 28 | 0.2 | 393.4 | 0.19
186.8 | 0 | 28 | 0.2 | 394 | 0.19
188.1 | 0 | 28 | 0.19 | 379.2 | 0.19
189.3 | 0 | 28 | 0.17 | 341.3 | 0.17
190.4 | 0 | 28 | 0.19 | 375.2 | 0.19
191.7 | 0 | 28 | 0.18 | 355.5 | 0.18
192.9 | 0 | 28 | 0.18 | 354.7 | 0.17
194.1 | 0 | 28 | 0.13 | 252 | 0.13
195.3 | 0 | 28 | 0.15 | 291 | 0.15
196.5 | 0 | 28 | 0.14 | 289.3 | 0.14
197.7 | 0 | 28 | 0.12 | 241.3 | 0.12
198.9 | 0 | 28 | 0.12 | 235 | 0.12
200 | 0 | 28 | 0.12 | 238.5 | 0.12