## Rendements détecteurs
Measurement time: | 29/09/2022 19:18:35
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 1.24 | 7053.7 | 1.11
0 | 0 | 149.1 | 1.1 | 6288.6 | 1.07
0 | 0 | 147.6 | 1.25 | 7138.1 | 1.1
0 | 0 | 146.6 | 1.14 | 6513.3 | 1.09
0 | 0 | 145.3 | 1.12 | 6370 | 1.11
0 | 0 | 144.1 | 1.24 | 7059.8 | 1.12
0 | 0 | 142.9 | 1.2 | 6842.5 | 1.14
0 | 0 | 141.7 | 1.26 | 7162.8 | 1.15
0 | 0 | 140.5 | 1.17 | 6650.7 | 1.15
0 | 0 | 139.3 | 1.17 | 6655.2 | 1.16
0 | 0 | 138.1 | 1.32 | 7495.2 | 1.19
0 | 0 | 136.9 | 1.23 | 6990.5 | 1.19
0 | 0 | 135.7 | 1.23 | 7016.4 | 1.18
0 | 0 | 134.5 | 1.33 | 7565.9 | 1.2
0 | 0 | 133.3 | 1.22 | 6943.3 | 1.2
0 | 0 | 132.1 | 1.33 | 7588.3 | 1.22
0 | 0 | 130.9 | 1.3 | 7375.1 | 1.22
0 | 0 | 129.7 | 1.27 | 7216.7 | 1.23
0 | 0 | 128.5 | 1.38 | 7864 | 1.22
0 | 0 | 127.3 | 1.28 | 7276.8 | 1.25
0 | 0 | 126 | 1.29 | 7356.1 | 1.25
0 | 0 | 124.9 | 1.3 | 7425.9 | 1.27
0 | 0 | 123.6 | 1.42 | 8066.2 | 1.29
0 | 0 | 122.4 | 1.38 | 7828.8 | 1.29
0 | 0 | 121.3 | 1.34 | 7606.8 | 1.3
0 | 0 | 120 | 1.42 | 8100 | 1.29
0 | 0 | 118.9 | 1.38 | 7860.3 | 1.32
0 | 0 | 117.6 | 1.36 | 7751.1 | 1.31
0 | 0 | 116.4 | 1.44 | 8198.9 | 1.35
0 | 0 | 115.2 | 1.37 | 7781.6 | 1.34
0 | 0 | 114 | 1.5 | 8516.1 | 1.35
0 | 0 | 112.8 | 1.39 | 7911.4 | 1.35
0 | 0 | 111.6 | 1.44 | 8208.1 | 1.4
0 | 0 | 110.4 | 1.49 | 8459.5 | 1.39
0 | 0 | 109.2 | 1.45 | 8239.2 | 1.42
0 | 0 | 108 | 1.54 | 8778.8 | 1.41
0 | 0 | 106.8 | 1.48 | 8403.7 | 1.43
0 | 0 | 105.6 | 1.48 | 8421.6 | 1.45
0 | 0 | 104.4 | 1.56 | 8884 | 1.45
0 | 0 | 103.2 | 1.52 | 8633.4 | 1.48
0 | 0 | 102 | 1.62 | 9226.3 | 1.49
0 | 0 | 100.8 | 1.53 | 8697.6 | 1.48
0 | 0 | 99.6 | 1.52 | 8636.7 | 1.5
0 | 0 | 98.4 | 1.66 | 9462.9 | 1.52
0 | 0 | 97.2 | 1.57 | 8946.2 | 1.56
0 | 0 | 95.9 | 1.68 | 9536.3 | 1.54
0 | 0 | 94.8 | 1.58 | 8974.1 | 1.53
0 | 0 | 93.5 | 1.6 | 9117.1 | 1.57
0 | 0 | 92.3 | 1.73 | 9821.1 | 1.59
0 | 0 | 91.2 | 1.63 | 9249.9 | 1.62
0 | 0 | 89.9 | 1.73 | 9845.2 | 1.6
0 | 0 | 88.7 | 1.65 | 9366.3 | 1.61
0 | 0 | 87.5 | 1.67 | 9522.9 | 1.65
0 | 0 | 86.3 | 1.82 | 10383 | 1.67
0 | 0 | 85.1 | 1.67 | 9497.1 | 1.66
0 | 0 | 83.9 | 1.81 | 10277.1 | 1.69
0 | 0 | 82.7 | 1.76 | 9996.7 | 1.69
0 | 0 | 81.5 | 1.73 | 9820.7 | 1.69
0 | 0 | 80.3 | 1.84 | 10461.6 | 1.71
0 | 0 | 79.1 | 1.76 | 10022.4 | 1.77
0 | 0 | 77.9 | 1.83 | 10397.9 | 1.73
0 | 0 | 76.7 | 1.8 | 10254.8 | 1.75
0 | 0 | 75.4 | 1.82 | 10378.2 | 1.78
0 | 0 | 74.2 | 1.91 | 10846.9 | 1.78
0 | 0 | 73.1 | 1.83 | 10426.2 | 1.81
0 | 0 | 71.8 | 1.86 | 10606.3 | 1.81
0 | 0 | 70.6 | 1.88 | 10707 | 1.84
0 | 0 | 69.4 | 1.96 | 11140 | 1.88
0 | 0 | 68.2 | 2.11 | 12027 | 1.94
0 | 0 | 67 | 2.06 | 11746.2 | 2.07
0 | 0 | 65.8 | 2.37 | 13495.9 | 2.3
0 | 0 | 64.6 | 2.78 | 15828.8 | 2.74
0 | 0 | 63.4 | 3.63 | 20657 | 3.43
0 | 0 | 62.2 | 4.9 | 27869.8 | 4.52
0 | 0 | 61 | 6.45 | 36712.8 | 6.39
0 | 0 | 59.7 | 8.24 | 46912 | 8.06
0 | 0 | 58.6 | 11.14 | 63397.9 | 10.97
0 | 0 | 57.3 | 15.39 | 87572 | 14.5
0 | 0 | 56.1 | 19.73 | 112317 | 18.16
0 | 0 | 54.9 | 23.56 | 134097.7 | 23.16
0 | 0 | 53.7 | 28.38 | 161511.3 | 27.72
0 | 0 | 52.5 | 33.19 | 188911.6 | 32.4
0 | 0 | 51.3 | 41.7 | 237335.8 | 38.69
0 | 0 | 50.1 | 48.06 | 273529.8 | 43.66 | 43.66 | 50.1
0 | 0 | 48.9 | 50.91 | 289741.6 | 49.68 | 49.68 | 48.9
0 | 0 | 47.7 | 57.53 | 327471.1 | 55.74 | 44.1617 | 50
0 | 0 | 46.5 | 62.01 | 352948.7 | 59.65
0 | 0 | 45.3 | 72.67 | 413639.8 | 66.31
0 | 0 | 44.1 | 77.28 | 439847.2 | 70.47
0 | 0 | 42.9 | 79.14 | 450472.8 | 75.02
0 | 0 | 41.7 | 82.92 | 471942.6 | 79.78
0 | 0 | 40.5 | 86.96 | 494954.2 | 82.54 | 40.5 | 82.54
0 | 0 | 39.2 | 94.94 | 540384.4 | 86.34 | 39.2 | 86.34
0 | 0 | 38.1 | 98.6 | 561191.7 | 89.2 | 39.6584 | 85
0 | 0 | 36.8 | 96.86 | 551312.5 | 91.72
0 | 0 | 35.6 | 99 | 563489 | 94.09
0 | 0 | 34.4 | 102.01 | 580594.8 | 95.7
0 | 0 | 33.2 | 108 | 614686 | 96.78
0 | 0 | 32 | 111.5 | 634613.4 | 98.34
0 | 0 | 30.8 | 103.76 | 590566.5 | 99.05
0 | 0 | 29.6 | 113.27 | 644725.4 | 99.75
0 | 0 | 28.4 | 107.98 | 614622.2 | 99.92
0 | 0 | 27.2 | 111.85 | 636599.7 | 100
0 | 0 | 26 | 113.89 | 648235.3 | 100
0 | 0 | 24.8 | 105.99 | 603240 | 100
0 | 0 | 23.6 | 112.96 | 642938 | 99.73
0 | 0 | 22.3 | 106.71 | 607350.3 | 99.21
0 | 0 | 21.1 | 111.46 | 634379.1 | 98.96
0 | 0 | 19.9 | 112.95 | 642855.8 | 98.41
0 | 0 | 18.7 | 104.1 | 592538.5 | 98.11
0 | 0 | 17.5 | 111.43 | 634228.9 | 97.59
0 | 0 | 16.3 | 105.04 | 597872.9 | 97.12
0 | 0 | 15.1 | 111.25 | 633205.7 | 96.61
0 | 0 | 13.9 | 111.55 | 634920.3 | 96.14
0 | 0 | 12.7 | 102.78 | 584993.3 | 95.56
0 | 0 | 11.5 | 107.88 | 613998.7 | 95.21
0 | 0 | 10.3 | 106.39 | 605570.4 | 94.61
0 | 0 | 9.1 | 110.57 | 629313.5 | 93.95
0 | 0 | 7.9 | 110.18 | 627098.9 | 93.71
0 | 0 | 6.7 | 102.13 | 581302.6 | 93.1
0 | 0 | 5.4 | 105.35 | 599601.3 | 92.42
0 | 0 | 4.2 | 102.76 | 584904.4 | 91.59
0 | 0 | 3 | 108.8 | 619288.9 | 90.85
0 | 0 | 1.8 | 100.12 | 569879.3 | 90.04
0 | 0 | 0.6 | 98.99 | 563413.5 | 89.42
0 | 0 | -0.5 | 105.85 | 602447 | 89.15
Measurement time: | 02/09/2022 18:34:57
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 1.04 | 2038.3 | 0.7
0 | 0 | 149.3 | 0.95 | 1863.1 | 0.95
0 | 0 | 147.9 | 0.97 | 1887.5 | 1
0 | 0 | 146.7 | 0.94 | 1839.7 | 0.95
0 | 0 | 145.6 | 0.97 | 1892.5 | 1
0 | 0 | 144.3 | 1.01 | 1983.2 | 1
0 | 0 | 143.1 | 0.95 | 1854.5 | 0.98
0 | 0 | 141.9 | 0.96 | 1873.7 | 0.96
0 | 0 | 140.7 | 0.99 | 1938.2 | 0.99
0 | 0 | 139.5 | 1.01 | 1977.1 | 1.01
0 | 0 | 138.3 | 1.03 | 2023.8 | 1.04
0 | 0 | 137.1 | 1.04 | 2038.4 | 1.05
0 | 0 | 135.9 | 1.07 | 2098.8 | 1.08
0 | 0 | 134.7 | 1.05 | 2059.5 | 1.05
0 | 0 | 133.5 | 1.04 | 2033.7 | 1.05
0 | 0 | 132.3 | 1.06 | 2064 | 1.09
0 | 0 | 131.1 | 1.14 | 2236.3 | 1.14
0 | 0 | 129.9 | 1.09 | 2131.7 | 1.1
0 | 0 | 128.7 | 1.09 | 2135.5 | 1.13
0 | 0 | 127.5 | 1.18 | 2306.9 | 1.17
0 | 0 | 126.3 | 1.13 | 2205.2 | 1.14
0 | 0 | 125.1 | 1.14 | 2220.6 | 1.18
0 | 0 | 123.9 | 1.15 | 2247.8 | 1.14
0 | 0 | 122.7 | 1.2 | 2340.9 | 1.2
0 | 0 | 121.5 | 1.11 | 2180.2 | 1.17
0 | 0 | 120.3 | 1.17 | 2285.3 | 1.17
0 | 0 | 119.1 | 1.17 | 2280.7 | 1.17
0 | 0 | 117.9 | 1.19 | 2331.2 | 1.2
0 | 0 | 116.7 | 1.2 | 2351.8 | 1.22
0 | 0 | 115.5 | 1.21 | 2374.7 | 1.23
0 | 0 | 114.3 | 1.24 | 2423.1 | 1.26
0 | 0 | 113.1 | 1.22 | 2393.2 | 1.24
0 | 0 | 111.9 | 1.21 | 2372.1 | 1.23
0 | 0 | 110.7 | 1.26 | 2464.1 | 1.28
0 | 0 | 109.5 | 1.25 | 2448.3 | 1.27
0 | 0 | 108.3 | 1.23 | 2403.2 | 1.24
0 | 0 | 107.1 | 1.29 | 2523.7 | 1.3
0 | 0 | 105.9 | 1.29 | 2526.6 | 1.3
0 | 0 | 104.7 | 1.33 | 2598.1 | 1.34
0 | 0 | 103.5 | 1.32 | 2577.6 | 1.33
0 | 0 | 102.3 | 1.36 | 2658.6 | 1.36
0 | 0 | 101.1 | 1.36 | 2657 | 1.38
0 | 0 | 99.9 | 1.34 | 2613.5 | 1.35
0 | 0 | 98.7 | 1.31 | 2570.5 | 1.39
0 | 0 | 97.5 | 1.39 | 2726.6 | 1.41
0 | 0 | 96.3 | 1.42 | 2774.4 | 1.43
0 | 0 | 95.1 | 1.44 | 2822.9 | 1.45
0 | 0 | 93.9 | 1.43 | 2792.5 | 1.45
0 | 0 | 92.7 | 1.41 | 2749.4 | 1.48
0 | 0 | 91.5 | 1.46 | 2854.8 | 1.46
0 | 0 | 90.3 | 1.48 | 2896.9 | 1.5
0 | 0 | 89.1 | 1.42 | 2786 | 1.49
0 | 0 | 87.9 | 1.45 | 2834.8 | 1.53
0 | 0 | 86.6 | 1.54 | 3003.9 | 1.56
0 | 0 | 85.4 | 1.47 | 2872.1 | 1.53
0 | 0 | 84.2 | 1.57 | 3073.7 | 1.58
0 | 0 | 83 | 1.59 | 3104.1 | 1.61
0 | 0 | 81.8 | 1.59 | 3111.7 | 1.61
0 | 0 | 80.6 | 1.5 | 2937.7 | 1.6
0 | 0 | 79.4 | 1.63 | 3177.8 | 1.63
0 | 0 | 78.2 | 1.59 | 3100.3 | 1.62
0 | 0 | 77 | 1.62 | 3167.6 | 1.64
0 | 0 | 75.8 | 1.63 | 3180.1 | 1.67
0 | 0 | 74.6 | 1.66 | 3250.7 | 1.7
0 | 0 | 73.4 | 1.7 | 3332.9 | 1.75
0 | 0 | 72.2 | 1.73 | 3377.6 | 1.77
0 | 0 | 71 | 1.73 | 3387.7 | 1.78
0 | 0 | 69.7 | 1.8 | 3516.1 | 1.83
0 | 0 | 68.5 | 1.9 | 3716.1 | 1.94
0 | 0 | 67.3 | 2.14 | 4188.2 | 2.18
0 | 0 | 66.1 | 2.48 | 4852.6 | 2.54
0 | 0 | 64.9 | 3.07 | 6009.5 | 3.14
0 | 0 | 63.7 | 4.08 | 7978.7 | 4.1
0 | 0 | 62.5 | 5.28 | 10318.5 | 5.6
0 | 0 | 61.3 | 7.06 | 13804.1 | 7.36
0 | 0 | 60.1 | 9.76 | 19088.6 | 10.02
0 | 0 | 58.9 | 13.05 | 25517.2 | 13.06
0 | 0 | 57.7 | 15.99 | 31270.9 | 16.66
0 | 0 | 56.5 | 20.43 | 39943.1 | 20.85
0 | 0 | 55.3 | 25.44 | 49743.5 | 25.6
0 | 0 | 54.1 | 30.58 | 59799.4 | 30.83
0 | 0 | 52.9 | 36.02 | 70431.6 | 36.1
0 | 0 | 51.7 | 39.18 | 76612.7 | 41.68
0 | 0 | 50.5 | 47.15 | 92197.7 | 47.68 | 50.5 | 47.68
0 | 0 | 49.3 | 52.37 | 102412.4 | 53.21 | 49.3 | 53.21
0 | 0 | 48.1 | 58.01 | 113434.2 | 58.65 |  | 50
0 | 0 | 46.9 | 63.22 | 123622 | 64.29
0 | 0 | 45.7 | 68.06 | 133091.5 | 69.1
0 | 0 | 44.5 | 74.31 | 145322.1 | 74.18
0 | 0 | 43.2 | 78.39 | 153290.1 | 78.37
0 | 0 | 42.1 | 82.87 | 162062.6 | 82.39 | 42.1 | 82.39
0 | 0 | 40.8 | 87.66 | 171421.8 | 86.1 | 40.8 | 86.1
0 | 0 | 39.6 | 87.68 | 171462.3 | 88.93 | 41.1854 | 85
0 | 0 | 38.4 | 92.85 | 181574.5 | 91.73
0 | 0 | 37.2 | 96.83 | 189347.2 | 94.06
0 | 0 | 36 | 99.41 | 194409.3 | 95.76
0 | 0 | 34.8 | 96.08 | 187892.1 | 97.1
0 | 0 | 33.6 | 101.53 | 198537.3 | 98.33
0 | 0 | 32.4 | 103.36 | 202126.6 | 99.1
0 | 0 | 31.2 | 104.12 | 203602.4 | 99.65
0 | 0 | 30 | 99.52 | 194620.3 | 99.92
0 | 0 | 28.8 | 104.46 | 204270.3 | 100
0 | 0 | 27.6 | 103.94 | 203264.3 | 99.98
0 | 0 | 26.4 | 104.41 | 204184.6 | 99.75
0 | 0 | 25.2 | 103.71 | 202800.6 | 99.29
0 | 0 | 24 | 102.16 | 199783 | 98.81
0 | 0 | 22.8 | 102.02 | 199501.4 | 98.5
0 | 0 | 21.5 | 103.79 | 202962.3 | 97.91
0 | 0 | 20.3 | 99.71 | 194982.6 | 97.16
0 | 0 | 19.2 | 102.2 | 199861.4 | 96.64
0 | 0 | 18 | 103.2 | 201818 | 95.75
0 | 0 | 16.7 | 101.1 | 197701 | 95.43
0 | 0 | 15.5 | 100.47 | 196468.9 | 94.88
0 | 0 | 14.3 | 102.19 | 199832.5 | 94.22
0 | 0 | 13.1 | 98.39 | 192401 | 93.63
0 | 0 | 11.9 | 97.15 | 189988.8 | 93.14
0 | 0 | 10.7 | 100.68 | 196893.2 | 92.45
0 | 0 | 9.5 | 94.49 | 184774.8 | 91.96
0 | 0 | 8.3 | 97.56 | 190780.5 | 91.44
0 | 0 | 7.1 | 99.6 | 194764 | 90.83
0 | 0 | 5.9 | 95.35 | 186461.1 | 89.81
0 | 0 | 4.7 | 93.36 | 182564.5 | 89.12
0 | 0 | 3.5 | 93.19 | 182235.8 | 88.24
0 | 0 | 2.3 | 93.86 | 183551.4 | 86.82
0 | 0 | 1.1 | 88.68 | 173421.3 | 85.97
0 | 0 | -0.5 | 92.94 | 181753.5 | 85.28