## Profil 15x15 6 MeV
Measurement time: | 02/09/2022 19:43:02
Radiation device: | Clinac3 Accelerator
Energy: | 6 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
176 | 0 | 13 | 1.74 | 2546.6 | 1.58
175.3 | 0 | 13 | 1.68 | 2467.9 | 1.98
173.9 | 0 | 13 | 1.7 | 2485.9 | 2.01
172.8 | 0 | 13 | 1.66 | 2432.5 | 1.97
171.6 | 0 | 13 | 1.68 | 2463.7 | 1.93
170.4 | 0 | 13 | 1.72 | 2513.9 | 2.05
169.1 | 0 | 13 | 1.68 | 2460.1 | 1.99
168 | 0 | 13 | 1.72 | 2515.3 | 1.98
166.8 | 0 | 13 | 1.69 | 2469.5 | 2.01
165.5 | 0 | 13 | 1.67 | 2439.3 | 1.98
164.3 | 0 | 13 | 1.68 | 2454.4 | 1.93
163.1 | 0 | 13 | 1.7 | 2489.9 | 2.03
161.9 | 0 | 13 | 1.68 | 2457.5 | 1.98
160.7 | 0 | 13 | 1.71 | 2505.8 | 1.98
159.5 | 0 | 13 | 1.73 | 2527.9 | 2.09
158.3 | 0 | 13 | 1.71 | 2505.5 | 1.98
157.1 | 0 | 13 | 1.71 | 2504.3 | 2.01
155.9 | 0 | 13 | 1.72 | 2524.3 | 2
154.7 | 0 | 13 | 1.67 | 2439.6 | 1.9
153.5 | 0 | 13 | 1.68 | 2462.1 | 1.98
152.3 | 0 | 13 | 1.74 | 2544.1 | 2.08
151.1 | 0 | 13 | 1.69 | 2472.5 | 1.96
149.9 | 0 | 13 | 1.73 | 2537.9 | 2.04
148.7 | 0 | 13 | 1.71 | 2501.8 | 2.05
147.5 | 0 | 13 | 1.73 | 2529.1 | 1.99
146.3 | 0 | 13 | 1.74 | 2555.7 | 2.06
145.1 | 0 | 13 | 1.71 | 2501.3 | 2.04
143.9 | 0 | 13 | 1.72 | 2515.3 | 1.98
142.7 | 0 | 13 | 1.74 | 2547.1 | 2.06
141.5 | 0 | 13 | 1.73 | 2538.6 | 2.06
140.3 | 0 | 13 | 1.71 | 2500.6 | 1.95
139.1 | 0 | 13 | 1.7 | 2493.9 | 2.02
137.9 | 0 | 13 | 1.69 | 2476.4 | 2.02
136.7 | 0 | 13 | 1.69 | 2479.3 | 1.93
135.5 | 0 | 13 | 1.65 | 2423.2 | 1.96
134.3 | 0 | 13 | 1.67 | 2449.8 | 1.99
133 | 0 | 13 | 1.7 | 2486 | 1.96
131.9 | 0 | 13 | 1.67 | 2452.3 | 1.98
130.7 | 0 | 13 | 1.63 | 2384 | 1.94
129.5 | 0 | 13 | 1.68 | 2458.7 | 1.94
128.3 | 0 | 13 | 1.62 | 2374.8 | 1.91
127 | 0 | 13 | 1.63 | 2386.6 | 1.96
125.9 | 0 | 13 | 1.62 | 2368.1 | 1.86
124.7 | 0 | 13 | 1.64 | 2395.9 | 1.95
123.5 | 0 | 13 | 1.66 | 2438.1 | 1.99
122.3 | 0 | 13 | 1.6 | 2341.6 | 1.82
121 | 0 | 13 | 1.6 | 2347 | 1.91
119.9 | 0 | 13 | 1.57 | 2299.1 | 1.88
118.7 | 0 | 13 | 1.55 | 2269.3 | 1.77
117.4 | 0 | 13 | 1.6 | 2345 | 1.92
116.3 | 0 | 13 | 1.54 | 2253 | 1.81
115 | 0 | 13 | 1.52 | 2219.5 | 1.74
113.9 | 0 | 13 | 1.47 | 2159.1 | 1.74
112.6 | 0 | 13 | 1.44 | 2112.9 | 1.72
111.4 | 0 | 13 | 1.44 | 2112.9 | 1.65
110.3 | 0 | 13 | 1.44 | 2111.6 | 1.7
109 | 0 | 13 | 1.44 | 2108.4 | 1.72
107.9 | 0 | 13 | 1.33 | 1954.7 | 1.54
106.6 | 0 | 13 | 1.35 | 1979.2 | 1.6
105.4 | 0 | 13 | 1.34 | 1956.3 | 1.61
104.3 | 0 | 13 | 1.24 | 1822.5 | 1.41
103 | 0 | 13 | 1.25 | 1831.7 | 1.45
101.8 | 0 | 13 | 1.17 | 1712.1 | 1.38
100.6 | 0 | 13 | 1.1 | 1618.5 | 1.26
99.4 | 0 | 13 | 1.01 | 1474.1 | 1.16
98.3 | 0 | 13 | 0.91 | 1336.1 | 1.08
97 | 0 | 13 | 0.72 | 1049.2 | 0.82
95.8 | 0 | 13 | 0.43 | 627.7 | 0.49
94.6 | 0 | 13 | 0.16 | 241.4 | 0.2
93.4 | 0 | 13 | 0.38 | 561.1 | 0.44
92.2 | 0 | 13 | 0.91 | 1331.9 | 1.07
91 | 0 | 13 | 1.58 | 2309.6 | 1.91
89.8 | 0 | 13 | 2.84 | 4156.8 | 3.26
88.6 | 0 | 13 | 3.97 | 5818.3 | 4.7
87.4 | 0 | 13 | 5.5 | 8055.9 | 6.62
86.2 | 0 | 13 | 7.86 | 11511.6 | 9.09
85 | 0 | 13 | 10.3 | 15086.4 | 12.19
83.8 | 0 | 13 | 13.37 | 19583.6 | 15.91
82.6 | 0 | 13 | 18.32 | 26839.9 | 20.95
81.4 | 0 | 13 | 22.86 | 33479.7 | 27.25
80.2 | 0 | 13 | 29.08 | 42590.3 | 34.48
79 | 0 | 13 | 37.22 | 54515.5 | 43.29
77.8 | 0 | 13 | 44.86 | 65705.5 | 53.14
76.6 | 0 | 13 | 53.95 | 79026.3 | 64.14
75.4 | 0 | 13 | 66.4 | 97267.1 | 75.97
74.2 | 0 | 13 | 73.61 | 107829.4 | 87.19
73 | 0 | 13 | 81.01 | 118671.2 | 97.92
71.8 | 0 | 13 | 93.47 | 136920.6 | 107.28
70.6 | 0 | 13 | 97.39 | 142655.8 | 115.43
69.4 | 0 | 13 | 102.46 | 150078.6 | 122.57
68.2 | 0 | 13 | 110.39 | 161698.1 | 128.13
67 | 0 | 13 | 113.68 | 166522.3 | 132.85
65.8 | 0 | 13 | 114.32 | 167462.9 | 135.86
64.6 | 0 | 13 | 119.58 | 175156 | 138.86
63.4 | 0 | 13 | 119.58 | 175168.3 | 141.16
62.2 | 0 | 13 | 121.77 | 178368.2 | 142.81
61 | 0 | 13 | 124.34 | 182133.2 | 144.45
59.8 | 0 | 13 | 124.13 | 181829 | 145.94
58.6 | 0 | 13 | 123.54 | 180965.9 | 146.68
57.4 | 0 | 13 | 125.35 | 183618.1 | 147.37
56.2 | 0 | 13 | 124.14 | 181840.1 | 148.1
54.9 | 0 | 13 | 128.42 | 188112.9 | 148.74
53.7 | 0 | 13 | 127.28 | 186442.2 | 149.27
52.5 | 0 | 13 | 125.37 | 183646.1 | 149.61
51.3 | 0 | 13 | 131.25 | 192260.5 | 150.33
50.1 | 0 | 13 | 127.98 | 187469 | 150.54
48.9 | 0 | 13 | 125.45 | 183768.4 | 150.78
47.7 | 0 | 13 | 133.19 | 195094.8 | 151.46
46.5 | 0 | 13 | 127.25 | 186393.9 | 151.37
45.3 | 0 | 13 | 126.36 | 185087.3 | 151.51
44.1 | 0 | 13 | 132.72 | 194404.2 | 151.92
42.9 | 0 | 13 | 128.17 | 187741.1 | 151.84
41.7 | 0 | 13 | 133.66 | 195790.8 | 152.55
40.5 | 0 | 13 | 127.59 | 186894 | 152.33
39.3 | 0 | 13 | 127.45 | 186696.6 | 152.72
38.1 | 0 | 13 | 129.8 | 190127.4 | 153.09
36.9 | 0 | 13 | 131.77 | 193014.3 | 152.83
35.7 | 0 | 13 | 128.98 | 188934 | 152.94
34.5 | 0 | 13 | 129.06 | 189055 | 153.08
33.3 | 0 | 13 | 134.06 | 196380.1 | 153.35
32.1 | 0 | 13 | 128.25 | 187856.9 | 152.77
30.9 | 0 | 13 | 128.87 | 188765.7 | 153.26
29.7 | 0 | 13 | 134.16 | 196519.5 | 153.81
28.5 | 0 | 13 | 129.41 | 189558.7 | 153.68
27.3 | 0 | 13 | 128.99 | 188949 | 153.42
26.1 | 0 | 13 | 135.19 | 198023.6 | 154.27
24.9 | 0 | 13 | 128.99 | 188952.1 | 154.08
23.7 | 0 | 13 | 129.42 | 189570 | 153.93
22.5 | 0 | 13 | 135.59 | 198615.7 | 154.53
21.3 | 0 | 13 | 129.13 | 189148.9 | 154.13
20.1 | 0 | 13 | 129.61 | 189852.6 | 154.21
18.9 | 0 | 13 | 134.46 | 196955.1 | 154.48
17.7 | 0 | 13 | 130.88 | 191720.5 | 154.36
16.5 | 0 | 13 | 129.27 | 189361.6 | 154.27
15.3 | 0 | 13 | 134.23 | 196617.1 | 154.57
14.1 | 0 | 13 | 130.68 | 191415.9 | 154.5
12.9 | 0 | 13 | 129.4 | 189554.2 | 154.47
11.7 | 0 | 13 | 135.36 | 198271 | 154.63
10.5 | 0 | 13 | 129.44 | 189607.8 | 154.25
9.3 | 0 | 13 | 129.89 | 190260.6 | 154.34
8.1 | 0 | 13 | 132.99 | 194809.1 | 154.51
6.9 | 0 | 13 | 131.83 | 193102.8 | 154.33
5.7 | 0 | 13 | 129.76 | 190074.4 | 154.1
4.5 | 0 | 13 | 133.27 | 195214 | 154.49
3.2 | 0 | 13 | 129.91 | 190290.4 | 154.18
2.1 | 0 | 13 | 133.94 | 196195 | 154.67
0.8 | 0 | 13 | 130.08 | 190538.5 | 154.44
-0.3 | 0 | 13 | 128.77 | 188619.8 | 154.24
-1.6 | 0 | 13 | 130.92 | 191771.1 | 154.26
-2.8 | 0 | 13 | 134.66 | 197258.2 | 154.72
-4 | 0 | 13 | 127.31 | 186486.9 | 154.52
-5.2 | 0 | 13 | 132.35 | 193873.7 | 154.37
-6.3 | 0 | 13 | 133.13 | 195008.9 | 154.66
-7.6 | 0 | 13 | 128.99 | 188949.1 | 154.45
-8.8 | 0 | 13 | 133.05 | 194894.4 | 154.52
-10 | 0 | 13 | 132.28 | 193765.9 | 154.65
-11.2 | 0 | 13 | 130.43 | 191058.1 | 154.52
-12.4 | 0 | 13 | 132.49 | 194076.4 | 154.41
-13.6 | 0 | 13 | 132.05 | 193433.6 | 154.53
-14.8 | 0 | 13 | 130.49 | 191139.3 | 154.38
-16 | 0 | 13 | 132.71 | 194393.6 | 154.47
-17.2 | 0 | 13 | 132.6 | 194230.4 | 154.35
-18.4 | 0 | 13 | 127.48 | 186736 | 154.14
-19.6 | 0 | 13 | 135.61 | 198638.4 | 154.41
-20.8 | 0 | 13 | 128.92 | 188838 | 154.07
-22 | 0 | 13 | 128.56 | 188312.3 | 153.74
-23.2 | 0 | 13 | 133.47 | 195505.8 | 154
-24.4 | 0 | 13 | 131.42 | 192504.7 | 154.13
-25.6 | 0 | 13 | 129.09 | 189096 | 153.7
-26.8 | 0 | 13 | 133.9 | 196143.5 | 154.06
-28 | 0 | 13 | 131.07 | 191994.2 | 154.02
-29.2 | 0 | 13 | 128.56 | 188310.6 | 153.38
-30.4 | 0 | 13 | 133.75 | 195917.2 | 153.67
-31.6 | 0 | 13 | 129.7 | 189984.8 | 153.48
-32.8 | 0 | 13 | 128.7 | 188522 | 153.65
-34 | 0 | 13 | 134.84 | 197517.1 | 153.68
-35.2 | 0 | 13 | 129.37 | 189500.3 | 153.42
-36.4 | 0 | 13 | 128.11 | 187663.1 | 152.82
-37.6 | 0 | 13 | 134.34 | 196783.2 | 153.26
-38.8 | 0 | 13 | 127.21 | 186337.1 | 152.85
-40 | 0 | 13 | 128.52 | 188253.5 | 152.8
-41.2 | 0 | 13 | 133.52 | 195576.1 | 152.66
-42.4 | 0 | 13 | 127.47 | 186724.3 | 152.26
-43.6 | 0 | 13 | 127.21 | 186339.8 | 152.41
-44.8 | 0 | 13 | 132.52 | 194125 | 152.37
-46 | 0 | 13 | 128.69 | 188509.4 | 151.99
-47.2 | 0 | 13 | 126.26 | 184947.8 | 151.19
-48.4 | 0 | 13 | 132.73 | 194425.4 | 151.7
-49.6 | 0 | 13 | 127.23 | 186368.8 | 151.08
-50.8 | 0 | 13 | 126.24 | 184925.8 | 150.54
-52 | 0 | 13 | 131.73 | 192967.1 | 150.91
-53.2 | 0 | 13 | 126.49 | 185281.9 | 150.16
-54.4 | 0 | 13 | 125.88 | 184394.1 | 149.43
-55.7 | 0 | 13 | 131.37 | 192428.5 | 149.77
-56.9 | 0 | 13 | 125.46 | 183777.1 | 149.02
-58 | 0 | 13 | 124.2 | 181927.5 | 148.08
-59.3 | 0 | 13 | 129.4 | 189549.2 | 147.88
-60.4 | 0 | 13 | 124.16 | 181876.4 | 147.19
-61.7 | 0 | 13 | 121.33 | 177719.6 | 145.66
-62.9 | 0 | 13 | 127.24 | 186386 | 145.38
-64.1 | 0 | 13 | 121.23 | 177573.4 | 143.83
-65.3 | 0 | 13 | 117.15 | 171602.9 | 141.81
-66.5 | 0 | 13 | 121.79 | 178399.4 | 140.52
-67.7 | 0 | 13 | 117.3 | 171821.4 | 137.92
-68.9 | 0 | 13 | 112.59 | 164920.2 | 134.37
-70.1 | 0 | 13 | 112.37 | 164599.7 | 131.06
-71.3 | 0 | 13 | 107.79 | 157898.4 | 125.68
-72.5 | 0 | 13 | 100.82 | 147682.6 | 119.54
-73.7 | 0 | 13 | 96.5 | 141361.6 | 111.9
-74.9 | 0 | 13 | 86.24 | 126330.1 | 102.45
-76.1 | 0 | 13 | 79.48 | 116428.2 | 92.17
-77.3 | 0 | 13 | 70.16 | 102770.9 | 81.08
-78.5 | 0 | 13 | 57.57 | 84324.6 | 69.62
-79.7 | 0 | 13 | 50.54 | 74024.9 | 58.39
-80.9 | 0 | 13 | 41.32 | 60520.4 | 48.16
-82.1 | 0 | 13 | 33.42 | 48956.8 | 38.6
-83.3 | 0 | 13 | 25.28 | 37033 | 30.37
-84.5 | 0 | 13 | 20.59 | 30164.9 | 24.12
-85.7 | 0 | 13 | 16.15 | 23654.8 | 18.59
-86.9 | 0 | 13 | 11.73 | 17184 | 14.06
-88.1 | 0 | 13 | 9.01 | 13195.9 | 10.61
-89.3 | 0 | 13 | 6.58 | 9642.2 | 7.64
-90.5 | 0 | 13 | 4.75 | 6959.2 | 5.58
-91.7 | 0 | 13 | 3.15 | 4621.4 | 3.75
-92.9 | 0 | 13 | 2.19 | 3209 | 2.52
-94.1 | 0 | 13 | 1.17 | 1718.3 | 1.4
-95.3 | 0 | 13 | 0.49 | 716.1 | 0.58
-96.5 | 0 | 13 | 0.11 | 162.4 | 0.13
-97.7 | 0 | 13 | 0.34 | 495.3 | 0.4
-98.9 | 0 | 13 | 0.55 | 806.9 | 0.65
-100.1 | 0 | 13 | 0.71 | 1044.4 | 0.81
-101.4 | 0 | 13 | 0.86 | 1261.1 | 1.04
-102.5 | 0 | 13 | 1 | 1460.5 | 1.17
-103.8 | 0 | 13 | 1 | 1463.9 | 1.15
-105 | 0 | 13 | 1.1 | 1611.8 | 1.31
-106.2 | 0 | 13 | 1.12 | 1640.5 | 1.32
-107.4 | 0 | 13 | 1.18 | 1727.5 | 1.36
-108.6 | 0 | 13 | 1.21 | 1777.8 | 1.45
-109.8 | 0 | 13 | 1.23 | 1805 | 1.46
-111 | 0 | 13 | 1.27 | 1856.1 | 1.46
-112.2 | 0 | 13 | 1.25 | 1830.9 | 1.48
-113.4 | 0 | 13 | 1.23 | 1794.7 | 1.46
-114.6 | 0 | 13 | 1.35 | 1981 | 1.55
-115.8 | 0 | 13 | 1.34 | 1966.6 | 1.62
-117 | 0 | 13 | 1.29 | 1891.8 | 1.47
-118.2 | 0 | 13 | 1.34 | 1964.7 | 1.59
-119.4 | 0 | 13 | 1.33 | 1945.2 | 1.53
-120.6 | 0 | 13 | 1.33 | 1947.9 | 1.51
-121.8 | 0 | 13 | 1.33 | 1950.5 | 1.59
-123 | 0 | 13 | 1.36 | 1995.7 | 1.56
-124.2 | 0 | 13 | 1.37 | 2003 | 1.63
-125.4 | 0 | 13 | 1.43 | 2095 | 1.65
-126.6 | 0 | 13 | 1.44 | 2105 | 1.7
-127.8 | 0 | 13 | 1.48 | 2173.4 | 1.77
-129 | 0 | 13 | 1.49 | 2189.7 | 1.72
-130.2 | 0 | 13 | 1.52 | 2230.5 | 1.81
-131.4 | 0 | 13 | 1.49 | 2187.5 | 1.71
-132.6 | 0 | 13 | 1.52 | 2223.6 | 1.79
-133.8 | 0 | 13 | 1.44 | 2103.5 | 1.65
-135 | 0 | 13 | 1.52 | 2225.4 | 1.73
-136.2 | 0 | 13 | 1.51 | 2216.1 | 1.8
-137.4 | 0 | 13 | 1.51 | 2205.1 | 1.74
-138.6 | 0 | 13 | 1.5 | 2191.4 | 1.76
-139.8 | 0 | 13 | 1.48 | 2173.2 | 1.71
-141 | 0 | 13 | 1.47 | 2159.7 | 1.76
-142.2 | 0 | 13 | 1.41 | 2068.1 | 1.67
-143.4 | 0 | 13 | 1.44 | 2112.3 | 1.66
-144.6 | 0 | 13 | 1.37 | 2010.1 | 1.65
-145.8 | 0 | 13 | 1.36 | 1985.4 | 1.6
-147 | 0 | 13 | 1.31 | 1920.8 | 1.5
-148.2 | 0 | 13 | 1.36 | 1992 | 1.65
-149.4 | 0 | 13 | 1.31 | 1912.4 | 1.53
-150.6 | 0 | 13 | 1.31 | 1914.5 | 1.52
-151.8 | 0 | 13 | 1.3 | 1902 | 1.53
-153 | 0 | 13 | 1.27 | 1864.3 | 1.51
-154.2 | 0 | 13 | 1.26 | 1844.9 | 1.47
-155.4 | 0 | 13 | 1.34 | 1960.4 | 1.58
-156.6 | 0 | 13 | 1.24 | 1815.2 | 1.46
-157.9 | 0 | 13 | 1.27 | 1865.6 | 1.49
-159 | 0 | 13 | 1.34 | 1957.6 | 1.61
-160.2 | 0 | 13 | 1.3 | 1910.6 | 1.52
-161.5 | 0 | 13 | 1.24 | 1821.9 | 1.45
-162.6 | 0 | 13 | 1.34 | 1958.6 | 1.61
-163.9 | 0 | 13 | 1.3 | 1902.1 | 1.51
-165.1 | 0 | 13 | 1.29 | 1896.8 | 1.52
-166.3 | 0 | 13 | 1.32 | 1929.1 | 1.58
-167.5 | 0 | 13 | 1.35 | 1974 | 1.57
-168.7 | 0 | 13 | 1.33 | 1951.4 | 1.56
-169.9 | 0 | 13 | 1.29 | 1890.2 | 1.55
-171.1 | 0 | 13 | 1.35 | 1975.7 | 1.55
-172.3 | 0 | 13 | 1.36 | 1986.2 | 1.6
-173.5 | 0 | 13 | 1.38 | 2020.9 | 1.65
-174.7 | 0 | 13 | 1.35 | 1979.6 | 1.55
-176 | 0 | 13 | 1.38 | 2018 | 1.63