## Profils orientation 6 et 18 MeV
Measurement time: | 02/09/2022 18:46:49
Radiation device: | Clinac3 Accelerator
Energy: | 6 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-150 | 0 | 13 | 0.11 | 207.2 | 0.11
-148.5 | 0 | 13 | 0.07 | 137.3 | 0.07
-147.4 | 0 | 13 | 0.09 | 168.5 | 0.08
-146.2 | 0 | 13 | 0.08 | 163.5 | 0.09
-145 | 0 | 13 | 0.04 | 71.5 | 0.04
-143.7 | 0 | 13 | 0.08 | 156.2 | 0.08
-142.6 | 0 | 13 | 0.03 | 51.5 | 0.03
-141.3 | 0 | 13 | 0.03 | 56 | 0.03
-140.1 | 0 | 13 | 0.07 | 134.4 | 0.07
-138.9 | 0 | 13 | 0.07 | 141.7 | 0.07
-137.7 | 0 | 13 | 0.08 | 147.6 | 0.08
-136.5 | 0 | 13 | 0.08 | 153.9 | 0.08
-135.3 | 0 | 13 | 0.04 | 69.7 | 0.03
-134.1 | 0 | 13 | 0.04 | 68.6 | 0.04
-132.9 | 0 | 13 | 0.03 | 52 | 0.03
-131.7 | 0 | 13 | 0.02 | 36.5 | 0.02
-130.5 | 0 | 13 | 0 | 7 | 0
-129.3 | 0 | 13 | 0.04 | 77.5 | 0.04
-128.1 | 0 | 13 | 0.01 | 22.2 | 0.01
-126.9 | 0 | 13 | 0.01 | 13.9 | 0.01
-125.7 | 0 | 13 | 0.02 | 36.8 | 0.02
-124.5 | 0 | 13 | 0.04 | 73.5 | 0.04
-123.3 | 0 | 13 | 0 | 7.3 | 0
-122.1 | 0 | 13 | 0.02 | 48.6 | 0.03
-120.9 | 0 | 13 | 0.05 | 104.6 | 0.05
-119.7 | 0 | 13 | 0.01 | 25.4 | 0.01
-118.5 | 0 | 13 | 0.02 | 35.8 | 0.02
-117.3 | 0 | 13 | 0.02 | 37.5 | 0.02
-116.1 | 0 | 13 | 0.05 | 88.1 | 0.05
-114.9 | 0 | 13 | 0.03 | 65.5 | 0.03
-113.7 | 0 | 13 | 0.04 | 78.6 | 0.04
-112.5 | 0 | 13 | 0.05 | 92.7 | 0.05
-111.3 | 0 | 13 | 0.08 | 146.9 | 0.08
-110.1 | 0 | 13 | 0.05 | 101 | 0.05
-108.9 | 0 | 13 | 0 | 6.9 | 0
-107.7 | 0 | 13 | 0.01 | 23 | 0.01
-106.5 | 0 | 13 | 0.04 | 83.6 | 0.04
-105.3 | 0 | 13 | 0.04 | 74 | 0.04
-104.1 | 0 | 13 | 0.05 | 95.8 | 0.05
-102.9 | 0 | 13 | 0.06 | 115.6 | 0.06
-101.7 | 0 | 13 | 0.07 | 142.7 | 0.08
-100.5 | 0 | 13 | 0.05 | 91.5 | 0.04
-99.3 | 0 | 13 | 0.04 | 83.4 | 0.04
-98.1 | 0 | 13 | 0.06 | 115.6 | 0.06
-96.8 | 0 | 13 | 0 | 7.4 | 0
-95.6 | 0 | 13 | 0.04 | 75.7 | 0.04
-94.4 | 0 | 13 | 0.03 | 49.5 | 0.03
-93.2 | 0 | 13 | 0 | 6.2 | 0
-92 | 0 | 13 | 0.06 | 120.4 | 0.06
-90.8 | 0 | 13 | 0.01 | 13.6 | 0.01
-89.6 | 0 | 13 | 0.01 | 27.3 | 0.01
-88.4 | 0 | 13 | 0.01 | 10.6 | 0.01
-87.2 | 0 | 13 | 0.04 | 79 | 0.04
-86 | 0 | 13 | 0.01 | 15.9 | 0.01
-84.8 | 0 | 13 | 0.03 | 49.1 | 0.03
-83.6 | 0 | 13 | 0.05 | 91.2 | 0.05
-82.4 | 0 | 13 | 0.1 | 200.1 | 0.1
-81.2 | 0 | 13 | 0.1 | 202.7 | 0.1
-80 | 0 | 13 | 0.15 | 297.1 | 0.15
-78.8 | 0 | 13 | 0.18 | 348.3 | 0.17
-77.6 | 0 | 13 | 0.24 | 476.8 | 0.24
-76.4 | 0 | 13 | 0.29 | 562 | 0.29
-75.2 | 0 | 13 | 0.44 | 859.4 | 0.43
-74 | 0 | 13 | 0.5 | 985.7 | 0.51
-72.8 | 0 | 13 | 0.65 | 1271.4 | 0.66
-71.6 | 0 | 13 | 0.99 | 1929.7 | 0.97
-70.4 | 0 | 13 | 1.26 | 2466 | 1.28
-69.2 | 0 | 13 | 1.72 | 3372.5 | 1.75
-68 | 0 | 13 | 2.46 | 4815.1 | 2.41
-66.8 | 0 | 13 | 3.15 | 6167 | 3.22
-65.6 | 0 | 13 | 4.26 | 8332.1 | 4.32
-64.4 | 0 | 13 | 5.91 | 11549.5 | 5.75
-63.2 | 0 | 13 | 7.38 | 14438 | 7.49
-62 | 0 | 13 | 9.46 | 18495.6 | 9.64
-60.8 | 0 | 13 | 12.82 | 25063.8 | 12.48
-59.6 | 0 | 13 | 15.52 | 30357 | 15.72
-58.4 | 0 | 13 | 19.55 | 38223.4 | 19.9
-57.2 | 0 | 13 | 25.24 | 49367.7 | 24.63
-56 | 0 | 13 | 30.01 | 58679.1 | 30.36
-54.8 | 0 | 13 | 35.86 | 70117.4 | 36.58
-53.6 | 0 | 13 | 44.2 | 86432 | 43.19
-52.3 | 0 | 13 | 49.59 | 96977.6 | 50.37
-51.1 | 0 | 13 | 55.95 | 109405.5 | 57.34
-49.9 | 0 | 13 | 65.77 | 128617 | 64.16
-48.7 | 0 | 13 | 69.19 | 135301.6 | 70.35
-47.5 | 0 | 13 | 74.34 | 145371.3 | 75.35
-46.3 | 0 | 13 | 82.15 | 160646.2 | 80.34
-45.1 | 0 | 13 | 83.15 | 162604.6 | 83.92
-43.9 | 0 | 13 | 85.3 | 166815.9 | 86.78
-42.7 | 0 | 13 | 90.65 | 177274.7 | 89.29
-41.5 | 0 | 13 | 89.28 | 174589.5 | 91.36
-40.3 | 0 | 13 | 93.2 | 182264 | 92.7
-39.1 | 0 | 13 | 94.98 | 185743.6 | 94.1
-37.9 | 0 | 13 | 93.42 | 182690.1 | 95.06
-36.7 | 0 | 13 | 96.34 | 188394.1 | 95.94
-35.5 | 0 | 13 | 96.49 | 188693.5 | 96.52
-34.3 | 0 | 13 | 97.92 | 191482.2 | 97.16
-33.1 | 0 | 13 | 95.48 | 186714.9 | 97.65
-31.9 | 0 | 13 | 98.15 | 191942.2 | 97.85
-30.7 | 0 | 13 | 99.09 | 193783.4 | 98.09
-29.5 | 0 | 13 | 97.42 | 190509.6 | 98.52
-28.3 | 0 | 13 | 98.1 | 191833.3 | 98.75
-27.1 | 0 | 13 | 97.18 | 190039.4 | 98.55
-25.9 | 0 | 13 | 99.93 | 195414 | 98.98
-24.7 | 0 | 13 | 99.9 | 195354.1 | 98.96
-23.5 | 0 | 13 | 97.18 | 190045.4 | 98.96
-22.3 | 0 | 13 | 100.39 | 196308 | 99.11
-21.1 | 0 | 13 | 99.4 | 194375.2 | 99.19
-19.9 | 0 | 13 | 100.14 | 195830.8 | 99.31
-18.7 | 0 | 13 | 98.06 | 191768.1 | 99.5
-17.5 | 0 | 13 | 98.76 | 193126.8 | 99.49
-16.3 | 0 | 13 | 101.19 | 197884.9 | 99.55
-15.1 | 0 | 13 | 97.65 | 190966.9 | 99.44
-13.9 | 0 | 13 | 98.94 | 193489.8 | 99.59
-12.7 | 0 | 13 | 100.98 | 197476.8 | 99.64
-11.5 | 0 | 13 | 98.21 | 192050.4 | 99.51
-10.3 | 0 | 13 | 99.09 | 193778.8 | 99.63
-9.1 | 0 | 13 | 98.04 | 191726.4 | 99.74
-7.9 | 0 | 13 | 101.72 | 198919.5 | 99.71
-6.7 | 0 | 13 | 98.98 | 193560.5 | 99.53
-5.5 | 0 | 13 | 102.02 | 199514.3 | 99.93
-4.3 | 0 | 13 | 97.89 | 191421.7 | 99.92
-3.1 | 0 | 13 | 98.37 | 192376.4 | 99.74
-1.9 | 0 | 13 | 102.76 | 200951.2 | 99.8
-0.6 | 0 | 13 | 97.7 | 191053.8 | 99.86
0.6 | 0 | 13 | 98.67 | 192950.7 | 99.92
1.8 | 0 | 13 | 102.54 | 200513.7 | 99.73
2.9 | 0 | 13 | 97.99 | 191620.3 | 99.94
4.2 | 0 | 13 | 98.06 | 191762.5 | 99.93
5.3 | 0 | 13 | 102.92 | 201262.1 | 99.88
6.5 | 0 | 13 | 97.96 | 191574.8 | 99.94
7.8 | 0 | 13 | 98.69 | 192983.7 | 100
9 | 0 | 13 | 98.12 | 191883.1 | 99.8
10.2 | 0 | 13 | 102.64 | 200713 | 99.74
11.4 | 0 | 13 | 98.53 | 192674.3 | 99.56
12.5 | 0 | 13 | 98.53 | 192685.6 | 99.94
13.8 | 0 | 13 | 101.64 | 198753.9 | 99.55
15 | 0 | 13 | 97.9 | 191457.1 | 99.55
16.2 | 0 | 13 | 97.6 | 190862.4 | 99.73
17.4 | 0 | 13 | 102.33 | 200104.5 | 99.52
18.6 | 0 | 13 | 98.35 | 192319.7 | 99.35
19.8 | 0 | 13 | 101.89 | 199252.2 | 99.58
21 | 0 | 13 | 97.49 | 190652 | 99.52
22.2 | 0 | 13 | 99.21 | 194009.9 | 99.47
23.4 | 0 | 13 | 97.39 | 190456.5 | 99.67
24.6 | 0 | 13 | 101.32 | 198127.3 | 99.53
25.8 | 0 | 13 | 98.58 | 192777.9 | 99.28
27 | 0 | 13 | 98.17 | 191979.4 | 99.02
28.2 | 0 | 13 | 101.08 | 197676 | 99
29.4 | 0 | 13 | 98.18 | 191995.4 | 98.7
30.6 | 0 | 13 | 97.3 | 190281.8 | 98.5
31.8 | 0 | 13 | 100.14 | 195826 | 98.07
33 | 0 | 13 | 96.57 | 188845.9 | 97.55
34.2 | 0 | 13 | 98.96 | 193529.4 | 96.84
35.4 | 0 | 13 | 94.52 | 184836.2 | 96.17
36.6 | 0 | 13 | 93.9 | 183624.8 | 95.24
37.8 | 0 | 13 | 95.15 | 186072.5 | 93.9
39 | 0 | 13 | 90.32 | 176634.7 | 92.07
40.2 | 0 | 13 | 89.18 | 174389.2 | 89.89
41.4 | 0 | 13 | 88.79 | 173632.7 | 87.54
42.6 | 0 | 13 | 81.97 | 160305.7 | 83.98
43.8 | 0 | 13 | 79.92 | 156288.2 | 80.15
45 | 0 | 13 | 76.97 | 150515.7 | 75.68
46.2 | 0 | 13 | 67.98 | 132942.3 | 70
47.4 | 0 | 13 | 64.01 | 125167.8 | 63.52
48.6 | 0 | 13 | 57.37 | 112191.8 | 56.96
49.8 | 0 | 13 | 49.12 | 96064.9 | 49.74
51 | 0 | 13 | 42.75 | 83604.9 | 42.61
52.2 | 0 | 13 | 35.32 | 69069.2 | 35.96
53.4 | 0 | 13 | 29.77 | 58217.8 | 29.56
54.6 | 0 | 13 | 24.17 | 47262.4 | 23.94
55.9 | 0 | 13 | 18.66 | 36490.8 | 19.07
57 | 0 | 13 | 15.16 | 29648.7 | 15.16
58.3 | 0 | 13 | 12.01 | 23483.1 | 11.83
59.5 | 0 | 13 | 9.19 | 17966.7 | 9.22
60.6 | 0 | 13 | 6.98 | 13641.7 | 7.14
61.9 | 0 | 13 | 5.61 | 10963.4 | 5.52
63.1 | 0 | 13 | 4.18 | 8174.1 | 4.23
64.3 | 0 | 13 | 3.09 | 6050.3 | 3.15
65.5 | 0 | 13 | 2.43 | 4757 | 2.38
66.7 | 0 | 13 | 1.7 | 3323.8 | 1.73
67.9 | 0 | 13 | 1.22 | 2393.9 | 1.26
69.1 | 0 | 13 | 0.96 | 1885.6 | 0.94
70.3 | 0 | 13 | 0.64 | 1250.2 | 0.65
71.5 | 0 | 13 | 0.41 | 806.6 | 0.42
72.7 | 0 | 13 | 0.33 | 638.7 | 0.32
73.9 | 0 | 13 | 0.21 | 404.6 | 0.21
75.1 | 0 | 13 | 0.1 | 192.8 | 0.1
76.3 | 0 | 13 | 0.05 | 103.1 | 0.05
77.5 | 0 | 13 | 0.01 | 9.9 | 0.01
78.7 | 0 | 13 | 0.02 | 35.5 | 0.02
79.9 | 0 | 13 | 0.02 | 42 | 0.02
81.1 | 0 | 13 | 0.07 | 146.4 | 0.08
82.3 | 0 | 13 | 0.05 | 104.2 | 0.05
83.5 | 0 | 13 | 0.12 | 241.8 | 0.12
84.7 | 0 | 13 | 0.13 | 244.6 | 0.13
85.9 | 0 | 13 | 0.15 | 293 | 0.15
87.1 | 0 | 13 | 0.13 | 247.5 | 0.12
88.3 | 0 | 13 | 0.14 | 282.3 | 0.15
89.5 | 0 | 13 | 0.22 | 420.5 | 0.22
90.7 | 0 | 13 | 0.16 | 315.5 | 0.16
91.9 | 0 | 13 | 0.18 | 361.8 | 0.19
93.1 | 0 | 13 | 0.17 | 331 | 0.17
94.4 | 0 | 13 | 0.17 | 327.6 | 0.16
95.5 | 0 | 13 | 0.19 | 379.4 | 0.2
96.7 | 0 | 13 | 0.23 | 454.8 | 0.24
97.9 | 0 | 13 | 0.19 | 371 | 0.19
99.2 | 0 | 13 | 0.17 | 331.2 | 0.17
100.4 | 0 | 13 | 0.22 | 431.2 | 0.23
101.5 | 0 | 13 | 0.19 | 380.5 | 0.19
102.8 | 0 | 13 | 0.18 | 358.5 | 0.19
104 | 0 | 13 | 0.14 | 274.6 | 0.14
105.2 | 0 | 13 | 0.16 | 321.4 | 0.16
106.4 | 0 | 13 | 0.19 | 372.1 | 0.19
107.5 | 0 | 13 | 0.16 | 309.9 | 0.16
108.8 | 0 | 13 | 0.19 | 369.3 | 0.19
110 | 0 | 13 | 0.19 | 378.4 | 0.2
111.2 | 0 | 13 | 0.18 | 354.7 | 0.18
112.4 | 0 | 13 | 0.22 | 430.1 | 0.22
113.6 | 0 | 13 | 0.16 | 316.6 | 0.16
114.8 | 0 | 13 | 0.18 | 356.1 | 0.18
116 | 0 | 13 | 0.22 | 426 | 0.22
117.2 | 0 | 13 | 0.2 | 385.9 | 0.2
118.4 | 0 | 13 | 0.2 | 387.8 | 0.2
119.6 | 0 | 13 | 0.23 | 455.1 | 0.23
120.8 | 0 | 13 | 0.18 | 348.4 | 0.18
122 | 0 | 13 | 0.22 | 424.2 | 0.22
123.2 | 0 | 13 | 0.24 | 469.3 | 0.24
124.4 | 0 | 13 | 0.25 | 487.9 | 0.25
125.6 | 0 | 13 | 0.29 | 561.9 | 0.28
126.8 | 0 | 13 | 0.27 | 526 | 0.27
128 | 0 | 13 | 0.28 | 554.8 | 0.29
129.2 | 0 | 13 | 0.29 | 566.1 | 0.28
130.4 | 0 | 13 | 0.33 | 635.7 | 0.33
131.6 | 0 | 13 | 0.31 | 610.8 | 0.31
132.8 | 0 | 13 | 0.31 | 613.8 | 0.32
134 | 0 | 13 | 0.3 | 590.7 | 0.31
135.2 | 0 | 13 | 0.3 | 586.8 | 0.3
136.4 | 0 | 13 | 0.27 | 532.5 | 0.28
137.6 | 0 | 13 | 0.3 | 590.8 | 0.3
138.8 | 0 | 13 | 0.28 | 544.3 | 0.27
140 | 0 | 13 | 0.3 | 588.1 | 0.31
141.2 | 0 | 13 | 0.28 | 547.8 | 0.28
142.4 | 0 | 13 | 0.31 | 605.1 | 0.32
143.6 | 0 | 13 | 0.27 | 522.4 | 0.26
144.8 | 0 | 13 | 0.3 | 595.3 | 0.31
146 | 0 | 13 | 0.28 | 552.4 | 0.29
147.2 | 0 | 13 | 0.24 | 460.6 | 0.23
148.4 | 0 | 13 | 0.33 | 650.9 | 0.34
149.6 | 0 | 13 | 0.31 | 606.9 | 0.31
150.7 | 0 | 13 | 0.25 | 498.4 | 0.25
Measurement time: | 02/09/2022 18:48:53
Radiation device: | Clinac3 Accelerator
Energy: | 6 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -150.7 | 13 | 0.46 | 902 | 1.03
0 | -150 | 13 | 0.28 | 551.4 | 0.28
0 | -148.6 | 13 | 0.24 | 477.5 | 0.25
0 | -147.4 | 13 | 0.29 | 571.2 | 0.3
0 | -146.2 | 13 | 0.26 | 512 | 0.26
0 | -145 | 13 | 0.26 | 499.6 | 0.26
0 | -143.8 | 13 | 0.27 | 532.1 | 0.28
0 | -142.6 | 13 | 0.3 | 594.1 | 0.3
0 | -141.4 | 13 | 0.28 | 546 | 0.28
0 | -140.2 | 13 | 0.28 | 557 | 0.29
0 | -139 | 13 | 0.26 | 510.5 | 0.26
0 | -137.8 | 13 | 0.24 | 462.2 | 0.24
0 | -136.6 | 13 | 0.22 | 432.9 | 0.23
0 | -135.4 | 13 | 0.25 | 490.4 | 0.25
0 | -134.2 | 13 | 0.27 | 532.6 | 0.28
0 | -133 | 13 | 0.23 | 440.3 | 0.23
0 | -131.8 | 13 | 0.21 | 419.5 | 0.21
0 | -130.6 | 13 | 0.19 | 380.3 | 0.2
0 | -129.4 | 13 | 0.21 | 408.7 | 0.21
0 | -128.2 | 13 | 0.18 | 348.6 | 0.17
0 | -127 | 13 | 0.25 | 487.5 | 0.25
0 | -125.8 | 13 | 0.21 | 403 | 0.21
0 | -124.6 | 13 | 0.25 | 492.2 | 0.25
0 | -123.4 | 13 | 0.24 | 467.3 | 0.24
0 | -122.2 | 13 | 0.24 | 476.6 | 0.25
0 | -121 | 13 | 0.19 | 364.4 | 0.18
0 | -119.8 | 13 | 0.22 | 424 | 0.21
0 | -118.6 | 13 | 0.21 | 413.3 | 0.21
0 | -117.4 | 13 | 0.24 | 471.9 | 0.24
0 | -116.2 | 13 | 0.24 | 472.5 | 0.24
0 | -114.9 | 13 | 0.28 | 544.2 | 0.28
0 | -113.8 | 13 | 0.23 | 456.2 | 0.23
0 | -112.6 | 13 | 0.21 | 409.9 | 0.21
0 | -111.3 | 13 | 0.2 | 384.8 | 0.2
0 | -110.1 | 13 | 0.23 | 448.1 | 0.22
0 | -109 | 13 | 0.21 | 409.8 | 0.21
0 | -107.8 | 13 | 0.24 | 464.1 | 0.24
0 | -106.5 | 13 | 0.21 | 400.9 | 0.2
0 | -105.3 | 13 | 0.19 | 380.4 | 0.2
0 | -104.1 | 13 | 0.22 | 437.2 | 0.23
0 | -102.9 | 13 | 0.23 | 451.9 | 0.23
0 | -101.7 | 13 | 0.2 | 399.4 | 0.2
0 | -100.5 | 13 | 0.2 | 395.7 | 0.2
0 | -99.3 | 13 | 0.21 | 411.2 | 0.21
0 | -98.1 | 13 | 0.18 | 346.8 | 0.18
0 | -96.9 | 13 | 0.2 | 393.4 | 0.2
0 | -95.7 | 13 | 0.16 | 321.6 | 0.16
0 | -94.5 | 13 | 0.19 | 363.1 | 0.19
0 | -93.3 | 13 | 0.19 | 363.8 | 0.19
0 | -92.1 | 13 | 0.17 | 333.8 | 0.17
0 | -90.9 | 13 | 0.18 | 350.1 | 0.18
0 | -89.7 | 13 | 0.16 | 313.9 | 0.16
0 | -88.5 | 13 | 0.18 | 347.8 | 0.18
0 | -87.3 | 13 | 0.13 | 249.4 | 0.13
0 | -86.1 | 13 | 0.11 | 209.9 | 0.11
0 | -84.9 | 13 | 0.13 | 244.5 | 0.12
0 | -83.7 | 13 | 0.08 | 159.4 | 0.08
0 | -82.5 | 13 | 0.07 | 128.2 | 0.06
0 | -81.3 | 13 | 0.05 | 88.2 | 0.05
0 | -80.1 | 13 | 0 | 8.6 | 0
0 | -78.9 | 13 | 0.01 | 17 | 0.01
0 | -77.7 | 13 | 0.09 | 178.4 | 0.09
0 | -76.5 | 13 | 0.11 | 209.7 | 0.11
0 | -75.3 | 13 | 0.24 | 466.6 | 0.24
0 | -74.1 | 13 | 0.33 | 645.3 | 0.33
0 | -72.9 | 13 | 0.49 | 950.2 | 0.48
0 | -71.7 | 13 | 0.7 | 1360.4 | 0.7
0 | -70.5 | 13 | 1 | 1955.8 | 1.01
0 | -69.3 | 13 | 1.46 | 2862.5 | 1.44
0 | -68.1 | 13 | 1.92 | 3756.7 | 1.94
0 | -66.9 | 13 | 2.64 | 5162.5 | 2.66
0 | -65.7 | 13 | 3.52 | 6879.6 | 3.57
0 | -64.5 | 13 | 4.78 | 9349.7 | 4.67
0 | -63.3 | 13 | 6.07 | 11867 | 6.1
0 | -62.1 | 13 | 7.66 | 14986.2 | 7.78
0 | -60.9 | 13 | 10.43 | 20395.9 | 10.13
0 | -59.7 | 13 | 12.93 | 25289.6 | 13.08
0 | -58.5 | 13 | 16.23 | 31732.8 | 16.48
0 | -57.3 | 13 | 21.35 | 41747.3 | 20.87
0 | -56.1 | 13 | 25.78 | 50413.8 | 25.85
0 | -54.9 | 13 | 32.75 | 64043.5 | 32.1
0 | -53.7 | 13 | 38.35 | 74994.6 | 38.82
0 | -52.5 | 13 | 45.46 | 88903.4 | 45.89
0 | -51.2 | 13 | 51.99 | 101677.4 | 53.2
0 | -50 | 13 | 62.62 | 122452.3 | 60.63
0 | -48.8 | 13 | 66.52 | 130084.4 | 67.27
0 | -47.6 | 13 | 71.99 | 140776.6 | 73.07
0 | -46.4 | 13 | 81.08 | 158565.4 | 78.57
0 | -45.2 | 13 | 81.87 | 160104.9 | 82.69
0 | -44 | 13 | 84.87 | 165973.6 | 86.1
0 | -42.8 | 13 | 91.72 | 179366.5 | 89.2
0 | -41.6 | 13 | 90.12 | 176238.9 | 91.12
0 | -40.4 | 13 | 95.73 | 187209.8 | 93.11
0 | -39.2 | 13 | 92.59 | 181061 | 94.43
0 | -38 | 13 | 94.32 | 184439 | 95.34
0 | -36.8 | 13 | 95.15 | 186075.8 | 96.49
0 | -35.6 | 13 | 99.36 | 194300.5 | 97.29
0 | -34.4 | 13 | 97.15 | 189981.2 | 97.7
0 | -33.2 | 13 | 97.03 | 189745 | 98.15
0 | -32 | 13 | 101.59 | 198658 | 98.7
0 | -30.8 | 13 | 98.03 | 191696.9 | 98.93
0 | -29.6 | 13 | 97.5 | 190657.2 | 98.94
0 | -28.4 | 13 | 102.53 | 200505.6 | 99.46
0 | -27.2 | 13 | 98.41 | 192438.9 | 99.49
0 | -26 | 13 | 100.97 | 197456.5 | 99.54
0 | -24.8 | 13 | 98.38 | 192380.2 | 99.67
0 | -23.6 | 13 | 98.82 | 193243.2 | 99.87
0 | -22.4 | 13 | 102.54 | 200521.5 | 99.89
0 | -21.2 | 13 | 97.92 | 191490 | 99.77
0 | -20 | 13 | 98.87 | 193337.6 | 99.87
0 | -18.8 | 13 | 102.27 | 199998.4 | 99.98
0 | -17.6 | 13 | 97.78 | 191207.4 | 99.93
0 | -16.4 | 13 | 99.57 | 194718.8 | 99.65
0 | -15.2 | 13 | 98.64 | 192893.5 | 99.91
0 | -13.9 | 13 | 102.77 | 200973.9 | 100
0 | -12.8 | 13 | 98.47 | 192568.3 | 99.71
0 | -11.6 | 13 | 98.15 | 191929.6 | 99.82
0 | -10.4 | 13 | 102.01 | 199494.6 | 99.9
0 | -9.2 | 13 | 100.02 | 195597.1 | 99.42
0 | -8 | 13 | 99.85 | 195256.2 | 99.62
0 | -6.8 | 13 | 99.14 | 193875.8 | 99.89
0 | -5.6 | 13 | 100.34 | 196228.1 | 99.97
0 | -4.4 | 13 | 100.9 | 197308.7 | 99.68
0 | -3.1 | 13 | 97.87 | 191388.5 | 99.9
0 | -1.9 | 13 | 101.56 | 198606.3 | 99.84
0 | -0.7 | 13 | 99.1 | 193802.5 | 99.72
0 | 0.5 | 13 | 99.81 | 195192.4 | 99.82
0 | 1.7 | 13 | 101.84 | 199154.2 | 99.94
0 | 2.9 | 13 | 97.75 | 191152.6 | 99.43
0 | 4.1 | 13 | 99.24 | 194076 | 99.83
0 | 5.3 | 13 | 101.88 | 199238.6 | 99.79
0 | 6.5 | 13 | 97.89 | 191432.7 | 99.62
0 | 7.7 | 13 | 99.59 | 194750.5 | 99.69
0 | 8.9 | 13 | 102.52 | 200474 | 99.61
0 | 10.1 | 13 | 98.32 | 192271.3 | 99.62
0 | 11.3 | 13 | 98.3 | 192225.3 | 99.56
0 | 12.5 | 13 | 102.95 | 201315.1 | 99.52
0 | 13.7 | 13 | 96.68 | 189056.4 | 99.32
0 | 14.9 | 13 | 99.3 | 194183.5 | 99.62
0 | 16.1 | 13 | 102.08 | 199631.9 | 99.28
0 | 17.3 | 13 | 98.49 | 192607.6 | 99.3
0 | 18.5 | 13 | 97.38 | 190436.2 | 99.39
0 | 19.7 | 13 | 101.92 | 199316.5 | 99.06
0 | 20.9 | 13 | 97.48 | 190617 | 99.24
0 | 22.1 | 13 | 97.79 | 191232.6 | 99.16
0 | 23.3 | 13 | 101.37 | 198230.4 | 99.15
0 | 24.5 | 13 | 98.06 | 191754 | 98.74
0 | 25.7 | 13 | 97.59 | 190846 | 98.94
0 | 26.9 | 13 | 101.91 | 199296.2 | 98.81
0 | 28.1 | 13 | 96.61 | 188919 | 98.28
0 | 29.3 | 13 | 97.37 | 190402.9 | 98.28
0 | 30.5 | 13 | 100.93 | 197379.1 | 97.98
0 | 31.7 | 13 | 96.28 | 188274.1 | 97.59
0 | 32.9 | 13 | 95.79 | 187316.4 | 97.18
0 | 34.1 | 13 | 99.73 | 195026.6 | 96.63
0 | 35.3 | 13 | 93.51 | 182863.2 | 95.89
0 | 36.5 | 13 | 94.62 | 185028.6 | 95.2
0 | 37.7 | 13 | 96.77 | 189233.7 | 94.15
0 | 38.9 | 13 | 90.23 | 176450.8 | 92.72
0 | 40.1 | 13 | 90.85 | 177660.7 | 91
0 | 41.3 | 13 | 91.15 | 178251.9 | 88.84
0 | 42.5 | 13 | 84.49 | 165219.2 | 85.94
0 | 43.7 | 13 | 81.72 | 159812.5 | 82.81
0 | 44.9 | 13 | 80.55 | 157522.7 | 78.3
0 | 46.2 | 13 | 72.12 | 141028.4 | 73.39
0 | 47.4 | 13 | 67.02 | 131066.8 | 67.84
0 | 48.6 | 13 | 62.68 | 122574.6 | 60.88
0 | 49.8 | 13 | 53.32 | 104262.1 | 53.99
0 | 51 | 13 | 46.34 | 90612.1 | 46.87
0 | 52.2 | 13 | 40.88 | 79951 | 39.7
0 | 53.4 | 13 | 32.14 | 62857.2 | 32.87
0 | 54.6 | 13 | 26.85 | 52507.5 | 26.67
0 | 55.8 | 13 | 21.96 | 42948.7 | 21.45
0 | 57 | 13 | 16.5 | 32269.5 | 17.07
0 | 58.2 | 13 | 13.76 | 26905.3 | 13.55
0 | 59.4 | 13 | 10.66 | 20847.7 | 10.63
0 | 60.6 | 13 | 8.08 | 15804.1 | 8.25
0 | 61.8 | 13 | 6.53 | 12766.1 | 6.42
0 | 63 | 13 | 5.05 | 9885.2 | 5.01
0 | 64.2 | 13 | 3.65 | 7137.1 | 3.78
0 | 65.4 | 13 | 2.9 | 5670.6 | 2.87
0 | 66.6 | 13 | 2.16 | 4218.2 | 2.15
0 | 67.8 | 13 | 1.55 | 3039.2 | 1.59
0 | 69 | 13 | 1.16 | 2262.3 | 1.14
0 | 70.2 | 13 | 0.81 | 1588 | 0.81
0 | 71.4 | 13 | 0.54 | 1060.1 | 0.55
0 | 72.6 | 13 | 0.34 | 657.1 | 0.33
0 | 73.8 | 13 | 0.25 | 484.1 | 0.25
0 | 75 | 13 | 0.12 | 229.5 | 0.12
0 | 76.2 | 13 | 0.08 | 162.6 | 0.08
0 | 77.4 | 13 | 0.02 | 48.8 | 0.02
0 | 78.6 | 13 | 0.05 | 93 | 0.05
0 | 79.8 | 13 | 0.07 | 145.9 | 0.07
0 | 81.1 | 13 | 0.08 | 151.5 | 0.08
0 | 82.2 | 13 | 0.12 | 241.6 | 0.13
0 | 83.4 | 13 | 0.09 | 174.9 | 0.09
0 | 84.6 | 13 | 0.14 | 277.1 | 0.14
0 | 85.8 | 13 | 0.15 | 291.9 | 0.15
0 | 87 | 13 | 0.16 | 322.1 | 0.16
0 | 88.2 | 13 | 0.17 | 326.1 | 0.17
0 | 89.4 | 13 | 0.18 | 350.6 | 0.18
0 | 90.6 | 13 | 0.15 | 299 | 0.15
0 | 91.8 | 13 | 0.19 | 363 | 0.19
0 | 93 | 13 | 0.19 | 375.5 | 0.2
0 | 94.2 | 13 | 0.19 | 376.2 | 0.19
0 | 95.5 | 13 | 0.22 | 431 | 0.22
0 | 96.7 | 13 | 0.19 | 365.3 | 0.19
0 | 97.9 | 13 | 0.22 | 437 | 0.22
0 | 99.1 | 13 | 0.23 | 440.8 | 0.23
0 | 100.2 | 13 | 0.19 | 379.8 | 0.2
0 | 101.5 | 13 | 0.26 | 514.3 | 0.25
0 | 102.7 | 13 | 0.21 | 416.4 | 0.21
0 | 103.8 | 13 | 0.22 | 424.2 | 0.22
0 | 105.1 | 13 | 0.21 | 405.4 | 0.2
0 | 106.3 | 13 | 0.23 | 448.5 | 0.23
0 | 107.5 | 13 | 0.22 | 420.7 | 0.22
0 | 108.6 | 13 | 0.23 | 444.4 | 0.22
0 | 109.9 | 13 | 0.2 | 389.3 | 0.2
0 | 111 | 13 | 0.23 | 453.5 | 0.24
0 | 112.2 | 13 | 0.23 | 446.2 | 0.22
0 | 113.4 | 13 | 0.25 | 488.3 | 0.25
0 | 114.7 | 13 | 0.27 | 523.3 | 0.27
0 | 115.9 | 13 | 0.26 | 510.9 | 0.25
0 | 117 | 13 | 0.27 | 531.3 | 0.28
0 | 118.2 | 13 | 0.26 | 500.5 | 0.26
0 | 119.4 | 13 | 0.24 | 462.6 | 0.23
0 | 120.7 | 13 | 0.24 | 463.9 | 0.24
0 | 121.9 | 13 | 0.25 | 480 | 0.25
0 | 123 | 13 | 0.24 | 478.3 | 0.24
0 | 124.2 | 13 | 0.23 | 453.8 | 0.24
0 | 125.5 | 13 | 0.22 | 436 | 0.23
0 | 126.7 | 13 | 0.23 | 446.5 | 0.22
0 | 127.9 | 13 | 0.2 | 391.3 | 0.2
0 | 129.1 | 13 | 0.21 | 413.3 | 0.21
0 | 130.3 | 13 | 0.19 | 380.4 | 0.19
0 | 131.4 | 13 | 0.22 | 438.9 | 0.23
0 | 132.7 | 13 | 0.23 | 445.7 | 0.23
0 | 133.9 | 13 | 0.25 | 482.2 | 0.24
0 | 135.1 | 13 | 0.23 | 453.8 | 0.23
0 | 136.3 | 13 | 0.25 | 497.2 | 0.26
0 | 137.5 | 13 | 0.15 | 291.3 | 0.15
0 | 138.6 | 13 | 0.25 | 484.1 | 0.25
0 | 139.9 | 13 | 0.27 | 530.4 | 0.27
0 | 141.1 | 13 | 0.24 | 478.9 | 0.24
0 | 142.3 | 13 | 0.26 | 516 | 0.27
0 | 143.5 | 13 | 0.26 | 502.2 | 0.26
0 | 144.7 | 13 | 0.24 | 461.6 | 0.23
0 | 145.9 | 13 | 0.28 | 538.5 | 0.28
0 | 147.1 | 13 | 0.3 | 585.5 | 0.3
0 | 148.3 | 13 | 0.31 | 598.9 | 0.3
0 | 149.5 | 13 | 0.29 | 561.8 | 0.29
0 | 150.7 | 13 | 0.33 | 637.3 | 0.32
Measurement time: | 02/09/2022 18:51:13
Radiation device: | Clinac3 Accelerator
Energy: | 18 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
151.1 | 0 | 21 | 0.05 | 97.9 | 0.06
150.4 | 0 | 21 | 0.13 | 264 | 0.14
149.1 | 0 | 21 | 0.15 | 293.6 | 0.15
147.9 | 0 | 21 | 0.14 | 276.7 | 0.14
146.7 | 0 | 21 | 0.13 | 251.7 | 0.13
145.5 | 0 | 21 | 0.23 | 444.9 | 0.22
144.3 | 0 | 21 | 0.14 | 276.3 | 0.14
143.1 | 0 | 21 | 0.15 | 292.1 | 0.15
141.9 | 0 | 21 | 0.15 | 290.9 | 0.14
140.6 | 0 | 21 | 0.14 | 271.7 | 0.14
139.4 | 0 | 21 | 0.14 | 283.3 | 0.14
138.2 | 0 | 21 | 0.18 | 344.8 | 0.17
137 | 0 | 21 | 0.18 | 342.9 | 0.18
135.8 | 0 | 21 | 0.2 | 395.5 | 0.2
134.6 | 0 | 21 | 0.21 | 418.8 | 0.21
133.4 | 0 | 21 | 0.18 | 349 | 0.18
132.2 | 0 | 21 | 0.21 | 412.6 | 0.21
131 | 0 | 21 | 0.26 | 510.4 | 0.25
129.8 | 0 | 21 | 0.2 | 382.6 | 0.2
128.7 | 0 | 21 | 0.21 | 412.7 | 0.21
127.4 | 0 | 21 | 0.24 | 477.1 | 0.24
126.2 | 0 | 21 | 0.2 | 397.6 | 0.2
125 | 0 | 21 | 0.21 | 414.3 | 0.21
123.8 | 0 | 21 | 0.28 | 539 | 0.26
122.6 | 0 | 21 | 0.23 | 447.9 | 0.23
121.4 | 0 | 21 | 0.23 | 448.2 | 0.23
120.2 | 0 | 21 | 0.24 | 472.4 | 0.23
119 | 0 | 21 | 0.24 | 471.7 | 0.24
117.8 | 0 | 21 | 0.25 | 480.4 | 0.25
116.6 | 0 | 21 | 0.32 | 634.5 | 0.31
115.4 | 0 | 21 | 0.28 | 544.4 | 0.28
114.2 | 0 | 21 | 0.3 | 589.2 | 0.3
113 | 0 | 21 | 0.33 | 644.4 | 0.32
111.8 | 0 | 21 | 0.31 | 603.7 | 0.31
110.6 | 0 | 21 | 0.33 | 652.7 | 0.33
109.4 | 0 | 21 | 0.37 | 730.6 | 0.36
108.2 | 0 | 21 | 0.37 | 731 | 0.37
107 | 0 | 21 | 0.39 | 756.2 | 0.39
105.8 | 0 | 21 | 0.43 | 840 | 0.41
104.6 | 0 | 21 | 0.44 | 868.6 | 0.44
103.4 | 0 | 21 | 0.47 | 912.9 | 0.47
102.2 | 0 | 21 | 0.51 | 1002.8 | 0.49
101 | 0 | 21 | 0.52 | 1011.1 | 0.52
99.8 | 0 | 21 | 0.52 | 1019 | 0.53
98.6 | 0 | 21 | 0.62 | 1220.5 | 0.6
97.4 | 0 | 21 | 0.59 | 1154.9 | 0.6
96.2 | 0 | 21 | 0.63 | 1241.2 | 0.64
95 | 0 | 21 | 0.73 | 1430.7 | 0.7
93.8 | 0 | 21 | 0.69 | 1358.3 | 0.71
92.6 | 0 | 21 | 0.84 | 1642.3 | 0.84
91.4 | 0 | 21 | 0.95 | 1852 | 0.92
90.2 | 0 | 21 | 1 | 1954.1 | 1.01
89 | 0 | 21 | 1.05 | 2058.6 | 1.06
87.8 | 0 | 21 | 1.25 | 2440.8 | 1.2
86.6 | 0 | 21 | 1.31 | 2552.4 | 1.3
85.3 | 0 | 21 | 1.45 | 2831.4 | 1.46
84.2 | 0 | 21 | 1.67 | 3272.5 | 1.6
83 | 0 | 21 | 1.71 | 3344.7 | 1.72
81.8 | 0 | 21 | 1.81 | 3548.1 | 1.83
80.6 | 0 | 21 | 2.05 | 4009.7 | 1.96
79.3 | 0 | 21 | 1.99 | 3897.6 | 2.03
78.2 | 0 | 21 | 2.13 | 4158.3 | 2.13
76.9 | 0 | 21 | 2.36 | 4607.4 | 2.27
75.7 | 0 | 21 | 2.31 | 4525.2 | 2.35
74.6 | 0 | 21 | 2.46 | 4805.4 | 2.47
73.3 | 0 | 21 | 2.69 | 5263.2 | 2.61
72.1 | 0 | 21 | 2.75 | 5380.5 | 2.77
70.9 | 0 | 21 | 2.89 | 5644.6 | 2.92
69.7 | 0 | 21 | 3.23 | 6323.2 | 3.13
68.5 | 0 | 21 | 3.31 | 6479.3 | 3.32
67.3 | 0 | 21 | 3.54 | 6921.9 | 3.59
66.1 | 0 | 21 | 4.09 | 7991.7 | 3.91
64.9 | 0 | 21 | 4.23 | 8265.8 | 4.23
63.7 | 0 | 21 | 4.55 | 8892.9 | 4.6
62.5 | 0 | 21 | 5.18 | 10135 | 4.96
61.3 | 0 | 21 | 5.16 | 10098.8 | 5.36
60.1 | 0 | 21 | 6.17 | 12072.9 | 6
58.9 | 0 | 21 | 7.25 | 14172.3 | 6.91
57.7 | 0 | 21 | 8.19 | 16013.2 | 8.25
56.5 | 0 | 21 | 10.4 | 20334 | 10.37
55.3 | 0 | 21 | 14.76 | 28863.7 | 14.16
54.1 | 0 | 21 | 19.69 | 38504.4 | 19.67
52.9 | 0 | 21 | 27.6 | 53971.2 | 27.66
51.7 | 0 | 21 | 39.93 | 78094.4 | 38.52
50.5 | 0 | 21 | 49.89 | 97568.5 | 50.19
49.3 | 0 | 21 | 61.87 | 120984.6 | 62.6
48.1 | 0 | 21 | 76.4 | 149404 | 73.27
46.9 | 0 | 21 | 81.31 | 158997.5 | 82.07
45.6 | 0 | 21 | 87.19 | 170510.5 | 88.11
44.5 | 0 | 21 | 94.29 | 184387.2 | 91.66
43.2 | 0 | 21 | 94.55 | 184895.9 | 94.23
42.1 | 0 | 21 | 94.96 | 185690 | 95.7
40.9 | 0 | 21 | 99.57 | 194720.2 | 96.57
39.6 | 0 | 21 | 97 | 189688 | 97.07
38.5 | 0 | 21 | 97.49 | 190649 | 97.57
37.2 | 0 | 21 | 101.72 | 198921 | 97.9
36 | 0 | 21 | 96.78 | 189263.8 | 97.98
34.8 | 0 | 21 | 99.43 | 194431.4 | 98.37
33.6 | 0 | 21 | 102.47 | 200392.9 | 98.62
32.4 | 0 | 21 | 96.01 | 187755.6 | 98.52
31.2 | 0 | 21 | 100.34 | 196219.7 | 98.79
30 | 0 | 21 | 101.03 | 197572.6 | 98.92
28.8 | 0 | 21 | 100.2 | 195953.2 | 99.1
27.6 | 0 | 21 | 98.69 | 192991.2 | 99.16
26.4 | 0 | 21 | 102.33 | 200112.5 | 99.3
25.2 | 0 | 21 | 99.22 | 194028.4 | 99.1
24 | 0 | 21 | 99.21 | 194014.9 | 99.37
22.8 | 0 | 21 | 102.58 | 200594.7 | 99.5
21.6 | 0 | 21 | 98.02 | 191683.2 | 99.27
20.4 | 0 | 21 | 99.39 | 194356.8 | 99.47
19.2 | 0 | 21 | 102.27 | 200001.5 | 99.29
18 | 0 | 21 | 97.32 | 190304.1 | 99.49
16.8 | 0 | 21 | 100.4 | 196344.5 | 99.57
15.6 | 0 | 21 | 102.56 | 200560.8 | 99.66
14.4 | 0 | 21 | 96.94 | 189573.2 | 99.4
13.2 | 0 | 21 | 101.57 | 198615.5 | 99.79
12 | 0 | 21 | 101.85 | 199169.9 | 99.49
10.8 | 0 | 21 | 98.64 | 192893.4 | 99.59
9.6 | 0 | 21 | 100.86 | 197242.8 | 99.81
8.4 | 0 | 21 | 101.89 | 199246.3 | 99.75
7.2 | 0 | 21 | 100.38 | 196293.4 | 99.59
6 | 0 | 21 | 99.9 | 195368.5 | 99.91
4.8 | 0 | 21 | 100.94 | 197395.9 | 99.98
3.6 | 0 | 21 | 101.35 | 198189.8 | 99.81
2.4 | 0 | 21 | 99.26 | 194106.1 | 99.89
1.2 | 0 | 21 | 102.14 | 199737.1 | 99.93
0 | 0 | 21 | 97.57 | 190809.4 | 99.72
-1.2 | 0 | 21 | 99.35 | 194274.1 | 100
-2.4 | 0 | 21 | 105 | 205337.6 | 99.92
-3.6 | 0 | 21 | 98.16 | 191963.9 | 99.88
-4.8 | 0 | 21 | 101.08 | 197673.4 | 99.87
-6 | 0 | 21 | 102.67 | 200770.4 | 99.77
-7.2 | 0 | 21 | 98.96 | 193521.8 | 99.73
-8.4 | 0 | 21 | 102.09 | 199637.8 | 99.71
-9.6 | 0 | 21 | 100.06 | 195663.1 | 99.79
-10.8 | 0 | 21 | 99.88 | 195329.6 | 99.67
-12 | 0 | 21 | 101.95 | 199361.2 | 99.79
-13.2 | 0 | 21 | 100.24 | 196025.9 | 99.48
-14.4 | 0 | 21 | 99.2 | 193982.2 | 99.54
-15.6 | 0 | 21 | 99.08 | 193745.8 | 99.58
-16.8 | 0 | 21 | 103.6 | 202595.8 | 99.47
-18 | 0 | 21 | 98.78 | 193168.1 | 99.26
-19.3 | 0 | 21 | 102.06 | 199573.5 | 99.49
-20.4 | 0 | 21 | 99.72 | 195007.5 | 99.21
-21.6 | 0 | 21 | 97.73 | 191108 | 99.2
-22.9 | 0 | 21 | 101.62 | 198726.3 | 99.27
-24.1 | 0 | 21 | 98.95 | 193510.4 | 99.2
-25.3 | 0 | 21 | 98.25 | 192132.5 | 98.86
-26.5 | 0 | 21 | 102 | 199458.2 | 99.09
-27.7 | 0 | 21 | 99.42 | 194415 | 98.89
-28.9 | 0 | 21 | 98.52 | 192664.4 | 98.8
-30.1 | 0 | 21 | 102.12 | 199698.5 | 98.76
-31.3 | 0 | 21 | 98.49 | 192594.4 | 98.65
-32.5 | 0 | 21 | 97.24 | 190151.1 | 98.62
-33.7 | 0 | 21 | 101.38 | 198253.7 | 98.43
-34.9 | 0 | 21 | 97.58 | 190813.2 | 98.28
-36.1 | 0 | 21 | 100.98 | 197477.7 | 98.2
-37.3 | 0 | 21 | 96.7 | 189108.3 | 97.93
-38.5 | 0 | 21 | 97.27 | 190217 | 97.52
-39.7 | 0 | 21 | 96.95 | 189589.9 | 97.44
-40.9 | 0 | 21 | 95.88 | 187498.1 | 97.06
-42.1 | 0 | 21 | 100.07 | 195700.8 | 96.59
-43.3 | 0 | 21 | 95.93 | 187594.1 | 95.9
-44.5 | 0 | 21 | 98.63 | 192867.9 | 95.34
-45.7 | 0 | 21 | 93.29 | 182438.2 | 93.68
-46.9 | 0 | 21 | 90.83 | 177621.7 | 91.3
-48.1 | 0 | 21 | 90.19 | 176375.3 | 87.65
-49.3 | 0 | 21 | 81.54 | 159450.2 | 81.46
-50.5 | 0 | 21 | 72.42 | 141628.8 | 72.69
-51.7 | 0 | 21 | 62.2 | 121635.4 | 62.35
-52.9 | 0 | 21 | 51.71 | 101120.6 | 49.7
-54.1 | 0 | 21 | 37.59 | 73515 | 37.77
-55.3 | 0 | 21 | 27.28 | 53337.8 | 27.42
-56.5 | 0 | 21 | 20.04 | 39180.3 | 19.31
-57.7 | 0 | 21 | 13.65 | 26694.6 | 13.74
-58.9 | 0 | 21 | 10.42 | 20384.6 | 10.1
-60.1 | 0 | 21 | 7.94 | 15521.6 | 7.95
-61.3 | 0 | 21 | 6.44 | 12599.7 | 6.53
-62.5 | 0 | 21 | 5.65 | 11056.3 | 5.67
-63.7 | 0 | 21 | 5.15 | 10071.7 | 5.08
-64.9 | 0 | 21 | 4.52 | 8838.5 | 4.44
-66.1 | 0 | 21 | 4.23 | 8268.2 | 4.16
-67.3 | 0 | 21 | 3.81 | 7445.2 | 3.81
-68.5 | 0 | 21 | 3.51 | 6864.8 | 3.52
-69.7 | 0 | 21 | 3.35 | 6541.7 | 3.28
-71 | 0 | 21 | 3.04 | 5951.6 | 3.05
-72.1 | 0 | 21 | 2.83 | 5542.9 | 2.82
-73.4 | 0 | 21 | 2.68 | 5246.1 | 2.66
-74.6 | 0 | 21 | 2.53 | 4941.9 | 2.53
-75.8 | 0 | 21 | 2.42 | 4733.7 | 2.38
-77 | 0 | 21 | 2.24 | 4388 | 2.23
-78.2 | 0 | 21 | 2.06 | 4032.5 | 2.08
-79.4 | 0 | 21 | 2.05 | 4017.6 | 2.01
-80.6 | 0 | 21 | 1.98 | 3875.7 | 1.97
-81.8 | 0 | 21 | 1.86 | 3637.4 | 1.87
-83 | 0 | 21 | 1.82 | 3559.1 | 1.76
-84.2 | 0 | 21 | 1.7 | 3330 | 1.7
-85.4 | 0 | 21 | 1.57 | 3066.5 | 1.58
-86.6 | 0 | 21 | 1.55 | 3022.8 | 1.51
-87.8 | 0 | 21 | 1.4 | 2738.7 | 1.4
-89 | 0 | 21 | 1.28 | 2502.6 | 1.29
-90.2 | 0 | 21 | 1.28 | 2493.7 | 1.24
-91.4 | 0 | 21 | 1.05 | 2053.5 | 1.05
-92.6 | 0 | 21 | 1.02 | 1996.4 | 1.03
-93.8 | 0 | 21 | 0.99 | 1932 | 0.96
-95 | 0 | 21 | 0.85 | 1671.2 | 0.86
-96.2 | 0 | 21 | 0.78 | 1528.8 | 0.79
-97.4 | 0 | 21 | 0.81 | 1584.7 | 0.79
-98.6 | 0 | 21 | 0.68 | 1337.7 | 0.69
-99.8 | 0 | 21 | 0.65 | 1277.4 | 0.66
-101 | 0 | 21 | 0.65 | 1278.3 | 0.63
-102.2 | 0 | 21 | 0.57 | 1112 | 0.57
-103.4 | 0 | 21 | 0.58 | 1139.8 | 0.56
-104.6 | 0 | 21 | 0.52 | 1013.9 | 0.52
-105.8 | 0 | 21 | 0.53 | 1029.4 | 0.53
-107 | 0 | 21 | 0.5 | 986.7 | 0.51
-108.2 | 0 | 21 | 0.5 | 987.4 | 0.49
-109.5 | 0 | 21 | 0.43 | 849.5 | 0.44
-110.6 | 0 | 21 | 0.42 | 816.7 | 0.42
-111.8 | 0 | 21 | 0.49 | 949 | 0.47
-113.1 | 0 | 21 | 0.45 | 881.2 | 0.45
-114.2 | 0 | 21 | 0.45 | 879.2 | 0.45
-115.4 | 0 | 21 | 0.48 | 941.9 | 0.46
-116.6 | 0 | 21 | 0.39 | 770.3 | 0.4
-117.8 | 0 | 21 | 0.42 | 824.4 | 0.41
-119.1 | 0 | 21 | 0.36 | 713.5 | 0.36
-120.2 | 0 | 21 | 0.41 | 802.2 | 0.41
-121.5 | 0 | 21 | 0.36 | 708 | 0.36
-122.7 | 0 | 21 | 0.41 | 804.3 | 0.4
-123.9 | 0 | 21 | 0.35 | 683 | 0.35
-125.1 | 0 | 21 | 0.42 | 818.9 | 0.42
-126.3 | 0 | 21 | 0.41 | 793 | 0.39
-127.5 | 0 | 21 | 0.35 | 686.5 | 0.35
-128.7 | 0 | 21 | 0.36 | 695.6 | 0.36
-129.9 | 0 | 21 | 0.34 | 661.1 | 0.33
-131.1 | 0 | 21 | 0.4 | 774.1 | 0.4
-132.3 | 0 | 21 | 0.36 | 706.6 | 0.36
-133.5 | 0 | 21 | 0.41 | 805.7 | 0.4
-134.7 | 0 | 21 | 0.36 | 699.7 | 0.36
-135.9 | 0 | 21 | 0.38 | 746.4 | 0.38
-137.1 | 0 | 21 | 0.42 | 824.6 | 0.41
-138.3 | 0 | 21 | 0.4 | 781.3 | 0.4
-139.5 | 0 | 21 | 0.38 | 749.4 | 0.39
-140.7 | 0 | 21 | 0.39 | 755.7 | 0.37
-141.9 | 0 | 21 | 0.35 | 685.1 | 0.35
-143.1 | 0 | 21 | 0.36 | 704.9 | 0.36
-144.3 | 0 | 21 | 0.4 | 779.5 | 0.38
-145.5 | 0 | 21 | 0.4 | 773.2 | 0.4
-146.7 | 0 | 21 | 0.35 | 679.4 | 0.35
-147.9 | 0 | 21 | 0.42 | 816.4 | 0.4
-149.1 | 0 | 21 | 0.3 | 589.6 | 0.3
-151.1 | 0 | 21 | 0.37 | 729.2 | 0.37
Measurement time: | 02/09/2022 18:53:11
Radiation device: | Clinac3 Accelerator
Energy: | 18 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -151.1 | 21 | 0.29 | 568.1 | 0.19
0 | -150.4 | 21 | 0.05 | 100.6 | 0.05
0 | -149.1 | 21 | 0.13 | 250.3 | 0.13
0 | -147.9 | 21 | 0.09 | 176.5 | 0.09
0 | -146.7 | 21 | 0.13 | 253.6 | 0.13
0 | -145.5 | 21 | 0.1 | 194.6 | 0.1
0 | -144.3 | 21 | 0.08 | 159.7 | 0.08
0 | -143.1 | 21 | 0.16 | 317.6 | 0.16
0 | -141.8 | 21 | 0.16 | 313.6 | 0.16
0 | -140.7 | 21 | 0.13 | 255.9 | 0.13
0 | -139.4 | 21 | 0.2 | 385.5 | 0.19
0 | -138.2 | 21 | 0.18 | 346.4 | 0.18
0 | -137 | 21 | 0.16 | 314.6 | 0.16
0 | -135.8 | 21 | 0.23 | 443 | 0.22
0 | -134.6 | 21 | 0.23 | 446.8 | 0.23
0 | -133.4 | 21 | 0.19 | 371.5 | 0.19
0 | -132.2 | 21 | 0.25 | 483.2 | 0.24
0 | -131 | 21 | 0.18 | 355.7 | 0.18
0 | -129.8 | 21 | 0.21 | 416.3 | 0.22
0 | -128.6 | 21 | 0.23 | 453.9 | 0.22
0 | -127.4 | 21 | 0.22 | 426.4 | 0.22
0 | -126.2 | 21 | 0.2 | 400.1 | 0.2
0 | -125 | 21 | 0.25 | 495.2 | 0.24
0 | -123.8 | 21 | 0.24 | 478.4 | 0.25
0 | -122.6 | 21 | 0.22 | 439.5 | 0.22
0 | -121.4 | 21 | 0.29 | 563.9 | 0.28
0 | -120.2 | 21 | 0.26 | 517.5 | 0.27
0 | -119 | 21 | 0.2 | 398.7 | 0.21
0 | -117.8 | 21 | 0.32 | 624.5 | 0.31
0 | -116.6 | 21 | 0.29 | 564.8 | 0.29
0 | -115.4 | 21 | 0.27 | 521.4 | 0.27
0 | -114.2 | 21 | 0.28 | 540.7 | 0.27
0 | -113 | 21 | 0.28 | 554 | 0.28
0 | -111.8 | 21 | 0.31 | 599.4 | 0.31
0 | -110.6 | 21 | 0.34 | 674.4 | 0.33
0 | -109.4 | 21 | 0.35 | 693.4 | 0.36
0 | -108.2 | 21 | 0.35 | 689.8 | 0.36
0 | -107 | 21 | 0.41 | 803.8 | 0.4
0 | -105.8 | 21 | 0.42 | 820.5 | 0.43
0 | -104.6 | 21 | 0.41 | 794.4 | 0.41
0 | -103.4 | 21 | 0.47 | 926.7 | 0.46
0 | -102.2 | 21 | 0.48 | 945.1 | 0.49
0 | -101 | 21 | 0.5 | 969 | 0.5
0 | -99.8 | 21 | 0.59 | 1150.1 | 0.56
0 | -98.6 | 21 | 0.61 | 1200.8 | 0.61
0 | -97.4 | 21 | 0.62 | 1205.8 | 0.62
0 | -96.2 | 21 | 0.72 | 1414.5 | 0.69
0 | -94.9 | 21 | 0.71 | 1388.4 | 0.71
0 | -93.7 | 21 | 0.75 | 1470.9 | 0.76
0 | -92.5 | 21 | 0.86 | 1689.2 | 0.83
0 | -91.3 | 21 | 0.93 | 1811.8 | 0.93
0 | -90.1 | 21 | 1.01 | 1971.8 | 1.02
0 | -88.9 | 21 | 1.12 | 2187.7 | 1.08
0 | -87.7 | 21 | 1.13 | 2206.5 | 1.14
0 | -86.5 | 21 | 1.18 | 2306.7 | 1.19
0 | -85.4 | 21 | 1.49 | 2914.7 | 1.44
0 | -84.1 | 21 | 1.54 | 3011.2 | 1.55
0 | -82.9 | 21 | 1.66 | 3252 | 1.68
0 | -81.7 | 21 | 1.9 | 3723.5 | 1.84
0 | -80.5 | 21 | 1.94 | 3795.3 | 1.96
0 | -79.3 | 21 | 2.01 | 3937.9 | 2.02
0 | -78.1 | 21 | 2.26 | 4423.6 | 2.17
0 | -76.9 | 21 | 2.26 | 4421 | 2.27
0 | -75.7 | 21 | 2.35 | 4587.3 | 2.35
0 | -74.5 | 21 | 2.63 | 5134.5 | 2.52
0 | -73.3 | 21 | 2.64 | 5164.6 | 2.66
0 | -72.1 | 21 | 2.79 | 5457.1 | 2.8
0 | -70.9 | 21 | 3.09 | 6040.5 | 2.98
0 | -69.7 | 21 | 3.15 | 6151 | 3.18
0 | -68.5 | 21 | 3.4 | 6653.6 | 3.43
0 | -67.3 | 21 | 3.85 | 7522.9 | 3.71
0 | -66.1 | 21 | 4 | 7819.7 | 4.05
0 | -64.9 | 21 | 4.37 | 8549.5 | 4.37
0 | -63.7 | 21 | 5 | 9775.8 | 4.83
0 | -62.5 | 21 | 5.3 | 10360.3 | 5.35
0 | -61.3 | 21 | 6.01 | 11758.3 | 6.04
0 | -60.1 | 21 | 7.4 | 14469.2 | 7.12
0 | -58.9 | 21 | 8.81 | 17234.3 | 8.84
0 | -57.7 | 21 | 11.52 | 22527.2 | 11.6
0 | -56.5 | 21 | 16.97 | 33180.6 | 16.29
0 | -55.3 | 21 | 22.98 | 44934.8 | 23.2
0 | -54.1 | 21 | 32.34 | 63251.1 | 32.62
0 | -52.9 | 21 | 45.45 | 88873.3 | 44.05
0 | -51.7 | 21 | 56.56 | 110604 | 56.97
0 | -50.5 | 21 | 67.95 | 132882.8 | 68.78
0 | -49.3 | 21 | 81.39 | 159159.9 | 78.43
0 | -48.1 | 21 | 84.41 | 165062.5 | 85.92
0 | -46.8 | 21 | 89.83 | 175673.7 | 90.56
0 | -45.7 | 21 | 97.4 | 190470.2 | 93.2
0 | -44.5 | 21 | 94.62 | 185028.6 | 95.4
0 | -43.3 | 21 | 95.89 | 187517.9 | 96.45
0 | -42 | 21 | 100.79 | 197101.4 | 96.95
0 | -40.9 | 21 | 96.61 | 188922.4 | 97.59
0 | -39.7 | 21 | 97.79 | 191242.4 | 97.94
0 | -38.4 | 21 | 101.8 | 199076.2 | 98.16
0 | -37.2 | 21 | 97.82 | 191285.4 | 98.49
0 | -36 | 21 | 98.37 | 192362.6 | 98.72
0 | -34.8 | 21 | 101.81 | 199087.5 | 99
0 | -33.6 | 21 | 97.14 | 189964.4 | 98.85
0 | -32.4 | 21 | 99.11 | 193805.6 | 99
0 | -31.2 | 21 | 102.22 | 199886.6 | 99.24
0 | -30 | 21 | 98.31 | 192247.7 | 99.19
0 | -28.8 | 21 | 99.49 | 194549.4 | 99.44
0 | -27.6 | 21 | 102.72 | 200879.5 | 99.23
0 | -26.4 | 21 | 97.29 | 190264.5 | 99.38
0 | -25.2 | 21 | 101.44 | 198365.7 | 99.47
0 | -24 | 21 | 103.79 | 202962.1 | 99.49
0 | -22.8 | 21 | 96.88 | 189450.5 | 99.42
0 | -21.6 | 21 | 102.1 | 199656.1 | 99.76
0 | -20.4 | 21 | 103.07 | 201567 | 99.58
0 | -19.2 | 21 | 98.55 | 192721.4 | 99.71
0 | -18 | 21 | 100.82 | 197157.7 | 99.81
0 | -16.8 | 21 | 101.08 | 197673.8 | 99.58
0 | -15.6 | 21 | 97.82 | 191294.7 | 99.9
0 | -14.4 | 21 | 101.72 | 198923.4 | 100
0 | -13.2 | 21 | 102.94 | 201313.7 | 99.86
0 | -12 | 21 | 96.69 | 189072.2 | 99.91
0 | -10.8 | 21 | 101.81 | 199086.1 | 100
0 | -9.6 | 21 | 102.9 | 201222.4 | 99.92
0 | -8.4 | 21 | 96.67 | 189050.8 | 99.92
0 | -7.2 | 21 | 101.86 | 199192.1 | 99.98
0 | -6 | 21 | 100.7 | 196924.8 | 99.97
0 | -4.8 | 21 | 99.17 | 193926.8 | 99.89
0 | -3.6 | 21 | 101.73 | 198936.7 | 99.88
0 | -2.4 | 21 | 100.92 | 197351.9 | 99.9
0 | -1.2 | 21 | 100.5 | 196537.5 | 99.84
0 | 0 | 21 | 101.41 | 198308.6 | 99.95
0 | 1.2 | 21 | 99.43 | 194432.2 | 99.64
0 | 2.4 | 21 | 99.43 | 194436.5 | 99.62
0 | 3.6 | 21 | 102.98 | 201388.3 | 99.86
0 | 4.8 | 21 | 100.06 | 195664.2 | 99.57
0 | 6 | 21 | 98.71 | 193028.2 | 99.55
0 | 7.2 | 21 | 102.67 | 200777.3 | 99.5
0 | 8.4 | 21 | 99.84 | 195233.4 | 99.35
0 | 9.6 | 21 | 98.7 | 193010.1 | 99.45
0 | 10.8 | 21 | 102.76 | 200944.4 | 99.49
0 | 12 | 21 | 99.58 | 194738.9 | 99.16
0 | 13.2 | 21 | 98.44 | 192511.3 | 99.32
0 | 14.4 | 21 | 102.23 | 199915 | 99.2
0 | 15.6 | 21 | 99 | 193600.2 | 99.35
0 | 16.8 | 21 | 97.52 | 190701.8 | 99.16
0 | 18.1 | 21 | 101.87 | 199206.9 | 99.18
0 | 19.2 | 21 | 98.72 | 193042.3 | 98.89
0 | 20.4 | 21 | 97.94 | 191528 | 99.06
0 | 21.6 | 21 | 101.92 | 199308.6 | 98.95
0 | 22.9 | 21 | 98.14 | 191910.9 | 98.78
0 | 24 | 21 | 101.62 | 198731.3 | 98.9
0 | 25.2 | 21 | 97.47 | 190613.5 | 98.73
0 | 26.5 | 21 | 98.4 | 192419.4 | 98.72
0 | 27.7 | 21 | 97.84 | 191332.9 | 98.66
0 | 28.9 | 21 | 101.58 | 198650 | 98.49
0 | 30.1 | 21 | 98.13 | 191904.1 | 98.35
0 | 31.3 | 21 | 98.28 | 192183.6 | 98.33
0 | 32.5 | 21 | 97.34 | 190351.3 | 98.31
0 | 33.7 | 21 | 101.83 | 199138.7 | 97.98
0 | 34.9 | 21 | 98.26 | 192145.8 | 98.09
0 | 36.1 | 21 | 100.4 | 196341.5 | 98.18
0 | 37.3 | 21 | 96.72 | 189134.7 | 97.67
0 | 38.5 | 21 | 98.16 | 191962.2 | 97.58
0 | 39.7 | 21 | 100.59 | 196713.1 | 97.58
0 | 40.9 | 21 | 95.99 | 187716.2 | 96.74
0 | 42.1 | 21 | 96.24 | 188192.9 | 96.25
0 | 43.3 | 21 | 97.1 | 189892.8 | 95.27
0 | 44.5 | 21 | 94.88 | 185547.7 | 93.36
0 | 45.7 | 21 | 89.93 | 175867.8 | 90.16
0 | 46.9 | 21 | 86.88 | 169902.3 | 84.83
0 | 48.1 | 21 | 77.45 | 151452.8 | 76.76
0 | 49.3 | 21 | 66.9 | 130831.2 | 66.8
0 | 50.5 | 21 | 56.1 | 109714.1 | 54.47
0 | 51.7 | 21 | 41.84 | 81814.6 | 42.24
0 | 52.9 | 21 | 31.08 | 60783.6 | 31
0 | 54.1 | 21 | 22.17 | 43348.8 | 21.81
0 | 55.3 | 21 | 15.48 | 30279.2 | 15.34
0 | 56.5 | 21 | 11.07 | 21645.6 | 11.19
0 | 57.7 | 21 | 8.65 | 16916.8 | 8.61
0 | 58.9 | 21 | 6.92 | 13531.6 | 6.99
0 | 60.1 | 21 | 6.11 | 11946.9 | 5.98
0 | 61.3 | 21 | 5.36 | 10490.6 | 5.32
0 | 62.5 | 21 | 4.7 | 9182.1 | 4.73
0 | 63.7 | 21 | 4.42 | 8647.9 | 4.32
0 | 64.9 | 21 | 4.02 | 7868.8 | 4
0 | 66.2 | 21 | 3.63 | 7092.9 | 3.62
0 | 67.4 | 21 | 3.49 | 6825.2 | 3.4
0 | 68.5 | 21 | 3.14 | 6144.5 | 3.12
0 | 69.7 | 21 | 2.98 | 5827.8 | 2.91
0 | 71 | 21 | 2.69 | 5255.9 | 2.72
0 | 72.2 | 21 | 2.59 | 5061.3 | 2.59
0 | 73.4 | 21 | 2.45 | 4799.8 | 2.49
0 | 74.6 | 21 | 2.43 | 4758.7 | 2.38
0 | 75.8 | 21 | 2.23 | 4367.7 | 2.24
0 | 77 | 21 | 2.04 | 3998.2 | 2.07
0 | 78.2 | 21 | 2.07 | 4046.5 | 2.02
0 | 79.4 | 21 | 1.81 | 3549.1 | 1.84
0 | 80.6 | 21 | 1.77 | 3456.8 | 1.72
0 | 81.8 | 21 | 1.58 | 3098.9 | 1.6
0 | 83 | 21 | 1.42 | 2775.1 | 1.41
0 | 84.2 | 21 | 1.27 | 2476.2 | 1.27
0 | 85.4 | 21 | 1.23 | 2399.9 | 1.18
0 | 86.6 | 21 | 1.06 | 2077.1 | 1.06
0 | 87.8 | 21 | 0.99 | 1931.2 | 1
0 | 89 | 21 | 0.98 | 1922.7 | 0.95
0 | 90.2 | 21 | 0.84 | 1635 | 0.84
0 | 91.4 | 21 | 0.87 | 1707.8 | 0.84
0 | 92.6 | 21 | 0.69 | 1346.3 | 0.7
0 | 93.8 | 21 | 0.66 | 1292.6 | 0.66
0 | 95 | 21 | 0.66 | 1281.7 | 0.66
0 | 96.2 | 21 | 0.64 | 1243 | 0.61
0 | 97.4 | 21 | 0.53 | 1032.6 | 0.53
0 | 98.6 | 21 | 0.56 | 1103.2 | 0.56
0 | 99.8 | 21 | 0.51 | 989.4 | 0.49
0 | 101 | 21 | 0.46 | 908.3 | 0.46
0 | 102.2 | 21 | 0.44 | 861.1 | 0.44
0 | 103.4 | 21 | 0.39 | 764.1 | 0.38
0 | 104.6 | 21 | 0.38 | 744.3 | 0.38
0 | 105.8 | 21 | 0.35 | 688 | 0.34
0 | 107 | 21 | 0.29 | 567 | 0.29
0 | 108.2 | 21 | 0.31 | 601.6 | 0.31
0 | 109.4 | 21 | 0.23 | 450.9 | 0.23
0 | 110.6 | 21 | 0.3 | 577.6 | 0.29
0 | 111.8 | 21 | 0.2 | 385.7 | 0.2
0 | 113 | 21 | 0.18 | 359 | 0.19
0 | 114.2 | 21 | 0.2 | 399.4 | 0.2
0 | 115.4 | 21 | 0.23 | 447.2 | 0.23
0 | 116.6 | 21 | 0.15 | 289.8 | 0.15
0 | 117.8 | 21 | 0.2 | 395.6 | 0.2
0 | 119 | 21 | 0.16 | 313.9 | 0.16
0 | 120.2 | 21 | 0.12 | 238.3 | 0.12
0 | 121.4 | 21 | 0.16 | 310.4 | 0.15
0 | 122.6 | 21 | 0.11 | 214.7 | 0.11
0 | 123.8 | 21 | 0.1 | 190.2 | 0.1
0 | 125 | 21 | 0.19 | 367.7 | 0.18
0 | 126.2 | 21 | 0.13 | 252.5 | 0.13
0 | 127.4 | 21 | 0.12 | 234.4 | 0.12
0 | 128.6 | 21 | 0.08 | 159.1 | 0.08
0 | 129.8 | 21 | 0.13 | 259.5 | 0.13
0 | 131 | 21 | 0.12 | 225.9 | 0.11
0 | 132.2 | 21 | 0.06 | 126 | 0.07
0 | 133.4 | 21 | 0.07 | 137.3 | 0.07
0 | 134.6 | 21 | 0.12 | 239.8 | 0.12
0 | 135.8 | 21 | 0.13 | 250.9 | 0.13
0 | 137 | 21 | 0.05 | 106 | 0.05
0 | 138.2 | 21 | 0.05 | 89.7 | 0.05
0 | 139.4 | 21 | 0.06 | 124.7 | 0.06
0 | 140.6 | 21 | 0.08 | 157.4 | 0.08
0 | 141.8 | 21 | 0.02 | 41.3 | 0.02
0 | 143 | 21 | 0.05 | 88.4 | 0.04
0 | 144.2 | 21 | 0.01 | 18.6 | 0.01
0 | 145.4 | 21 | 0.01 | 20.8 | 0.01
0 | 146.6 | 21 | 0.09 | 176 | 0.09
0 | 147.8 | 21 | 0.02 | 37.1 | 0.02
0 | 149 | 21 | 0.02 | 42 | 0.02
0 | 150.2 | 21 | 0 | 6.6 | 0
0 | 151.1 | 21 | 0.03 | 53.2 | 0.03