## Profils taile de champ 5x5 et 2
Measurement time: | 02/09/2022 20:14:29
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 250 x 250 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 200 | 0.16 | 230.9 | 0.12
0 | 0 | 199.3 | 0.68 | 990.2 | 0.7
0 | 0 | 198 | 0.7 | 1027.8 | 0.75
0 | 0 | 196.8 | 0.68 | 996.3 | 0.71
0 | 0 | 195.6 | 0.64 | 930.7 | 0.66
0 | 0 | 194.4 | 0.66 | 971 | 0.69
0 | 0 | 193.2 | 0.63 | 922.7 | 0.66
0 | 0 | 192 | 0.66 | 968 | 0.69
0 | 0 | 190.8 | 0.6 | 871.9 | 0.62
0 | 0 | 189.6 | 0.63 | 917.3 | 0.65
0 | 0 | 188.4 | 0.65 | 946.9 | 0.68
0 | 0 | 187.2 | 0.6 | 882.2 | 0.62
0 | 0 | 186 | 0.64 | 941.6 | 0.69
0 | 0 | 184.8 | 0.62 | 910.4 | 0.65
0 | 0 | 183.6 | 0.55 | 811.8 | 0.57
0 | 0 | 182.4 | 0.53 | 776.3 | 0.56
0 | 0 | 181.2 | 0.6 | 881.3 | 0.65
0 | 0 | 180 | 0.5 | 735.4 | 0.51
0 | 0 | 178.8 | 0.6 | 879.7 | 0.63
0 | 0 | 177.6 | 0.59 | 868.4 | 0.65
0 | 0 | 176.4 | 0.54 | 789.6 | 0.55
0 | 0 | 175.2 | 0.51 | 746.4 | 0.54
0 | 0 | 174 | 0.57 | 832.2 | 0.62
0 | 0 | 172.8 | 0.51 | 746 | 0.51
0 | 0 | 171.6 | 0.57 | 831.4 | 0.63
0 | 0 | 170.4 | 0.5 | 733.1 | 0.52
0 | 0 | 169.2 | 0.44 | 646.2 | 0.44
0 | 0 | 167.9 | 0.55 | 801.8 | 0.6
0 | 0 | 166.7 | 0.51 | 741.1 | 0.53
0 | 0 | 165.5 | 0.44 | 638.2 | 0.45
0 | 0 | 164.3 | 0.43 | 632.3 | 0.45
0 | 0 | 163.1 | 0.43 | 622.9 | 0.45
0 | 0 | 161.9 | 0.46 | 676.1 | 0.48
0 | 0 | 160.7 | 0.34 | 498.4 | 0.34
0 | 0 | 159.5 | 0.46 | 677.2 | 0.51
0 | 0 | 158.3 | 0.38 | 549.8 | 0.39
0 | 0 | 157.1 | 0.4 | 581.5 | 0.41
0 | 0 | 155.9 | 0.35 | 515.2 | 0.37
0 | 0 | 154.7 | 0.41 | 597.5 | 0.43
0 | 0 | 153.5 | 0.33 | 477.3 | 0.34
0 | 0 | 152.3 | 0.42 | 612.3 | 0.44
0 | 0 | 151.1 | 0.39 | 576.5 | 0.41
0 | 0 | 149.9 | 0.33 | 489.1 | 0.35
0 | 0 | 148.7 | 0.31 | 448 | 0.32
0 | 0 | 147.5 | 0.26 | 387.5 | 0.26
0 | 0 | 146.3 | 0.38 | 554.3 | 0.41
0 | 0 | 145.1 | 0.27 | 390.6 | 0.28
0 | 0 | 143.9 | 0.28 | 406.1 | 0.29
0 | 0 | 142.7 | 0.3 | 445.8 | 0.31
0 | 0 | 141.5 | 0.28 | 413.6 | 0.29
0 | 0 | 140.3 | 0.26 | 374.3 | 0.27
0 | 0 | 139.1 | 0.21 | 312.2 | 0.22
0 | 0 | 137.9 | 0.18 | 269.7 | 0.19
0 | 0 | 136.7 | 0.16 | 234.4 | 0.17
0 | 0 | 135.5 | 0.14 | 208 | 0.15
0 | 0 | 134.3 | 0.24 | 347.9 | 0.25
0 | 0 | 133.1 | 0.19 | 280.3 | 0.2
0 | 0 | 131.9 | 0.22 | 315.8 | 0.23
0 | 0 | 130.7 | 0.21 | 305.8 | 0.22
0 | 0 | 129.5 | 0.11 | 165 | 0.12
0 | 0 | 128.3 | 0.19 | 274.7 | 0.2
0 | 0 | 127.1 | 0.22 | 322.6 | 0.24
0 | 0 | 125.9 | 0.14 | 202.9 | 0.14
0 | 0 | 124.7 | 0.2 | 287.3 | 0.21
0 | 0 | 123.5 | 0.17 | 244.7 | 0.18
0 | 0 | 122.3 | 0.11 | 163 | 0.12
0 | 0 | 121.1 | 0.18 | 258.3 | 0.19
0 | 0 | 119.9 | 0.12 | 172.6 | 0.12
0 | 0 | 118.7 | 0.14 | 211.2 | 0.15
0 | 0 | 117.4 | 0.11 | 168.1 | 0.12
0 | 0 | 116.3 | 0.06 | 82.6 | 0.06
0 | 0 | 115.1 | 0.02 | 36.1 | 0.03
0 | 0 | 113.8 | 0.07 | 98.2 | 0.07
0 | 0 | 112.6 | 0.02 | 31.9 | 0.02
0 | 0 | 111.5 | 0.05 | 72 | 0.05
0 | 0 | 110.2 | 0.1 | 143.4 | 0.1
0 | 0 | 109 | 0.07 | 109.2 | 0.08
0 | 0 | 107.8 | 0.08 | 113.6 | 0.08
0 | 0 | 106.6 | 0.15 | 214.4 | 0.15
0 | 0 | 105.4 | 0.09 | 132.8 | 0.09
0 | 0 | 104.2 | 0.13 | 193.4 | 0.14
0 | 0 | 103 | 0.11 | 161.1 | 0.12
0 | 0 | 101.8 | 0.18 | 270.1 | 0.19
0 | 0 | 100.6 | 0.15 | 220.2 | 0.16
0 | 0 | 99.4 | 0.2 | 296.4 | 0.21
0 | 0 | 98.2 | 0.18 | 267 | 0.19
0 | 0 | 97 | 0.18 | 258.9 | 0.19
0 | 0 | 95.8 | 0.25 | 373.1 | 0.26
0 | 0 | 94.6 | 0.2 | 293.8 | 0.21
0 | 0 | 93.4 | 0.2 | 299.4 | 0.21
0 | 0 | 92.2 | 0.2 | 299.8 | 0.21
0 | 0 | 91 | 0.24 | 349.8 | 0.25
0 | 0 | 89.8 | 0.26 | 379.8 | 0.27
0 | 0 | 88.6 | 0.3 | 437.5 | 0.31
0 | 0 | 87.4 | 0.28 | 404.3 | 0.29
0 | 0 | 86.2 | 0.41 | 605.3 | 0.42
0 | 0 | 85 | 0.3 | 443.3 | 0.32
0 | 0 | 83.8 | 0.33 | 487.2 | 0.35
0 | 0 | 82.5 | 0.41 | 598.8 | 0.42
0 | 0 | 81.3 | 0.35 | 508.2 | 0.37
0 | 0 | 80.1 | 0.42 | 615.1 | 0.44
0 | 0 | 78.9 | 0.44 | 641.6 | 0.44
0 | 0 | 77.7 | 0.4 | 584.9 | 0.43
0 | 0 | 76.5 | 0.43 | 633.6 | 0.45
0 | 0 | 75.3 | 0.51 | 749.6 | 0.52
0 | 0 | 74.1 | 0.42 | 611.8 | 0.45
0 | 0 | 72.9 | 0.5 | 737 | 0.52
0 | 0 | 71.7 | 0.62 | 910.7 | 0.63
0 | 0 | 70.5 | 0.56 | 816.6 | 0.6
0 | 0 | 69.3 | 0.79 | 1151.2 | 0.82
0 | 0 | 68.1 | 1.17 | 1710.2 | 1.18
0 | 0 | 66.9 | 1.41 | 2072 | 1.53
0 | 0 | 65.7 | 2.29 | 3353.7 | 2.4
0 | 0 | 64.4 | 3.71 | 5433.8 | 3.76
0 | 0 | 63.2 | 5.21 | 7626.4 | 5.47
0 | 0 | 62 | 7.6 | 11132.9 | 7.99
0 | 0 | 60.8 | 10.44 | 15286.5 | 11.01
0 | 0 | 59.6 | 14.4 | 21094.9 | 15.08
0 | 0 | 58.4 | 18.21 | 26673.4 | 19.82
0 | 0 | 57.2 | 25.08 | 36744.2 | 25.54
0 | 0 | 56 | 30.38 | 44503 | 31.98
0 | 0 | 54.8 | 36.46 | 53405.4 | 38.52
0 | 0 | 53.6 | 43.07 | 63091.6 | 45.17
0 | 0 | 52.4 | 50.06 | 73322.3 | 52.9
0 | 0 | 51.2 | 54.87 | 80379.8 | 60
0 | 0 | 50 | 66.65 | 97627.3 | 67.98
0 | 0 | 48.8 | 72.03 | 105508.4 | 75.53
0 | 0 | 47.6 | 75.95 | 111256.9 | 82.29
0 | 0 | 46.4 | 88.77 | 130032.7 | 89.67
0 | 0 | 45.2 | 88.58 | 129756.2 | 95.83
0 | 0 | 44 | 96.64 | 141558.2 | 101.34
0 | 0 | 42.8 | 106.28 | 155682.1 | 107.45
0 | 0 | 41.6 | 103.48 | 151586.4 | 112
0 | 0 | 40.4 | 112.26 | 164436.7 | 116.34
0 | 0 | 39.1 | 120.48 | 176479.6 | 120.6
0 | 0 | 37.9 | 113.53 | 166306.1 | 123.17
0 | 0 | 36.7 | 121.08 | 177362.4 | 126.34
0 | 0 | 35.5 | 128.46 | 188170.9 | 128.49
0 | 0 | 34.3 | 119.2 | 174604.9 | 129.88
0 | 0 | 33.1 | 125.63 | 184028.1 | 131.42
0 | 0 | 31.9 | 131.31 | 192341.8 | 132.83
0 | 0 | 30.7 | 127.93 | 187395.8 | 133.24
0 | 0 | 29.5 | 122.11 | 178874.5 | 133.13
0 | 0 | 28.3 | 131.43 | 192518.6 | 133.9
0 | 0 | 27.1 | 127.1 | 186174 | 133.66
0 | 0 | 25.9 | 127.3 | 186475.7 | 133.89
0 | 0 | 24.7 | 126.02 | 184594.7 | 133.36
0 | 0 | 23.5 | 124.77 | 182764.6 | 132.89
0 | 0 | 22.3 | 125.1 | 183246.5 | 132.71
0 | 0 | 21.1 | 124.63 | 182555.1 | 132.06
0 | 0 | 19.9 | 124.15 | 181858.1 | 131.58
0 | 0 | 18.7 | 123.05 | 180244.2 | 130.95
0 | 0 | 17.5 | 129.48 | 189663.4 | 130.86
0 | 0 | 16.2 | 121.88 | 178535.1 | 130.19
0 | 0 | 15 | 123.38 | 180726.2 | 129.83
0 | 0 | 13.9 | 121.52 | 178000.4 | 128.92
0 | 0 | 12.6 | 123.09 | 180300.4 | 128.58
0 | 0 | 11.5 | 121.28 | 177656 | 128.13
0 | 0 | 10.2 | 120.41 | 176374 | 127.35
0 | 0 | 9 | 121.03 | 177281 | 126.68
0 | 0 | 7.8 | 119.13 | 174497.8 | 126.5
0 | 0 | 6.6 | 119.62 | 175227.5 | 125.53
0 | 0 | 5.4 | 118.73 | 173922.6 | 124.88
0 | 0 | 4.2 | 115.39 | 169023.3 | 123.39
0 | 0 | 3 | 121.39 | 177810.8 | 122.46
0 | 0 | 1.8 | 112.51 | 164801.8 | 120.99
0 | 0 | 0.6 | 115.45 | 169106.3 | 120.09
0 | 0 | -0.5 | 119.22 | 174633.9 | 120.03
Measurement time: | 02/09/2022 20:15:50
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 250 x 250 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
228.5 | 0 | 28 | 1.78 | 2610.7 | 2.5
227.8 | 0 | 28 | 1.58 | 2318.6 | 1.63
226.3 | 0 | 28 | 1.63 | 2382.4 | 1.66
225.2 | 0 | 28 | 1.57 | 2297.9 | 1.66
224 | 0 | 28 | 1.6 | 2344.7 | 1.68
222.8 | 0 | 28 | 1.55 | 2273.4 | 1.6
221.6 | 0 | 28 | 1.53 | 2240.1 | 1.61
220.4 | 0 | 28 | 1.53 | 2244.6 | 1.6
219.2 | 0 | 28 | 1.56 | 2287.6 | 1.63
218 | 0 | 28 | 1.57 | 2302.5 | 1.65
216.8 | 0 | 28 | 1.56 | 2281 | 1.65
215.6 | 0 | 28 | 1.55 | 2277 | 1.64
214.4 | 0 | 28 | 1.52 | 2222.9 | 1.6
213.2 | 0 | 28 | 1.6 | 2341.5 | 1.69
212 | 0 | 28 | 1.56 | 2290.4 | 1.66
210.8 | 0 | 28 | 1.49 | 2187.5 | 1.56
209.6 | 0 | 28 | 1.58 | 2312 | 1.67
208.4 | 0 | 28 | 1.49 | 2186.5 | 1.56
207.2 | 0 | 28 | 1.49 | 2187.9 | 1.56
206 | 0 | 28 | 1.54 | 2260.2 | 1.65
204.8 | 0 | 28 | 1.52 | 2225.2 | 1.56
203.6 | 0 | 28 | 1.59 | 2330.7 | 1.69
202.3 | 0 | 28 | 1.49 | 2189.4 | 1.53
201.2 | 0 | 28 | 1.46 | 2145.5 | 1.52
200 | 0 | 28 | 1.52 | 2219.5 | 1.61
198.7 | 0 | 28 | 1.43 | 2089.3 | 1.44
197.6 | 0 | 28 | 1.47 | 2156.1 | 1.56
196.4 | 0 | 28 | 1.42 | 2079.5 | 1.47
195.2 | 0 | 28 | 1.37 | 2007.9 | 1.43
194 | 0 | 28 | 1.47 | 2158.9 | 1.58
192.7 | 0 | 28 | 1.41 | 2069.4 | 1.47
191.6 | 0 | 28 | 1.52 | 2222.8 | 1.59
190.3 | 0 | 28 | 1.41 | 2070.5 | 1.54
189.1 | 0 | 28 | 1.37 | 2012.7 | 1.42
187.9 | 0 | 28 | 1.42 | 2086.4 | 1.54
186.7 | 0 | 28 | 1.39 | 2033.7 | 1.43
185.6 | 0 | 28 | 1.42 | 2081.1 | 1.51
184.3 | 0 | 28 | 1.41 | 2066.8 | 1.47
183.1 | 0 | 28 | 1.37 | 2001 | 1.43
181.9 | 0 | 28 | 1.36 | 1993.3 | 1.43
180.7 | 0 | 28 | 1.35 | 1977.7 | 1.4
179.5 | 0 | 28 | 1.35 | 1983.6 | 1.42
178.3 | 0 | 28 | 1.33 | 1951.2 | 1.39
177.1 | 0 | 28 | 1.31 | 1920.2 | 1.37
175.9 | 0 | 28 | 1.31 | 1920.6 | 1.37
174.7 | 0 | 28 | 1.32 | 1940.4 | 1.38
173.5 | 0 | 28 | 1.24 | 1823.3 | 1.26
172.3 | 0 | 28 | 1.23 | 1797.8 | 1.34
171.1 | 0 | 28 | 1.11 | 1621.3 | 1.13
169.9 | 0 | 28 | 1.14 | 1676.5 | 1.21
168.7 | 0 | 28 | 1.09 | 1598.4 | 1.18
167.5 | 0 | 28 | 1 | 1465.7 | 1.06
166.3 | 0 | 28 | 0.97 | 1427.6 | 1.01
165.1 | 0 | 28 | 0.87 | 1277.2 | 0.95
163.9 | 0 | 28 | 0.76 | 1119.3 | 0.78
162.7 | 0 | 28 | 0.72 | 1052 | 0.77
161.5 | 0 | 28 | 0.53 | 769.2 | 0.57
160.3 | 0 | 28 | 0.29 | 424.5 | 0.29
159.1 | 0 | 28 | 0.16 | 235.5 | 0.17
157.9 | 0 | 28 | 0.03 | 38.6 | 0.03
156.7 | 0 | 28 | 0.22 | 319.4 | 0.23
155.5 | 0 | 28 | 0.53 | 777 | 0.55
154.3 | 0 | 28 | 0.82 | 1204.3 | 0.86
153.1 | 0 | 28 | 1.19 | 1739.6 | 1.24
151.9 | 0 | 28 | 1.67 | 2450.9 | 1.73
150.6 | 0 | 28 | 2.12 | 3109.5 | 2.23
149.5 | 0 | 28 | 2.71 | 3971.1 | 2.82
148.2 | 0 | 28 | 3.37 | 4932 | 3.55
147.1 | 0 | 28 | 4.09 | 5987.7 | 4.29
145.8 | 0 | 28 | 5.3 | 7763.7 | 5.37
144.6 | 0 | 28 | 6.14 | 8986.7 | 6.47
143.4 | 0 | 28 | 7 | 10249.6 | 7.64
142.2 | 0 | 28 | 9.24 | 13540 | 9.44
141 | 0 | 28 | 10.8 | 15817.5 | 11.39
139.8 | 0 | 28 | 12.63 | 18503.6 | 13.67
138.6 | 0 | 28 | 16.76 | 24550.3 | 16.91
137.4 | 0 | 28 | 18.55 | 27166.4 | 20.17
136.2 | 0 | 28 | 23.46 | 34366.9 | 24.48
135 | 0 | 28 | 28.65 | 41962.9 | 29.82
133.8 | 0 | 28 | 34.72 | 50865.2 | 35.93
132.6 | 0 | 28 | 41.49 | 60771 | 43.25
131.4 | 0 | 28 | 49.5 | 72510.7 | 51.05
130.2 | 0 | 28 | 57.07 | 83596.6 | 59.78
129 | 0 | 28 | 66.14 | 96876.9 | 68.78
127.8 | 0 | 28 | 76.14 | 111536.2 | 77.71
126.6 | 0 | 28 | 82.56 | 120930.5 | 86.71
125.4 | 0 | 28 | 85.38 | 125060.3 | 93.99
124.2 | 0 | 28 | 100.03 | 146530.8 | 101.41
123 | 0 | 28 | 97.96 | 143498.9 | 106.77
121.8 | 0 | 28 | 106.46 | 155948.1 | 111.57
120.6 | 0 | 28 | 114.3 | 167427.1 | 115.95
119.4 | 0 | 28 | 110.42 | 161742.7 | 118.58
118.2 | 0 | 28 | 115.42 | 169066.4 | 120.97
117 | 0 | 28 | 119.58 | 175166.4 | 123.28
115.8 | 0 | 28 | 119.94 | 175691.7 | 124.48
114.6 | 0 | 28 | 121.04 | 177296.6 | 125.69
113.4 | 0 | 28 | 122.37 | 179246.6 | 127.24
112.2 | 0 | 28 | 120.89 | 177080.9 | 127.75
111 | 0 | 28 | 123.85 | 181412.5 | 128.64
109.8 | 0 | 28 | 125.28 | 183516.7 | 129.47
108.6 | 0 | 28 | 123.72 | 181231.8 | 129.71
107.4 | 0 | 28 | 125.1 | 183250 | 130.19
106.2 | 0 | 28 | 124.85 | 182881 | 130.54
105 | 0 | 28 | 124.56 | 182454.6 | 130.81
103.8 | 0 | 28 | 123.65 | 181131 | 131.09
102.6 | 0 | 28 | 126.42 | 185183.5 | 131.54
101.4 | 0 | 28 | 125 | 183102.7 | 131.49
100.2 | 0 | 28 | 124.15 | 181863.1 | 131.85
99 | 0 | 28 | 126.78 | 185713.4 | 131.93
97.8 | 0 | 28 | 128.25 | 187865.4 | 132.39
96.5 | 0 | 28 | 126.24 | 184922.6 | 132.5
95.4 | 0 | 28 | 128.08 | 187609.3 | 132.31
94.1 | 0 | 28 | 125.93 | 184465.3 | 133.04
93 | 0 | 28 | 126.75 | 185661.8 | 132.69
91.8 | 0 | 28 | 126.97 | 185986.6 | 132.66
90.5 | 0 | 28 | 124.8 | 182808.5 | 132.73
89.3 | 0 | 28 | 125.41 | 183702.3 | 132.78
88.1 | 0 | 28 | 125.95 | 184496 | 132.87
86.9 | 0 | 28 | 126.86 | 185829.6 | 132.67
85.7 | 0 | 28 | 126.37 | 185107.1 | 132.94
84.5 | 0 | 28 | 126.48 | 185270.8 | 132.99
83.3 | 0 | 28 | 127.69 | 187044.4 | 133.11
82.1 | 0 | 28 | 126.82 | 185766.1 | 132.88
80.9 | 0 | 28 | 128.66 | 188466.5 | 133.27
79.7 | 0 | 28 | 127.58 | 186874.4 | 133.16
78.5 | 0 | 28 | 128.19 | 187770.9 | 133.02
77.3 | 0 | 28 | 128.57 | 188337.4 | 133.17
76.1 | 0 | 28 | 127.56 | 186849.7 | 133.23
74.9 | 0 | 28 | 128.1 | 187637.5 | 133.23
73.7 | 0 | 28 | 128.47 | 188190.9 | 133.29
72.5 | 0 | 28 | 127.91 | 187371.8 | 133.36
71.3 | 0 | 28 | 124.67 | 182616.8 | 133.22
70.1 | 0 | 28 | 130.3 | 190862.8 | 133.18
68.9 | 0 | 28 | 126.96 | 185974.4 | 133.37
67.7 | 0 | 28 | 125.21 | 183409.3 | 133.12
66.5 | 0 | 28 | 125.86 | 184366.4 | 133.1
65.3 | 0 | 28 | 127.02 | 186064.5 | 133.46
64.1 | 0 | 28 | 125.19 | 183382.1 | 133.33
62.9 | 0 | 28 | 126.8 | 185740.3 | 133.16
61.7 | 0 | 28 | 126.82 | 185764.3 | 132.96
60.5 | 0 | 28 | 121.48 | 177948.8 | 132.79
59.3 | 0 | 28 | 131.06 | 191984.5 | 133.31
58 | 0 | 28 | 125.69 | 184119.9 | 132.77
56.8 | 0 | 28 | 121.41 | 177848.9 | 132.67
55.6 | 0 | 28 | 133.21 | 195132.9 | 133.53
54.4 | 0 | 28 | 121.52 | 177997.5 | 132.87
53.2 | 0 | 28 | 127.76 | 187144.9 | 133.26
52 | 0 | 28 | 131.55 | 192698 | 133.61
50.8 | 0 | 28 | 122.91 | 180046.6 | 133.12
49.6 | 0 | 28 | 127.1 | 186183.3 | 133.36
48.4 | 0 | 28 | 132.52 | 194118.1 | 133.51
47.2 | 0 | 28 | 123.36 | 180701.2 | 133.22
46 | 0 | 28 | 127.42 | 186641.5 | 133.41
44.8 | 0 | 28 | 133.4 | 195412.8 | 133.44
43.6 | 0 | 28 | 122.16 | 178946.7 | 133.06
42.4 | 0 | 28 | 127.53 | 186806.1 | 133.53
41.2 | 0 | 28 | 132.94 | 194739.8 | 133.71
40 | 0 | 28 | 122.84 | 179938.6 | 133.12
38.8 | 0 | 28 | 127.22 | 186356.8 | 133.49
37.6 | 0 | 28 | 133.52 | 195578.8 | 133.51
36.4 | 0 | 28 | 121.8 | 178414.5 | 133.19
35.2 | 0 | 28 | 128.42 | 188118.4 | 133.42
34 | 0 | 28 | 131.78 | 193033.3 | 133.42
32.8 | 0 | 28 | 123.07 | 180280.7 | 132.98
31.6 | 0 | 28 | 126.96 | 185966.3 | 133.05
30.4 | 0 | 28 | 133.04 | 194880.7 | 133.73
29.2 | 0 | 28 | 125.12 | 183276.4 | 133.09
28 | 0 | 28 | 126.79 | 185726.3 | 132.96
26.8 | 0 | 28 | 125.89 | 184403.9 | 133.39
25.6 | 0 | 28 | 125.5 | 183839.2 | 133.27
24.4 | 0 | 28 | 130.84 | 191650.7 | 133.53
23.2 | 0 | 28 | 124.5 | 182363.6 | 133.41
22 | 0 | 28 | 126.9 | 185889.6 | 133.47
20.8 | 0 | 28 | 126.02 | 184602.2 | 133.26
19.6 | 0 | 28 | 126.27 | 184956.3 | 133.57
18.4 | 0 | 28 | 126.07 | 184673.9 | 133.52
17.2 | 0 | 28 | 123.5 | 180912.1 | 133.37
16 | 0 | 28 | 127.26 | 186417.9 | 133.31
14.8 | 0 | 28 | 125.58 | 183947.8 | 133.42
13.6 | 0 | 28 | 125.59 | 183959.3 | 133.5
12.4 | 0 | 28 | 127.9 | 187344.8 | 133.42
11.2 | 0 | 28 | 124.85 | 182882.1 | 133.59
9.9 | 0 | 28 | 125 | 183104.5 | 133.69
8.8 | 0 | 28 | 130.42 | 191036.4 | 133.52
7.6 | 0 | 28 | 124.97 | 183057.2 | 133.83
6.3 | 0 | 28 | 126.94 | 185946.5 | 133.83
5.2 | 0 | 28 | 132.2 | 193646 | 134.32
3.9 | 0 | 28 | 126.42 | 185189.5 | 133.84
2.7 | 0 | 28 | 128.46 | 188172.3 | 133.7
1.5 | 0 | 28 | 131.33 | 192372.8 | 133.89
0.3 | 0 | 28 | 125.54 | 183891.8 | 133.66
-0.9 | 0 | 28 | 129.72 | 190010.8 | 133.74
-2.1 | 0 | 28 | 131.77 | 193019.3 | 134
-3.2 | 0 | 28 | 124 | 181640.9 | 133.7
-4.5 | 0 | 28 | 128.92 | 188842.8 | 133.96
-5.7 | 0 | 28 | 127.85 | 187278.6 | 134.04
-6.9 | 0 | 28 | 125.52 | 183863.7 | 133.73
-8.1 | 0 | 28 | 124.6 | 182523.6 | 133.84
-9.3 | 0 | 28 | 129.13 | 189147.5 | 134.14
-10.5 | 0 | 28 | 127.09 | 186160.3 | 133.7
-11.7 | 0 | 28 | 127.26 | 186416.9 | 133.82
-12.9 | 0 | 28 | 127.4 | 186622.7 | 133.65
-14.1 | 0 | 28 | 130.48 | 191134.1 | 134.07
-15.3 | 0 | 28 | 124.61 | 182527.6 | 133.85
-16.5 | 0 | 28 | 130.24 | 190770.9 | 133.79
-17.7 | 0 | 28 | 125.03 | 183152.5 | 133.69
-18.9 | 0 | 28 | 128.69 | 188502.5 | 133.86
-20.1 | 0 | 28 | 131.68 | 192891.2 | 133.85
-21.3 | 0 | 28 | 125.37 | 183640.9 | 133.75
-22.5 | 0 | 28 | 129.47 | 189643.3 | 133.88
-23.7 | 0 | 28 | 131.1 | 192035.3 | 133.98
-24.9 | 0 | 28 | 126.86 | 185825.7 | 133.44
-26.1 | 0 | 28 | 128.02 | 187523.3 | 133.71
-27.3 | 0 | 28 | 131.68 | 192891.7 | 133.89
-28.5 | 0 | 28 | 126.87 | 185843.5 | 133.42
-29.7 | 0 | 28 | 127.75 | 187128.7 | 133.28
-30.9 | 0 | 28 | 130.29 | 190849.8 | 133.89
-32.1 | 0 | 28 | 125.65 | 184060.3 | 133.5
-33.3 | 0 | 28 | 128.47 | 188181.8 | 133.37
-34.5 | 0 | 28 | 129.47 | 189655.1 | 133.74
-35.7 | 0 | 28 | 125.66 | 184062.3 | 133.29
-36.9 | 0 | 28 | 128.42 | 188111.8 | 133.64
-38.1 | 0 | 28 | 128.82 | 188704 | 133.5
-39.3 | 0 | 28 | 126.24 | 184923.8 | 133.34
-40.5 | 0 | 28 | 125.16 | 183336.4 | 133.43
-41.7 | 0 | 28 | 128.25 | 187864.6 | 133.23
-42.9 | 0 | 28 | 126.72 | 185617.7 | 133.55
-44.2 | 0 | 28 | 126.58 | 185412.6 | 133.44
-45.3 | 0 | 28 | 126.25 | 184928.4 | 133.16
-46.5 | 0 | 28 | 126.26 | 184941.8 | 133.38
-47.7 | 0 | 28 | 126.27 | 184965 | 133.28
-48.9 | 0 | 28 | 127.28 | 186444.4 | 133.33
-50.2 | 0 | 28 | 126.68 | 185561.6 | 133.48
-51.3 | 0 | 28 | 125.66 | 184072.3 | 133.42
-52.6 | 0 | 28 | 126.85 | 185809 | 133.04
-53.7 | 0 | 28 | 127.21 | 186345 | 133.01
-54.9 | 0 | 28 | 126.42 | 185178.5 | 133.3
-56.2 | 0 | 28 | 127.59 | 186901 | 133.23
-57.4 | 0 | 28 | 125.9 | 184413.7 | 133.22
-58.6 | 0 | 28 | 126.13 | 184763 | 133
-59.8 | 0 | 28 | 128 | 187500.7 | 133.4
-61 | 0 | 28 | 128.22 | 187814.7 | 133.43
-62.2 | 0 | 28 | 127.11 | 186195.6 | 133.33
-63.4 | 0 | 28 | 129.89 | 190262.3 | 133.66
-64.6 | 0 | 28 | 127.28 | 186445.5 | 133.48
-65.8 | 0 | 28 | 126.76 | 185673.5 | 133.48
-67 | 0 | 28 | 129.55 | 189763.4 | 133.63
-68.2 | 0 | 28 | 127.18 | 186292.4 | 133.46
-69.4 | 0 | 28 | 128.81 | 188678 | 133.47
-70.6 | 0 | 28 | 128.78 | 188645.8 | 133.58
-71.8 | 0 | 28 | 127.54 | 186819.4 | 133.52
-73 | 0 | 28 | 126.13 | 184762.6 | 133.14
-74.2 | 0 | 28 | 130.28 | 190831.7 | 133.43
-75.4 | 0 | 28 | 126.65 | 185522.2 | 133.45
-76.6 | 0 | 28 | 129.39 | 189529.3 | 133.84
-77.8 | 0 | 28 | 128.87 | 188771.9 | 133.55
-79 | 0 | 28 | 126.39 | 185133.8 | 133.52
-80.2 | 0 | 28 | 129.57 | 189796.2 | 133.43
-81.4 | 0 | 28 | 127.12 | 186214.1 | 133.43
-82.6 | 0 | 28 | 128.68 | 188485.7 | 133.26
-83.8 | 0 | 28 | 129.07 | 189069.4 | 133.28
-85 | 0 | 28 | 126.31 | 185019 | 133.38
-86.2 | 0 | 28 | 129.88 | 190255.6 | 133.06
-87.4 | 0 | 28 | 127.04 | 186096.3 | 133.06
-88.6 | 0 | 28 | 128 | 187500.2 | 133.23
-89.8 | 0 | 28 | 128.01 | 187509.1 | 133.01
-91 | 0 | 28 | 126.45 | 185232.4 | 132.97
-92.2 | 0 | 28 | 129.79 | 190120.3 | 133.13
-93.4 | 0 | 28 | 125.86 | 184364.5 | 132.93
-94.6 | 0 | 28 | 127.52 | 186790.9 | 132.74
-95.8 | 0 | 28 | 127.47 | 186722.4 | 132.81
-97 | 0 | 28 | 126.54 | 185353.5 | 132.57
-98.2 | 0 | 28 | 127.11 | 186186.2 | 132.66
-99.4 | 0 | 28 | 127.2 | 186327.3 | 132.31
-100.6 | 0 | 28 | 126.42 | 185178.7 | 132.06
-101.9 | 0 | 28 | 126.76 | 185674.2 | 131.95
-103 | 0 | 28 | 126.62 | 185480.1 | 131.71
-104.3 | 0 | 28 | 126.12 | 184741.7 | 131.66
-105.4 | 0 | 28 | 125.56 | 183929.6 | 131.48
-106.7 | 0 | 28 | 126.4 | 185159.1 | 131.2
-107.9 | 0 | 28 | 126.6 | 185442.4 | 130.89
-109.1 | 0 | 28 | 124.79 | 182800.4 | 130.38
-110.3 | 0 | 28 | 125.75 | 184194.2 | 129.98
-111.5 | 0 | 28 | 123.49 | 180889.3 | 129.35
-112.7 | 0 | 28 | 123.72 | 181232.7 | 128.97
-113.9 | 0 | 28 | 123.01 | 180184.5 | 128.47
-115.1 | 0 | 28 | 121.3 | 177688.7 | 127.23
-116.3 | 0 | 28 | 121.54 | 178036.9 | 126.52
-117.5 | 0 | 28 | 118.67 | 173825.9 | 124.83
-118.7 | 0 | 28 | 122.58 | 179554.1 | 123.7
-119.9 | 0 | 28 | 115.43 | 169085.1 | 120.84
-121.1 | 0 | 28 | 108.72 | 159250.6 | 118.22
-122.3 | 0 | 28 | 112.51 | 164801.5 | 114.98
-123.5 | 0 | 28 | 105.11 | 153971 | 110.18
-124.7 | 0 | 28 | 94.4 | 138271.9 | 104.3
-125.9 | 0 | 28 | 96.4 | 141207.4 | 98.32
-127.1 | 0 | 28 | 86.11 | 126142.8 | 90.54
-128.3 | 0 | 28 | 73.9 | 108245.1 | 81.92
-129.5 | 0 | 28 | 72.19 | 105739.2 | 73.59
-130.7 | 0 | 28 | 60 | 87884.5 | 64.34
-131.9 | 0 | 28 | 50.57 | 74074.4 | 55.65
-133.1 | 0 | 28 | 47.05 | 68926.6 | 47.68
-134.3 | 0 | 28 | 37.18 | 54456.9 | 39.63
-135.5 | 0 | 28 | 30.55 | 44746.6 | 33.34
-136.7 | 0 | 28 | 27.58 | 40397.7 | 27.8
-137.9 | 0 | 28 | 21.37 | 31305.2 | 22.79
-139.1 | 0 | 28 | 17.54 | 25690.7 | 19.03
-140.3 | 0 | 28 | 15.66 | 22933.7 | 15.87
-141.5 | 0 | 28 | 11.93 | 17470.9 | 12.96
-142.7 | 0 | 28 | 10.41 | 15245.7 | 10.9
-143.9 | 0 | 28 | 8.58 | 12571.4 | 8.96
-145.1 | 0 | 28 | 7.08 | 10377.3 | 7.38
-146.3 | 0 | 28 | 5.76 | 8443.3 | 6.08
-147.5 | 0 | 28 | 4.81 | 7046.8 | 5.01
-148.7 | 0 | 28 | 3.85 | 5637.8 | 4.03
-149.9 | 0 | 28 | 3.03 | 4439.8 | 3.16
-151.1 | 0 | 28 | 2.39 | 3505 | 2.48
-152.4 | 0 | 28 | 1.76 | 2576.1 | 1.85
-153.5 | 0 | 28 | 1.34 | 1968.2 | 1.39
-154.8 | 0 | 28 | 0.9 | 1313.9 | 0.94
-156 | 0 | 28 | 0.61 | 896.4 | 0.64
-157.2 | 0 | 28 | 0.4 | 592.9 | 0.42
-158.4 | 0 | 28 | 0.13 | 189.8 | 0.14
-159.6 | 0 | 28 | 0.07 | 109 | 0.08
-160.8 | 0 | 28 | 0.28 | 411.6 | 0.29
-162 | 0 | 28 | 0.41 | 604.5 | 0.43
-163.2 | 0 | 28 | 0.53 | 770.9 | 0.55
-164.4 | 0 | 28 | 0.59 | 861.4 | 0.61
-165.6 | 0 | 28 | 0.67 | 984.6 | 0.7
-166.8 | 0 | 28 | 0.72 | 1050.6 | 0.74
-168 | 0 | 28 | 0.77 | 1133.5 | 0.8
-169.2 | 0 | 28 | 0.86 | 1262.3 | 0.9
-170.4 | 0 | 28 | 0.89 | 1306 | 0.93
-171.6 | 0 | 28 | 0.88 | 1285.2 | 0.91
-172.8 | 0 | 28 | 0.95 | 1389 | 0.99
-174 | 0 | 28 | 0.95 | 1394.1 | 0.99
-175.2 | 0 | 28 | 1.02 | 1490.9 | 1.06
-176.4 | 0 | 28 | 0.95 | 1393.4 | 1
-177.6 | 0 | 28 | 0.99 | 1448.4 | 1.03
-178.8 | 0 | 28 | 1.01 | 1479.5 | 1.07
-180 | 0 | 28 | 1.02 | 1496.4 | 1.05
-181.2 | 0 | 28 | 1.02 | 1488.3 | 1.08
-182.4 | 0 | 28 | 1.04 | 1529.2 | 1.13
-183.6 | 0 | 28 | 1.01 | 1483.2 | 1.05
-184.8 | 0 | 28 | 1.05 | 1542.5 | 1.11
-186 | 0 | 28 | 1.06 | 1546.6 | 1.12
-187.2 | 0 | 28 | 1.04 | 1527.2 | 1.1
-188.4 | 0 | 28 | 1.09 | 1590.8 | 1.15
-189.6 | 0 | 28 | 1.05 | 1539.6 | 1.12
-190.8 | 0 | 28 | 1.1 | 1611.9 | 1.15
-192 | 0 | 28 | 1.12 | 1641.9 | 1.2
-193.2 | 0 | 28 | 1.14 | 1668 | 1.21
-194.4 | 0 | 28 | 1.12 | 1639.2 | 1.18
-195.6 | 0 | 28 | 1.16 | 1696.6 | 1.22
-196.8 | 0 | 28 | 1.15 | 1679.4 | 1.22
-198 | 0 | 28 | 1.17 | 1708.6 | 1.21
-199.2 | 0 | 28 | 1.14 | 1675.4 | 1.21
-200.4 | 0 | 28 | 1.15 | 1684.6 | 1.21
-201.6 | 0 | 28 | 1.17 | 1713.6 | 1.21
-202.8 | 0 | 28 | 1.23 | 1802.2 | 1.3
-204 | 0 | 28 | 1.19 | 1737.4 | 1.24
-205.2 | 0 | 28 | 1.18 | 1724.6 | 1.23
-206.4 | 0 | 28 | 1.2 | 1761.4 | 1.26
-207.6 | 0 | 28 | 1.2 | 1760.5 | 1.25
-208.8 | 0 | 28 | 1.2 | 1763.6 | 1.24
-210 | 0 | 28 | 1.27 | 1860.8 | 1.34
-211.2 | 0 | 28 | 1.24 | 1811.6 | 1.29
-212.4 | 0 | 28 | 1.27 | 1854.9 | 1.33
-213.6 | 0 | 28 | 1.26 | 1849.8 | 1.31
-214.8 | 0 | 28 | 1.26 | 1847.3 | 1.32
-216 | 0 | 28 | 1.3 | 1902 | 1.35
-217.3 | 0 | 28 | 1.29 | 1884.7 | 1.34
-218.5 | 0 | 28 | 1.33 | 1948.3 | 1.39
-219.7 | 0 | 28 | 1.39 | 2030.4 | 1.44
-220.9 | 0 | 28 | 1.31 | 1915.9 | 1.36
-222.1 | 0 | 28 | 1.36 | 1989.5 | 1.42
-223.3 | 0 | 28 | 1.35 | 1974.6 | 1.4
-224.5 | 0 | 28 | 1.37 | 2011.5 | 1.43
-225.7 | 0 | 28 | 1.37 | 2007.6 | 1.44
-226.9 | 0 | 28 | 1.37 | 2002.2 | 1.41
-228.5 | 0 | 28 | 1.35 | 1970.5 | 1.4
Measurement time: | 02/09/2022 20:25:48
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 196.7 | 1.1 | 1608.6 | 1.05
0 | 0 | 195.5 | 1.15 | 1681.5 | 1.13
0 | 0 | 194.3 | 1.07 | 1562 | 1.03
0 | 0 | 193.1 | 1.08 | 1584.1 | 1.04
0 | 0 | 191.9 | 1.18 | 1728.6 | 1.16
0 | 0 | 190.7 | 1.12 | 1639.6 | 1.08
0 | 0 | 189.5 | 1.14 | 1663.5 | 1.09
0 | 0 | 188.3 | 1.1 | 1611.9 | 1.06
0 | 0 | 187.1 | 1.06 | 1556.5 | 1.03
0 | 0 | 185.9 | 1.04 | 1521.4 | 0.99
0 | 0 | 184.7 | 1.07 | 1562.9 | 1.03
0 | 0 | 183.5 | 1.06 | 1549.3 | 1
0 | 0 | 182.3 | 1 | 1459.2 | 0.95
0 | 0 | 181.1 | 1.04 | 1529 | 1.02
0 | 0 | 179.9 | 1.01 | 1483.5 | 0.97
0 | 0 | 178.7 | 0.94 | 1370.2 | 0.89
0 | 0 | 177.5 | 1.04 | 1516.5 | 1
0 | 0 | 176.3 | 0.99 | 1444.2 | 0.95
0 | 0 | 175.1 | 1.03 | 1512 | 0.99
0 | 0 | 173.9 | 1 | 1464.6 | 0.97
0 | 0 | 172.7 | 0.96 | 1409 | 0.93
0 | 0 | 171.5 | 0.98 | 1436.4 | 0.95
0 | 0 | 170.3 | 0.99 | 1447.1 | 0.96
0 | 0 | 169.1 | 0.9 | 1319.2 | 0.87
0 | 0 | 167.8 | 0.91 | 1337.2 | 0.89
0 | 0 | 166.6 | 0.94 | 1374.7 | 0.9
0 | 0 | 165.5 | 0.97 | 1419.4 | 0.95
0 | 0 | 164.3 | 0.94 | 1371.2 | 0.91
0 | 0 | 163 | 0.87 | 1268.8 | 0.83
0 | 0 | 161.9 | 0.91 | 1335.5 | 0.88
0 | 0 | 160.6 | 0.89 | 1302.6 | 0.86
0 | 0 | 159.4 | 0.88 | 1292.4 | 0.85
0 | 0 | 158.3 | 0.84 | 1223.2 | 0.8
0 | 0 | 157 | 0.85 | 1245.7 | 0.81
0 | 0 | 155.8 | 0.84 | 1226.3 | 0.81
0 | 0 | 154.6 | 0.84 | 1236.4 | 0.81
0 | 0 | 153.4 | 0.78 | 1137.2 | 0.73
0 | 0 | 152.2 | 0.84 | 1234.6 | 0.82
0 | 0 | 151 | 0.77 | 1130.8 | 0.73
0 | 0 | 149.8 | 0.79 | 1155 | 0.76
0 | 0 | 148.6 | 0.75 | 1105.7 | 0.73
0 | 0 | 147.4 | 0.8 | 1173.2 | 0.77
0 | 0 | 146.2 | 0.78 | 1143.9 | 0.77
0 | 0 | 145 | 0.69 | 1008.4 | 0.66
0 | 0 | 143.8 | 0.75 | 1104.6 | 0.73
0 | 0 | 142.6 | 0.7 | 1023.5 | 0.68
0 | 0 | 141.4 | 0.76 | 1106.8 | 0.73
0 | 0 | 140.2 | 0.7 | 1030.1 | 0.69
0 | 0 | 139 | 0.66 | 961.1 | 0.62
0 | 0 | 137.8 | 0.68 | 991.8 | 0.67
0 | 0 | 136.6 | 0.65 | 947.6 | 0.6
0 | 0 | 135.4 | 0.68 | 993.5 | 0.68
0 | 0 | 134.2 | 0.63 | 919.8 | 0.6
0 | 0 | 133 | 0.65 | 948.1 | 0.65
0 | 0 | 131.8 | 0.63 | 921.5 | 0.61
0 | 0 | 130.6 | 0.58 | 852.4 | 0.56
0 | 0 | 129.4 | 0.53 | 780.3 | 0.51
0 | 0 | 128.2 | 0.56 | 824.4 | 0.54
0 | 0 | 127 | 0.57 | 835.5 | 0.55
0 | 0 | 125.8 | 0.47 | 687 | 0.45
0 | 0 | 124.6 | 0.54 | 795.2 | 0.52
0 | 0 | 123.4 | 0.53 | 771.3 | 0.51
0 | 0 | 122.2 | 0.52 | 764.1 | 0.5
0 | 0 | 121 | 0.45 | 652.6 | 0.42
0 | 0 | 119.8 | 0.47 | 686.1 | 0.45
0 | 0 | 118.6 | 0.44 | 643 | 0.42
0 | 0 | 117.4 | 0.47 | 690.5 | 0.45
0 | 0 | 116.2 | 0.42 | 609.2 | 0.4
0 | 0 | 115 | 0.35 | 519 | 0.34
0 | 0 | 113.8 | 0.38 | 551.2 | 0.36
0 | 0 | 112.6 | 0.32 | 468.5 | 0.31
0 | 0 | 111.4 | 0.35 | 514 | 0.34
0 | 0 | 110.2 | 0.38 | 557.6 | 0.37
0 | 0 | 109 | 0.39 | 570.8 | 0.37
0 | 0 | 107.8 | 0.36 | 528.6 | 0.35
0 | 0 | 106.6 | 0.32 | 464.6 | 0.31
0 | 0 | 105.4 | 0.3 | 445.7 | 0.29
0 | 0 | 104.2 | 0.33 | 476.5 | 0.32
0 | 0 | 103 | 0.37 | 542.8 | 0.38
0 | 0 | 101.7 | 0.23 | 332 | 0.21
0 | 0 | 100.5 | 0.37 | 537.6 | 0.37
0 | 0 | 99.3 | 0.21 | 314.5 | 0.21
0 | 0 | 98.1 | 0.21 | 302 | 0.2
0 | 0 | 96.9 | 0.23 | 331.7 | 0.22
0 | 0 | 95.7 | 0.19 | 278 | 0.18
0 | 0 | 94.5 | 0.14 | 204.1 | 0.13
0 | 0 | 93.3 | 0.07 | 106 | 0.07
0 | 0 | 92.1 | 0.13 | 185.7 | 0.12
0 | 0 | 90.9 | 0.08 | 123.5 | 0.08
0 | 0 | 89.7 | 0.09 | 131.3 | 0.09
0 | 0 | 88.5 | 0.03 | 48.4 | 0.03
0 | 0 | 87.3 | 0.05 | 80 | 0.05
0 | 0 | 86.1 | 0.07 | 98 | 0.06
0 | 0 | 84.9 | 0.04 | 58.8 | 0.04
0 | 0 | 83.7 | 0.02 | 25.4 | 0.02
0 | 0 | 82.5 | 0.02 | 22.3 | 0.01
0 | 0 | 81.3 | 0.02 | 30.3 | 0.02
0 | 0 | 80.1 | 0.06 | 90.7 | 0.06
0 | 0 | 78.9 | 0.08 | 124 | 0.08
0 | 0 | 77.7 | 0.12 | 168.7 | 0.11
0 | 0 | 76.4 | 0.08 | 124.2 | 0.08
0 | 0 | 75.2 | 0.17 | 243.8 | 0.16
0 | 0 | 74 | 0.16 | 231.4 | 0.15
0 | 0 | 72.8 | 0.24 | 354.1 | 0.23
0 | 0 | 71.6 | 0.2 | 296 | 0.2
0 | 0 | 70.4 | 0.28 | 406.4 | 0.27
0 | 0 | 69.2 | 0.53 | 779.7 | 0.51
0 | 0 | 68 | 0.79 | 1151.5 | 0.77
0 | 0 | 66.8 | 1.29 | 1890.2 | 1.26
0 | 0 | 65.6 | 2.17 | 3182.3 | 2.1
0 | 0 | 64.4 | 3.45 | 5060.6 | 3.32
0 | 0 | 63.2 | 5.34 | 7819.4 | 5.18
0 | 0 | 62 | 7.76 | 11370 | 7.46
0 | 0 | 60.8 | 10.92 | 15994.5 | 10.65
0 | 0 | 59.6 | 14.6 | 21386.7 | 14.46
0 | 0 | 58.4 | 19.76 | 28940.7 | 19.03
0 | 0 | 57.2 | 25.82 | 37820.1 | 24.58
0 | 0 | 56 | 31.82 | 46608.1 | 30.4
0 | 0 | 54.8 | 38.96 | 57065.7 | 36.99
0 | 0 | 53.5 | 45.96 | 67323.4 | 43.92
0 | 0 | 52.3 | 53.12 | 77811.9 | 50.88
0 | 0 | 51.1 | 60.57 | 88717.5 | 57.89
0 | 0 | 49.9 | 68.2 | 99900.4 | 65.22
0 | 0 | 48.7 | 75.71 | 110896.5 | 72.23
0 | 0 | 47.5 | 82.85 | 121361.7 | 79.02
0 | 0 | 46.3 | 89.27 | 130764.2 | 85.27
0 | 0 | 45.1 | 95.37 | 139702.1 | 91.34
0 | 0 | 43.9 | 101.87 | 149215.9 | 97.4
0 | 0 | 42.7 | 107.68 | 157731.4 | 102.66
0 | 0 | 41.5 | 111.7 | 163625.7 | 107.16
0 | 0 | 40.3 | 116.58 | 170764.7 | 111.4
0 | 0 | 39.1 | 119.09 | 174450.2 | 115.14
0 | 0 | 37.9 | 123.66 | 181135.2 | 118.33
0 | 0 | 36.7 | 125.92 | 184449.9 | 121.2
0 | 0 | 35.4 | 127.81 | 187225.2 | 123.14
0 | 0 | 34.2 | 129.48 | 189659.3 | 125.23
0 | 0 | 33 | 131.68 | 192890.2 | 126.53
0 | 0 | 31.8 | 131.42 | 192513.5 | 127.45
0 | 0 | 30.6 | 133.85 | 196069 | 128.53
0 | 0 | 29.4 | 133.85 | 196068 | 129
0 | 0 | 28.2 | 132.93 | 194718.9 | 129.15
0 | 0 | 27 | 130.88 | 191721.1 | 129.28
0 | 0 | 25.8 | 137.5 | 201406.3 | 129.32
0 | 0 | 24.6 | 130.15 | 190643.1 | 128.9
0 | 0 | 23.4 | 136.69 | 200223.2 | 128.74
0 | 0 | 22.2 | 130.86 | 191688.1 | 127.95
0 | 0 | 21 | 130.91 | 191760.5 | 127.6
0 | 0 | 19.8 | 127.34 | 186526.3 | 126.79
0 | 0 | 18.6 | 132.75 | 194451.8 | 126.63
0 | 0 | 17.4 | 128.16 | 187725.5 | 125.9
0 | 0 | 16.2 | 125.96 | 184505.6 | 125.16
0 | 0 | 14.9 | 133.02 | 194855.4 | 125.2
0 | 0 | 13.7 | 126.12 | 184743.8 | 124.15
0 | 0 | 12.5 | 130.41 | 191032.1 | 124.17
0 | 0 | 11.3 | 130.52 | 191185.3 | 123.52
0 | 0 | 10.1 | 124.23 | 181968.9 | 122.61
0 | 0 | 8.9 | 127.6 | 186903.8 | 121.86
0 | 0 | 7.7 | 129.43 | 189592.9 | 121.42
0 | 0 | 6.5 | 120.8 | 176952.5 | 120.21
0 | 0 | 5.3 | 124.38 | 182200.2 | 119.59
0 | 0 | 4.1 | 127.18 | 186288.4 | 118.44
0 | 0 | 2.9 | 118.03 | 172890.4 | 116.54
0 | 0 | 1.7 | 120.03 | 175817.1 | 115.37
0 | 0 | 0.5 | 122.96 | 180117.3 | 114.81
0 | 0 | -0.5 | 114.96 | 168389 | 113.85
Measurement time: | 02/09/2022 20:27:12
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-125.7 | 0 | 28 | 1.37 | 2008.8 | 1.58
-125 | 0 | 28 | 1.67 | 2446.6 | 1.6
-123.6 | 0 | 28 | 1.65 | 2414.4 | 1.6
-122.4 | 0 | 28 | 1.66 | 2433.9 | 1.64
-121.2 | 0 | 28 | 1.64 | 2395.5 | 1.6
-120 | 0 | 28 | 1.62 | 2367.3 | 1.56
-118.8 | 0 | 28 | 1.65 | 2410.5 | 1.62
-117.6 | 0 | 28 | 1.74 | 2547.4 | 1.73
-116.4 | 0 | 28 | 1.64 | 2404.2 | 1.55
-115.2 | 0 | 28 | 1.69 | 2482.8 | 1.69
-114 | 0 | 28 | 1.66 | 2430.2 | 1.6
-112.8 | 0 | 28 | 1.61 | 2362.5 | 1.51
-111.6 | 0 | 28 | 1.67 | 2447.7 | 1.67
-110.4 | 0 | 28 | 1.68 | 2453.8 | 1.6
-109.2 | 0 | 28 | 1.58 | 2321.1 | 1.48
-108 | 0 | 28 | 1.64 | 2408 | 1.63
-106.8 | 0 | 28 | 1.7 | 2491.2 | 1.63
-105.6 | 0 | 28 | 1.67 | 2440.9 | 1.56
-104.4 | 0 | 28 | 1.66 | 2438 | 1.64
-103.2 | 0 | 28 | 1.75 | 2563.9 | 1.68
-102 | 0 | 28 | 1.75 | 2561.8 | 1.63
-100.8 | 0 | 28 | 1.77 | 2596.8 | 1.77
-99.6 | 0 | 28 | 1.73 | 2530.9 | 1.67
-98.4 | 0 | 28 | 1.7 | 2489.6 | 1.57
-97.2 | 0 | 28 | 1.71 | 2504.3 | 1.71
-96 | 0 | 28 | 1.72 | 2521.6 | 1.66
-94.7 | 0 | 28 | 1.7 | 2490.5 | 1.58
-93.6 | 0 | 28 | 1.71 | 2503.3 | 1.71
-92.3 | 0 | 28 | 1.67 | 2442.8 | 1.62
-91.1 | 0 | 28 | 1.61 | 2364.5 | 1.5
-89.9 | 0 | 28 | 1.7 | 2489.6 | 1.71
-88.7 | 0 | 28 | 1.7 | 2491.7 | 1.66
-87.5 | 0 | 28 | 1.63 | 2381.3 | 1.52
-86.3 | 0 | 28 | 1.66 | 2425.2 | 1.61
-85.1 | 0 | 28 | 1.64 | 2403.3 | 1.68
-83.9 | 0 | 28 | 1.6 | 2338.8 | 1.5
-82.7 | 0 | 28 | 1.64 | 2399.3 | 1.6
-81.5 | 0 | 28 | 1.62 | 2369.4 | 1.59
-80.3 | 0 | 28 | 1.59 | 2327 | 1.57
-79.1 | 0 | 28 | 1.52 | 2231.6 | 1.49
-77.9 | 0 | 28 | 1.54 | 2255.3 | 1.52
-76.7 | 0 | 28 | 1.51 | 2213.5 | 1.47
-75.5 | 0 | 28 | 1.45 | 2129.6 | 1.44
-74.3 | 0 | 28 | 1.41 | 2066.9 | 1.38
-73.1 | 0 | 28 | 1.3 | 1898 | 1.27
-71.9 | 0 | 28 | 1.32 | 1929.3 | 1.31
-70.7 | 0 | 28 | 1.19 | 1749.5 | 1.16
-69.5 | 0 | 28 | 1.17 | 1718.2 | 1.15
-68.3 | 0 | 28 | 1.08 | 1575.2 | 1.06
-67.1 | 0 | 28 | 0.91 | 1340.1 | 0.88
-65.9 | 0 | 28 | 0.9 | 1317.4 | 0.89
-64.7 | 0 | 28 | 0.74 | 1083.9 | 0.72
-63.5 | 0 | 28 | 0.69 | 1014.2 | 0.69
-62.3 | 0 | 28 | 0.44 | 650.7 | 0.42
-61.1 | 0 | 28 | 0.4 | 579 | 0.39
-59.9 | 0 | 28 | 0.17 | 247.9 | 0.16
-58.6 | 0 | 28 | 0.06 | 89.2 | 0.06
-57.4 | 0 | 28 | 0.27 | 397.8 | 0.27
-56.2 | 0 | 28 | 0.63 | 925.9 | 0.61
-55 | 0 | 28 | 1.09 | 1594 | 1.03
-53.8 | 0 | 28 | 1.46 | 2141.5 | 1.42
-52.6 | 0 | 28 | 1.99 | 2908.9 | 1.91
-51.4 | 0 | 28 | 2.67 | 3907.3 | 2.54
-50.2 | 0 | 28 | 3.33 | 4879.5 | 3.23
-49 | 0 | 28 | 4.24 | 6204.7 | 4.06
-47.8 | 0 | 28 | 5.24 | 7680.4 | 4.99
-46.6 | 0 | 28 | 6.26 | 9163.8 | 6.15
-45.4 | 0 | 28 | 7.86 | 11506.9 | 7.6
-44.2 | 0 | 28 | 9.58 | 14029.5 | 9.38
-43 | 0 | 28 | 11.72 | 17168.6 | 11.51
-41.8 | 0 | 28 | 14.69 | 21519.5 | 14.28
-40.6 | 0 | 28 | 18.16 | 26600 | 17.58
-39.4 | 0 | 28 | 22.25 | 32598.1 | 21.65
-38.2 | 0 | 28 | 27.52 | 40308.3 | 26.52
-37 | 0 | 28 | 34.04 | 49863.6 | 32.35
-35.8 | 0 | 28 | 40.8 | 59771.4 | 39.13
-34.6 | 0 | 28 | 49.2 | 72070.3 | 47
-33.4 | 0 | 28 | 58.19 | 85244 | 55.36
-32.2 | 0 | 28 | 66.58 | 97532.7 | 64.32
-31 | 0 | 28 | 76.3 | 111764.5 | 73.26
-29.8 | 0 | 28 | 85.53 | 125282.2 | 81.88
-28.6 | 0 | 28 | 92.47 | 135458 | 89.91
-27.4 | 0 | 28 | 100.51 | 147222.8 | 96.91
-26.2 | 0 | 28 | 107.22 | 157062.5 | 103.29
-25 | 0 | 28 | 110.98 | 162565.6 | 108.32
-23.8 | 0 | 28 | 116.72 | 170974.6 | 112.48
-22.6 | 0 | 28 | 120.09 | 175916.5 | 115.82
-21.4 | 0 | 28 | 121.54 | 178027.9 | 118.45
-20.2 | 0 | 28 | 122.99 | 180162.4 | 120.46
-19 | 0 | 28 | 125.33 | 183580.4 | 122.11
-17.7 | 0 | 28 | 127.01 | 186051.8 | 123.41
-16.5 | 0 | 28 | 127.77 | 187161.7 | 124.46
-15.4 | 0 | 28 | 129.26 | 189348.8 | 125.55
-14.1 | 0 | 28 | 129.29 | 189381.7 | 125.74
-12.9 | 0 | 28 | 131.87 | 193162.6 | 126.5
-11.7 | 0 | 28 | 132.04 | 193409.4 | 126.95
-10.5 | 0 | 28 | 133.1 | 194974.1 | 127.28
-9.3 | 0 | 28 | 133.25 | 195186.2 | 127.88
-8.1 | 0 | 28 | 133.91 | 196155 | 128.08
-6.9 | 0 | 28 | 132.99 | 194812.1 | 128.24
-5.7 | 0 | 28 | 133.99 | 196265.2 | 128.63
-4.5 | 0 | 28 | 134.48 | 196994.4 | 128.86
-3.3 | 0 | 28 | 133.68 | 195819.5 | 128.82
-2.1 | 0 | 28 | 132.64 | 194292.5 | 129.02
-0.9 | 0 | 28 | 134.3 | 196721.9 | 129.01
0.3 | 0 | 28 | 132.49 | 194073.7 | 128.99
1.5 | 0 | 28 | 134.09 | 196420.4 | 129.04
2.7 | 0 | 28 | 131.7 | 192910 | 128.93
3.9 | 0 | 28 | 133.27 | 195211.8 | 128.69
5.1 | 0 | 28 | 137.14 | 200886.5 | 128.92
6.3 | 0 | 28 | 135.58 | 198606.6 | 128.57
7.5 | 0 | 28 | 131.93 | 193250.5 | 128.22
8.7 | 0 | 28 | 134.04 | 196339.9 | 127.76
9.9 | 0 | 28 | 128.72 | 188545.9 | 127.42
11.1 | 0 | 28 | 129.89 | 190262.6 | 126.68
12.3 | 0 | 28 | 126.48 | 185266.5 | 126.16
13.5 | 0 | 28 | 130.83 | 191648.8 | 125.53
14.7 | 0 | 28 | 129.49 | 189675.5 | 124.72
15.9 | 0 | 28 | 128.51 | 188237.1 | 123.43
17.1 | 0 | 28 | 127.74 | 187121.5 | 122.02
18.3 | 0 | 28 | 123.53 | 180954.5 | 120.43
19.5 | 0 | 28 | 123.16 | 180412.7 | 118.48
20.7 | 0 | 28 | 119.27 | 174711.3 | 115.93
21.9 | 0 | 28 | 117.05 | 171461.4 | 112.46
23.1 | 0 | 28 | 113.27 | 165920.5 | 108.84
24.3 | 0 | 28 | 108.72 | 159260.7 | 103.81
25.5 | 0 | 28 | 102.06 | 149502.7 | 98.17
26.7 | 0 | 28 | 94.99 | 139141.5 | 91.28
27.9 | 0 | 28 | 87.07 | 127543.6 | 83.53
29.1 | 0 | 28 | 78.02 | 114286.6 | 74.89
30.3 | 0 | 28 | 67.86 | 99397.4 | 65.85
31.5 | 0 | 28 | 58.62 | 85865.2 | 56.81
32.7 | 0 | 28 | 50.43 | 73871 | 47.98
33.9 | 0 | 28 | 40.62 | 59497.2 | 39.61
35.1 | 0 | 28 | 32.48 | 47572 | 32.63
36.3 | 0 | 28 | 27.2 | 39845.2 | 26.42
37.5 | 0 | 28 | 21.96 | 32171.1 | 21.3
38.7 | 0 | 28 | 17.68 | 25904.9 | 17.18
39.9 | 0 | 28 | 14.25 | 20878.7 | 13.84
41.1 | 0 | 28 | 11.48 | 16812.3 | 11.17
42.4 | 0 | 28 | 9.26 | 13561.9 | 8.97
43.5 | 0 | 28 | 7.43 | 10888.6 | 7.16
44.8 | 0 | 28 | 6.15 | 9005 | 5.93
46 | 0 | 28 | 4.87 | 7132.1 | 4.69
47.2 | 0 | 28 | 3.89 | 5698.8 | 3.73
48.4 | 0 | 28 | 3.13 | 4585.4 | 3
49.6 | 0 | 28 | 2.52 | 3694.8 | 2.42
50.8 | 0 | 28 | 1.85 | 2713.5 | 1.78
52 | 0 | 28 | 1.46 | 2139 | 1.4
53.2 | 0 | 28 | 1.07 | 1566 | 1.03
54.4 | 0 | 28 | 0.74 | 1085.6 | 0.71
55.6 | 0 | 28 | 0.43 | 630.7 | 0.41
56.8 | 0 | 28 | 0.15 | 222.8 | 0.15
58 | 0 | 28 | 0.05 | 74.5 | 0.05
59.2 | 0 | 28 | 0.28 | 411.3 | 0.27
60.4 | 0 | 28 | 0.49 | 719.4 | 0.47
61.6 | 0 | 28 | 0.63 | 922 | 0.61
62.8 | 0 | 28 | 0.68 | 1001.4 | 0.66
64 | 0 | 28 | 0.84 | 1233.8 | 0.81
65.2 | 0 | 28 | 1.04 | 1518.4 | 1
66.4 | 0 | 28 | 1.07 | 1567.6 | 1.04
67.6 | 0 | 28 | 1.24 | 1810.6 | 1.2
68.8 | 0 | 28 | 1.26 | 1844.4 | 1.2
70 | 0 | 28 | 1.3 | 1902.1 | 1.26
71.2 | 0 | 28 | 1.41 | 2068.2 | 1.41
72.5 | 0 | 28 | 1.42 | 2082 | 1.35
73.7 | 0 | 28 | 1.54 | 2258.6 | 1.52
74.8 | 0 | 28 | 1.62 | 2375.6 | 1.55
76.1 | 0 | 28 | 1.6 | 2347.4 | 1.56
77.2 | 0 | 28 | 1.65 | 2412 | 1.59
78.5 | 0 | 28 | 1.65 | 2421.4 | 1.59
79.6 | 0 | 28 | 1.64 | 2399.3 | 1.57
80.8 | 0 | 28 | 1.7 | 2487 | 1.62
82.1 | 0 | 28 | 1.7 | 2488.2 | 1.62
83.3 | 0 | 28 | 1.66 | 2438.5 | 1.6
84.5 | 0 | 28 | 1.72 | 2515.5 | 1.65
85.7 | 0 | 28 | 1.76 | 2579.3 | 1.7
86.9 | 0 | 28 | 1.73 | 2537.1 | 1.66
88.1 | 0 | 28 | 1.77 | 2585.8 | 1.7
89.3 | 0 | 28 | 1.75 | 2559.3 | 1.69
90.5 | 0 | 28 | 1.76 | 2582.8 | 1.69
91.7 | 0 | 28 | 1.75 | 2563.5 | 1.69
92.9 | 0 | 28 | 1.8 | 2630.2 | 1.73
94.1 | 0 | 28 | 1.86 | 2717.7 | 1.81
95.3 | 0 | 28 | 1.84 | 2692.8 | 1.78
96.5 | 0 | 28 | 1.84 | 2699.5 | 1.8
97.7 | 0 | 28 | 1.84 | 2692.7 | 1.79
98.9 | 0 | 28 | 1.81 | 2654.6 | 1.77
100.1 | 0 | 28 | 1.81 | 2651.8 | 1.81
101.3 | 0 | 28 | 1.78 | 2601.8 | 1.67
102.5 | 0 | 28 | 1.85 | 2709.7 | 1.8
103.7 | 0 | 28 | 1.82 | 2666.7 | 1.8
104.9 | 0 | 28 | 1.8 | 2642.4 | 1.73
106.1 | 0 | 28 | 1.88 | 2754.5 | 1.86
107.3 | 0 | 28 | 1.8 | 2631.7 | 1.75
108.5 | 0 | 28 | 1.75 | 2569.1 | 1.69
109.7 | 0 | 28 | 1.82 | 2664.1 | 1.78
110.9 | 0 | 28 | 1.77 | 2589.6 | 1.7
112.1 | 0 | 28 | 1.83 | 2683.1 | 1.76
113.3 | 0 | 28 | 1.83 | 2682.5 | 1.76
114.5 | 0 | 28 | 1.86 | 2719.3 | 1.79
115.7 | 0 | 28 | 1.79 | 2629 | 1.73
116.9 | 0 | 28 | 1.83 | 2683.1 | 1.78
118.1 | 0 | 28 | 1.86 | 2728.6 | 1.78
119.3 | 0 | 28 | 1.83 | 2677.6 | 1.75
120.5 | 0 | 28 | 1.83 | 2683.8 | 1.76
121.7 | 0 | 28 | 1.85 | 2705.6 | 1.76
122.9 | 0 | 28 | 1.87 | 2735.5 | 1.78
124.2 | 0 | 28 | 1.89 | 2762.3 | 1.83
125.7 | 0 | 28 | 1.87 | 2740.4 | 1.81