## Taille de champ 10x10 12 MeV
Measurement time: | 02/09/2022 20:43:26
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 200 | 1.1 | 1604.3 | 0.93
0 | 0 | 199.3 | 1.05 | 1542.5 | 0.94
0 | 0 | 197.9 | 1.09 | 1600.2 | 1.01
0 | 0 | 196.7 | 1.07 | 1566.2 | 0.95
0 | 0 | 195.5 | 1.06 | 1557.7 | 0.97
0 | 0 | 194.3 | 1.02 | 1499.1 | 0.92
0 | 0 | 193.1 | 1.07 | 1573.1 | 1
0 | 0 | 191.9 | 1.06 | 1555.9 | 0.97
0 | 0 | 190.7 | 1.01 | 1485.5 | 0.92
0 | 0 | 189.5 | 0.98 | 1439.7 | 0.88
0 | 0 | 188.3 | 1.11 | 1622.3 | 1.06
0 | 0 | 187.1 | 0.98 | 1435.2 | 0.87
0 | 0 | 185.9 | 0.99 | 1453.1 | 0.91
0 | 0 | 184.7 | 0.96 | 1405.1 | 0.86
0 | 0 | 183.5 | 0.94 | 1375.6 | 0.84
0 | 0 | 182.3 | 0.99 | 1454.7 | 0.9
0 | 0 | 181.1 | 0.92 | 1343.5 | 0.83
0 | 0 | 179.9 | 0.92 | 1353.5 | 0.83
0 | 0 | 178.7 | 0.96 | 1407.1 | 0.87
0 | 0 | 177.5 | 0.94 | 1380.5 | 0.85
0 | 0 | 176.3 | 0.94 | 1384.1 | 0.86
0 | 0 | 175.1 | 0.96 | 1402.8 | 0.86
0 | 0 | 173.9 | 0.9 | 1324.4 | 0.82
0 | 0 | 172.7 | 0.94 | 1376 | 0.85
0 | 0 | 171.5 | 0.93 | 1366 | 0.84
0 | 0 | 170.3 | 0.9 | 1316 | 0.81
0 | 0 | 169.1 | 0.86 | 1253.2 | 0.77
0 | 0 | 167.9 | 0.87 | 1273.1 | 0.79
0 | 0 | 166.7 | 0.84 | 1233.5 | 0.76
0 | 0 | 165.5 | 0.88 | 1293.2 | 0.8
0 | 0 | 164.3 | 0.87 | 1270.9 | 0.78
0 | 0 | 163.1 | 0.84 | 1227.7 | 0.76
0 | 0 | 161.9 | 0.77 | 1130.2 | 0.69
0 | 0 | 160.7 | 0.79 | 1151.5 | 0.71
0 | 0 | 159.5 | 0.83 | 1210.5 | 0.75
0 | 0 | 158.3 | 0.84 | 1228.7 | 0.75
0 | 0 | 157.1 | 0.77 | 1129.6 | 0.7
0 | 0 | 155.9 | 0.78 | 1137.4 | 0.7
0 | 0 | 154.7 | 0.76 | 1119.1 | 0.69
0 | 0 | 153.5 | 0.74 | 1077.2 | 0.66
0 | 0 | 152.3 | 0.72 | 1056.5 | 0.65
0 | 0 | 151.1 | 0.76 | 1117.1 | 0.69
0 | 0 | 149.8 | 0.75 | 1104 | 0.68
0 | 0 | 148.7 | 0.69 | 1013.4 | 0.62
0 | 0 | 147.4 | 0.73 | 1070.4 | 0.66
0 | 0 | 146.3 | 0.64 | 941.9 | 0.58
0 | 0 | 145 | 0.66 | 972 | 0.6
0 | 0 | 143.8 | 0.72 | 1058.8 | 0.68
0 | 0 | 142.6 | 0.62 | 910.7 | 0.56
0 | 0 | 141.4 | 0.68 | 998.9 | 0.61
0 | 0 | 140.3 | 0.68 | 995.6 | 0.64
0 | 0 | 139 | 0.6 | 881.3 | 0.54
0 | 0 | 137.8 | 0.57 | 836.9 | 0.52
0 | 0 | 136.6 | 0.62 | 908.5 | 0.56
0 | 0 | 135.4 | 0.58 | 855.7 | 0.53
0 | 0 | 134.3 | 0.62 | 911.9 | 0.58
0 | 0 | 133 | 0.6 | 883.6 | 0.54
0 | 0 | 131.8 | 0.5 | 738.4 | 0.46
0 | 0 | 130.6 | 0.51 | 750.9 | 0.46
0 | 0 | 129.4 | 0.62 | 903.6 | 0.58
0 | 0 | 128.2 | 0.46 | 680.1 | 0.42
0 | 0 | 127 | 0.44 | 650.4 | 0.4
0 | 0 | 125.8 | 0.46 | 670.7 | 0.41
0 | 0 | 124.6 | 0.43 | 634 | 0.39
0 | 0 | 123.4 | 0.47 | 691.9 | 0.42
0 | 0 | 122.2 | 0.43 | 629.5 | 0.39
0 | 0 | 121 | 0.4 | 590.9 | 0.36
0 | 0 | 119.8 | 0.4 | 581.8 | 0.36
0 | 0 | 118.6 | 0.41 | 607 | 0.37
0 | 0 | 117.4 | 0.4 | 579.2 | 0.36
0 | 0 | 116.2 | 0.29 | 425.9 | 0.26
0 | 0 | 115 | 0.37 | 548.2 | 0.34
0 | 0 | 113.8 | 0.34 | 503.2 | 0.31
0 | 0 | 112.6 | 0.32 | 466.4 | 0.29
0 | 0 | 111.4 | 0.27 | 399.2 | 0.25
0 | 0 | 110.2 | 0.27 | 399.8 | 0.25
0 | 0 | 109 | 0.28 | 414.7 | 0.26
0 | 0 | 107.8 | 0.29 | 428 | 0.27
0 | 0 | 106.6 | 0.3 | 433.7 | 0.27
0 | 0 | 105.4 | 0.3 | 444.1 | 0.28
0 | 0 | 104.2 | 0.22 | 318.3 | 0.2
0 | 0 | 103 | 0.21 | 309.5 | 0.19
0 | 0 | 101.8 | 0.2 | 292.2 | 0.18
0 | 0 | 100.6 | 0.18 | 259.1 | 0.16
0 | 0 | 99.4 | 0.13 | 185.4 | 0.11
0 | 0 | 98.2 | 0.18 | 270.2 | 0.17
0 | 0 | 97 | 0.16 | 227.6 | 0.14
0 | 0 | 95.8 | 0.16 | 239.6 | 0.15
0 | 0 | 94.6 | 0.11 | 163.9 | 0.1
0 | 0 | 93.4 | 0.18 | 260.9 | 0.17
0 | 0 | 92.2 | 0.06 | 86.9 | 0.05
0 | 0 | 91 | 0.07 | 107.1 | 0.07
0 | 0 | 89.8 | 0.12 | 173.7 | 0.11
0 | 0 | 88.6 | 0.02 | 25.8 | 0.02
0 | 0 | 87.4 | 0.07 | 95.7 | 0.06
0 | 0 | 86.2 | 0.05 | 77.3 | 0.05
0 | 0 | 84.9 | 0.02 | 25 | 0.02
0 | 0 | 83.7 | 0.05 | 71.4 | 0.05
0 | 0 | 82.5 | 0.03 | 41 | 0.03
0 | 0 | 81.3 | 0.1 | 152.8 | 0.09
0 | 0 | 80.1 | 0.08 | 115.1 | 0.07
0 | 0 | 78.9 | 0.05 | 71.9 | 0.05
0 | 0 | 77.7 | 0.17 | 254.6 | 0.16
0 | 0 | 76.5 | 0.17 | 251.3 | 0.16
0 | 0 | 75.3 | 0.16 | 237.6 | 0.15
0 | 0 | 74.1 | 0.17 | 243.3 | 0.15
0 | 0 | 72.9 | 0.21 | 300.5 | 0.19
0 | 0 | 71.7 | 0.29 | 426.9 | 0.27
0 | 0 | 70.5 | 0.34 | 499.2 | 0.31
0 | 0 | 69.3 | 0.46 | 673.1 | 0.42
0 | 0 | 68.1 | 0.79 | 1158.5 | 0.72
0 | 0 | 66.9 | 1.38 | 2026.8 | 1.25
0 | 0 | 65.7 | 2.12 | 3109.8 | 1.92
0 | 0 | 64.5 | 3.37 | 4931.4 | 3.04
0 | 0 | 63.3 | 5.19 | 7608.3 | 4.69
0 | 0 | 62.1 | 7.55 | 11066.4 | 6.85
0 | 0 | 60.8 | 10.98 | 16080.2 | 9.91
0 | 0 | 59.7 | 14.27 | 20904.7 | 13.38
0 | 0 | 58.4 | 19.79 | 28992.6 | 17.9
0 | 0 | 57.2 | 25.22 | 36949.4 | 22.79
0 | 0 | 56 | 30.31 | 44392.6 | 28.44
0 | 0 | 54.8 | 38.59 | 56527.6 | 34.94
0 | 0 | 53.6 | 45.51 | 66657 | 40.99
0 | 0 | 52.4 | 53.63 | 78558.7 | 48.16
0 | 0 | 51.2 | 61.45 | 90013.5 | 55.19
0 | 0 | 50 | 65.69 | 96218.7 | 61.85
0 | 0 | 48.8 | 76.44 | 111968.5 | 68.96
0 | 0 | 47.6 | 80.1 | 117331.1 | 75.18
0 | 0 | 46.4 | 90.47 | 132517.1 | 81.57
0 | 0 | 45.2 | 97.67 | 143070.4 | 87.92
0 | 0 | 44 | 103.27 | 151265.4 | 93.16
0 | 0 | 42.8 | 109.41 | 160262.5 | 98.42
0 | 0 | 41.6 | 114.12 | 167167.8 | 103.01
0 | 0 | 40.4 | 117.56 | 172205.4 | 107.08
0 | 0 | 39.1 | 123.98 | 181601.6 | 110.68
0 | 0 | 38 | 121.58 | 178097.1 | 113.72
0 | 0 | 36.7 | 126.83 | 185776.9 | 116.37
0 | 0 | 35.5 | 130.88 | 191719.4 | 118.56
0 | 0 | 34.3 | 136.7 | 200234.5 | 120.15
0 | 0 | 33.1 | 131.26 | 192270 | 121.43
0 | 0 | 31.9 | 135.74 | 198841.1 | 122.48
0 | 0 | 30.7 | 134.35 | 196805.6 | 122.96
0 | 0 | 29.5 | 132.75 | 194453.1 | 123.32
0 | 0 | 28.3 | 136.89 | 200519.1 | 123.6
0 | 0 | 27.1 | 136.4 | 199805 | 123.46
0 | 0 | 25.9 | 131.65 | 192836.3 | 123.16
0 | 0 | 24.7 | 137.14 | 200883 | 122.97
0 | 0 | 23.5 | 134.76 | 197393.9 | 122.56
0 | 0 | 22.3 | 130.53 | 191209.7 | 122.06
0 | 0 | 21.1 | 135.69 | 198755.3 | 121.82
0 | 0 | 19.9 | 132.96 | 194768.5 | 121.19
0 | 0 | 18.7 | 129.28 | 189372.4 | 120.27
0 | 0 | 17.5 | 133.46 | 195500.1 | 120.03
0 | 0 | 16.3 | 137.14 | 200889.6 | 119.7
0 | 0 | 15 | 127.61 | 186924.4 | 118.61
0 | 0 | 13.9 | 131.65 | 192841.4 | 118.49
0 | 0 | 12.6 | 134.14 | 196488 | 118.05
0 | 0 | 11.5 | 127.06 | 186119.7 | 117.25
0 | 0 | 10.2 | 129 | 188955.1 | 116.62
0 | 0 | 9 | 132.81 | 194537.2 | 116.36
0 | 0 | 7.8 | 123.52 | 180937.2 | 115.35
0 | 0 | 6.6 | 127.04 | 186090 | 114.53
0 | 0 | 5.4 | 130.24 | 190784.1 | 113.83
0 | 0 | 4.2 | 119.43 | 174941.7 | 112.26
0 | 0 | 3 | 122.51 | 179456.8 | 111.07
0 | 0 | 1.8 | 124.94 | 183021.5 | 110
0 | 0 | 0.6 | 114.8 | 168157.9 | 108.8
0 | 0 | -0.5 | 118.85 | 174099.7 | 108.57
Measurement time: | 02/09/2022 20:44:47
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-151.4 | 0 | 28 | 1.76 | 2581.1 | 1.94
-150.6 | 0 | 28 | 1.72 | 2526.5 | 1.58
-149.3 | 0 | 28 | 1.71 | 2502 | 1.56
-148.1 | 0 | 28 | 1.73 | 2530.3 | 1.51
-146.9 | 0 | 28 | 1.73 | 2530.4 | 1.61
-145.7 | 0 | 28 | 1.71 | 2505.3 | 1.55
-144.5 | 0 | 28 | 1.75 | 2566 | 1.51
-143.3 | 0 | 28 | 1.83 | 2679.8 | 1.73
-142.1 | 0 | 28 | 1.76 | 2571.7 | 1.65
-140.9 | 0 | 28 | 1.7 | 2495.2 | 1.47
-139.7 | 0 | 28 | 1.77 | 2588.1 | 1.66
-138.5 | 0 | 28 | 1.77 | 2589.5 | 1.68
-137.3 | 0 | 28 | 1.67 | 2450.5 | 1.44
-136.1 | 0 | 28 | 1.72 | 2524.8 | 1.63
-134.9 | 0 | 28 | 1.77 | 2591.1 | 1.67
-133.7 | 0 | 28 | 1.74 | 2547.9 | 1.51
-132.5 | 0 | 28 | 1.74 | 2549.7 | 1.62
-131.3 | 0 | 28 | 1.72 | 2519.9 | 1.56
-130.1 | 0 | 28 | 1.78 | 2606.6 | 1.54
-128.9 | 0 | 28 | 1.77 | 2589.6 | 1.67
-127.7 | 0 | 28 | 1.79 | 2627.4 | 1.62
-126.5 | 0 | 28 | 1.73 | 2535.8 | 1.51
-125.3 | 0 | 28 | 1.74 | 2544.3 | 1.65
-124.1 | 0 | 28 | 1.69 | 2482.2 | 1.55
-122.9 | 0 | 28 | 1.73 | 2533.2 | 1.52
-121.7 | 0 | 28 | 1.74 | 2551.5 | 1.66
-120.5 | 0 | 28 | 1.73 | 2538.4 | 1.53
-119.3 | 0 | 28 | 1.73 | 2540.7 | 1.61
-118.1 | 0 | 28 | 1.71 | 2504.9 | 1.55
-116.9 | 0 | 28 | 1.71 | 2509.6 | 1.63
-115.7 | 0 | 28 | 1.69 | 2473.8 | 1.49
-114.5 | 0 | 28 | 1.8 | 2631.3 | 1.67
-113.3 | 0 | 28 | 1.68 | 2465.4 | 1.58
-112.1 | 0 | 28 | 1.67 | 2441 | 1.48
-110.9 | 0 | 28 | 1.68 | 2456.4 | 1.54
-109.6 | 0 | 28 | 1.71 | 2504.8 | 1.6
-108.4 | 0 | 28 | 1.62 | 2373.4 | 1.45
-107.2 | 0 | 28 | 1.6 | 2349.4 | 1.48
-106 | 0 | 28 | 1.63 | 2381.6 | 1.52
-104.9 | 0 | 28 | 1.58 | 2307.3 | 1.4
-103.6 | 0 | 28 | 1.57 | 2299.4 | 1.45
-102.4 | 0 | 28 | 1.6 | 2341.7 | 1.44
-101.2 | 0 | 28 | 1.54 | 2249.7 | 1.36
-100 | 0 | 28 | 1.59 | 2329.3 | 1.46
-98.8 | 0 | 28 | 1.48 | 2166 | 1.33
-97.6 | 0 | 28 | 1.42 | 2082.6 | 1.25
-96.4 | 0 | 28 | 1.42 | 2075.7 | 1.31
-95.2 | 0 | 28 | 1.42 | 2079.8 | 1.29
-94 | 0 | 28 | 1.42 | 2079.3 | 1.32
-92.8 | 0 | 28 | 1.3 | 1906.7 | 1.19
-91.6 | 0 | 28 | 1.22 | 1793.7 | 1.1
-90.4 | 0 | 28 | 1.13 | 1651.3 | 1
-89.2 | 0 | 28 | 1.07 | 1565.8 | 0.98
-88 | 0 | 28 | 0.9 | 1319.6 | 0.81
-86.8 | 0 | 28 | 0.8 | 1164.9 | 0.71
-85.6 | 0 | 28 | 0.71 | 1046.1 | 0.65
-84.4 | 0 | 28 | 0.67 | 978.5 | 0.63
-83.2 | 0 | 28 | 0.44 | 650.2 | 0.4
-82 | 0 | 28 | 0.24 | 345.7 | 0.21
-80.8 | 0 | 28 | 0.01 | 15.2 | 0.01
-79.6 | 0 | 28 | 0.21 | 307.1 | 0.19
-78.4 | 0 | 28 | 0.54 | 787.5 | 0.49
-77.2 | 0 | 28 | 0.86 | 1255 | 0.77
-76 | 0 | 28 | 1.27 | 1860.9 | 1.15
-74.8 | 0 | 28 | 1.74 | 2554.7 | 1.59
-73.6 | 0 | 28 | 2.33 | 3418.8 | 2.15
-72.4 | 0 | 28 | 3.01 | 4415.3 | 2.74
-71.2 | 0 | 28 | 3.69 | 5404.3 | 3.39
-70 | 0 | 28 | 4.73 | 6928.3 | 4.34
-68.8 | 0 | 28 | 5.76 | 8443.7 | 5.29
-67.6 | 0 | 28 | 7.23 | 10583.4 | 6.61
-66.4 | 0 | 28 | 8.86 | 12980.4 | 8.12
-65.2 | 0 | 28 | 10.97 | 16075.8 | 9.97
-63.9 | 0 | 28 | 13.35 | 19553 | 12.15
-62.8 | 0 | 28 | 16.1 | 23577.8 | 14.68
-61.5 | 0 | 28 | 20.37 | 29837.9 | 18.21
-60.3 | 0 | 28 | 24.25 | 35515.2 | 22.12
-59.1 | 0 | 28 | 29.9 | 43803.2 | 27
-57.9 | 0 | 28 | 36.28 | 53146.8 | 32.75
-56.7 | 0 | 28 | 43.08 | 63106.3 | 39.29
-55.5 | 0 | 28 | 52.41 | 76767.7 | 46.75
-54.3 | 0 | 28 | 58 | 84962.2 | 54.44
-53.1 | 0 | 28 | 69.25 | 101444.7 | 62.91
-51.9 | 0 | 28 | 79.02 | 115743.6 | 71.06
-50.7 | 0 | 28 | 86.8 | 127141.9 | 78.95
-49.5 | 0 | 28 | 95.39 | 139734.3 | 86.22
-48.3 | 0 | 28 | 100.92 | 147834.9 | 92.47
-47.1 | 0 | 28 | 107.55 | 157547.6 | 98.13
-45.9 | 0 | 28 | 112.24 | 164407.3 | 102.62
-44.7 | 0 | 28 | 115.96 | 169859.7 | 106.14
-43.5 | 0 | 28 | 119.52 | 175068.4 | 109.36
-42.3 | 0 | 28 | 121.97 | 178662.5 | 111.54
-41.1 | 0 | 28 | 124.96 | 183044 | 113.63
-39.9 | 0 | 28 | 128.95 | 188887.2 | 115.3
-38.7 | 0 | 28 | 126.97 | 185986.1 | 116.45
-37.5 | 0 | 28 | 131.39 | 192459.4 | 117.6
-36.3 | 0 | 28 | 131.67 | 192877.3 | 118.29
-35.1 | 0 | 28 | 130.81 | 191611.9 | 119.08
-33.9 | 0 | 28 | 126.7 | 185596.5 | 119.18
-32.7 | 0 | 28 | 132.69 | 194362.4 | 120.21
-31.5 | 0 | 28 | 135.43 | 198380.1 | 120.61
-30.3 | 0 | 28 | 131.52 | 192648.6 | 120.76
-29.1 | 0 | 28 | 133.63 | 195742.4 | 121.27
-27.9 | 0 | 28 | 133.33 | 195301.7 | 121.35
-26.7 | 0 | 28 | 132.89 | 194654.7 | 121.62
-25.5 | 0 | 28 | 136.53 | 199995.8 | 122.07
-24.3 | 0 | 28 | 133.36 | 195342.4 | 122.05
-23.1 | 0 | 28 | 130.51 | 191170.9 | 122.13
-21.9 | 0 | 28 | 134.22 | 196604.4 | 122.41
-20.7 | 0 | 28 | 134.82 | 197493.1 | 122.8
-19.5 | 0 | 28 | 134.3 | 196732 | 122.79
-18.3 | 0 | 28 | 135.67 | 198726.8 | 122.92
-17 | 0 | 28 | 135.52 | 198509.6 | 122.81
-15.8 | 0 | 28 | 136.51 | 199966.6 | 122.86
-14.6 | 0 | 28 | 129.76 | 190072 | 123.01
-13.4 | 0 | 28 | 137.47 | 201362.9 | 123.06
-12.2 | 0 | 28 | 136.58 | 200059.9 | 122.97
-11 | 0 | 28 | 130.46 | 191098 | 123.08
-9.8 | 0 | 28 | 137.42 | 201295 | 123.23
-8.6 | 0 | 28 | 137.13 | 200874.9 | 123.19
-7.4 | 0 | 28 | 136.26 | 199592 | 123.32
-6.2 | 0 | 28 | 137.02 | 200711.1 | 123.44
-5 | 0 | 28 | 137.34 | 201171.9 | 123.13
-3.8 | 0 | 28 | 130.05 | 190500.8 | 123.17
-2.6 | 0 | 28 | 137.63 | 201597.9 | 123.46
-1.4 | 0 | 28 | 136.49 | 199935.9 | 123.21
-0.2 | 0 | 28 | 136.8 | 200386.7 | 123.36
1 | 0 | 28 | 136.64 | 200154.9 | 123.25
2.2 | 0 | 28 | 135.82 | 198954.9 | 123.36
3.4 | 0 | 28 | 135.96 | 199159.9 | 123.52
4.6 | 0 | 28 | 135.68 | 198749.5 | 123.37
5.8 | 0 | 28 | 135.01 | 197769.5 | 123.37
7 | 0 | 28 | 135.8 | 198923.4 | 123.37
8.2 | 0 | 28 | 134.75 | 197379.3 | 123.23
9.4 | 0 | 28 | 134.72 | 197337.5 | 123.17
10.6 | 0 | 28 | 134.49 | 196997.1 | 123.27
11.8 | 0 | 28 | 139.41 | 204215.9 | 123.14
13 | 0 | 28 | 133.72 | 195882.7 | 122.8
14.2 | 0 | 28 | 127.98 | 187471.3 | 122.66
15.4 | 0 | 28 | 139.75 | 204704.5 | 122.98
16.6 | 0 | 28 | 133.45 | 195479.2 | 122.59
17.8 | 0 | 28 | 127.96 | 187434.1 | 122.56
19 | 0 | 28 | 139.58 | 204460 | 122.81
20.2 | 0 | 28 | 132.73 | 194429.6 | 122.25
21.4 | 0 | 28 | 131.15 | 192112.8 | 122.31
22.6 | 0 | 28 | 133.38 | 195381 | 122.13
23.8 | 0 | 28 | 132.62 | 194262.1 | 121.79
25 | 0 | 28 | 128.1 | 187643.9 | 121.56
26.2 | 0 | 28 | 136.81 | 200405.7 | 121.43
27.4 | 0 | 28 | 127.93 | 187396.9 | 120.77
28.6 | 0 | 28 | 131.63 | 192811.3 | 120.81
29.8 | 0 | 28 | 135.84 | 198979.3 | 120.51
31 | 0 | 28 | 130.28 | 190834.3 | 119.79
32.2 | 0 | 28 | 125.7 | 184131.4 | 119.02
33.4 | 0 | 28 | 133.01 | 194831.8 | 118.55
34.6 | 0 | 28 | 129.35 | 189476.7 | 117.55
35.8 | 0 | 28 | 123.04 | 180229.1 | 116.33
37 | 0 | 28 | 128.87 | 188770.2 | 115.21
38.2 | 0 | 28 | 120.54 | 176573.1 | 113.27
39.4 | 0 | 28 | 120.95 | 177175.5 | 111.89
40.6 | 0 | 28 | 123.67 | 181155.8 | 109.43
41.8 | 0 | 28 | 112.19 | 164335.3 | 106.03
43 | 0 | 28 | 112.45 | 164723.3 | 102.69
44.2 | 0 | 28 | 107.69 | 157750.8 | 97.82
45.4 | 0 | 28 | 102.23 | 149752.8 | 92.56
46.7 | 0 | 28 | 94.94 | 139073.2 | 86.45
47.9 | 0 | 28 | 87.18 | 127706.4 | 79.47
49 | 0 | 28 | 79.03 | 115758.9 | 71.64
50.3 | 0 | 28 | 70.06 | 102621.9 | 63.5
51.5 | 0 | 28 | 61.22 | 89682 | 55.27
52.7 | 0 | 28 | 51.94 | 76086.8 | 47
53.9 | 0 | 28 | 44.1 | 64595 | 39.68
55.1 | 0 | 28 | 36.25 | 53099.6 | 32.75
56.3 | 0 | 28 | 29.9 | 43795.1 | 26.91
57.5 | 0 | 28 | 24.34 | 35657.7 | 21.93
58.7 | 0 | 28 | 19.78 | 28970.7 | 17.84
59.9 | 0 | 28 | 16.01 | 23455.8 | 14.45
61.1 | 0 | 28 | 12.25 | 17948 | 11.57
62.3 | 0 | 28 | 10.65 | 15602.3 | 9.59
63.5 | 0 | 28 | 8.69 | 12735.1 | 7.79
64.7 | 0 | 28 | 6.54 | 9573.4 | 6.22
65.9 | 0 | 28 | 5.73 | 8387 | 5.15
67.1 | 0 | 28 | 4.58 | 6711.7 | 4.14
68.3 | 0 | 28 | 3.7 | 5420.7 | 3.35
69.5 | 0 | 28 | 3 | 4394.6 | 2.71
70.7 | 0 | 28 | 2.32 | 3398 | 2.11
71.9 | 0 | 28 | 1.85 | 2708.1 | 1.68
73.1 | 0 | 28 | 1.34 | 1956.5 | 1.22
74.3 | 0 | 28 | 0.95 | 1385.7 | 0.86
75.5 | 0 | 28 | 0.71 | 1040.4 | 0.65
76.7 | 0 | 28 | 0.36 | 521.2 | 0.33
77.9 | 0 | 28 | 0.02 | 33.1 | 0.02
79.1 | 0 | 28 | 0.11 | 158.2 | 0.1
80.3 | 0 | 28 | 0.32 | 475.5 | 0.29
81.5 | 0 | 28 | 0.53 | 778 | 0.49
82.7 | 0 | 28 | 0.66 | 968.7 | 0.6
84 | 0 | 28 | 0.77 | 1121.3 | 0.69
85.2 | 0 | 28 | 1.01 | 1483.7 | 0.93
86.3 | 0 | 28 | 1.08 | 1578.4 | 0.98
87.6 | 0 | 28 | 1.12 | 1642.9 | 1
88.8 | 0 | 28 | 1.25 | 1825.8 | 1.15
90 | 0 | 28 | 1.36 | 1993.8 | 1.24
91.2 | 0 | 28 | 1.4 | 2047.8 | 1.24
92.4 | 0 | 28 | 1.5 | 2197.4 | 1.39
93.6 | 0 | 28 | 1.51 | 2215.4 | 1.35
94.8 | 0 | 28 | 1.59 | 2323.3 | 1.41
96 | 0 | 28 | 1.58 | 2309 | 1.45
97.2 | 0 | 28 | 1.67 | 2445.3 | 1.49
98.4 | 0 | 28 | 1.68 | 2454.9 | 1.49
99.6 | 0 | 28 | 1.7 | 2486.5 | 1.55
100.8 | 0 | 28 | 1.72 | 2526.8 | 1.64
102 | 0 | 28 | 1.72 | 2516.9 | 1.51
103.2 | 0 | 28 | 1.74 | 2552.4 | 1.61
104.4 | 0 | 28 | 1.78 | 2601.1 | 1.68
105.6 | 0 | 28 | 1.75 | 2556.5 | 1.53
106.8 | 0 | 28 | 1.8 | 2636 | 1.68
108 | 0 | 28 | 1.79 | 2628.3 | 1.63
109.2 | 0 | 28 | 1.77 | 2592.7 | 1.55
110.4 | 0 | 28 | 1.81 | 2647.1 | 1.71
111.6 | 0 | 28 | 1.84 | 2689.2 | 1.67
112.8 | 0 | 28 | 1.84 | 2696.7 | 1.63
114 | 0 | 28 | 1.82 | 2659.2 | 1.73
115.2 | 0 | 28 | 1.86 | 2719.4 | 1.7
116.4 | 0 | 28 | 1.79 | 2623.8 | 1.57
117.6 | 0 | 28 | 1.9 | 2778.8 | 1.8
118.8 | 0 | 28 | 1.88 | 2750.3 | 1.71
120 | 0 | 28 | 1.83 | 2683.6 | 1.6
121.3 | 0 | 28 | 1.85 | 2713.4 | 1.77
122.4 | 0 | 28 | 1.85 | 2703.6 | 1.68
123.6 | 0 | 28 | 1.84 | 2693.1 | 1.61
124.9 | 0 | 28 | 1.97 | 2887.6 | 1.89
126 | 0 | 28 | 1.89 | 2773.6 | 1.71
127.2 | 0 | 28 | 1.89 | 2762.5 | 1.64
128.4 | 0 | 28 | 1.9 | 2779.7 | 1.78
129.7 | 0 | 28 | 1.91 | 2802.2 | 1.72
130.8 | 0 | 28 | 1.93 | 2833.1 | 1.76
132.1 | 0 | 28 | 1.92 | 2808.4 | 1.8
133.3 | 0 | 28 | 1.96 | 2877.7 | 1.77
134.5 | 0 | 28 | 1.96 | 2864.8 | 1.7
135.7 | 0 | 28 | 1.96 | 2865 | 1.84
136.9 | 0 | 28 | 1.94 | 2849 | 1.75
138.1 | 0 | 28 | 1.86 | 2721.4 | 1.62
139.3 | 0 | 28 | 1.96 | 2877.7 | 1.84
140.5 | 0 | 28 | 1.96 | 2867.1 | 1.85
141.7 | 0 | 28 | 1.91 | 2802.6 | 1.65
142.9 | 0 | 28 | 2.01 | 2942.1 | 1.91
144.1 | 0 | 28 | 1.91 | 2799.4 | 1.74
145.3 | 0 | 28 | 1.92 | 2817.3 | 1.68
146.5 | 0 | 28 | 2.01 | 2937.9 | 1.92
147.7 | 0 | 28 | 1.95 | 2861.5 | 1.78
148.9 | 0 | 28 | 2.02 | 2958.4 | 1.78
150.1 | 0 | 28 | 1.97 | 2880.6 | 1.81
151.4 | 0 | 28 | 1.99 | 2922.2 | 1.84