## 1_Profils chambre de ref
Measurement time: | 07/09/2022 16:34:35 |  |  |  |  | Position de la chambre | Symétrie (%) | Homogénéité (%) | Pénombre G (cm) | Pénombre D (cm)
Radiation device: | Clinac3 Accelerator |  |  |  |  | Dans le coin | 1.76 | 3.56 | 1.38 | 1.39
Energy: | 12 MeV |  |  |  |  | Centrée (crossline) | 2.77 | 4.61 | 1.42 | 1.41
Controller: | CCU |  |  |  |  | Hors champ + normalisation | 8.92 | 8.35 | 1.45 | 1.47
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Désactivée | 7.27 | 5 | 1.37 | 1.33
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-200 | 0 | 28 | 0.25 | 508.3 | 0.21
-199.3 | 0 | 28 | 0.28 | 566.7 | 0.28
-197.9 | 0 | 28 | 0.3 | 589.7 | 0.3
-196.7 | 0 | 28 | 0.3 | 590.4 | 0.3
-195.5 | 0 | 28 | 0.26 | 523.9 | 0.26
-194.3 | 0 | 28 | 0.3 | 592.5 | 0.3
-193.1 | 0 | 28 | 0.31 | 613.9 | 0.3
-191.9 | 0 | 28 | 0.31 | 623 | 0.31
-190.7 | 0 | 28 | 0.32 | 649.1 | 0.32
-189.5 | 0 | 28 | 0.3 | 598.9 | 0.3
-188.3 | 0 | 28 | 0.31 | 616.5 | 0.3
-187.1 | 0 | 28 | 0.27 | 537.5 | 0.28
-185.9 | 0 | 28 | 0.34 | 674.2 | 0.33
-184.7 | 0 | 28 | 0.28 | 568.3 | 0.28
-183.5 | 0 | 28 | 0.33 | 649.8 | 0.32
-182.3 | 0 | 28 | 0.32 | 636.3 | 0.32
-181.1 | 0 | 28 | 0.32 | 648.5 | 0.32
-179.9 | 0 | 28 | 0.37 | 729.2 | 0.36
-178.7 | 0 | 28 | 0.34 | 681.3 | 0.35
-177.5 | 0 | 28 | 0.34 | 682.3 | 0.34
-176.3 | 0 | 28 | 0.38 | 767.3 | 0.39
-175.1 | 0 | 28 | 0.39 | 769.7 | 0.38
-173.9 | 0 | 28 | 0.35 | 699.1 | 0.35
-172.7 | 0 | 28 | 0.4 | 790.7 | 0.4
-171.5 | 0 | 28 | 0.39 | 784.2 | 0.38
-170.3 | 0 | 28 | 0.36 | 724.9 | 0.36
-169.1 | 0 | 28 | 0.34 | 685.3 | 0.34
-167.9 | 0 | 28 | 0.35 | 695.2 | 0.34
-166.7 | 0 | 28 | 0.35 | 701.2 | 0.36
-165.5 | 0 | 28 | 0.36 | 720.1 | 0.36
-164.3 | 0 | 28 | 0.34 | 685.4 | 0.34
-163.1 | 0 | 28 | 0.39 | 773.2 | 0.39
-161.9 | 0 | 28 | 0.37 | 732.2 | 0.38
-160.7 | 0 | 28 | 0.43 | 859.9 | 0.41
-159.5 | 0 | 28 | 0.37 | 730.2 | 0.37
-158.3 | 0 | 28 | 0.36 | 711.4 | 0.35
-157.1 | 0 | 28 | 0.37 | 748.1 | 0.36
-155.9 | 0 | 28 | 0.35 | 709 | 0.36
-154.7 | 0 | 28 | 0.38 | 755 | 0.38
-153.5 | 0 | 28 | 0.39 | 780.9 | 0.38
-152.2 | 0 | 28 | 0.36 | 718.5 | 0.37
-151 | 0 | 28 | 0.38 | 758.2 | 0.38
-149.8 | 0 | 28 | 0.4 | 808.7 | 0.39
-148.6 | 0 | 28 | 0.38 | 763.9 | 0.39
-147.4 | 0 | 28 | 0.4 | 791.4 | 0.39
-146.2 | 0 | 28 | 0.39 | 777.9 | 0.37
-145 | 0 | 28 | 0.35 | 692.4 | 0.36
-143.8 | 0 | 28 | 0.38 | 766.5 | 0.39
-142.6 | 0 | 28 | 0.41 | 824 | 0.4
-141.4 | 0 | 28 | 0.42 | 833.3 | 0.42
-140.2 | 0 | 28 | 0.39 | 770.5 | 0.39
-139 | 0 | 28 | 0.42 | 839.8 | 0.4
-137.8 | 0 | 28 | 0.39 | 781.2 | 0.4
-136.6 | 0 | 28 | 0.42 | 845.9 | 0.42
-135.4 | 0 | 28 | 0.45 | 903.2 | 0.43
-134.2 | 0 | 28 | 0.39 | 781.6 | 0.4
-133 | 0 | 28 | 0.45 | 900.9 | 0.45
-131.8 | 0 | 28 | 0.42 | 846.6 | 0.41
-130.6 | 0 | 28 | 0.43 | 857.2 | 0.43
-129.4 | 0 | 28 | 0.42 | 847.7 | 0.44
-128.2 | 0 | 28 | 0.48 | 967.6 | 0.47
-127 | 0 | 28 | 0.44 | 880.7 | 0.46
-125.8 | 0 | 28 | 0.5 | 998 | 0.49
-124.6 | 0 | 28 | 0.56 | 1126.5 | 0.53
-123.4 | 0 | 28 | 0.48 | 959.9 | 0.49
-122.2 | 0 | 28 | 0.48 | 958.4 | 0.47
-121 | 0 | 28 | 0.54 | 1084.4 | 0.52
-119.8 | 0 | 28 | 0.53 | 1066.9 | 0.55
-118.6 | 0 | 28 | 0.58 | 1161.4 | 0.58
-117.4 | 0 | 28 | 0.65 | 1288.6 | 0.62
-116.2 | 0 | 28 | 0.65 | 1306.4 | 0.68
-115 | 0 | 28 | 0.71 | 1416.5 | 0.7
-113.8 | 0 | 28 | 0.79 | 1580.9 | 0.77
-112.6 | 0 | 28 | 0.81 | 1616.8 | 0.81
-111.4 | 0 | 28 | 0.95 | 1893.3 | 0.93
-110.2 | 0 | 28 | 1.04 | 2087.1 | 1.01
-109 | 0 | 28 | 1.08 | 2154 | 1.11
-107.8 | 0 | 28 | 1.3 | 2587.3 | 1.27
-106.6 | 0 | 28 | 1.45 | 2899.3 | 1.45
-105.3 | 0 | 28 | 1.59 | 3179.9 | 1.63
-104.2 | 0 | 28 | 1.88 | 3759.1 | 1.86
-102.9 | 0 | 28 | 2.23 | 4457.6 | 2.13
-101.8 | 0 | 28 | 2.4 | 4803 | 2.45
-100.5 | 0 | 28 | 2.79 | 5572.5 | 2.75
-99.3 | 0 | 28 | 3.38 | 6744.4 | 3.24
-98.1 | 0 | 28 | 3.63 | 7253 | 3.71
-96.9 | 0 | 28 | 4.27 | 8528 | 4.29
-95.7 | 0 | 28 | 4.98 | 9958.2 | 4.98
-94.5 | 0 | 28 | 5.77 | 11518.4 | 5.85
-93.3 | 0 | 28 | 7.01 | 14005.7 | 6.78
-92.1 | 0 | 28 | 8.16 | 16302.6 | 7.94
-90.9 | 0 | 28 | 9.29 | 18553.1 | 9.39
-89.7 | 0 | 28 | 11.15 | 22269.1 | 11.03
-88.5 | 0 | 28 | 13.52 | 27016.5 | 13.1
-87.3 | 0 | 28 | 15.57 | 31098 | 15.63
-86.1 | 0 | 28 | 18.48 | 36910.1 | 18.52
-84.9 | 0 | 28 | 22.46 | 44869 | 22.19
-83.7 | 0 | 28 | 26.07 | 52076.3 | 26.21
-82.5 | 0 | 28 | 31.26 | 62442.3 | 31.02
-81.3 | 0 | 28 | 37.37 | 74667.2 | 36.51
-80.1 | 0 | 28 | 42.6 | 85104.2 | 42.42
-78.9 | 0 | 28 | 48.91 | 97712.4 | 48.65
-77.7 | 0 | 28 | 56.15 | 112176.4 | 54.9
-76.5 | 0 | 28 | 60.38 | 120633.6 | 60.98
-75.3 | 0 | 28 | 66.12 | 132095.8 | 66.76
-74.1 | 0 | 28 | 73.41 | 146655.7 | 71.85
-72.9 | 0 | 28 | 75.96 | 151751.3 | 76.51
-71.7 | 0 | 28 | 80.38 | 160575.8 | 80.29
-70.5 | 0 | 28 | 84.68 | 169180.9 | 83.62
-69.3 | 0 | 28 | 86.02 | 171853.8 | 86.3
-68.1 | 0 | 28 | 88.61 | 177025.8 | 88.2
-66.9 | 0 | 28 | 90.33 | 180465 | 90.28
-65.7 | 0 | 28 | 90.97 | 181735.6 | 91.64
-64.5 | 0 | 28 | 92.31 | 184427.9 | 92.85
-63.3 | 0 | 28 | 94.18 | 188163.4 | 93.63
-62.1 | 0 | 28 | 94.31 | 188415.3 | 94.46
-60.9 | 0 | 28 | 95.26 | 190302.4 | 95.03
-59.7 | 0 | 28 | 97.81 | 195408.3 | 95.57
-58.4 | 0 | 28 | 96.43 | 192646.7 | 96.27
-57.3 | 0 | 28 | 96.6 | 192988 | 96.46
-56.1 | 0 | 28 | 97.27 | 194329.4 | 96.93
-54.8 | 0 | 28 | 97.06 | 193916 | 97.22
-53.6 | 0 | 28 | 98.58 | 196935.7 | 97.61
-52.4 | 0 | 28 | 99.21 | 198204.7 | 97.77
-51.2 | 0 | 28 | 98.96 | 197697.9 | 98.03
-50 | 0 | 28 | 99.31 | 198399.8 | 98.29
-48.8 | 0 | 28 | 100.14 | 200055 | 98.44
-47.6 | 0 | 28 | 98.43 | 196636 | 98.33
-46.4 | 0 | 28 | 97.69 | 195158.4 | 98.73
-45.2 | 0 | 28 | 99.73 | 199247.5 | 98.75
-44 | 0 | 28 | 98.56 | 196914 | 99.04
-42.8 | 0 | 28 | 100.2 | 200180.7 | 98.78
-41.6 | 0 | 28 | 99.77 | 199318.7 | 99.08
-40.4 | 0 | 28 | 100.18 | 200148 | 99.28
-39.2 | 0 | 28 | 100.68 | 201132.2 | 99.22
-38 | 0 | 28 | 99.42 | 198625 | 99.28
-36.8 | 0 | 28 | 98.1 | 195985.1 | 99.49
-35.6 | 0 | 28 | 98.69 | 197169.8 | 99.4
-34.4 | 0 | 28 | 99.55 | 198875.1 | 99.55
-33.2 | 0 | 28 | 99.85 | 199476 | 99.5
-32 | 0 | 28 | 99.44 | 198654.8 | 99.67
-30.8 | 0 | 28 | 101.47 | 202719.6 | 99.49
-29.6 | 0 | 28 | 99.64 | 199058.1 | 99.75
-28.4 | 0 | 28 | 100.9 | 201585.3 | 99.66
-27.2 | 0 | 28 | 99.49 | 198766 | 99.8
-26 | 0 | 28 | 100.04 | 199853.8 | 99.66
-24.8 | 0 | 28 | 100.43 | 200646.7 | 99.74
-23.6 | 0 | 28 | 100.55 | 200872.5 | 99.87
-22.4 | 0 | 28 | 102.03 | 203843.7 | 99.61
-21.2 | 0 | 28 | 100.53 | 200830.6 | 99.97
-20 | 0 | 28 | 102.02 | 203824.6 | 99.83
-18.8 | 0 | 28 | 98.29 | 196362.4 | 99.84
-17.6 | 0 | 28 | 99 | 197774.5 | 99.68
-16.4 | 0 | 28 | 100.56 | 200894.3 | 99.66
-15.2 | 0 | 28 | 101.6 | 202971.6 | 99.77
-14 | 0 | 28 | 100.38 | 200534.5 | 99.7
-12.8 | 0 | 28 | 97.35 | 194496 | 99.8
-11.6 | 0 | 28 | 100.4 | 200579.5 | 99.87
-10.4 | 0 | 28 | 99.2 | 198176.4 | 99.54
-9.2 | 0 | 28 | 100.87 | 201514.6 | 99.8
-8 | 0 | 28 | 99.32 | 198416.8 | 99.91
-6.7 | 0 | 28 | 96.56 | 192910.7 | 99.91
-5.5 | 0 | 28 | 100.34 | 200468.4 | 99.83
-4.3 | 0 | 28 | 100.45 | 200682 | 99.9
-3.1 | 0 | 28 | 101.24 | 202255.3 | 99.94
-1.9 | 0 | 28 | 96.19 | 192173.7 | 99.76
-0.7 | 0 | 28 | 101.15 | 202086.2 | 99.96
0.5 | 0 | 28 | 99.8 | 199384.5 | 99.98
1.7 | 0 | 28 | 100.38 | 200532.4 | 99.93
2.9 | 0 | 28 | 100.47 | 200726.6 | 99.86
4.1 | 0 | 28 | 100.75 | 201271.3 | 100
5.3 | 0 | 28 | 100.81 | 201402 | 99.86
6.5 | 0 | 28 | 101.63 | 203044.8 | 99.73
7.7 | 0 | 28 | 101.4 | 202581 | 99.98
8.9 | 0 | 28 | 100 | 199774 | 99.78
10.1 | 0 | 28 | 103.54 | 206850.5 | 99.9
11.3 | 0 | 28 | 95.63 | 191056.2 | 99.8
12.5 | 0 | 28 | 101.02 | 201822.9 | 99.74
13.7 | 0 | 28 | 100.96 | 201700.5 | 99.86
14.9 | 0 | 28 | 100.74 | 201264.1 | 99.61
16.1 | 0 | 28 | 96.73 | 193253.8 | 99.68
17.3 | 0 | 28 | 99.28 | 198334.8 | 99.75
18.5 | 0 | 28 | 102.16 | 204098.3 | 99.55
19.7 | 0 | 28 | 95.25 | 190298.4 | 99.57
20.9 | 0 | 28 | 100.27 | 200320.2 | 99.74
22.1 | 0 | 28 | 101.32 | 202419.7 | 99.79
23.3 | 0 | 28 | 100.4 | 200574.9 | 99.55
24.5 | 0 | 28 | 101.04 | 201860.9 | 99.78
25.7 | 0 | 28 | 100.77 | 201327.8 | 99.59
26.9 | 0 | 28 | 100.06 | 199902.1 | 99.48
28.1 | 0 | 28 | 99.04 | 197869.4 | 99.58
29.3 | 0 | 28 | 100 | 199776.3 | 99.56
30.5 | 0 | 28 | 99.78 | 199348.1 | 99.41
31.7 | 0 | 28 | 100.11 | 199995.5 | 99.18
32.9 | 0 | 28 | 100.75 | 201280.4 | 99.37
34.1 | 0 | 28 | 101.28 | 202344.4 | 99.4
35.3 | 0 | 28 | 99.57 | 198920.7 | 99.04
36.5 | 0 | 28 | 98.89 | 197565.3 | 99.12
37.7 | 0 | 28 | 98.1 | 195994.9 | 98.97
38.9 | 0 | 28 | 102.08 | 203944.9 | 98.9
40.1 | 0 | 28 | 97.31 | 194400.7 | 98.81
41.4 | 0 | 28 | 95.28 | 190343.9 | 98.56
42.5 | 0 | 28 | 100.6 | 200983 | 98.7
43.8 | 0 | 28 | 99.44 | 198660.2 | 98.31
45 | 0 | 28 | 98.36 | 196508 | 98.19
46.2 | 0 | 28 | 98.33 | 196436.3 | 98.27
47.4 | 0 | 28 | 97.62 | 195032 | 97.93
48.6 | 0 | 28 | 96.25 | 192297.5 | 97.76
49.8 | 0 | 28 | 98.69 | 197165.6 | 97.6
51 | 0 | 28 | 98.36 | 196495.8 | 97.19
52.2 | 0 | 28 | 98.56 | 196895.1 | 97.04
53.4 | 0 | 28 | 94.48 | 188750.2 | 96.82
54.6 | 0 | 28 | 95.93 | 191646.6 | 96.41
55.8 | 0 | 28 | 96.59 | 192972.1 | 96.08
57 | 0 | 28 | 95.78 | 191346.3 | 95.57
58.2 | 0 | 28 | 96.15 | 192096.2 | 95.22
59.4 | 0 | 28 | 94.86 | 189516.4 | 94.47
60.6 | 0 | 28 | 95.94 | 191663.2 | 93.7
61.8 | 0 | 28 | 94.36 | 188504 | 92.87
63 | 0 | 28 | 92.55 | 184907 | 91.87
64.2 | 0 | 28 | 90.41 | 180625.8 | 90.78
65.4 | 0 | 28 | 88.66 | 177122.4 | 89.18
66.6 | 0 | 28 | 88.54 | 176880.3 | 87.33
67.8 | 0 | 28 | 85.91 | 171629.6 | 85.42
69 | 0 | 28 | 83.34 | 166489.5 | 82.47
70.2 | 0 | 28 | 80.33 | 160494.4 | 79.19
71.4 | 0 | 28 | 76.35 | 152532 | 75.57
72.6 | 0 | 28 | 71.44 | 142715.8 | 70.86
73.9 | 0 | 28 | 65.28 | 130422.8 | 65.63
75 | 0 | 28 | 59.52 | 118912.4 | 60.08
76.2 | 0 | 28 | 53.84 | 107553.1 | 53.93
77.4 | 0 | 28 | 47.49 | 94869.2 | 47.69
78.6 | 0 | 28 | 41.97 | 83848.7 | 41.61
79.9 | 0 | 28 | 35.98 | 71887.5 | 35.54
81.1 | 0 | 28 | 30.44 | 60817 | 30.2
82.3 | 0 | 28 | 25.67 | 51281.6 | 25.24
83.5 | 0 | 28 | 21.29 | 42530.5 | 21.21
84.7 | 0 | 28 | 17.49 | 34940.5 | 17.64
85.9 | 0 | 28 | 14.71 | 29389.3 | 14.62
87.1 | 0 | 28 | 12.26 | 24497.3 | 12.27
88.3 | 0 | 28 | 10.35 | 20673.5 | 10.22
89.5 | 0 | 28 | 8.82 | 17620.9 | 8.65
90.7 | 0 | 28 | 7.43 | 14834.9 | 7.34
91.9 | 0 | 28 | 6.03 | 12044.8 | 6.26
93.1 | 0 | 28 | 5.33 | 10649.3 | 5.3
94.3 | 0 | 28 | 4.54 | 9067.3 | 4.65
95.5 | 0 | 28 | 4.07 | 8131.9 | 4.01
96.7 | 0 | 28 | 3.57 | 7130.6 | 3.47
97.9 | 0 | 28 | 3.01 | 6010.9 | 3.04
99.1 | 0 | 28 | 2.68 | 5359.6 | 2.68
100.3 | 0 | 28 | 2.43 | 4857.4 | 2.37
101.5 | 0 | 28 | 2.04 | 4079.5 | 2.07
102.7 | 0 | 28 | 1.84 | 3685.6 | 1.86
103.9 | 0 | 28 | 1.7 | 3387.1 | 1.68
105.1 | 0 | 28 | 1.42 | 2838.9 | 1.47
106.3 | 0 | 28 | 1.32 | 2632.5 | 1.31
107.5 | 0 | 28 | 1.24 | 2473 | 1.18
108.7 | 0 | 28 | 1.04 | 2069.2 | 1.05
109.9 | 0 | 28 | 0.95 | 1892.2 | 0.95
111.1 | 0 | 28 | 0.9 | 1800.5 | 0.86
112.3 | 0 | 28 | 0.73 | 1466.9 | 0.76
113.5 | 0 | 28 | 0.7 | 1394.6 | 0.69
114.7 | 0 | 28 | 0.61 | 1214.1 | 0.61
115.9 | 0 | 28 | 0.57 | 1146.3 | 0.59
117.1 | 0 | 28 | 0.54 | 1073.3 | 0.54
118.3 | 0 | 28 | 0.52 | 1047.7 | 0.5
119.6 | 0 | 28 | 0.41 | 827.8 | 0.43
120.7 | 0 | 28 | 0.43 | 862.8 | 0.43
122 | 0 | 28 | 0.45 | 905.4 | 0.43
123.2 | 0 | 28 | 0.43 | 850.9 | 0.44
124.4 | 0 | 28 | 0.43 | 868.8 | 0.42
125.6 | 0 | 28 | 0.36 | 724 | 0.37
126.8 | 0 | 28 | 0.35 | 699.7 | 0.36
128 | 0 | 28 | 0.38 | 764 | 0.38
129.2 | 0 | 28 | 0.32 | 641.2 | 0.3
130.4 | 0 | 28 | 0.29 | 581.4 | 0.3
131.6 | 0 | 28 | 0.31 | 610.7 | 0.3
132.8 | 0 | 28 | 0.31 | 626.1 | 0.3
134 | 0 | 28 | 0.25 | 500.5 | 0.26
135.2 | 0 | 28 | 0.28 | 556.9 | 0.28
136.4 | 0 | 28 | 0.31 | 622.8 | 0.3
137.6 | 0 | 28 | 0.27 | 531.3 | 0.27
138.8 | 0 | 28 | 0.28 | 554 | 0.27
140 | 0 | 28 | 0.25 | 491.1 | 0.23
141.2 | 0 | 28 | 0.22 | 443.3 | 0.23
142.4 | 0 | 28 | 0.3 | 589.4 | 0.29
143.6 | 0 | 28 | 0.3 | 597.8 | 0.28
144.8 | 0 | 28 | 0.23 | 460.9 | 0.24
146 | 0 | 28 | 0.25 | 506.3 | 0.25
147.2 | 0 | 28 | 0.27 | 548.8 | 0.27
148.4 | 0 | 28 | 0.22 | 442 | 0.22
149.6 | 0 | 28 | 0.25 | 494.3 | 0.25
150.8 | 0 | 28 | 0.25 | 505.4 | 0.26
152 | 0 | 28 | 0.24 | 470.7 | 0.24
153.2 | 0 | 28 | 0.2 | 396.2 | 0.2
154.4 | 0 | 28 | 0.27 | 547.3 | 0.26
155.6 | 0 | 28 | 0.23 | 451.2 | 0.23
156.8 | 0 | 28 | 0.19 | 382.9 | 0.19
158 | 0 | 28 | 0.23 | 454.7 | 0.22
159.2 | 0 | 28 | 0.22 | 435.1 | 0.22
160.4 | 0 | 28 | 0.2 | 395 | 0.2
161.6 | 0 | 28 | 0.22 | 435.2 | 0.21
162.8 | 0 | 28 | 0.17 | 344 | 0.18
164 | 0 | 28 | 0.19 | 373.2 | 0.19
165.2 | 0 | 28 | 0.22 | 443.9 | 0.21
166.4 | 0 | 28 | 0.22 | 436.7 | 0.22
167.7 | 0 | 28 | 0.18 | 361.8 | 0.18
168.8 | 0 | 28 | 0.17 | 341.3 | 0.17
170 | 0 | 28 | 0.15 | 302.8 | 0.15
171.2 | 0 | 28 | 0.15 | 304.3 | 0.15
172.4 | 0 | 28 | 0.16 | 316.3 | 0.16
173.6 | 0 | 28 | 0.13 | 269.3 | 0.14
174.8 | 0 | 28 | 0.15 | 296.5 | 0.15
176 | 0 | 28 | 0.15 | 308.4 | 0.15
177.2 | 0 | 28 | 0.17 | 343.9 | 0.17
178.4 | 0 | 28 | 0.17 | 337.9 | 0.17
179.6 | 0 | 28 | 0.19 | 372.4 | 0.18
180.8 | 0 | 28 | 0.15 | 300.6 | 0.15
182 | 0 | 28 | 0.2 | 390.1 | 0.2
183.2 | 0 | 28 | 0.16 | 318.1 | 0.16
184.5 | 0 | 28 | 0.18 | 355.4 | 0.18
185.6 | 0 | 28 | 0.2 | 393.4 | 0.19
186.8 | 0 | 28 | 0.2 | 394 | 0.19
188.1 | 0 | 28 | 0.19 | 379.2 | 0.19
189.3 | 0 | 28 | 0.17 | 341.3 | 0.17
190.4 | 0 | 28 | 0.19 | 375.2 | 0.18
191.7 | 0 | 28 | 0.18 | 355.5 | 0.18
192.9 | 0 | 28 | 0.18 | 354.7 | 0.17
194.1 | 0 | 28 | 0.13 | 252 | 0.12
195.3 | 0 | 28 | 0.15 | 291 | 0.15
196.5 | 0 | 28 | 0.14 | 289.3 | 0.14
197.7 | 0 | 28 | 0.12 | 241.3 | 0.12
198.9 | 0 | 28 | 0.12 | 235 | 0.12
200 | 0 | 28 | 0.12 | 238.5 | 0.12
Measurement time: | 07/09/2022 16:38:34
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
200 | 0 | 28 | 0.11 | 219.9 | 0.11
198.9 | 0 | 28 | 0.14 | 288.4 | 0.14
197.7 | 0 | 28 | 0.18 | 355.2 | 0.16
196.5 | 0 | 28 | 0.13 | 262.9 | 0.13
195.3 | 0 | 28 | 0.14 | 287.1 | 0.14
194.1 | 0 | 28 | 0.11 | 218.6 | 0.1
192.9 | 0 | 28 | 0.12 | 247.9 | 0.12
191.7 | 0 | 28 | 0.13 | 265.2 | 0.13
190.5 | 0 | 28 | 0.11 | 225.6 | 0.11
189.3 | 0 | 28 | 0.13 | 249.7 | 0.12
188.1 | 0 | 28 | 0.21 | 428.4 | 0.2
186.9 | 0 | 28 | 0.18 | 360.8 | 0.16
185.7 | 0 | 28 | 0.15 | 296.7 | 0.15
184.5 | 0 | 28 | 0.14 | 271.3 | 0.13
183.3 | 0 | 28 | 0.15 | 295 | 0.14
182 | 0 | 28 | 0.14 | 277.3 | 0.13
180.8 | 0 | 28 | 0.18 | 351.7 | 0.17
179.7 | 0 | 28 | 0.21 | 411 | 0.19
178.4 | 0 | 28 | 0.21 | 414.6 | 0.2
177.3 | 0 | 28 | 0.18 | 356 | 0.17
176 | 0 | 28 | 0.18 | 364.9 | 0.17
174.8 | 0 | 28 | 0.16 | 326.4 | 0.16
173.6 | 0 | 28 | 0.19 | 386.6 | 0.19
172.4 | 0 | 28 | 0.14 | 288.1 | 0.13
171.2 | 0 | 28 | 0.17 | 334.4 | 0.17
170 | 0 | 28 | 0.18 | 352.4 | 0.17
168.8 | 0 | 28 | 0.18 | 354.9 | 0.16
167.6 | 0 | 28 | 0.16 | 315.8 | 0.16
166.4 | 0 | 28 | 0.2 | 401.5 | 0.2
165.3 | 0 | 28 | 0.2 | 396.6 | 0.19
164 | 0 | 28 | 0.17 | 338.7 | 0.17
162.8 | 0 | 28 | 0.19 | 372.4 | 0.18
161.6 | 0 | 28 | 0.23 | 462.6 | 0.22
160.4 | 0 | 28 | 0.21 | 421.5 | 0.2
159.2 | 0 | 28 | 0.19 | 376.1 | 0.18
158 | 0 | 28 | 0.16 | 312.4 | 0.15
156.8 | 0 | 28 | 0.19 | 373.6 | 0.18
155.6 | 0 | 28 | 0.22 | 434.3 | 0.21
154.4 | 0 | 28 | 0.17 | 331.3 | 0.16
153.2 | 0 | 28 | 0.21 | 424.2 | 0.2
152 | 0 | 28 | 0.19 | 386.1 | 0.19
150.8 | 0 | 28 | 0.21 | 409.8 | 0.19
149.6 | 0 | 28 | 0.21 | 424.5 | 0.21
148.4 | 0 | 28 | 0.19 | 385.8 | 0.19
147.2 | 0 | 28 | 0.24 | 482.8 | 0.23
146 | 0 | 28 | 0.29 | 571.6 | 0.28
144.8 | 0 | 28 | 0.27 | 545.5 | 0.26
143.6 | 0 | 28 | 0.26 | 517.7 | 0.25
142.4 | 0 | 28 | 0.25 | 492.8 | 0.24
141.2 | 0 | 28 | 0.23 | 458.2 | 0.22
139.9 | 0 | 28 | 0.25 | 502.7 | 0.24
138.8 | 0 | 28 | 0.29 | 569.7 | 0.27
137.6 | 0 | 28 | 0.26 | 514.2 | 0.24
136.3 | 0 | 28 | 0.27 | 544.7 | 0.26
135.2 | 0 | 28 | 0.27 | 532.1 | 0.26
133.9 | 0 | 28 | 0.31 | 609.7 | 0.3
132.7 | 0 | 28 | 0.25 | 509 | 0.24
131.5 | 0 | 28 | 0.29 | 580.9 | 0.28
130.3 | 0 | 28 | 0.32 | 638.3 | 0.31
129.1 | 0 | 28 | 0.31 | 629.1 | 0.3
127.9 | 0 | 28 | 0.32 | 636.8 | 0.3
126.7 | 0 | 28 | 0.32 | 648.4 | 0.31
125.5 | 0 | 28 | 0.37 | 730.2 | 0.35
124.3 | 0 | 28 | 0.42 | 837.5 | 0.4
123.1 | 0 | 28 | 0.39 | 787.8 | 0.38
121.9 | 0 | 28 | 0.4 | 796.7 | 0.38
120.7 | 0 | 28 | 0.43 | 851 | 0.41
119.5 | 0 | 28 | 0.46 | 914.6 | 0.44
118.3 | 0 | 28 | 0.5 | 1000.1 | 0.48
117.1 | 0 | 28 | 0.54 | 1078.2 | 0.52
115.9 | 0 | 28 | 0.57 | 1142.8 | 0.55
114.7 | 0 | 28 | 0.66 | 1317.2 | 0.63
113.5 | 0 | 28 | 0.68 | 1354.4 | 0.67
112.3 | 0 | 28 | 0.78 | 1559.6 | 0.74
111.2 | 0 | 28 | 0.9 | 1790.8 | 0.87
109.9 | 0 | 28 | 1.04 | 2067.9 | 0.99
108.7 | 0 | 28 | 1.14 | 2273.2 | 1.1
107.5 | 0 | 28 | 1.25 | 2506 | 1.23
106.3 | 0 | 28 | 1.51 | 3014.3 | 1.43
105.1 | 0 | 28 | 1.67 | 3330.3 | 1.62
103.9 | 0 | 28 | 1.94 | 3881.1 | 1.85
102.7 | 0 | 28 | 2.23 | 4457.7 | 2.12
101.5 | 0 | 28 | 2.4 | 4803.9 | 2.35
100.3 | 0 | 28 | 2.92 | 5827.8 | 2.75
99.1 | 0 | 28 | 3.26 | 6522.7 | 3.13
97.9 | 0 | 28 | 3.67 | 7329.8 | 3.58
96.7 | 0 | 28 | 4.32 | 8627 | 4.1
95.5 | 0 | 28 | 4.86 | 9708.4 | 4.7
94.3 | 0 | 28 | 5.89 | 11762.9 | 5.49
93.1 | 0 | 28 | 6.27 | 12524.2 | 6.33
91.9 | 0 | 28 | 7.76 | 15498.9 | 7.37
90.7 | 0 | 28 | 9.16 | 18302.1 | 8.66
89.5 | 0 | 28 | 10.14 | 20257 | 10.03
88.3 | 0 | 28 | 12.73 | 25433.3 | 11.95
87.1 | 0 | 28 | 14.49 | 28955.8 | 14.07
85.9 | 0 | 28 | 17.95 | 35854.5 | 16.66
84.7 | 0 | 28 | 20.59 | 41127.8 | 19.75
83.5 | 0 | 28 | 24.19 | 48325.9 | 23.25
82.3 | 0 | 28 | 29.16 | 58250.3 | 27.7
81.1 | 0 | 28 | 33.95 | 67835.3 | 32.39
79.9 | 0 | 28 | 38.43 | 76770 | 37.92
78.7 | 0 | 28 | 46.06 | 92014 | 43.65
77.5 | 0 | 28 | 52.19 | 104257.5 | 49.93
76.3 | 0 | 28 | 59.99 | 119853.6 | 56.27
75.1 | 0 | 28 | 64.14 | 128149.4 | 62.09
73.9 | 0 | 28 | 73.3 | 146443.2 | 68.06
72.7 | 0 | 28 | 71.54 | 142919.5 | 72.99
71.5 | 0 | 28 | 80.53 | 160892.5 | 77.22
70.3 | 0 | 28 | 86.12 | 172053.9 | 81.34
69.1 | 0 | 28 | 86.95 | 173701.5 | 84.44
67.9 | 0 | 28 | 93.22 | 186246 | 87.13
66.7 | 0 | 28 | 92.6 | 184994.8 | 89.12
65.5 | 0 | 28 | 95.8 | 191398.1 | 90.97
64.2 | 0 | 28 | 93.8 | 187398.4 | 92.5
63.1 | 0 | 28 | 96.46 | 192717.2 | 93.62
61.8 | 0 | 28 | 98.25 | 196286.7 | 94.74
60.6 | 0 | 28 | 98.95 | 197677.2 | 95.62
59.4 | 0 | 28 | 100.45 | 200679.1 | 96.1
58.2 | 0 | 28 | 101.92 | 203619.5 | 96.87
57 | 0 | 28 | 103.16 | 206100.8 | 97.34
55.8 | 0 | 28 | 102.74 | 205248.5 | 97.77
54.6 | 0 | 28 | 103.62 | 207011.8 | 98.26
53.4 | 0 | 28 | 103.84 | 207454.2 | 98.55
52.2 | 0 | 28 | 102.91 | 205589 | 98.77
51 | 0 | 28 | 102.81 | 205392.1 | 98.95
49.8 | 0 | 28 | 101.36 | 202495.8 | 99.36
48.6 | 0 | 28 | 107.99 | 215750.7 | 99.56
47.4 | 0 | 28 | 99.71 | 199194.7 | 99.67
46.2 | 0 | 28 | 102.2 | 204177.9 | 99.64
45 | 0 | 28 | 104.29 | 208350.4 | 99.76
43.8 | 0 | 28 | 104.16 | 208085.3 | 99.92
42.6 | 0 | 28 | 104.92 | 209607.9 | 99.9
41.4 | 0 | 28 | 105.14 | 210056.6 | 99.97
40.2 | 0 | 28 | 103.23 | 206237.2 | 100
39 | 0 | 28 | 104.44 | 208645.2 | 99.76
37.8 | 0 | 28 | 102.65 | 205079.9 | 99.89
36.6 | 0 | 28 | 104.64 | 209060.5 | 99.84
35.4 | 0 | 28 | 104.36 | 208500.1 | 99.85
34.2 | 0 | 28 | 104.04 | 207850.3 | 99.72
33 | 0 | 28 | 106.02 | 211799.5 | 99.98
31.8 | 0 | 28 | 104.29 | 208359.6 | 99.82
30.6 | 0 | 28 | 105.51 | 210796.5 | 99.65
29.4 | 0 | 28 | 104.45 | 208681.9 | 99.89
28.2 | 0 | 28 | 104.23 | 208228.7 | 99.74
27 | 0 | 28 | 104.19 | 208142.8 | 99.55
25.8 | 0 | 28 | 103.21 | 206188.7 | 99.68
24.6 | 0 | 28 | 105.57 | 210904 | 99.72
23.3 | 0 | 28 | 104.9 | 209573.1 | 99.68
22.1 | 0 | 28 | 104.66 | 209087.1 | 99.45
21 | 0 | 28 | 105.25 | 210279.5 | 99.69
19.7 | 0 | 28 | 103.35 | 206469.8 | 99.6
18.6 | 0 | 28 | 104.46 | 208698.9 | 99.6
17.4 | 0 | 28 | 103.59 | 206962.4 | 99.65
16.1 | 0 | 28 | 102.26 | 204294.1 | 99.59
15 | 0 | 28 | 104.47 | 208715.4 | 99.5
13.7 | 0 | 28 | 103.12 | 206008.1 | 99.72
12.6 | 0 | 28 | 105.16 | 210088.5 | 99.66
11.4 | 0 | 28 | 104.95 | 209668.1 | 99.8
10.1 | 0 | 28 | 105.22 | 210200.9 | 99.81
9 | 0 | 28 | 102.17 | 204123 | 99.52
7.7 | 0 | 28 | 103.88 | 207535.9 | 99.82
6.5 | 0 | 28 | 101.21 | 202200.4 | 99.59
5.3 | 0 | 28 | 103.69 | 207159.7 | 99.51
4.1 | 0 | 28 | 102.24 | 204248.2 | 99.65
2.9 | 0 | 28 | 105.68 | 211132.2 | 99.63
1.7 | 0 | 28 | 103.71 | 207197.5 | 99.48
0.5 | 0 | 28 | 105.69 | 211157.2 | 99.5
-0.7 | 0 | 28 | 103.8 | 207383.4 | 99.33
-1.9 | 0 | 28 | 102.09 | 203957.9 | 99.13
-3.1 | 0 | 28 | 103.59 | 206958.7 | 99.17
-4.3 | 0 | 28 | 100.52 | 200828.5 | 99.13
-5.5 | 0 | 28 | 104.69 | 209143 | 99.1
-6.7 | 0 | 28 | 103.83 | 207440.7 | 98.88
-7.9 | 0 | 28 | 104.15 | 208076.2 | 98.78
-9.1 | 0 | 28 | 104.52 | 208807.8 | 98.66
-10.3 | 0 | 28 | 102.48 | 204745.7 | 98.4
-11.5 | 0 | 28 | 102.37 | 204514.3 | 98.53
-12.7 | 0 | 28 | 100.93 | 201648.5 | 98.27
-13.9 | 0 | 28 | 101.32 | 202412.3 | 98.01
-15.1 | 0 | 28 | 101.91 | 203603 | 98.08
-16.3 | 0 | 28 | 101.41 | 202599.5 | 97.9
-17.5 | 0 | 28 | 102.43 | 204626.5 | 97.98
-18.7 | 0 | 28 | 102.8 | 205365.8 | 97.58
-19.9 | 0 | 28 | 102.48 | 204736.4 | 97.52
-21.1 | 0 | 28 | 102.98 | 205735.9 | 97.58
-22.3 | 0 | 28 | 102.14 | 204065.6 | 97.4
-23.5 | 0 | 28 | 101.85 | 203468 | 97.33
-24.7 | 0 | 28 | 100.71 | 201196.9 | 97.48
-26 | 0 | 28 | 101.6 | 202975.4 | 97.32
-27.2 | 0 | 28 | 101.81 | 203398.4 | 97.18
-28.4 | 0 | 28 | 101.39 | 202548.7 | 97.11
-29.5 | 0 | 28 | 101.43 | 202642.7 | 97.03
-30.8 | 0 | 28 | 102.75 | 205271.6 | 96.87
-32 | 0 | 28 | 101.66 | 203099.2 | 96.75
-33.2 | 0 | 28 | 96.05 | 191882.8 | 96.74
-34.4 | 0 | 28 | 102.88 | 205536.6 | 96.75
-35.5 | 0 | 28 | 100.63 | 201031.8 | 96.56
-36.8 | 0 | 28 | 99.61 | 198993 | 96.46
-38 | 0 | 28 | 100.67 | 201126.2 | 96.36
-39.2 | 0 | 28 | 100.49 | 200769.8 | 96.34
-40.4 | 0 | 28 | 97.73 | 195254.9 | 96.18
-41.6 | 0 | 28 | 101.42 | 202626.7 | 95.98
-42.8 | 0 | 28 | 98.75 | 197289.7 | 95.92
-44 | 0 | 28 | 102.5 | 204771.6 | 95.59
-45.1 | 0 | 28 | 101.25 | 202275.9 | 95.46
-46.4 | 0 | 28 | 99.36 | 198506 | 95.51
-47.6 | 0 | 28 | 99.6 | 198973.4 | 95.2
-48.8 | 0 | 28 | 101.82 | 203425.1 | 94.93
-50 | 0 | 28 | 98.08 | 195944.7 | 94.82
-51.2 | 0 | 28 | 96.95 | 193685 | 94.72
-52.4 | 0 | 28 | 99.8 | 199389.3 | 94.33
-53.6 | 0 | 28 | 97.43 | 194637.8 | 94.22
-54.8 | 0 | 28 | 96.32 | 192437.6 | 93.9
-56 | 0 | 28 | 96.67 | 193131.9 | 93.56
-57.2 | 0 | 28 | 94.46 | 188720.2 | 93.3
-58.4 | 0 | 28 | 99.16 | 198105.2 | 92.96
-59.6 | 0 | 28 | 98.31 | 196406.8 | 92.32
-60.8 | 0 | 28 | 93.84 | 187476.3 | 91.92
-62 | 0 | 28 | 96.9 | 193595.1 | 91.31
-63.2 | 0 | 28 | 97.9 | 195595 | 90.66
-64.4 | 0 | 28 | 92.27 | 184330 | 89.58
-65.6 | 0 | 28 | 91.23 | 182252.8 | 88.56
-66.8 | 0 | 28 | 93.55 | 186896.6 | 87.17
-68 | 0 | 28 | 86.37 | 172556.5 | 85.49
-69.2 | 0 | 28 | 87.16 | 174129.2 | 83.55
-70.4 | 0 | 28 | 87.13 | 174075.5 | 80.99
-71.6 | 0 | 28 | 79.8 | 159421.7 | 78.1
-72.8 | 0 | 28 | 78.18 | 156194.5 | 74.44
-74 | 0 | 28 | 75.63 | 151088.3 | 69.72
-75.2 | 0 | 28 | 66.16 | 132179.3 | 64.93
-76.4 | 0 | 28 | 63.23 | 126314.5 | 59.45
-77.6 | 0 | 28 | 57.65 | 115167.3 | 53.48
-78.8 | 0 | 28 | 47.79 | 95472.8 | 47.4
-80 | 0 | 28 | 43.14 | 86190.4 | 41.48
-81.2 | 0 | 28 | 38.71 | 77330.6 | 35.63
-82.4 | 0 | 28 | 30.82 | 61570.4 | 30.28
-83.6 | 0 | 28 | 25.86 | 51659 | 25.78
-84.8 | 0 | 28 | 23.6 | 47142.4 | 21.6
-86.1 | 0 | 28 | 18.45 | 36868.3 | 18.2
-87.2 | 0 | 28 | 15.38 | 30734.6 | 15.3
-88.5 | 0 | 28 | 13.88 | 27737.2 | 12.74
-89.6 | 0 | 28 | 10.84 | 21651.8 | 10.78
-90.9 | 0 | 28 | 9.44 | 18866.9 | 9.18
-92.1 | 0 | 28 | 8.4 | 16791.3 | 7.8
-93.3 | 0 | 28 | 6.98 | 13954.3 | 6.69
-94.5 | 0 | 28 | 5.61 | 11205.9 | 5.7
-95.7 | 0 | 28 | 5.3 | 10583.5 | 4.91
-96.9 | 0 | 28 | 4.2 | 8391.2 | 4.22
-98.1 | 0 | 28 | 3.76 | 7513.1 | 3.59
-99.3 | 0 | 28 | 3.4 | 6792.1 | 3.13
-100.5 | 0 | 28 | 2.78 | 5547.2 | 2.74
-101.7 | 0 | 28 | 2.43 | 4861.5 | 2.32
-102.9 | 0 | 28 | 2.23 | 4454.7 | 2.06
-104.1 | 0 | 28 | 1.87 | 3744.3 | 1.81
-105.3 | 0 | 28 | 1.55 | 3098.9 | 1.55
-106.5 | 0 | 28 | 1.53 | 3046.8 | 1.41
-107.7 | 0 | 28 | 1.2 | 2398.9 | 1.2
-108.9 | 0 | 28 | 1.1 | 2190.5 | 1.05
-110.1 | 0 | 28 | 1.07 | 2147.1 | 0.98
-111.3 | 0 | 28 | 0.88 | 1751.8 | 0.86
-112.5 | 0 | 28 | 0.69 | 1374.6 | 0.69
-113.7 | 0 | 28 | 0.74 | 1475.1 | 0.68
-114.9 | 0 | 28 | 0.6 | 1195.7 | 0.6
-116.1 | 0 | 28 | 0.65 | 1295.7 | 0.61
-117.3 | 0 | 28 | 0.5 | 1000.3 | 0.49
-118.5 | 0 | 28 | 0.47 | 947.8 | 0.46
-119.7 | 0 | 28 | 0.51 | 1027.9 | 0.49
-120.9 | 0 | 28 | 0.49 | 969.7 | 0.45
-122.1 | 0 | 28 | 0.43 | 852.7 | 0.42
-123.3 | 0 | 28 | 0.43 | 853.5 | 0.41
-124.5 | 0 | 28 | 0.44 | 879.1 | 0.41
-125.7 | 0 | 28 | 0.36 | 713.1 | 0.35
-126.9 | 0 | 28 | 0.39 | 789 | 0.38
-128.1 | 0 | 28 | 0.42 | 835.3 | 0.4
-129.3 | 0 | 28 | 0.42 | 831.3 | 0.4
-130.5 | 0 | 28 | 0.4 | 790.3 | 0.38
-131.7 | 0 | 28 | 0.37 | 731.1 | 0.34
-133 | 0 | 28 | 0.32 | 648.6 | 0.31
-134.1 | 0 | 28 | 0.35 | 696.6 | 0.33
-135.4 | 0 | 28 | 0.39 | 777.7 | 0.37
-136.6 | 0 | 28 | 0.31 | 617.9 | 0.3
-137.7 | 0 | 28 | 0.31 | 620.7 | 0.3
-139 | 0 | 28 | 0.35 | 691.7 | 0.33
-140.2 | 0 | 28 | 0.32 | 634.5 | 0.31
-141.3 | 0 | 28 | 0.34 | 679.7 | 0.32
-142.6 | 0 | 28 | 0.26 | 510.5 | 0.25
-143.7 | 0 | 28 | 0.28 | 567.6 | 0.28
-145 | 0 | 28 | 0.3 | 607.6 | 0.3
-146.2 | 0 | 28 | 0.3 | 591.9 | 0.28
-147.4 | 0 | 28 | 0.28 | 565.8 | 0.27
-148.6 | 0 | 28 | 0.29 | 573.9 | 0.27
-149.8 | 0 | 28 | 0.28 | 553.6 | 0.27
-151 | 0 | 28 | 0.28 | 562.7 | 0.27
-152.2 | 0 | 28 | 0.33 | 651.3 | 0.32
-153.4 | 0 | 28 | 0.31 | 623 | 0.29
-154.6 | 0 | 28 | 0.27 | 549.4 | 0.27
-155.8 | 0 | 28 | 0.27 | 533.8 | 0.25
-157 | 0 | 28 | 0.28 | 550.1 | 0.26
-158.2 | 0 | 28 | 0.29 | 583.8 | 0.28
-159.4 | 0 | 28 | 0.29 | 569.6 | 0.27
-160.6 | 0 | 28 | 0.28 | 549.5 | 0.26
-161.8 | 0 | 28 | 0.22 | 446.8 | 0.22
-163 | 0 | 28 | 0.29 | 573.7 | 0.28
-164.2 | 0 | 28 | 0.28 | 558.6 | 0.27
-165.4 | 0 | 28 | 0.23 | 457.5 | 0.22
-166.6 | 0 | 28 | 0.28 | 568 | 0.27
-167.8 | 0 | 28 | 0.26 | 520.7 | 0.25
-169 | 0 | 28 | 0.24 | 479.1 | 0.23
-170.2 | 0 | 28 | 0.23 | 463.6 | 0.23
-171.4 | 0 | 28 | 0.23 | 459.4 | 0.22
-172.6 | 0 | 28 | 0.25 | 509.2 | 0.24
-173.8 | 0 | 28 | 0.23 | 463.4 | 0.22
-175 | 0 | 28 | 0.25 | 508.9 | 0.24
-176.2 | 0 | 28 | 0.23 | 468.1 | 0.22
-177.4 | 0 | 28 | 0.22 | 435.2 | 0.21
-178.6 | 0 | 28 | 0.25 | 491 | 0.24
-179.8 | 0 | 28 | 0.26 | 511 | 0.25
-181 | 0 | 28 | 0.24 | 484.5 | 0.24
-182.2 | 0 | 28 | 0.26 | 523.5 | 0.26
-183.4 | 0 | 28 | 0.23 | 458.5 | 0.22
-184.6 | 0 | 28 | 0.19 | 387.6 | 0.19
-185.8 | 0 | 28 | 0.21 | 427.8 | 0.21
-187 | 0 | 28 | 0.2 | 401.2 | 0.19
-188.3 | 0 | 28 | 0.17 | 332.9 | 0.16
-189.4 | 0 | 28 | 0.18 | 365.9 | 0.18
-190.7 | 0 | 28 | 0.18 | 361.6 | 0.17
-191.8 | 0 | 28 | 0.17 | 348.2 | 0.17
-193.1 | 0 | 28 | 0.17 | 341.2 | 0.16
-194.3 | 0 | 28 | 0.16 | 319.5 | 0.16
-195.5 | 0 | 28 | 0.19 | 379.7 | 0.19
-196.7 | 0 | 28 | 0.16 | 324.8 | 0.15
-197.9 | 0 | 28 | 0.18 | 359.9 | 0.17
-199.1 | 0 | 28 | 0.18 | 352.3 | 0.17
-200 | 0 | 28 | 0.17 | 341.9 | 0.16
Measurement time: | 07/09/2022 16:42:41
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-199.1 | 0 | 28 | 0.13 | 256.6 | 0.14
-197.3 | 0 | 28 | 0.14 | 289.5 | 0.14
-196.1 | 0 | 28 | 0.15 | 298.3 | 0.15
-194.9 | 0 | 28 | 0.16 | 325.8 | 0.16
-193.7 | 0 | 28 | 0.19 | 388.5 | 0.19
-192.5 | 0 | 28 | 0.17 | 332.7 | 0.16
-191.3 | 0 | 28 | 0.19 | 377 | 0.18
-190.1 | 0 | 28 | 0.17 | 329.6 | 0.16
-188.9 | 0 | 28 | 0.15 | 306.8 | 0.15
-187.7 | 0 | 28 | 0.21 | 412.5 | 0.2
-186.5 | 0 | 28 | 0.2 | 402.5 | 0.2
-185.3 | 0 | 28 | 0.18 | 351.8 | 0.17
-184.1 | 0 | 28 | 0.18 | 355.2 | 0.17
-182.8 | 0 | 28 | 0.2 | 391.9 | 0.19
-181.6 | 0 | 28 | 0.17 | 344.2 | 0.16
-180.5 | 0 | 28 | 0.25 | 497 | 0.24
-179.2 | 0 | 28 | 0.22 | 431.7 | 0.21
-178.1 | 0 | 28 | 0.26 | 511.1 | 0.24
-176.9 | 0 | 28 | 0.24 | 474.4 | 0.23
-175.7 | 0 | 28 | 0.28 | 554.1 | 0.27
-174.4 | 0 | 28 | 0.22 | 431.3 | 0.21
-173.3 | 0 | 28 | 0.27 | 533.9 | 0.26
-172 | 0 | 28 | 0.2 | 408.8 | 0.2
-170.9 | 0 | 28 | 0.21 | 416.8 | 0.2
-169.6 | 0 | 28 | 0.24 | 487.1 | 0.23
-168.4 | 0 | 28 | 0.25 | 491.2 | 0.24
-167.2 | 0 | 28 | 0.25 | 492.7 | 0.24
-166 | 0 | 28 | 0.22 | 434.7 | 0.21
-164.8 | 0 | 28 | 0.23 | 450.1 | 0.22
-163.6 | 0 | 28 | 0.21 | 423.5 | 0.2
-162.4 | 0 | 28 | 0.24 | 485.3 | 0.24
-161.2 | 0 | 28 | 0.24 | 471.2 | 0.23
-160 | 0 | 28 | 0.23 | 458.3 | 0.22
-158.8 | 0 | 28 | 0.26 | 525 | 0.25
-157.6 | 0 | 28 | 0.23 | 459 | 0.22
-156.4 | 0 | 28 | 0.25 | 494.4 | 0.23
-155.2 | 0 | 28 | 0.24 | 482.5 | 0.24
-154 | 0 | 28 | 0.22 | 448.5 | 0.21
-152.8 | 0 | 28 | 0.25 | 504.5 | 0.24
-151.6 | 0 | 28 | 0.24 | 485 | 0.23
-150.4 | 0 | 28 | 0.24 | 481.8 | 0.23
-149.2 | 0 | 28 | 0.25 | 492.5 | 0.24
-148 | 0 | 28 | 0.24 | 483.1 | 0.24
-146.8 | 0 | 28 | 0.23 | 457.2 | 0.22
-145.6 | 0 | 28 | 0.24 | 473.6 | 0.23
-144.4 | 0 | 28 | 0.25 | 503.9 | 0.25
-143.2 | 0 | 28 | 0.24 | 472.2 | 0.23
-142 | 0 | 28 | 0.28 | 555.7 | 0.27
-140.8 | 0 | 28 | 0.3 | 589.6 | 0.28
-139.6 | 0 | 28 | 0.29 | 576.8 | 0.28
-138.4 | 0 | 28 | 0.27 | 539.1 | 0.26
-137.2 | 0 | 28 | 0.29 | 572.2 | 0.28
-136 | 0 | 28 | 0.29 | 587.4 | 0.28
-134.8 | 0 | 28 | 0.31 | 613.1 | 0.3
-133.6 | 0 | 28 | 0.34 | 676.2 | 0.33
-132.4 | 0 | 28 | 0.36 | 712.4 | 0.34
-131.2 | 0 | 28 | 0.37 | 734.8 | 0.36
-130 | 0 | 28 | 0.37 | 737.7 | 0.35
-128.8 | 0 | 28 | 0.38 | 766.8 | 0.37
-127.6 | 0 | 28 | 0.38 | 766.6 | 0.38
-126.4 | 0 | 28 | 0.42 | 840.7 | 0.4
-125.2 | 0 | 28 | 0.4 | 795.7 | 0.39
-124 | 0 | 28 | 0.43 | 849.1 | 0.41
-122.8 | 0 | 28 | 0.46 | 918.8 | 0.45
-121.6 | 0 | 28 | 0.42 | 838.6 | 0.41
-120.3 | 0 | 28 | 0.46 | 927.5 | 0.44
-119.2 | 0 | 28 | 0.53 | 1064.1 | 0.51
-118 | 0 | 28 | 0.53 | 1053 | 0.5
-116.7 | 0 | 28 | 0.56 | 1127 | 0.54
-115.6 | 0 | 28 | 0.58 | 1148.9 | 0.58
-114.3 | 0 | 28 | 0.68 | 1354.1 | 0.65
-113.2 | 0 | 28 | 0.78 | 1552.7 | 0.75
-112 | 0 | 28 | 0.81 | 1619.7 | 0.77
-110.8 | 0 | 28 | 0.82 | 1643 | 0.8
-109.5 | 0 | 28 | 1.06 | 2115.3 | 1.01
-108.3 | 0 | 28 | 1.07 | 2130.5 | 1.08
-107.1 | 0 | 28 | 1.36 | 2725 | 1.27
-105.9 | 0 | 28 | 1.47 | 2934.7 | 1.47
-104.7 | 0 | 28 | 1.74 | 3474.7 | 1.68
-103.5 | 0 | 28 | 1.97 | 3935.8 | 1.9
-102.3 | 0 | 28 | 2.24 | 4473.5 | 2.13
-101.1 | 0 | 28 | 2.59 | 5168.6 | 2.48
-99.9 | 0 | 28 | 2.98 | 5955.6 | 2.87
-98.7 | 0 | 28 | 3.34 | 6679.7 | 3.32
-97.5 | 0 | 28 | 4.06 | 8114.3 | 3.86
-96.3 | 0 | 28 | 4.66 | 9314.7 | 4.49
-95.1 | 0 | 28 | 5.5 | 10979.7 | 5.25
-93.9 | 0 | 28 | 6.32 | 12628 | 6.11
-92.7 | 0 | 28 | 7.41 | 14800.9 | 7.16
-91.5 | 0 | 28 | 8.67 | 17321.8 | 8.38
-90.3 | 0 | 28 | 10.23 | 20428.3 | 9.84
-89.1 | 0 | 28 | 12.23 | 24423.6 | 11.72
-87.9 | 0 | 28 | 14.43 | 28834.1 | 13.86
-86.7 | 0 | 28 | 17.36 | 34672.3 | 16.52
-85.5 | 0 | 28 | 20.65 | 41249.6 | 19.72
-84.3 | 0 | 28 | 24.62 | 49180.2 | 23.48
-83.1 | 0 | 28 | 29.45 | 58836.4 | 27.88
-81.9 | 0 | 28 | 34.35 | 68631.5 | 32.89
-80.7 | 0 | 28 | 38.66 | 77239.7 | 38.51
-79.5 | 0 | 28 | 46.65 | 93203.2 | 44.34
-78.3 | 0 | 28 | 52.42 | 104719.2 | 50.52
-77.1 | 0 | 28 | 58.93 | 117723.9 | 56.46
-75.9 | 0 | 28 | 64.17 | 128205.5 | 62.27
-74.7 | 0 | 28 | 70.35 | 140538.3 | 67.61
-73.5 | 0 | 28 | 74.55 | 148933.6 | 72.28
-72.3 | 0 | 28 | 79.17 | 158166 | 76.39
-71.1 | 0 | 28 | 82.98 | 165772 | 79.94
-69.8 | 0 | 28 | 85.9 | 171619.2 | 82.61
-68.7 | 0 | 28 | 88.84 | 177488.9 | 84.97
-67.4 | 0 | 28 | 90.82 | 181441.6 | 86.91
-66.2 | 0 | 28 | 92.68 | 185151.5 | 88.38
-65 | 0 | 28 | 89.54 | 178880.6 | 89.49
-63.8 | 0 | 28 | 93.57 | 186944.8 | 90.5
-62.6 | 0 | 28 | 95.46 | 190714.6 | 91.37
-61.4 | 0 | 28 | 95.03 | 189861.3 | 91.91
-60.2 | 0 | 28 | 95.87 | 191530.6 | 92.73
-59 | 0 | 28 | 95.94 | 191664 | 93.13
-57.8 | 0 | 28 | 98.58 | 196945.9 | 93.52
-56.6 | 0 | 28 | 97.21 | 194199.7 | 93.83
-55.4 | 0 | 28 | 97.21 | 194209.7 | 94.34
-54.2 | 0 | 28 | 98.55 | 196883.9 | 94.7
-53 | 0 | 28 | 98.4 | 196589.3 | 94.89
-51.8 | 0 | 28 | 99.09 | 197963.5 | 95.19
-50.6 | 0 | 28 | 100.37 | 200519.3 | 95.37
-49.4 | 0 | 28 | 99.87 | 199528.9 | 95.7
-48.2 | 0 | 28 | 100.36 | 200504.4 | 95.95
-47 | 0 | 28 | 100.39 | 200562.3 | 96.04
-45.8 | 0 | 28 | 99.67 | 199124.1 | 96.25
-44.6 | 0 | 28 | 99.61 | 199008 | 96.31
-43.4 | 0 | 28 | 99.58 | 198942.9 | 96.4
-42.2 | 0 | 28 | 99.45 | 198681.2 | 96.6
-41 | 0 | 28 | 99.84 | 199464.6 | 96.84
-39.8 | 0 | 28 | 99.84 | 199467.7 | 96.9
-38.6 | 0 | 28 | 100.73 | 201249.3 | 96.91
-37.4 | 0 | 28 | 101.44 | 202666.1 | 97.09
-36.2 | 0 | 28 | 101.13 | 202036.5 | 97.1
-35 | 0 | 28 | 102.25 | 204283.8 | 97.14
-33.8 | 0 | 28 | 102.14 | 204065.8 | 97.29
-32.6 | 0 | 28 | 97.01 | 193810.1 | 97.54
-31.4 | 0 | 28 | 101.3 | 202381.4 | 97.34
-30.2 | 0 | 28 | 100.53 | 200850.2 | 97.49
-29 | 0 | 28 | 100.76 | 201299.2 | 97.73
-27.8 | 0 | 28 | 101.13 | 202034 | 97.49
-26.6 | 0 | 28 | 99.88 | 199547.1 | 97.86
-25.4 | 0 | 28 | 102.15 | 204078.1 | 97.96
-24.2 | 0 | 28 | 101.6 | 202983.6 | 97.95
-23 | 0 | 28 | 102.28 | 204327 | 98.08
-21.8 | 0 | 28 | 102.85 | 205472.1 | 98.21
-20.6 | 0 | 28 | 99.2 | 198174.4 | 98.16
-19.3 | 0 | 28 | 101.42 | 202626.3 | 98.18
-18.2 | 0 | 28 | 103.34 | 206450.6 | 98.25
-17 | 0 | 28 | 101.32 | 202426.6 | 98.31
-15.7 | 0 | 28 | 101.55 | 202873.5 | 98.34
-14.6 | 0 | 28 | 100.17 | 200116.7 | 98.52
-13.3 | 0 | 28 | 100.61 | 201002.7 | 98.59
-12.2 | 0 | 28 | 100.28 | 200337.9 | 98.42
-10.9 | 0 | 28 | 101.66 | 203090.9 | 98.77
-9.7 | 0 | 28 | 102.31 | 204397.8 | 98.78
-8.5 | 0 | 28 | 102.99 | 205758.2 | 98.73
-7.3 | 0 | 28 | 102.92 | 205620.6 | 98.86
-6.1 | 0 | 28 | 103.85 | 207465.9 | 98.78
-4.9 | 0 | 28 | 103.42 | 206619.8 | 98.86
-3.7 | 0 | 28 | 103.75 | 207276.8 | 99.12
-2.5 | 0 | 28 | 103.41 | 206594.4 | 98.82
-1.3 | 0 | 28 | 101.91 | 203604.6 | 99.3
-0.1 | 0 | 28 | 103.09 | 205957 | 99.38
1.1 | 0 | 28 | 103.77 | 207318.5 | 99.08
2.3 | 0 | 28 | 101.58 | 202946.3 | 99.55
3.5 | 0 | 28 | 103.21 | 206189.8 | 99.42
4.7 | 0 | 28 | 104.78 | 209333.4 | 99.53
5.9 | 0 | 28 | 101.9 | 203573.4 | 99.65
7.1 | 0 | 28 | 103.45 | 206678 | 99.72
8.3 | 0 | 28 | 107.33 | 214429.3 | 99.63
9.5 | 0 | 28 | 101.5 | 202769.6 | 99.8
10.7 | 0 | 28 | 103.25 | 206277.7 | 99.82
11.9 | 0 | 28 | 104.89 | 209543.3 | 99.66
13.1 | 0 | 28 | 103.04 | 205845.5 | 99.87
14.3 | 0 | 28 | 98.81 | 197406.2 | 99.82
15.5 | 0 | 28 | 107.49 | 214735.5 | 99.87
16.7 | 0 | 28 | 100.76 | 201302.8 | 100
17.9 | 0 | 28 | 103.89 | 207546.2 | 99.83
19.1 | 0 | 28 | 108.04 | 215837.6 | 99.98
20.3 | 0 | 28 | 101.03 | 201848.2 | 99.81
21.5 | 0 | 28 | 104.43 | 208630.6 | 99.89
22.7 | 0 | 28 | 107.6 | 214969.5 | 99.75
23.9 | 0 | 28 | 99.32 | 198428.5 | 99.73
25.1 | 0 | 28 | 106.07 | 211912.8 | 99.76
26.3 | 0 | 28 | 101.48 | 202739.3 | 99.8
27.5 | 0 | 28 | 102.61 | 204994.7 | 99.92
28.7 | 0 | 28 | 99.04 | 197873.1 | 99.75
29.9 | 0 | 28 | 106.55 | 212867.5 | 99.78
31.1 | 0 | 28 | 103.75 | 207282 | 99.85
32.3 | 0 | 28 | 99.77 | 199313.3 | 99.92
33.5 | 0 | 28 | 107.98 | 215721.3 | 99.76
34.7 | 0 | 28 | 100.84 | 201460.3 | 99.62
35.9 | 0 | 28 | 104.48 | 208731.6 | 99.62
37.1 | 0 | 28 | 107.2 | 214171.9 | 99.64
38.3 | 0 | 28 | 99.75 | 199283.7 | 99.35
39.6 | 0 | 28 | 103.05 | 205882.7 | 99.43
40.8 | 0 | 28 | 105.84 | 211451.2 | 99.27
42 | 0 | 28 | 98.8 | 197381.3 | 99.08
43.2 | 0 | 28 | 101.84 | 203455 | 98.99
44.4 | 0 | 28 | 106.46 | 212697.2 | 99
45.6 | 0 | 28 | 97.88 | 195548.4 | 98.54
46.8 | 0 | 28 | 102.66 | 205102.5 | 98.61
48 | 0 | 28 | 107.18 | 214119.7 | 98.36
49.2 | 0 | 28 | 98.84 | 197463.4 | 98.09
50.4 | 0 | 28 | 101.68 | 203137.3 | 97.91
51.6 | 0 | 28 | 107.37 | 214512.6 | 97.38
52.8 | 0 | 28 | 97.47 | 194733.2 | 97.09
54 | 0 | 28 | 95.7 | 191195.6 | 96.68
55.2 | 0 | 28 | 104.97 | 209706.8 | 96.21
56.4 | 0 | 28 | 95.18 | 190153.5 | 95.99
57.6 | 0 | 28 | 98.6 | 196977.6 | 95.55
58.8 | 0 | 28 | 101.93 | 203638.2 | 94.67
60 | 0 | 28 | 96.91 | 193617.5 | 94.12
61.2 | 0 | 28 | 92.63 | 185055.4 | 93.32
62.4 | 0 | 28 | 99.69 | 199168.2 | 92.27
63.6 | 0 | 28 | 91.2 | 182201.2 | 91.32
64.8 | 0 | 28 | 93.49 | 186772.6 | 89.84
66 | 0 | 28 | 94.96 | 189718.7 | 88.15
67.2 | 0 | 28 | 87.78 | 175364.7 | 86.16
68.4 | 0 | 28 | 87.54 | 174896.3 | 83.9
69.6 | 0 | 28 | 86.95 | 173710.6 | 81.04
70.8 | 0 | 28 | 79.01 | 157849.7 | 77.31
72 | 0 | 28 | 76.93 | 153690.3 | 73.2
73.2 | 0 | 28 | 69.59 | 139030.3 | 68.49
74.5 | 0 | 28 | 64.25 | 128359.1 | 62.9
75.6 | 0 | 28 | 59.51 | 118896.1 | 57.37
76.8 | 0 | 28 | 54.81 | 109502.3 | 51.18
78 | 0 | 28 | 45.22 | 90350.1 | 45.04
79.2 | 0 | 28 | 39.91 | 79728.6 | 39.1
80.5 | 0 | 28 | 35.33 | 70574.6 | 33.57
81.6 | 0 | 28 | 29.29 | 58512 | 28.52
82.9 | 0 | 28 | 24.36 | 48664.5 | 24.08
84 | 0 | 28 | 21.23 | 42422.5 | 20.23
85.2 | 0 | 28 | 17.45 | 34855.1 | 17
86.4 | 0 | 28 | 14.52 | 29013.8 | 14.33
87.7 | 0 | 28 | 12.85 | 25665.8 | 12.13
88.9 | 0 | 28 | 10.53 | 21045.5 | 10.27
90.1 | 0 | 28 | 9.09 | 18152.1 | 8.71
91.3 | 0 | 28 | 7.99 | 15972.4 | 7.47
92.5 | 0 | 28 | 6.54 | 13066 | 6.4
93.7 | 0 | 28 | 5.77 | 11536 | 5.5
94.9 | 0 | 28 | 5.03 | 10054.1 | 4.76
96.1 | 0 | 28 | 4.21 | 8416.5 | 4.12
97.3 | 0 | 28 | 3.74 | 7470.9 | 3.61
98.5 | 0 | 28 | 3.3 | 6585.2 | 3.16
99.7 | 0 | 28 | 2.82 | 5637.9 | 2.75
100.9 | 0 | 28 | 2.45 | 4899.9 | 2.41
102.1 | 0 | 28 | 2.19 | 4379.9 | 2.12
103.3 | 0 | 28 | 1.89 | 3773.3 | 1.83
104.5 | 0 | 28 | 1.66 | 3310.7 | 1.6
105.7 | 0 | 28 | 1.49 | 2972.8 | 1.4
106.9 | 0 | 28 | 1.27 | 2543 | 1.24
108.1 | 0 | 28 | 1.16 | 2323.4 | 1.1
109.3 | 0 | 28 | 0.99 | 1977.2 | 0.94
110.5 | 0 | 28 | 0.84 | 1678.7 | 0.81
111.7 | 0 | 28 | 0.74 | 1485.3 | 0.71
112.9 | 0 | 28 | 0.66 | 1326.1 | 0.63
114.1 | 0 | 28 | 0.54 | 1080.2 | 0.52
115.3 | 0 | 28 | 0.49 | 982.8 | 0.48
116.5 | 0 | 28 | 0.41 | 820.3 | 0.4
117.7 | 0 | 28 | 0.39 | 770.5 | 0.37
118.9 | 0 | 28 | 0.36 | 721.6 | 0.35
120.1 | 0 | 28 | 0.32 | 649.1 | 0.31
121.3 | 0 | 28 | 0.28 | 568.5 | 0.28
122.5 | 0 | 28 | 0.27 | 548.6 | 0.26
123.7 | 0 | 28 | 0.26 | 526.2 | 0.25
124.9 | 0 | 28 | 0.2 | 403.9 | 0.19
126.1 | 0 | 28 | 0.2 | 395.7 | 0.19
127.3 | 0 | 28 | 0.24 | 485.4 | 0.23
128.5 | 0 | 28 | 0.2 | 393.4 | 0.19
129.7 | 0 | 28 | 0.17 | 348 | 0.17
130.9 | 0 | 28 | 0.18 | 364.9 | 0.18
132.1 | 0 | 28 | 0.16 | 328 | 0.16
133.4 | 0 | 28 | 0.17 | 343 | 0.17
134.6 | 0 | 28 | 0.14 | 284.6 | 0.14
135.7 | 0 | 28 | 0.09 | 179 | 0.09
137 | 0 | 28 | 0.12 | 231.9 | 0.11
138.2 | 0 | 28 | 0.11 | 227.7 | 0.11
139.4 | 0 | 28 | 0.09 | 173.7 | 0.08
140.6 | 0 | 28 | 0.1 | 205.3 | 0.1
141.8 | 0 | 28 | 0.09 | 180.8 | 0.09
143 | 0 | 28 | 0.14 | 283.1 | 0.13
144.2 | 0 | 28 | 0.11 | 213.3 | 0.1
145.4 | 0 | 28 | 0.11 | 214.5 | 0.1
146.6 | 0 | 28 | 0.09 | 185.5 | 0.09
147.8 | 0 | 28 | 0.1 | 193.2 | 0.09
149 | 0 | 28 | 0.1 | 201 | 0.1
150.2 | 0 | 28 | 0.09 | 182.4 | 0.09
151.4 | 0 | 28 | 0.11 | 222.6 | 0.11
152.6 | 0 | 28 | 0.09 | 176.3 | 0.08
153.8 | 0 | 28 | 0.07 | 135.4 | 0.07
155 | 0 | 28 | 0.11 | 221.2 | 0.11
156.2 | 0 | 28 | 0.06 | 113.2 | 0.05
157.4 | 0 | 28 | 0.03 | 57.2 | 0.03
158.6 | 0 | 28 | 0.03 | 63.9 | 0.03
159.8 | 0 | 28 | 0.02 | 43.4 | 0.02
161 | 0 | 28 | 0.07 | 144.6 | 0.07
162.2 | 0 | 28 | 0.05 | 99.1 | 0.05
163.4 | 0 | 28 | 0.01 | 23.6 | 0.01
164.6 | 0 | 28 | 0.05 | 92 | 0.04
165.8 | 0 | 28 | 0.05 | 92.3 | 0.04
167 | 0 | 28 | 0.04 | 77.4 | 0.04
168.2 | 0 | 28 | 0.05 | 100 | 0.05
169.4 | 0 | 28 | 0.04 | 69.9 | 0.03
170.6 | 0 | 28 | 0.03 | 68.9 | 0.03
171.8 | 0 | 28 | 0.07 | 142.4 | 0.07
173 | 0 | 28 | 0.04 | 75.3 | 0.04
174.2 | 0 | 28 | 0.02 | 49.5 | 0.02
175.4 | 0 | 28 | 0.02 | 35.6 | 0.02
176.6 | 0 | 28 | 0.04 | 84.2 | 0.04
177.8 | 0 | 28 | 0.07 | 134.5 | 0.07
179 | 0 | 28 | 0.03 | 69.1 | 0.03
180.2 | 0 | 28 | 0.02 | 32.4 | 0.02
181.4 | 0 | 28 | 0.02 | 30.3 | 0.01
182.7 | 0 | 28 | 0.04 | 71.5 | 0.03
183.8 | 0 | 28 | 0.03 | 56.3 | 0.03
185.1 | 0 | 28 | 0 | 9 | 0
186.3 | 0 | 28 | 0.04 | 72.4 | 0.03
187.5 | 0 | 28 | 0.04 | 75.7 | 0.04
188.7 | 0 | 28 | 0.05 | 109.6 | 0.05
189.9 | 0 | 28 | 0.07 | 137.6 | 0.07
191.1 | 0 | 28 | 0.04 | 84.9 | 0.04
192.3 | 0 | 28 | 0.03 | 53.4 | 0.03
193.5 | 0 | 28 | 0.02 | 33 | 0.02
194.7 | 0 | 28 | 0 | 2.4 | 0
195.9 | 0 | 28 | 0.01 | 19.4 | 0.01
197.1 | 0 | 28 | 0.02 | 31.3 | 0.01
198.3 | 0 | 28 | 0.01 | 15.9 | 0.01
200 | 0 | 28 | 0.01 | 18.2 | 0.01
Measurement time: | 07/09/2022 16:48:51
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-200 | 0 | 28 | 0.08 | 153.1 | 0.08
-199.3 | 0 | 28 | 0.1 | 200 | 0.1
-198 | 0 | 28 | 0.12 | 240.8 | 0.11
-196.8 | 0 | 28 | 0.07 | 137.5 | 0.07
-195.6 | 0 | 28 | 0.09 | 172 | 0.08
-194.4 | 0 | 28 | 0.1 | 208 | 0.1
-193.2 | 0 | 28 | 0.1 | 210.1 | 0.1
-192 | 0 | 28 | 0.09 | 181.3 | 0.09
-190.8 | 0 | 28 | 0.11 | 221.4 | 0.11
-189.6 | 0 | 28 | 0.13 | 262.4 | 0.13
-188.4 | 0 | 28 | 0.12 | 236.6 | 0.12
-187.2 | 0 | 28 | 0.13 | 257.6 | 0.13
-186 | 0 | 28 | 0.1 | 192.7 | 0.09
-184.7 | 0 | 28 | 0.13 | 251.4 | 0.12
-183.5 | 0 | 28 | 0.14 | 279.9 | 0.13
-182.3 | 0 | 28 | 0.15 | 294.6 | 0.14
-181.1 | 0 | 28 | 0.1 | 204.5 | 0.1
-180 | 0 | 28 | 0.15 | 300.2 | 0.14
-178.7 | 0 | 28 | 0.15 | 291.2 | 0.14
-177.5 | 0 | 28 | 0.16 | 314.4 | 0.15
-176.3 | 0 | 28 | 0.15 | 308.4 | 0.14
-175.1 | 0 | 28 | 0.16 | 311 | 0.15
-173.9 | 0 | 28 | 0.13 | 262.6 | 0.13
-172.7 | 0 | 28 | 0.17 | 336.6 | 0.16
-171.5 | 0 | 28 | 0.19 | 381.2 | 0.19
-170.4 | 0 | 28 | 0.17 | 333.1 | 0.17
-169.1 | 0 | 28 | 0.18 | 365.2 | 0.17
-167.9 | 0 | 28 | 0.12 | 247.1 | 0.12
-166.7 | 0 | 28 | 0.17 | 342.3 | 0.17
-165.5 | 0 | 28 | 0.15 | 306.6 | 0.15
-164.3 | 0 | 28 | 0.16 | 315.9 | 0.15
-163.1 | 0 | 28 | 0.14 | 289.8 | 0.14
-161.9 | 0 | 28 | 0.13 | 252.7 | 0.12
-160.7 | 0 | 28 | 0.16 | 310.9 | 0.15
-159.5 | 0 | 28 | 0.16 | 323.6 | 0.16
-158.3 | 0 | 28 | 0.16 | 325.4 | 0.15
-157.1 | 0 | 28 | 0.15 | 306.7 | 0.15
-155.9 | 0 | 28 | 0.16 | 326.9 | 0.17
-154.7 | 0 | 28 | 0.22 | 434.3 | 0.2
-153.5 | 0 | 28 | 0.19 | 377.1 | 0.18
-152.3 | 0 | 28 | 0.16 | 322.8 | 0.15
-151.1 | 0 | 28 | 0.18 | 366.8 | 0.18
-149.9 | 0 | 28 | 0.16 | 310.5 | 0.15
-148.7 | 0 | 28 | 0.16 | 328.1 | 0.16
-147.5 | 0 | 28 | 0.2 | 405.6 | 0.19
-146.3 | 0 | 28 | 0.15 | 293.7 | 0.14
-145.1 | 0 | 28 | 0.17 | 345.7 | 0.17
-143.9 | 0 | 28 | 0.18 | 361.1 | 0.17
-142.7 | 0 | 28 | 0.15 | 302.5 | 0.15
-141.5 | 0 | 28 | 0.2 | 396.9 | 0.18
-140.3 | 0 | 28 | 0.18 | 369 | 0.17
-139.1 | 0 | 28 | 0.19 | 380 | 0.18
-137.9 | 0 | 28 | 0.22 | 443.9 | 0.21
-136.7 | 0 | 28 | 0.22 | 449.6 | 0.2
-135.5 | 0 | 28 | 0.19 | 371.1 | 0.18
-134.3 | 0 | 28 | 0.21 | 427.6 | 0.22
-133.1 | 0 | 28 | 0.26 | 526.2 | 0.25
-131.9 | 0 | 28 | 0.22 | 434 | 0.21
-130.7 | 0 | 28 | 0.2 | 406.5 | 0.19
-129.5 | 0 | 28 | 0.22 | 445.5 | 0.2
-128.3 | 0 | 28 | 0.24 | 476.7 | 0.23
-127.1 | 0 | 28 | 0.26 | 513.4 | 0.24
-125.9 | 0 | 28 | 0.32 | 649.1 | 0.3
-124.7 | 0 | 28 | 0.29 | 572.9 | 0.28
-123.4 | 0 | 28 | 0.36 | 727.5 | 0.37
-122.3 | 0 | 28 | 0.36 | 712.7 | 0.34
-121.1 | 0 | 28 | 0.35 | 694.7 | 0.32
-119.9 | 0 | 28 | 0.38 | 770.2 | 0.36
-118.7 | 0 | 28 | 0.48 | 956 | 0.44
-117.4 | 0 | 28 | 0.4 | 802.8 | 0.4
-116.3 | 0 | 28 | 0.52 | 1037.9 | 0.5
-115 | 0 | 28 | 0.58 | 1164.5 | 0.54
-113.8 | 0 | 28 | 0.54 | 1087 | 0.56
-112.7 | 0 | 28 | 0.68 | 1363.6 | 0.64
-111.4 | 0 | 28 | 0.71 | 1425.3 | 0.67
-110.2 | 0 | 28 | 0.83 | 1656.5 | 0.84
-109 | 0 | 28 | 0.99 | 1984.9 | 0.89
-107.8 | 0 | 28 | 1.12 | 2233.2 | 1.05
-106.6 | 0 | 28 | 1.21 | 2414.7 | 1.15
-105.4 | 0 | 28 | 1.43 | 2859.2 | 1.34
-104.2 | 0 | 28 | 1.72 | 3448.7 | 1.6
-103 | 0 | 28 | 1.9 | 3792.4 | 1.81
-101.8 | 0 | 28 | 2.27 | 4549 | 2.12
-100.6 | 0 | 28 | 2.75 | 5509.5 | 2.49
-99.4 | 0 | 28 | 2.94 | 5874.4 | 2.78
-98.2 | 0 | 28 | 3.54 | 7089 | 3.31
-97 | 0 | 28 | 4.31 | 8628.2 | 3.92
-95.8 | 0 | 28 | 4.63 | 9267.5 | 4.65
-94.6 | 0 | 28 | 5.67 | 11351.6 | 5.48
-93.4 | 0 | 28 | 6.99 | 13986 | 6.36
-92.2 | 0 | 28 | 7.52 | 15041.7 | 7.26
-91 | 0 | 28 | 9.19 | 18392.5 | 8.64
-89.8 | 0 | 28 | 11.19 | 22393.7 | 10.32
-88.6 | 0 | 28 | 12.6 | 25223.7 | 12.49
-87.4 | 0 | 28 | 15.38 | 30785.5 | 14.7
-86.2 | 0 | 28 | 18.6 | 37221.1 | 18.25
-85 | 0 | 28 | 21.51 | 43044 | 21.87
-83.8 | 0 | 28 | 26.23 | 52487.5 | 25.44
-82.6 | 0 | 28 | 31.28 | 62603.8 | 29.94
-81.4 | 0 | 28 | 35.73 | 71498.8 | 35.1
-80.2 | 0 | 28 | 42.55 | 85139 | 40.28
-79 | 0 | 28 | 49.08 | 98213.1 | 45.81
-77.8 | 0 | 28 | 53.96 | 107972.5 | 49.22
-76.6 | 0 | 28 | 61.32 | 122699.3 | 57.59
-75.4 | 0 | 28 | 67.92 | 135916 | 63.24
-74.2 | 0 | 28 | 71.58 | 143243.1 | 69.49
-73 | 0 | 28 | 76.67 | 153421.3 | 72.29
-71.8 | 0 | 28 | 81.64 | 163378.5 | 74.32
-70.6 | 0 | 28 | 83.84 | 167781.1 | 80.58
-69.4 | 0 | 28 | 86.89 | 173885.2 | 82.63
-68.2 | 0 | 28 | 90.15 | 180405.6 | 84.88
-66.9 | 0 | 28 | 89.87 | 179843.2 | 83.58
-65.7 | 0 | 28 | 92.33 | 184761.8 | 83.89
-64.5 | 0 | 28 | 94.55 | 189202.5 | 85.94
-63.3 | 0 | 28 | 93.45 | 187011.7 | 84.38
-62.1 | 0 | 28 | 95.82 | 191748.8 | 86.69
-60.9 | 0 | 28 | 96.45 | 193001.2 | 91.6
-59.7 | 0 | 28 | 96.24 | 192592.8 | 89.71
-58.5 | 0 | 28 | 96.68 | 193473.6 | 87.87
-57.3 | 0 | 28 | 97.97 | 196056.4 | 84.59
-56.1 | 0 | 28 | 97.15 | 194403.4 | 90.27
-54.9 | 0 | 28 | 98.39 | 196893.1 | 88.54
-53.7 | 0 | 28 | 97.97 | 196049.9 | 89.6
-52.5 | 0 | 28 | 98.02 | 196140.7 | 91.3
-51.3 | 0 | 28 | 99.84 | 199788.1 | 89.77
-50.1 | 0 | 28 | 99.34 | 198782.8 | 91.12
-48.9 | 0 | 28 | 99.5 | 199102.2 | 91.48
-47.7 | 0 | 28 | 99.26 | 198630.6 | 92.42
-46.5 | 0 | 28 | 99.8 | 199705.2 | 90.61
-45.3 | 0 | 28 | 99.95 | 200016.2 | 89.68
-44.1 | 0 | 28 | 99.72 | 199547 | 89.2
-42.9 | 0 | 28 | 95.67 | 191440.9 | 89.6
-41.7 | 0 | 28 | 99.95 | 200009.9 | 94.46
-40.5 | 0 | 28 | 99.75 | 199607.5 | 95.24
-39.3 | 0 | 28 | 101.03 | 202182.4 | 94.22
-38.1 | 0 | 28 | 100.32 | 200747.7 | 91.06
-36.9 | 0 | 28 | 95.87 | 191855.7 | 89.53
-35.7 | 0 | 28 | 101.34 | 202802.9 | 91.99
-34.5 | 0 | 28 | 96.29 | 192685.6 | 90.8
-33.3 | 0 | 28 | 99.23 | 198569.6 | 86.98
-32.1 | 0 | 28 | 99.73 | 199565.5 | 90.61
-30.9 | 0 | 28 | 100.73 | 201570.1 | 89.96
-29.7 | 0 | 28 | 99.79 | 199691 | 89.96
-28.5 | 0 | 28 | 100.32 | 200756.7 | 93.21
-27.3 | 0 | 28 | 99.88 | 199876.5 | 94.79
-26.1 | 0 | 28 | 99.14 | 198394.6 | 92.7
-24.9 | 0 | 28 | 101.9 | 203920.3 | 97.58
-23.7 | 0 | 28 | 98.31 | 196734.4 | 100
-22.5 | 0 | 28 | 100.66 | 201441 | 92.44
-21.3 | 0 | 28 | 98.68 | 197466.5 | 93.69
-20 | 0 | 28 | 99.59 | 199284.6 | 92.37
-18.9 | 0 | 28 | 97.96 | 196021.1 | 89.57
-17.7 | 0 | 28 | 98.55 | 197211.3 | 93.32
-16.4 | 0 | 28 | 102.27 | 204647.3 | 92.63
-15.2 | 0 | 28 | 96.69 | 193487.5 | 92.19
-14 | 0 | 28 | 98.94 | 197985.7 | 91.9
-12.8 | 0 | 28 | 99.91 | 199930.5 | 94.99
-11.6 | 0 | 28 | 99.77 | 199655.3 | 92.21
-10.4 | 0 | 28 | 100.2 | 200513.2 | 94.33
-9.2 | 0 | 28 | 100.35 | 200805.9 | 90.67
-8 | 0 | 28 | 100.39 | 200895.8 | 92.35
-6.8 | 0 | 28 | 100.63 | 201378.6 | 92.62
-5.6 | 0 | 28 | 100.53 | 201162.8 | 91.11
-4.4 | 0 | 28 | 100.54 | 201196.5 | 92.44
-3.2 | 0 | 28 | 101.55 | 203208.2 | 95.32
-2 | 0 | 28 | 100.85 | 201806.9 | 96.23
-0.8 | 0 | 28 | 100.6 | 201312.4 | 92.9
0.4 | 0 | 28 | 101.32 | 202763.6 | 95.9
1.6 | 0 | 28 | 100.73 | 201566.2 | 93.54
2.8 | 0 | 28 | 100.06 | 200240.6 | 91.11
4 | 0 | 28 | 100.49 | 201090.1 | 90.51
5.2 | 0 | 28 | 99.27 | 198655.9 | 91.33
6.4 | 0 | 28 | 99.95 | 200009.9 | 90.27
7.6 | 0 | 28 | 99.21 | 198527.6 | 98.32
8.8 | 0 | 28 | 99.04 | 198191.3 | 98.44
10 | 0 | 28 | 102.7 | 205508.7 | 93.39
11.2 | 0 | 28 | 95 | 190106 | 92.51
12.4 | 0 | 28 | 100.2 | 200509.9 | 94.01
13.6 | 0 | 28 | 100 | 200111.4 | 92.72
14.8 | 0 | 28 | 100.57 | 201257.4 | 93.41
16 | 0 | 28 | 101.02 | 202151.4 | 89.94
17.2 | 0 | 28 | 99.72 | 199557.6 | 93.6
18.4 | 0 | 28 | 100.29 | 200695.5 | 93.08
19.6 | 0 | 28 | 100.92 | 201943.8 | 94.3
20.8 | 0 | 28 | 100.95 | 202005.1 | 91.91
22 | 0 | 28 | 96.49 | 193089.3 | 93.57
23.3 | 0 | 28 | 100.2 | 200516.1 | 91.08
24.5 | 0 | 28 | 101 | 202107 | 95.62
25.6 | 0 | 28 | 95.97 | 192043.4 | 92.43
26.8 | 0 | 28 | 100.15 | 200408.9 | 95.38
28 | 0 | 28 | 100.41 | 200926.9 | 92.18
29.2 | 0 | 28 | 98.99 | 198099.8 | 96.64
30.5 | 0 | 28 | 99.12 | 198353.5 | 97.59
31.7 | 0 | 28 | 98.77 | 197649.1 | 94.23
32.9 | 0 | 28 | 98.18 | 196463 | 90.49
34.1 | 0 | 28 | 102.28 | 204667.3 | 95.66
35.3 | 0 | 28 | 96.64 | 193394.4 | 92.14
36.5 | 0 | 28 | 99.63 | 199362.7 | 98.44
37.7 | 0 | 28 | 99.51 | 199136.2 | 88.81
38.9 | 0 | 28 | 98.26 | 196633 | 88.04
40.1 | 0 | 28 | 98.76 | 197628.7 | 93.92
41.3 | 0 | 28 | 97.97 | 196058.6 | 89.44
42.5 | 0 | 28 | 99.03 | 198164 | 92.74
43.7 | 0 | 28 | 99.52 | 199149.5 | 89.54
44.9 | 0 | 28 | 99.05 | 198209.9 | 87.73
46.1 | 0 | 28 | 99.57 | 199251.6 | 93.46
47.3 | 0 | 28 | 97.99 | 196083.9 | 91.38
48.5 | 0 | 28 | 98.71 | 197539.7 | 92.02
49.7 | 0 | 28 | 97.16 | 194435.3 | 90.58
50.9 | 0 | 28 | 96.26 | 192621.5 | 87.32
52.1 | 0 | 28 | 99.41 | 198925.6 | 91.21
53.3 | 0 | 28 | 95.57 | 191246.5 | 90.75
54.5 | 0 | 28 | 91.82 | 183741 | 88.59
55.7 | 0 | 28 | 99.14 | 198385.3 | 89.16
56.9 | 0 | 28 | 92.21 | 184526.1 | 91.12
58.1 | 0 | 28 | 94.28 | 188665.8 | 86.2
59.3 | 0 | 28 | 94.21 | 188530.1 | 90.02
60.5 | 0 | 28 | 95.12 | 190338.9 | 88.6
61.7 | 0 | 28 | 93.29 | 186690.4 | 85.79
62.9 | 0 | 28 | 92.34 | 184773.6 | 84.56
64.1 | 0 | 28 | 92.58 | 185254.7 | 82.95
65.3 | 0 | 28 | 89.38 | 178862.1 | 80.46
66.5 | 0 | 28 | 88.19 | 176476.5 | 82.4
67.7 | 0 | 28 | 84.56 | 169208.8 | 78.65
68.9 | 0 | 28 | 81.93 | 163949.8 | 76.31
70.1 | 0 | 28 | 79.56 | 159202.9 | 73.96
71.3 | 0 | 28 | 74.72 | 149529 | 69.19
72.5 | 0 | 28 | 70.53 | 141144.7 | 67.88
73.8 | 0 | 28 | 65.77 | 131619.3 | 59.71
75 | 0 | 28 | 60.55 | 121165.3 | 56.65
76.1 | 0 | 28 | 54.8 | 109652 | 51.92
77.4 | 0 | 28 | 47.98 | 96004.6 | 45.31
78.5 | 0 | 28 | 42.26 | 84574.9 | 39.81
79.8 | 0 | 28 | 34.51 | 69050.4 | 34.97
81 | 0 | 28 | 30.36 | 60753.9 | 28.83
82.2 | 0 | 28 | 25.48 | 50980 | 24.42
83.4 | 0 | 28 | 21.22 | 42458.2 | 20.74
84.6 | 0 | 28 | 17.49 | 35006.7 | 16.64
85.8 | 0 | 28 | 14.5 | 29013.9 | 13.84
87 | 0 | 28 | 12.13 | 24269.5 | 11.03
88.2 | 0 | 28 | 10.22 | 20442.2 | 9.69
89.4 | 0 | 28 | 8.56 | 17133.6 | 8.1
90.6 | 0 | 28 | 7.39 | 14788.1 | 6.81
91.8 | 0 | 28 | 5.92 | 11842.2 | 5.95
93 | 0 | 28 | 5.25 | 10509.9 | 4.87
94.2 | 0 | 28 | 4.47 | 8939.7 | 4.07
95.4 | 0 | 28 | 3.68 | 7359.5 | 3.59
96.6 | 0 | 28 | 3.34 | 6675.7 | 2.98
97.8 | 0 | 28 | 2.88 | 5772.5 | 2.69
99 | 0 | 28 | 2.49 | 4974.8 | 2.37
100.2 | 0 | 28 | 2.17 | 4348.4 | 2.04
101.4 | 0 | 28 | 1.92 | 3848.7 | 1.81
102.6 | 0 | 28 | 1.64 | 3288.3 | 1.59
103.8 | 0 | 28 | 1.46 | 2923.3 | 1.4
105 | 0 | 28 | 1.25 | 2499 | 1.23
106.2 | 0 | 28 | 1.1 | 2196 | 1.04
107.4 | 0 | 28 | 0.97 | 1937.4 | 0.93
108.6 | 0 | 28 | 0.83 | 1657.5 | 0.77
109.8 | 0 | 28 | 0.76 | 1525 | 0.7
111 | 0 | 28 | 0.62 | 1234 | 0.59
112.2 | 0 | 28 | 0.59 | 1182.6 | 0.53
113.4 | 0 | 28 | 0.51 | 1017.1 | 0.46
114.6 | 0 | 28 | 0.45 | 901.5 | 0.42
115.8 | 0 | 28 | 0.36 | 717.3 | 0.33
117 | 0 | 28 | 0.34 | 678.9 | 0.3
118.2 | 0 | 28 | 0.31 | 628 | 0.29
119.4 | 0 | 28 | 0.29 | 576.8 | 0.26
120.6 | 0 | 28 | 0.29 | 579.5 | 0.25
121.8 | 0 | 28 | 0.24 | 488.8 | 0.23
123 | 0 | 28 | 0.18 | 361.1 | 0.17
124.2 | 0 | 28 | 0.24 | 485.2 | 0.21
125.4 | 0 | 28 | 0.18 | 356.5 | 0.16
126.6 | 0 | 28 | 0.15 | 295.4 | 0.14
127.8 | 0 | 28 | 0.16 | 327.4 | 0.14
129 | 0 | 28 | 0.12 | 243.3 | 0.11
130.2 | 0 | 28 | 0.17 | 334.9 | 0.15
131.4 | 0 | 28 | 0.13 | 268.6 | 0.12
132.6 | 0 | 28 | 0.09 | 186 | 0.09
133.8 | 0 | 28 | 0.08 | 169 | 0.08
135 | 0 | 28 | 0.06 | 122.8 | 0.06
136.2 | 0 | 28 | 0.05 | 109.3 | 0.05
137.4 | 0 | 28 | 0.05 | 106.2 | 0.05
138.6 | 0 | 28 | 0.09 | 177.3 | 0.08
139.9 | 0 | 28 | 0.11 | 213 | 0.1
141 | 0 | 28 | 0.1 | 196.5 | 0.09
142.3 | 0 | 28 | 0.08 | 161.6 | 0.07
143.5 | 0 | 28 | 0.07 | 135.3 | 0.06
144.7 | 0 | 28 | 0.05 | 107 | 0.05
145.9 | 0 | 28 | 0.07 | 149.9 | 0.07
147.1 | 0 | 28 | 0.02 | 49.9 | 0.02
148.3 | 0 | 28 | 0.06 | 117.5 | 0.05
149.5 | 0 | 28 | 0.03 | 54.9 | 0.02
150.7 | 0 | 28 | 0.04 | 87.1 | 0.04
151.9 | 0 | 28 | 0.04 | 77.5 | 0.04
153.1 | 0 | 28 | 0.02 | 38.2 | 0.02
154.3 | 0 | 28 | 0.04 | 80.2 | 0.04
155.5 | 0 | 28 | 0.01 | 14.9 | 0.01
156.7 | 0 | 28 | 0.01 | 27.1 | 0.01
157.9 | 0 | 28 | 0.01 | 17.4 | 0.01
159.1 | 0 | 28 | 0.03 | 54.3 | 0.03
160.3 | 0 | 28 | 0.05 | 106.9 | 0.05
161.5 | 0 | 28 | 0.01 | 24.8 | 0.01
162.7 | 0 | 28 | 0.04 | 70.1 | 0.03
163.9 | 0 | 28 | 0.01 | 10.2 | 0
165.1 | 0 | 28 | 0.02 | 41.4 | 0.02
166.3 | 0 | 28 | 0.02 | 30.7 | 0.01
167.5 | 0 | 28 | 0.01 | 15.3 | 0.01
168.7 | 0 | 28 | 0.02 | 39.5 | 0.02
169.9 | 0 | 28 | 0.02 | 32.4 | 0.02
171.1 | 0 | 28 | 0.01 | 17.4 | 0.01
172.3 | 0 | 28 | 0.02 | 41 | 0.02
173.5 | 0 | 28 | 0 | 9.8 | 0
174.7 | 0 | 28 | 0.02 | 45.9 | 0.02
175.9 | 0 | 28 | 0.02 | 43.5 | 0.02
177.1 | 0 | 28 | 0.01 | 14.2 | 0.01
178.3 | 0 | 28 | 0.07 | 142.3 | 0.06
179.5 | 0 | 28 | 0.01 | 26.4 | 0.01
180.7 | 0 | 28 | 0 | 6.8 | 0
181.9 | 0 | 28 | 0 | 0.7 | 0
183.1 | 0 | 28 | 0.01 | 19.8 | 0.01
184.3 | 0 | 28 | 0.01 | 19.3 | 0.01
185.5 | 0 | 28 | 0.01 | 10.1 | 0
186.7 | 0 | 28 | 0.03 | 52.1 | 0.03
187.9 | 0 | 28 | 0.03 | 62.3 | 0.03
189.1 | 0 | 28 | 0.04 | 74 | 0.03
190.3 | 0 | 28 | 0.01 | 12.2 | 0.01
191.5 | 0 | 28 | 0.02 | 49 | 0.02
192.7 | 0 | 28 | 0.02 | 48.6 | 0.02
193.9 | 0 | 28 | 0.01 | 21.7 | 0.01
195.1 | 0 | 28 | 0.03 | 53.5 | 0.03
196.3 | 0 | 28 | 0.02 | 42.3 | 0.02
197.6 | 0 | 28 | 0.02 | 30.6 | 0.02
198.7 | 0 | 28 | 0.01 | 19.1 | 0.01
200 | 0 | 28 | 0.01 | 16.2 | 0.01
Measurement time: | 07/09/2022 16:53:37
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-200 | 0 | 28 | 0.14 | 286.3 | 0
-199.3 | 0 | 28 | 0.06 | 127.1 | 0
-197.9 | 0 | 28 | 0.07 | 133.3 | 0
-196.7 | 0 | 28 | 0.06 | 132.8 | 0
-195.5 | 0 | 28 | 0.03 | 57.6 | 0
-194.3 | 0 | 28 | 0.09 | 193.7 | 0
-193.1 | 0 | 28 | 0.1 | 210.3 | 0
-191.9 | 0 | 28 | 0.09 | 190.2 | 0
-190.7 | 0 | 28 | 0.07 | 137.3 | 0
-189.5 | 0 | 28 | 0.1 | 207.5 | 0
-188.3 | 0 | 28 | 0.09 | 190.1 | 0
-187.1 | 0 | 28 | 0.07 | 137.6 | 0
-185.9 | 0 | 28 | 0.09 | 179.3 | 0
-184.7 | 0 | 28 | 0.14 | 291.1 | 0
-183.5 | 0 | 28 | 0.11 | 217.4 | 0
-182.3 | 0 | 28 | 0.1 | 202.7 | 0
-181.1 | 0 | 28 | 0.08 | 167.7 | 0
-179.9 | 0 | 28 | 0.11 | 215.4 | 0
-178.7 | 0 | 28 | 0.11 | 218.3 | 0
-177.5 | 0 | 28 | 0.15 | 306.3 | 0
-176.3 | 0 | 28 | 0.11 | 215.8 | 0
-175.1 | 0 | 28 | 0.11 | 230.3 | 0
-173.9 | 0 | 28 | 0.09 | 188.3 | 0
-172.7 | 0 | 28 | 0.06 | 116 | 0
-171.5 | 0 | 28 | 0.12 | 238.9 | 0
-170.3 | 0 | 28 | 0.14 | 283.2 | 0
-169.1 | 0 | 28 | 0.07 | 151.3 | 0
-167.9 | 0 | 28 | 0.12 | 254 | 0
-166.7 | 0 | 28 | 0.12 | 255.3 | 0
-165.5 | 0 | 28 | 0.08 | 173.1 | 0
-164.3 | 0 | 28 | 0.12 | 249.5 | 0
-163.1 | 0 | 28 | 0.13 | 271.7 | 0
-161.9 | 0 | 28 | 0.13 | 271.2 | 0
-160.7 | 0 | 28 | 0.13 | 265.4 | 0
-159.5 | 0 | 28 | 0.17 | 357.5 | 0
-158.3 | 0 | 28 | 0.13 | 266.2 | 0
-157.1 | 0 | 28 | 0.16 | 317.4 | 0
-155.9 | 0 | 28 | 0.14 | 280.8 | 0
-154.7 | 0 | 28 | 0.14 | 279.4 | 0
-153.4 | 0 | 28 | 0.13 | 257 | 0
-152.3 | 0 | 28 | 0.15 | 313.8 | 0
-151 | 0 | 28 | 0.12 | 240.3 | 0
-149.8 | 0 | 28 | 0.14 | 290.7 | 0
-148.6 | 0 | 28 | 0.14 | 284.3 | 0
-147.4 | 0 | 28 | 0.15 | 308.9 | 0
-146.2 | 0 | 28 | 0.17 | 339.4 | 0
-145 | 0 | 28 | 0.15 | 305.2 | 0
-143.8 | 0 | 28 | 0.12 | 245.2 | 0
-142.6 | 0 | 28 | 0.14 | 294.4 | 0
-141.4 | 0 | 28 | 0.18 | 366.9 | 0
-140.2 | 0 | 28 | 0.14 | 283.5 | 0
-139 | 0 | 28 | 0.15 | 304.6 | 0
-137.8 | 0 | 28 | 0.19 | 384 | 0
-136.6 | 0 | 28 | 0.15 | 311.2 | 0
-135.4 | 0 | 28 | 0.17 | 349.2 | 0
-134.2 | 0 | 28 | 0.22 | 455.6 | 0
-133 | 0 | 28 | 0.17 | 344.9 | 0
-131.8 | 0 | 28 | 0.19 | 387.2 | 0
-130.6 | 0 | 28 | 0.22 | 445.4 | 0
-129.4 | 0 | 28 | 0.17 | 358.2 | 0
-128.2 | 0 | 28 | 0.24 | 492.7 | 0
-127 | 0 | 28 | 0.24 | 496.1 | 0
-125.8 | 0 | 28 | 0.25 | 508.7 | 0
-124.6 | 0 | 28 | 0.28 | 567.9 | 0
-123.4 | 0 | 28 | 0.28 | 581.4 | 0
-122.2 | 0 | 28 | 0.3 | 606.3 | 0
-121 | 0 | 28 | 0.32 | 658 | 0
-119.8 | 0 | 28 | 0.37 | 766.1 | 0
-118.6 | 0 | 28 | 0.38 | 784.4 | 0
-117.4 | 0 | 28 | 0.41 | 832.5 | 0
-116.2 | 0 | 28 | 0.44 | 904.5 | 0
-115 | 0 | 28 | 0.47 | 958.7 | 0
-113.8 | 0 | 28 | 0.57 | 1168.5 | 0
-112.6 | 0 | 28 | 0.64 | 1303.3 | 0
-111.4 | 0 | 28 | 0.64 | 1306.7 | 0
-110.2 | 0 | 28 | 0.8 | 1647.5 | 0
-109 | 0 | 28 | 0.96 | 1966.5 | 0
-107.8 | 0 | 28 | 0.99 | 2021.6 | 0
-106.6 | 0 | 28 | 1.2 | 2466.5 | 0
-105.4 | 0 | 28 | 1.46 | 2990.5 | 0
-104.1 | 0 | 28 | 1.55 | 3171.3 | 0
-102.9 | 0 | 28 | 1.87 | 3827.1 | 0
-101.8 | 0 | 28 | 2.23 | 4571.6 | 0
-100.5 | 0 | 28 | 2.41 | 4933.4 | 0
-99.3 | 0 | 28 | 2.95 | 6034.5 | 0
-98.1 | 0 | 28 | 3.55 | 7267.9 | 0
-96.9 | 0 | 28 | 4 | 8179.8 | 0
-95.7 | 0 | 28 | 4.76 | 9745.1 | 0
-94.5 | 0 | 28 | 5.7 | 11679.5 | 0
-93.3 | 0 | 28 | 6.42 | 13137.4 | 0
-92.1 | 0 | 28 | 7.71 | 15784.2 | 0
-90.9 | 0 | 28 | 8.8 | 18015.2 | 0
-89.7 | 0 | 28 | 10.64 | 21783.7 | 0
-88.5 | 0 | 28 | 12.82 | 26240.9 | 0
-87.3 | 0 | 28 | 15.54 | 31807.5 | 0
-86.1 | 0 | 28 | 17.28 | 35372.6 | 0
-84.9 | 0 | 28 | 21.28 | 43569.4 | 0
-83.7 | 0 | 28 | 26.35 | 53943.2 | 0
-82.5 | 0 | 28 | 30.29 | 62006.2 | 0
-81.3 | 0 | 28 | 35.86 | 73412.3 | 0
-80.1 | 0 | 28 | 42.34 | 86691.2 | 0
-78.9 | 0 | 28 | 47.35 | 96928.8 | 0
-77.7 | 0 | 28 | 54.55 | 111680.8 | 0
-76.5 | 0 | 28 | 57.87 | 118468.1 | 0
-75.3 | 0 | 28 | 64.9 | 132868.1 | 0
-74.1 | 0 | 28 | 70.96 | 145273.7 | 0
-72.9 | 0 | 28 | 76.44 | 156500.7 | 0
-71.7 | 0 | 28 | 77.82 | 159309.9 | 0
-70.5 | 0 | 28 | 82.09 | 168060.9 | 0
-69.3 | 0 | 28 | 85.73 | 175511.6 | 0
-68.1 | 0 | 28 | 85.17 | 174364.2 | 0
-66.9 | 0 | 28 | 86.62 | 177324.3 | 0
-65.7 | 0 | 28 | 88.55 | 181292.1 | 0
-64.5 | 0 | 28 | 89.32 | 182852.9 | 0
-63.3 | 0 | 28 | 90.23 | 184721.4 | 0
-62.1 | 0 | 28 | 92.25 | 188865.5 | 0
-60.9 | 0 | 28 | 91.85 | 188036.5 | 0
-59.7 | 0 | 28 | 93.71 | 191847.1 | 0
-58.5 | 0 | 28 | 94.64 | 193762.1 | 0
-57.3 | 0 | 28 | 93.47 | 191354 | 0
-56.1 | 0 | 28 | 95.53 | 195581.1 | 0
-54.9 | 0 | 28 | 95.94 | 196412.3 | 0
-53.6 | 0 | 28 | 95.05 | 194583.9 | 0
-52.4 | 0 | 28 | 96.35 | 197252.7 | 0
-51.2 | 0 | 28 | 96.8 | 198180.1 | 0
-50.1 | 0 | 28 | 95.16 | 194826.7 | 0
-48.8 | 0 | 28 | 97.28 | 199166.7 | 0
-47.6 | 0 | 28 | 97.42 | 199446.9 | 0
-46.4 | 0 | 28 | 96.93 | 198445.6 | 0
-45.2 | 0 | 28 | 97.88 | 200390.1 | 0
-44 | 0 | 28 | 97.26 | 199110.1 | 0
-42.8 | 0 | 28 | 98.14 | 200909.4 | 0
-41.6 | 0 | 28 | 97.8 | 200224.1 | 0
-40.4 | 0 | 28 | 92.88 | 190141.1 | 0
-39.2 | 0 | 28 | 97.77 | 200163.5 | 0
-38 | 0 | 28 | 96.6 | 197772.2 | 0
-36.8 | 0 | 28 | 98.53 | 201715.3 | 0
-35.6 | 0 | 28 | 95.72 | 195965.9 | 0
-34.4 | 0 | 28 | 96.97 | 198513.8 | 0
-33.2 | 0 | 28 | 97.02 | 198624.8 | 0
-32 | 0 | 28 | 96.8 | 198185.5 | 0
-30.8 | 0 | 28 | 96.55 | 197658.6 | 0
-29.6 | 0 | 28 | 96.46 | 197489.4 | 0
-28.4 | 0 | 28 | 97.37 | 199343.1 | 0
-27.2 | 0 | 28 | 97.85 | 200328.5 | 0
-26 | 0 | 28 | 98.06 | 200746.8 | 0
-24.8 | 0 | 28 | 98.13 | 200891.1 | 0
-23.6 | 0 | 28 | 97.31 | 199219.5 | 0
-22.4 | 0 | 28 | 100 | 204726.6 | 0
-21.2 | 0 | 28 | 97.09 | 198771.7 | 0
-20 | 0 | 28 | 99.87 | 204461.3 | 0
-18.8 | 0 | 28 | 94.29 | 193035 | 0
-17.6 | 0 | 28 | 98.08 | 200794.3 | 0
-16.4 | 0 | 28 | 98.71 | 202089.7 | 0
-15.2 | 0 | 28 | 94.29 | 193041.9 | 0
-14 | 0 | 28 | 97.05 | 198693.3 | 0
-12.8 | 0 | 28 | 99.16 | 203009.5 | 0
-11.6 | 0 | 28 | 98.04 | 200723.3 | 0
-10.4 | 0 | 28 | 98.13 | 200900.7 | 0
-9.2 | 0 | 28 | 97.6 | 199803.5 | 0
-8 | 0 | 28 | 97.47 | 199537.3 | 0
-6.7 | 0 | 28 | 98.05 | 200726.5 | 0
-5.6 | 0 | 28 | 96.49 | 197542.9 | 0
-4.4 | 0 | 28 | 97.55 | 199711.2 | 0
-3.1 | 0 | 28 | 97.1 | 198782.6 | 0
-1.9 | 0 | 28 | 97.06 | 198714.9 | 0
-0.7 | 0 | 28 | 97.18 | 198962.1 | 0
0.5 | 0 | 28 | 97.26 | 199125.5 | 0
1.7 | 0 | 28 | 98.17 | 200988.4 | 0
2.9 | 0 | 28 | 97.84 | 200297.3 | 0
4.1 | 0 | 28 | 98.01 | 200661.8 | 0
5.3 | 0 | 28 | 98.09 | 200807.2 | 0
6.5 | 0 | 28 | 97.62 | 199849.5 | 0
7.7 | 0 | 28 | 97.75 | 200124.9 | 0
8.9 | 0 | 28 | 98.22 | 201085.9 | 0
10.1 | 0 | 28 | 93.84 | 192118.2 | 0
11.3 | 0 | 28 | 98.36 | 201369.8 | 0
12.5 | 0 | 28 | 98.65 | 201963.9 | 0
13.7 | 0 | 28 | 99.16 | 203015.6 | 0
14.9 | 0 | 28 | 97.78 | 200175.4 | 0
16.1 | 0 | 28 | 93.84 | 192118.7 | 0
17.3 | 0 | 28 | 98.37 | 201392 | 0
18.5 | 0 | 28 | 98.4 | 201457.5 | 0
19.7 | 0 | 28 | 98.11 | 200855.8 | 0
20.9 | 0 | 28 | 98.52 | 201705 | 0
22.1 | 0 | 28 | 98.25 | 201134.1 | 0
23.3 | 0 | 28 | 98.41 | 201466.5 | 0
24.5 | 0 | 28 | 98.46 | 201581.5 | 0
25.7 | 0 | 28 | 97.56 | 199729 | 0
26.9 | 0 | 28 | 98.11 | 200856.1 | 0
28.1 | 0 | 28 | 97.79 | 200201.7 | 0
29.3 | 0 | 28 | 97.54 | 199699.3 | 0
30.5 | 0 | 28 | 97.28 | 199148.6 | 0
31.7 | 0 | 28 | 97.41 | 199429.8 | 0
32.9 | 0 | 28 | 97.04 | 198674.4 | 0
34.1 | 0 | 28 | 96.33 | 197213.1 | 0
35.3 | 0 | 28 | 96.27 | 197081.6 | 0
36.5 | 0 | 28 | 99.48 | 203666.1 | 0
37.7 | 0 | 28 | 95.11 | 194710.7 | 0
38.9 | 0 | 28 | 91.3 | 186907 | 0
40.1 | 0 | 28 | 99.98 | 204691.7 | 0
41.3 | 0 | 28 | 95.08 | 194645.5 | 0
42.5 | 0 | 28 | 91.4 | 187115.4 | 0
43.7 | 0 | 28 | 99.74 | 204188.9 | 0
44.9 | 0 | 28 | 91.33 | 186967.1 | 0
46.1 | 0 | 28 | 96.09 | 196728.4 | 0
47.3 | 0 | 28 | 95.72 | 195955.4 | 0
48.5 | 0 | 28 | 95.47 | 195454.8 | 0
49.7 | 0 | 28 | 94.82 | 194115 | 0
50.9 | 0 | 28 | 94.82 | 194130 | 0
52.1 | 0 | 28 | 95.27 | 195050.4 | 0
53.3 | 0 | 28 | 95.12 | 194731.5 | 0
54.5 | 0 | 28 | 94.41 | 193290 | 0
55.7 | 0 | 28 | 94.66 | 193788.8 | 0
56.9 | 0 | 28 | 93.78 | 192001.4 | 0
58.1 | 0 | 28 | 93.21 | 190819.5 | 0
59.3 | 0 | 28 | 92.87 | 190136.3 | 0
60.5 | 0 | 28 | 92.64 | 189663.4 | 0
61.7 | 0 | 28 | 91.72 | 187765.8 | 0
62.9 | 0 | 28 | 89.7 | 183631.7 | 0
64.1 | 0 | 28 | 90.06 | 184383.7 | 0
65.4 | 0 | 28 | 87.38 | 178882.4 | 0
66.6 | 0 | 28 | 86.88 | 177868.4 | 0
67.8 | 0 | 28 | 83.92 | 171805.9 | 0
69 | 0 | 28 | 81.09 | 166015.3 | 0
70.2 | 0 | 28 | 78.68 | 161080.1 | 0
71.4 | 0 | 28 | 74.25 | 152000.4 | 0
72.6 | 0 | 28 | 70.32 | 143957.6 | 0
73.8 | 0 | 28 | 64.91 | 132881.6 | 0
75 | 0 | 28 | 59.15 | 121086.5 | 0
76.2 | 0 | 28 | 52.91 | 108324.7 | 0
77.4 | 0 | 28 | 46.71 | 95619.8 | 0
78.6 | 0 | 28 | 40.55 | 83020.4 | 0
79.8 | 0 | 28 | 34.58 | 70797.6 | 0
81 | 0 | 28 | 29.2 | 59771.5 | 0
82.2 | 0 | 28 | 24.66 | 50484 | 0
83.4 | 0 | 28 | 20.33 | 41617.2 | 0
84.6 | 0 | 28 | 17.09 | 34994.3 | 0
85.8 | 0 | 28 | 14.16 | 28985 | 0
87 | 0 | 28 | 11.82 | 24188.5 | 0
88.2 | 0 | 28 | 9.83 | 20123.8 | 0
89.4 | 0 | 28 | 8.24 | 16861.7 | 0
90.6 | 0 | 28 | 7.07 | 14467.3 | 0
91.8 | 0 | 28 | 5.88 | 12044.6 | 0
93 | 0 | 28 | 5.09 | 10426.5 | 0
94.2 | 0 | 28 | 4.26 | 8720.7 | 0
95.4 | 0 | 28 | 3.71 | 7590.6 | 0
96.6 | 0 | 28 | 3.23 | 6616.5 | 0
97.8 | 0 | 28 | 2.77 | 5680.9 | 0
99 | 0 | 28 | 2.42 | 4955.2 | 0
100.2 | 0 | 28 | 2.16 | 4419.4 | 0
101.4 | 0 | 28 | 1.74 | 3571.4 | 0
102.6 | 0 | 28 | 1.6 | 3276.7 | 0
103.8 | 0 | 28 | 1.41 | 2878.5 | 0
105 | 0 | 28 | 1.23 | 2516.3 | 0
106.2 | 0 | 28 | 1.08 | 2202 | 0
107.4 | 0 | 28 | 0.91 | 1865.8 | 0
108.6 | 0 | 28 | 0.82 | 1676.8 | 0
109.9 | 0 | 28 | 0.63 | 1292.9 | 0
111.1 | 0 | 28 | 0.55 | 1132.1 | 0
112.3 | 0 | 28 | 0.51 | 1044.4 | 0
113.5 | 0 | 28 | 0.4 | 812 | 0
114.7 | 0 | 28 | 0.32 | 652.3 | 0
115.9 | 0 | 28 | 0.33 | 685.4 | 0
117.1 | 0 | 28 | 0.27 | 559.2 | 0
118.3 | 0 | 28 | 0.26 | 524.8 | 0
119.5 | 0 | 28 | 0.29 | 586.4 | 0
120.7 | 0 | 28 | 0.18 | 377 | 0
121.9 | 0 | 28 | 0.17 | 349.2 | 0
123.1 | 0 | 28 | 0.15 | 303.7 | 0
124.3 | 0 | 28 | 0.12 | 239.8 | 0
125.5 | 0 | 28 | 0.14 | 278.1 | 0
126.7 | 0 | 28 | 0.1 | 206 | 0
127.9 | 0 | 28 | 0.07 | 145.9 | 0
129.1 | 0 | 28 | 0.1 | 201.2 | 0
130.3 | 0 | 28 | 0.08 | 156.1 | 0
131.5 | 0 | 28 | 0.1 | 209.8 | 0
132.7 | 0 | 28 | 0.06 | 123.5 | 0
133.9 | 0 | 28 | 0.1 | 208.5 | 0
135.1 | 0 | 28 | 0.06 | 126.5 | 0
136.3 | 0 | 28 | 0.07 | 147.2 | 0
137.5 | 0 | 28 | 0.08 | 170.5 | 0
138.7 | 0 | 28 | 0.07 | 152.2 | 0
139.9 | 0 | 28 | 0 | 8.2 | 0
141.1 | 0 | 28 | 0.06 | 121.6 | 0
142.3 | 0 | 28 | 0.05 | 96.8 | 0
143.5 | 0 | 28 | 0.02 | 44.6 | 0
144.7 | 0 | 28 | 0.03 | 53.5 | 0
145.9 | 0 | 28 | 0.02 | 34.3 | 0
147.1 | 0 | 28 | 0.03 | 58.1 | 0
148.3 | 0 | 28 | 0.01 | 11.4 | 0
149.5 | 0 | 28 | 0.06 | 126.8 | 0
150.7 | 0 | 28 | 0.04 | 76.6 | 0
151.9 | 0 | 28 | 0.04 | 89.4 | 0
153.1 | 0 | 28 | 0.03 | 53.8 | 0
154.3 | 0 | 28 | 0.01 | 23.9 | 0
155.5 | 0 | 28 | 0.01 | 13.6 | 0
156.7 | 0 | 28 | 0.03 | 64 | 0
158 | 0 | 28 | 0 | 4.3 | 0
159.1 | 0 | 28 | 0.02 | 38.2 | 0
160.3 | 0 | 28 | 0.04 | 73.3 | 0
161.6 | 0 | 28 | 0.04 | 87.5 | 0
162.7 | 0 | 28 | 0.03 | 55 | 0
164 | 0 | 28 | 0 | 8.7 | 0
165.1 | 0 | 28 | 0.02 | 34.4 | 0
166.3 | 0 | 28 | 0.01 | 27.1 | 0
167.6 | 0 | 28 | 0.01 | 16.6 | 0
168.7 | 0 | 28 | 0 | 3.6 | 0
169.9 | 0 | 28 | 0.04 | 74 | 0
171.2 | 0 | 28 | 0.04 | 77.6 | 0
172.4 | 0 | 28 | 0 | 0 | 0
173.6 | 0 | 28 | 0.02 | 38.2 | 0
174.8 | 0 | 28 | 0.04 | 78.1 | 0
176 | 0 | 28 | 0.05 | 98.1 | 0
177.2 | 0 | 28 | 0.06 | 128.5 | 0
178.4 | 0 | 28 | 0.08 | 173 | 0
179.6 | 0 | 28 | 0.08 | 160.1 | 0
180.8 | 0 | 28 | 0.04 | 79.7 | 0
182 | 0 | 28 | 0.07 | 137.2 | 0
183.2 | 0 | 28 | 0.03 | 62.6 | 0
184.4 | 0 | 28 | 0.05 | 102.1 | 0
185.6 | 0 | 28 | 0.03 | 68.3 | 0
186.8 | 0 | 28 | 0.06 | 115.3 | 0
188 | 0 | 28 | 0.08 | 167 | 0
189.2 | 0 | 28 | 0.06 | 118.2 | 0
190.4 | 0 | 28 | 0.07 | 150.3 | 0
191.6 | 0 | 28 | 0.02 | 33.7 | 0
192.8 | 0 | 28 | 0.04 | 88.7 | 0
194 | 0 | 28 | 0.06 | 128.3 | 0
195.2 | 0 | 28 | 0.04 | 71.8 | 0
196.4 | 0 | 28 | 0.05 | 102.3 | 0
197.6 | 0 | 28 | 0.09 | 191.4 | 0
198.8 | 0 | 28 | 0.01 | 27.9 | 0
200 | 0 | 28 | 0.04 | 90.2 | 0