## rendements énergie roos
Measurement time: | 07/09/2022 18:47:58
Radiation device: | Clinac3 Accelerator
Energy: | 6 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 200 | 2.73 | 1356.2 | 0.16
0 | 0 | 199.3 | 2.24 | 1113.3 | 0.19
0 | 0 | 198 | 2.08 | 1033.1 | 0.18
0 | 0 | 196.8 | 2.25 | 1116.7 | 0.18
0 | 0 | 195.6 | 2.15 | 1067.2 | 0.18
0 | 0 | 194.4 | 2.18 | 1084.5 | 0.19
0 | 0 | 193.1 | 2.23 | 1109.8 | 0.18
0 | 0 | 192 | 2.22 | 1103.9 | 0.19
0 | 0 | 190.8 | 2.15 | 1070.4 | 0.18
0 | 0 | 189.6 | 2.35 | 1169.7 | 0.19
0 | 0 | 188.4 | 2.45 | 1217.6 | 0.21
0 | 0 | 187.1 | 2.29 | 1138.4 | 0.2
0 | 0 | 185.9 | 2.64 | 1313.9 | 0.21
0 | 0 | 184.7 | 2.4 | 1190.7 | 0.2
0 | 0 | 183.5 | 2.45 | 1217.9 | 0.21
0 | 0 | 182.3 | 2.38 | 1183.5 | 0.2
0 | 0 | 181.1 | 2.49 | 1235.4 | 0.21
0 | 0 | 179.9 | 2.56 | 1270.7 | 0.21
0 | 0 | 178.7 | 2.62 | 1303.6 | 0.22
0 | 0 | 177.5 | 2.52 | 1252.3 | 0.21
0 | 0 | 176.3 | 2.58 | 1280.8 | 0.21
0 | 0 | 175.1 | 2.48 | 1229.7 | 0.21
0 | 0 | 173.9 | 2.5 | 1243.1 | 0.21
0 | 0 | 172.7 | 2.68 | 1330.4 | 0.22
0 | 0 | 171.5 | 2.74 | 1361 | 0.23
0 | 0 | 170.3 | 2.53 | 1258.2 | 0.22
0 | 0 | 169.1 | 2.7 | 1340.4 | 0.22
0 | 0 | 167.9 | 2.64 | 1311.6 | 0.22
0 | 0 | 166.7 | 2.73 | 1356 | 0.23
0 | 0 | 165.5 | 2.88 | 1428.5 | 0.24
0 | 0 | 164.3 | 2.82 | 1399 | 0.24
0 | 0 | 163.1 | 2.74 | 1360 | 0.23
0 | 0 | 161.9 | 2.8 | 1391.2 | 0.24
0 | 0 | 160.7 | 2.99 | 1486.8 | 0.25
0 | 0 | 159.5 | 3.04 | 1508.8 | 0.25
0 | 0 | 158.3 | 2.77 | 1378.4 | 0.24
0 | 0 | 157.1 | 2.86 | 1422.5 | 0.25
0 | 0 | 155.9 | 2.94 | 1461.2 | 0.24
0 | 0 | 154.7 | 2.95 | 1464.3 | 0.26
0 | 0 | 153.5 | 2.87 | 1427.5 | 0.23
0 | 0 | 152.3 | 2.99 | 1486.9 | 0.26
0 | 0 | 151.1 | 3.02 | 1502.4 | 0.25
0 | 0 | 149.9 | 3.1 | 1539.4 | 0.25
0 | 0 | 148.7 | 3.04 | 1509.8 | 0.27
0 | 0 | 147.5 | 2.93 | 1454.9 | 0.24
0 | 0 | 146.2 | 3.08 | 1529.2 | 0.27
0 | 0 | 145.1 | 3.2 | 1592 | 0.27
0 | 0 | 143.9 | 3.1 | 1541.4 | 0.25
0 | 0 | 142.7 | 3.03 | 1505.9 | 0.26
0 | 0 | 141.5 | 3.21 | 1592.5 | 0.27
0 | 0 | 140.3 | 3.23 | 1603.8 | 0.28
0 | 0 | 139 | 3.24 | 1607.8 | 0.27
0 | 0 | 137.8 | 2.93 | 1456.6 | 0.25
0 | 0 | 136.6 | 3.17 | 1577 | 0.26
0 | 0 | 135.4 | 3.12 | 1552.4 | 0.27
0 | 0 | 134.2 | 3.38 | 1677 | 0.28
0 | 0 | 133 | 3.02 | 1501.5 | 0.26
0 | 0 | 131.8 | 3.25 | 1617.1 | 0.28
0 | 0 | 130.6 | 3.36 | 1668.5 | 0.29
0 | 0 | 129.4 | 3.47 | 1722 | 0.29
0 | 0 | 128.2 | 3.38 | 1681.5 | 0.29
0 | 0 | 127 | 3.71 | 1845.2 | 0.31
0 | 0 | 125.8 | 3.46 | 1719 | 0.3
0 | 0 | 124.6 | 3.54 | 1756.7 | 0.3
0 | 0 | 123.4 | 3.56 | 1768.3 | 0.29
0 | 0 | 122.2 | 3.6 | 1788.6 | 0.32
0 | 0 | 121 | 3.57 | 1771.5 | 0.3
0 | 0 | 119.8 | 3.58 | 1780.4 | 0.31
0 | 0 | 118.6 | 3.78 | 1876.4 | 0.31
0 | 0 | 117.4 | 3.66 | 1817.6 | 0.32
0 | 0 | 116.2 | 3.79 | 1883.2 | 0.32
0 | 0 | 115 | 3.76 | 1868.9 | 0.33
0 | 0 | 113.8 | 3.69 | 1831.4 | 0.32
0 | 0 | 112.6 | 3.82 | 1898.1 | 0.32
0 | 0 | 111.4 | 4.09 | 2032.1 | 0.34
0 | 0 | 110.2 | 3.78 | 1878.4 | 0.33
0 | 0 | 109 | 4.17 | 2072.2 | 0.35
0 | 0 | 107.8 | 3.93 | 1951.1 | 0.34
0 | 0 | 106.6 | 4.23 | 2101.1 | 0.35
0 | 0 | 105.4 | 4.12 | 2046.6 | 0.35
0 | 0 | 104.2 | 4.04 | 2008.7 | 0.35
0 | 0 | 103 | 4.12 | 2046.1 | 0.36
0 | 0 | 101.8 | 4.33 | 2149.9 | 0.36
0 | 0 | 100.6 | 4.19 | 2083.6 | 0.37
0 | 0 | 99.4 | 4.11 | 2042.4 | 0.36
0 | 0 | 98.2 | 4.57 | 2268.5 | 0.38
0 | 0 | 97 | 4.34 | 2155.4 | 0.38
0 | 0 | 95.8 | 4.42 | 2196.9 | 0.39
0 | 0 | 94.6 | 4.34 | 2157.5 | 0.37
0 | 0 | 93.4 | 4.54 | 2256.8 | 0.39
0 | 0 | 92.1 | 4.74 | 2352.6 | 0.39
0 | 0 | 91 | 4.83 | 2399.8 | 0.41
0 | 0 | 89.7 | 4.82 | 2395.8 | 0.42
0 | 0 | 88.6 | 4.55 | 2262.7 | 0.4
0 | 0 | 87.3 | 4.93 | 2448.5 | 0.42
0 | 0 | 86.1 | 4.7 | 2336 | 0.41
0 | 0 | 84.9 | 4.79 | 2381.3 | 0.42
0 | 0 | 83.7 | 4.9 | 2432.9 | 0.41
0 | 0 | 82.5 | 4.77 | 2369.2 | 0.42
0 | 0 | 81.3 | 4.9 | 2436.5 | 0.43
0 | 0 | 80.1 | 5.29 | 2627.5 | 0.44
0 | 0 | 78.9 | 4.98 | 2476.2 | 0.44
0 | 0 | 77.7 | 5.02 | 2493.6 | 0.44
0 | 0 | 76.5 | 5.42 | 2691.3 | 0.45
0 | 0 | 75.3 | 5.13 | 2546.2 | 0.45
0 | 0 | 74.1 | 5.26 | 2614.7 | 0.46
0 | 0 | 72.9 | 5.55 | 2756.1 | 0.47
0 | 0 | 71.7 | 5.35 | 2658.6 | 0.47
0 | 0 | 70.5 | 5.3 | 2634 | 0.47
0 | 0 | 69.2 | 5.48 | 2720.9 | 0.46
0 | 0 | 68 | 5.51 | 2739.5 | 0.49
0 | 0 | 66.8 | 5.46 | 2711.7 | 0.48
0 | 0 | 65.6 | 5.87 | 2915.9 | 0.49
0 | 0 | 64.4 | 5.56 | 2762 | 0.48
0 | 0 | 63.2 | 5.39 | 2678.4 | 0.47
0 | 0 | 62 | 5.92 | 2941 | 0.5
0 | 0 | 60.8 | 5.67 | 2819.2 | 0.5
0 | 0 | 59.6 | 6 | 2982.2 | 0.53
0 | 0 | 58.4 | 6.14 | 3049.7 | 0.52
0 | 0 | 57.2 | 5.86 | 2911.6 | 0.51
0 | 0 | 56 | 6.07 | 3015.3 | 0.53
0 | 0 | 54.8 | 6.23 | 3094.8 | 0.52
0 | 0 | 53.6 | 6.16 | 3058.6 | 0.54
0 | 0 | 52.4 | 6.17 | 3064.4 | 0.54
0 | 0 | 51.2 | 6.52 | 3239.2 | 0.55
0 | 0 | 50 | 6.26 | 3108.9 | 0.55
0 | 0 | 48.8 | 6.23 | 3094.7 | 0.55
0 | 0 | 47.6 | 6.8 | 3377.4 | 0.57
0 | 0 | 46.4 | 6.61 | 3286.3 | 0.58
0 | 0 | 45.2 | 6.43 | 3193.4 | 0.57
0 | 0 | 44 | 6.76 | 3359.2 | 0.57
0 | 0 | 42.7 | 6.64 | 3296.4 | 0.58
0 | 0 | 41.5 | 6.86 | 3407.9 | 0.6
0 | 0 | 40.3 | 7.1 | 3526.7 | 0.6
0 | 0 | 39.1 | 7.09 | 3520.7 | 0.61
0 | 0 | 37.9 | 7.08 | 3515.8 | 0.62
0 | 0 | 36.7 | 7.02 | 3486.7 | 0.61
0 | 0 | 35.5 | 7.62 | 3787.6 | 0.65
0 | 0 | 34.3 | 8.65 | 4295 | 0.76
0 | 0 | 33.1 | 12.53 | 6225.6 | 1.09
0 | 0 | 31.9 | 23.82 | 11835.5 | 2.04
0 | 0 | 30.7 | 47.34 | 23520.8 | 4.18
0 | 0 | 29.5 | 86.88 | 43163.4 | 7.59
0 | 0 | 28.3 | 153.58 | 76300.7 | 13.15
0 | 0 | 27.1 | 234.32 | 116413.2 | 20.58
0 | 0 | 25.9 | 336.46 | 167157.4 | 29.27
0 | 0 | 24.7 | 472.55 | 234770.8 | 40.2
0 | 0 | 23.5 | 565.8 | 281094.5 | 50.07
0 | 0 | 22.3 | 701.07 | 348298.3 | 60.21
0 | 0 | 21.1 | 831.94 | 413319.2 | 70.38
0 | 0 | 19.9 | 910.74 | 452467.1 | 78.6
0 | 0 | 18.6 | 1006.22 | 499902.9 | 86.19
0 | 0 | 17.4 | 1105.78 | 549366.8 | 91.76
0 | 0 | 16.2 | 1106.87 | 549907.9 | 96.21
0 | 0 | 15 | 1184.5 | 588476.7 | 98.92
0 | 0 | 13.8 | 1209.92 | 601104.5 | 100
0 | 0 | 12.6 | 1180.48 | 586475.5 | 100
0 | 0 | 11.4 | 1186.86 | 589646.2 | 98.86
0 | 0 | 10.2 | 1191.52 | 591963.2 | 97.26
0 | 0 | 9 | 1134.86 | 563813.3 | 94.97
0 | 0 | 7.8 | 1129.48 | 561140.4 | 92.48
0 | 0 | 6.6 | 1113.34 | 553121.3 | 89.98
0 | 0 | 5.4 | 1060.06 | 526648.8 | 87.37
0 | 0 | 4.2 | 1039.68 | 516523.7 | 84.95
0 | 0 | 3 | 1039.33 | 516352 | 82.57
0 | 0 | 1.8 | 982.36 | 488049.2 | 80.91
0 | 0 | 0.6 | 983.85 | 488788.2 | 80.2
0 | 0 | -0.5 | 1017.72 | 505615.6 | 80.36
Measurement time: | 07/09/2022 18:49:40
Radiation device: | Clinac3 Accelerator
Energy: | 9 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 6.52 | 3240.8 | 0.6
0 | 0 | 149.3 | 6.63 | 3295.8 | 0.55
0 | 0 | 148 | 6.64 | 3297.2 | 0.56
0 | 0 | 146.7 | 6.88 | 3418.1 | 0.56
0 | 0 | 145.6 | 6.57 | 3264.7 | 0.56
0 | 0 | 144.4 | 6.79 | 3375.3 | 0.58
0 | 0 | 143.2 | 7.07 | 3513.2 | 0.57
0 | 0 | 141.9 | 6.76 | 3356.5 | 0.57
0 | 0 | 140.7 | 6.57 | 3266 | 0.56
0 | 0 | 139.5 | 7.24 | 3594.6 | 0.58
0 | 0 | 138.3 | 6.91 | 3431.8 | 0.59
0 | 0 | 137.2 | 7.26 | 3608.4 | 0.61
0 | 0 | 136 | 7.39 | 3671.8 | 0.6
0 | 0 | 134.7 | 6.87 | 3413.9 | 0.6
0 | 0 | 133.5 | 7.26 | 3604.7 | 0.6
0 | 0 | 132.3 | 7.68 | 3817 | 0.62
0 | 0 | 131.1 | 7.36 | 3654.2 | 0.62
0 | 0 | 129.9 | 7.3 | 3628 | 0.62
0 | 0 | 128.8 | 7.7 | 3825.7 | 0.62
0 | 0 | 127.5 | 7.24 | 3598.1 | 0.63
0 | 0 | 126.3 | 7.81 | 3878 | 0.64
0 | 0 | 125.1 | 7.74 | 3847.3 | 0.63
0 | 0 | 123.9 | 7.42 | 3684.6 | 0.65
0 | 0 | 122.7 | 7.9 | 3924.9 | 0.64
0 | 0 | 121.5 | 8.1 | 4022.6 | 0.65
0 | 0 | 120.3 | 7.6 | 3774.2 | 0.67
0 | 0 | 119.1 | 8.42 | 4182.7 | 0.68
0 | 0 | 117.9 | 8.46 | 4203.8 | 0.69
0 | 0 | 116.7 | 7.93 | 3937.3 | 0.67
0 | 0 | 115.5 | 8.03 | 3987.3 | 0.68
0 | 0 | 114.3 | 8.55 | 4248.5 | 0.7
0 | 0 | 113.1 | 8.02 | 3985.7 | 0.7
0 | 0 | 111.9 | 8.76 | 4353 | 0.71
0 | 0 | 110.7 | 8.81 | 4376.5 | 0.71
0 | 0 | 109.5 | 8.13 | 4041.5 | 0.72
0 | 0 | 108.3 | 9.05 | 4494.1 | 0.73
0 | 0 | 107.1 | 8.97 | 4456.7 | 0.73
0 | 0 | 105.9 | 8.43 | 4188 | 0.75
0 | 0 | 104.7 | 9.21 | 4577.1 | 0.75
0 | 0 | 103.5 | 9.42 | 4678.3 | 0.76
0 | 0 | 102.3 | 8.8 | 4369.8 | 0.78
0 | 0 | 101.1 | 9.39 | 4667.4 | 0.77
0 | 0 | 99.9 | 9.23 | 4588.1 | 0.77
0 | 0 | 98.7 | 9.1 | 4519.4 | 0.79
0 | 0 | 97.5 | 9.47 | 4704.5 | 0.78
0 | 0 | 96.3 | 9.27 | 4605.8 | 0.79
0 | 0 | 95.1 | 9.35 | 4644.8 | 0.8
0 | 0 | 93.8 | 9.72 | 4829 | 0.81
0 | 0 | 92.6 | 9.46 | 4698.9 | 0.82
0 | 0 | 91.4 | 10 | 4968.9 | 0.82
0 | 0 | 90.2 | 9.68 | 4810.7 | 0.83
0 | 0 | 89 | 10.05 | 4995.3 | 0.83
0 | 0 | 87.8 | 9.78 | 4858.5 | 0.84
0 | 0 | 86.6 | 9.8 | 4870.8 | 0.84
0 | 0 | 85.4 | 10.21 | 5073.2 | 0.86
0 | 0 | 84.2 | 10.22 | 5076.4 | 0.87
0 | 0 | 83 | 10.45 | 5192.1 | 0.88
0 | 0 | 81.8 | 10.92 | 5423.8 | 0.89
0 | 0 | 80.6 | 10.26 | 5097.5 | 0.88
0 | 0 | 79.4 | 10.37 | 5151.3 | 0.89
0 | 0 | 78.2 | 10.7 | 5314.7 | 0.91
0 | 0 | 77 | 10.73 | 5330.8 | 0.9
0 | 0 | 75.8 | 10.91 | 5421.2 | 0.93
0 | 0 | 74.6 | 11.41 | 5666.4 | 0.93
0 | 0 | 73.4 | 11 | 5462.5 | 0.95
0 | 0 | 72.1 | 10.98 | 5453.5 | 0.94
0 | 0 | 70.9 | 11.69 | 5806.2 | 0.96
0 | 0 | 69.7 | 11.28 | 5601.8 | 0.97
0 | 0 | 68.5 | 11.33 | 5629.7 | 0.98
0 | 0 | 67.3 | 12.2 | 6061.9 | 1
0 | 0 | 66.1 | 11.48 | 5705.5 | 0.99
0 | 0 | 64.9 | 11.78 | 5852.5 | 1.02
0 | 0 | 63.7 | 12.5 | 6211.1 | 1.03
0 | 0 | 62.5 | 11.87 | 5896.8 | 1.02
0 | 0 | 61.3 | 11.73 | 5828.8 | 1.02
0 | 0 | 60.1 | 12.87 | 6395.1 | 1.05
0 | 0 | 58.9 | 12.5 | 6209.9 | 1.07
0 | 0 | 57.7 | 12.43 | 6175.3 | 1.07
0 | 0 | 56.5 | 13.02 | 6468.8 | 1.07
0 | 0 | 55.3 | 12.48 | 6198.7 | 1.09
0 | 0 | 54.1 | 12.99 | 6451.5 | 1.11
0 | 0 | 52.8 | 13.46 | 6686.4 | 1.11
0 | 0 | 51.6 | 13.32 | 6619.2 | 1.15
0 | 0 | 50.5 | 14.41 | 7161 | 1.25
0 | 0 | 49.2 | 18.17 | 9025.6 | 1.5
0 | 0 | 48 | 23.24 | 11543.5 | 2.02
0 | 0 | 46.8 | 34.01 | 16897.3 | 2.93
0 | 0 | 45.6 | 54.84 | 27243.3 | 4.52
0 | 0 | 44.4 | 83.22 | 41343.7 | 7.15
0 | 0 | 43.2 | 120.28 | 59754.9 | 10.42
0 | 0 | 42 | 185.51 | 92161.6 | 15.2
0 | 0 | 40.8 | 242.97 | 120708.9 | 21.07
0 | 0 | 39.6 | 313.66 | 155830.1 | 26.84
0 | 0 | 38.4 | 398.95 | 198205.6 | 34.34
0 | 0 | 37.2 | 509.41 | 253078.8 | 41.74
0 | 0 | 36 | 570.61 | 283485 | 48.94
0 | 0 | 34.8 | 693.67 | 344622.7 | 56.8
0 | 0 | 33.6 | 753.54 | 374366.4 | 63.8
0 | 0 | 32.4 | 832.78 | 413733.6 | 70.55
0 | 0 | 31.2 | 946.38 | 470175.2 | 76.79
0 | 0 | 30 | 976.49 | 485133 | 81.86
0 | 0 | 28.7 | 1028.39 | 510917 | 86.87
0 | 0 | 27.6 | 1120.07 | 556464.6 | 90.59
0 | 0 | 26.3 | 1138.83 | 565785.3 | 93.72
0 | 0 | 25.1 | 1157.04 | 574833.6 | 96.38
0 | 0 | 23.9 | 1219.36 | 605793.4 | 98
0 | 0 | 22.7 | 1223.93 | 608065.6 | 99.42
0 | 0 | 21.5 | 1211.47 | 601871.3 | 100
0 | 0 | 20.3 | 1268.76 | 630337.5 | 99.99
0 | 0 | 19.1 | 1225.93 | 609055.3 | 99.73
0 | 0 | 17.9 | 1215.68 | 603962.7 | 99.16
0 | 0 | 16.7 | 1237.42 | 614768 | 98.22
0 | 0 | 15.5 | 1216.45 | 604348.3 | 97.23
0 | 0 | 14.3 | 1183.8 | 588124.2 | 96.16
0 | 0 | 13.1 | 1192.87 | 592631.5 | 94.93
0 | 0 | 11.9 | 1191.6 | 591999.9 | 93.65
0 | 0 | 10.7 | 1148.1 | 570389.4 | 92.57
0 | 0 | 9.5 | 1162.47 | 577528.1 | 91.29
0 | 0 | 8.3 | 1154.64 | 573638.3 | 90.1
0 | 0 | 7.1 | 1115.22 | 554054.2 | 89.18
0 | 0 | 5.9 | 1124.65 | 558742.3 | 88.02
0 | 0 | 4.6 | 1123.96 | 558396.4 | 86.92
0 | 0 | 3.5 | 1076.5 | 534818.8 | 85.63
0 | 0 | 2.3 | 1086.4 | 539739.4 | 84.57
0 | 0 | 1 | 1088.08 | 540569.5 | 83.79
0 | 0 | -0.5 | 1065.51 | 529358.9 | 83.94
Measurement time: | 07/09/2022 18:50:46
Radiation device: | Clinac3 Accelerator
Energy: | 15 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 11.39 | 5658.1 | 2.07
0 | 0 | 149.3 | 24.76 | 12299.6 | 2.03
0 | 0 | 147.9 | 24.38 | 12110.8 | 2.04
0 | 0 | 146.7 | 25.2 | 12521.5 | 2.06
0 | 0 | 145.6 | 25.61 | 12725.6 | 2.07
0 | 0 | 144.3 | 24.81 | 12324 | 2.09
0 | 0 | 143.1 | 25.72 | 12776.6 | 2.1
0 | 0 | 141.9 | 25.82 | 12826.6 | 2.09
0 | 0 | 140.8 | 25.32 | 12577.5 | 2.13
0 | 0 | 139.5 | 26.26 | 13048 | 2.14
0 | 0 | 138.3 | 25.88 | 12855.8 | 2.18
0 | 0 | 137.1 | 26.68 | 13254.2 | 2.16
0 | 0 | 135.9 | 26.68 | 13256.1 | 2.18
0 | 0 | 134.7 | 26.3 | 13066.5 | 2.21
0 | 0 | 133.5 | 27.23 | 13528.7 | 2.2
0 | 0 | 132.3 | 27.36 | 13594.5 | 2.23
0 | 0 | 131.1 | 26.57 | 13198.1 | 2.24
0 | 0 | 129.9 | 27.71 | 13766.3 | 2.24
0 | 0 | 128.7 | 27.99 | 13906.5 | 2.27
0 | 0 | 127.5 | 27.07 | 13449.5 | 2.28
0 | 0 | 126.3 | 28.47 | 14143.3 | 2.3
0 | 0 | 125.1 | 28.51 | 14165 | 2.32
0 | 0 | 123.9 | 27.93 | 13874.3 | 2.35
0 | 0 | 122.7 | 29.02 | 14417.4 | 2.35
0 | 0 | 121.5 | 29.22 | 14517 | 2.37
0 | 0 | 120.3 | 28.2 | 14008.6 | 2.38
0 | 0 | 119.1 | 29.88 | 14844.6 | 2.41
0 | 0 | 117.9 | 29.72 | 14766.4 | 2.43
0 | 0 | 116.7 | 28.81 | 14311.3 | 2.43
0 | 0 | 115.5 | 30.14 | 14975.4 | 2.46
0 | 0 | 114.3 | 30.68 | 15242.6 | 2.49
0 | 0 | 113.1 | 30.28 | 15042.7 | 2.51
0 | 0 | 111.9 | 30.02 | 14913.1 | 2.52
0 | 0 | 110.7 | 31.28 | 15542.6 | 2.55
0 | 0 | 109.5 | 29.88 | 14846.6 | 2.56
0 | 0 | 108.3 | 31.65 | 15724.4 | 2.58
0 | 0 | 107.1 | 31.5 | 15650.8 | 2.6
0 | 0 | 105.9 | 31.67 | 15735.4 | 2.62
0 | 0 | 104.7 | 31.98 | 15887.2 | 2.65
0 | 0 | 103.5 | 32.13 | 15962.3 | 2.67
0 | 0 | 102.3 | 33.01 | 16400 | 2.69
0 | 0 | 101.1 | 32.45 | 16123 | 2.72
0 | 0 | 99.9 | 32.93 | 16358.7 | 2.73
0 | 0 | 98.7 | 32.73 | 16262.5 | 2.76
0 | 0 | 97.5 | 33.76 | 16773.3 | 2.75
0 | 0 | 96.3 | 34.1 | 16940.8 | 2.8
0 | 0 | 95.1 | 32.97 | 16381.1 | 2.82
0 | 0 | 93.9 | 34.66 | 17217.9 | 2.84
0 | 0 | 92.7 | 35.19 | 17485.1 | 2.88
0 | 0 | 91.4 | 34.17 | 16974 | 2.89
0 | 0 | 90.2 | 34.73 | 17253.6 | 2.91
0 | 0 | 89 | 35.93 | 17852.2 | 2.95
0 | 0 | 87.8 | 35.24 | 17508 | 2.97
0 | 0 | 86.6 | 36.05 | 17910 | 3.03
0 | 0 | 85.4 | 38.17 | 18965.4 | 3.15
0 | 0 | 84.2 | 38.53 | 19143.6 | 3.27
0 | 0 | 83 | 41.82 | 20779.1 | 3.49
0 | 0 | 81.8 | 47.17 | 23434.2 | 3.89
0 | 0 | 80.6 | 51.55 | 25609.9 | 4.4
0 | 0 | 79.4 | 63.06 | 31329.4 | 5.21
0 | 0 | 78.2 | 74.49 | 37006.6 | 6.32
0 | 0 | 77 | 93.21 | 46308.1 | 7.66
0 | 0 | 75.8 | 116.71 | 57983.8 | 9.57
0 | 0 | 74.6 | 137.49 | 68305.9 | 11.63
0 | 0 | 73.4 | 172.51 | 85706.2 | 14.21
0 | 0 | 72.2 | 212.48 | 105562 | 17.25
0 | 0 | 71 | 241.89 | 120172.5 | 20.49
0 | 0 | 69.7 | 296.21 | 147161 | 24.23
0 | 0 | 68.5 | 345.98 | 171888.5 | 28.12
0 | 0 | 67.3 | 380.43 | 189000.5 | 32.18
0 | 0 | 66.1 | 448.68 | 222908.1 | 36.6
0 | 0 | 64.9 | 504.89 | 250835.9 | 40.88
0 | 0 | 63.7 | 535.91 | 266247.2 | 45.38
0 | 0 | 62.5 | 613.85 | 304970 | 49.75
0 | 0 | 61.3 | 670.28 | 333001.9 | 54.23
0 | 0 | 60.1 | 700.53 | 348031.6 | 58.62
0 | 0 | 58.9 | 777.77 | 386404.5 | 62.51
0 | 0 | 57.7 | 823.58 | 409166.4 | 66.7
0 | 0 | 56.5 | 847.69 | 421145 | 70.31
0 | 0 | 55.3 | 927.82 | 460953 | 73.87
0 | 0 | 54.1 | 961.82 | 477844.8 | 77.3
0 | 0 | 52.9 | 975.63 | 484706.9 | 80.07
0 | 0 | 51.7 | 1042.31 | 517832.9 | 82.97
0 | 0 | 50.5 | 1079.05 | 536085.5 | 85.53
0 | 0 | 49.3 | 1065.3 | 529253.6 | 87.81
0 | 0 | 48.1 | 1136.89 | 564823.1 | 89.62
0 | 0 | 46.9 | 1154.32 | 573478.4 | 91.46
0 | 0 | 45.7 | 1130.36 | 561579.3 | 93
0 | 0 | 44.5 | 1197.75 | 595056.3 | 94.33
0 | 0 | 43.2 | 1217.79 | 605013.3 | 95.48
0 | 0 | 42 | 1232.2 | 612172.7 | 96.57
0 | 0 | 40.8 | 1199.73 | 596043.1 | 97.43
0 | 0 | 39.6 | 1247.53 | 619787.2 | 97.98
0 | 0 | 38.4 | 1264.46 | 628198.6 | 98.55
0 | 0 | 37.2 | 1221.55 | 606879.7 | 98.92
0 | 0 | 36 | 1286.09 | 638947.4 | 99.05
0 | 0 | 34.8 | 1265.31 | 628623.4 | 99.61
0 | 0 | 33.6 | 1256.21 | 624101 | 99.78
0 | 0 | 32.4 | 1306.59 | 649131.4 | 99.86
0 | 0 | 31.2 | 1245.98 | 619018.9 | 100
0 | 0 | 30 | 1303.51 | 647598 | 99.98
0 | 0 | 28.8 | 1306.69 | 649178.8 | 99.94
0 | 0 | 27.6 | 1270.62 | 631258.1 | 99.82
0 | 0 | 26.4 | 1282.62 | 637223.1 | 99.8
0 | 0 | 25.2 | 1294.63 | 643187.9 | 99.69
0 | 0 | 24 | 1270.29 | 631095.4 | 99.59
0 | 0 | 22.7 | 1277.46 | 634658.7 | 99.52
0 | 0 | 21.6 | 1300.88 | 646293.4 | 99.39
0 | 0 | 20.4 | 1266.75 | 629337.8 | 99.29
0 | 0 | 19.1 | 1278.11 | 634982.2 | 99.03
0 | 0 | 17.9 | 1299.61 | 645663.6 | 98.68
0 | 0 | 16.7 | 1269.04 | 630474.6 | 98.63
0 | 0 | 15.5 | 1274.62 | 633248.7 | 98.31
0 | 0 | 14.3 | 1304.87 | 648277.9 | 98.02
0 | 0 | 13.1 | 1256.78 | 624384.8 | 97.97
0 | 0 | 11.9 | 1276.12 | 633992.4 | 97.6
0 | 0 | 10.7 | 1300.03 | 645873.4 | 97.36
0 | 0 | 9.5 | 1250.02 | 621023.1 | 97.16
0 | 0 | 8.3 | 1279.91 | 635875.8 | 96.75
0 | 0 | 7.1 | 1286.79 | 639292.3 | 96.41
0 | 0 | 5.9 | 1229.57 | 610868 | 95.97
0 | 0 | 4.7 | 1283.97 | 637890.4 | 95.42
0 | 0 | 3.5 | 1268.48 | 630198.3 | 94.55
0 | 0 | 2.3 | 1228.35 | 610258.7 | 93.9
0 | 0 | 1.1 | 1235.68 | 613903.2 | 93.28
0 | 0 | -0.5 | 1213.81 | 603036.4 | 93.51
Measurement time: | 07/09/2022 18:51:51
Radiation device: | Clinac3 Accelerator
Energy: | 18 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 37.11 | 18435.6 | 3.14
0 | 0 | 149.1 | 37.3 | 18528.8 | 3.2
0 | 0 | 147.6 | 38.96 | 19355.6 | 3.24
0 | 0 | 146.6 | 37.65 | 18703.4 | 3.27
0 | 0 | 145.4 | 37.73 | 18744.6 | 3.28
0 | 0 | 144.1 | 38.41 | 19080.8 | 3.29
0 | 0 | 143 | 40.33 | 20037.8 | 3.32
0 | 0 | 141.7 | 37.94 | 18851.4 | 3.32
0 | 0 | 140.6 | 40.69 | 20213.7 | 3.37
0 | 0 | 139.3 | 39.54 | 19642.5 | 3.38
0 | 0 | 138.1 | 39.8 | 19772 | 3.43
0 | 0 | 137 | 41.66 | 20696.4 | 3.43
0 | 0 | 135.7 | 39.3 | 19524.6 | 3.46
0 | 0 | 134.5 | 42.32 | 21025.3 | 3.49
0 | 0 | 133.3 | 40.45 | 20098.3 | 3.51
0 | 0 | 132.1 | 43.33 | 21526.7 | 3.52
0 | 0 | 130.9 | 40.59 | 20167.7 | 3.55
0 | 0 | 129.7 | 41.84 | 20788.9 | 3.57
0 | 0 | 128.5 | 41.48 | 20608.2 | 3.58
0 | 0 | 127.3 | 42.57 | 21149.6 | 3.6
0 | 0 | 126.1 | 44.56 | 22138.3 | 3.63
0 | 0 | 124.9 | 41.36 | 20548 | 3.64
0 | 0 | 123.7 | 42.47 | 21101.5 | 3.66
0 | 0 | 122.5 | 43.79 | 21754.1 | 3.71
0 | 0 | 121.3 | 45.15 | 22430.9 | 3.73
0 | 0 | 120.1 | 43.56 | 21639.2 | 3.76
0 | 0 | 118.9 | 46.77 | 23238.3 | 3.78
0 | 0 | 117.7 | 43.19 | 21458.7 | 3.78
0 | 0 | 116.5 | 45.23 | 22471.5 | 3.84
0 | 0 | 115.3 | 46.26 | 22980.6 | 3.85
0 | 0 | 114.1 | 44.52 | 22116.1 | 3.88
0 | 0 | 112.9 | 48.16 | 23925.6 | 3.92
0 | 0 | 111.7 | 44.74 | 22229.3 | 3.95
0 | 0 | 110.5 | 46.5 | 23101.9 | 3.98
0 | 0 | 109.2 | 48.21 | 23952.3 | 4.01
0 | 0 | 108.1 | 49.17 | 24427.1 | 4.04
0 | 0 | 106.9 | 46.76 | 23229.7 | 4.07
0 | 0 | 105.6 | 46.51 | 23107.4 | 4.13
0 | 0 | 104.5 | 48.49 | 24091 | 4.15
0 | 0 | 103.3 | 50.24 | 24957.5 | 4.21
0 | 0 | 102.1 | 48.92 | 24305.2 | 4.29
0 | 0 | 100.9 | 53.84 | 26750.3 | 4.39
0 | 0 | 99.6 | 51.27 | 25472.4 | 4.54
0 | 0 | 98.5 | 55.92 | 27780.7 | 4.8
0 | 0 | 97.3 | 61.87 | 30736.1 | 5.16
0 | 0 | 96 | 63.69 | 31642.7 | 5.59
0 | 0 | 94.8 | 75.42 | 37469.8 | 6.2
0 | 0 | 93.6 | 79.57 | 39531.4 | 7.02
0 | 0 | 92.4 | 92.96 | 46182.6 | 8.1
0 | 0 | 91.2 | 115.04 | 57153.1 | 9.54
0 | 0 | 90 | 128.74 | 63961.1 | 10.82
0 | 0 | 88.8 | 145.88 | 72476.9 | 12.88
0 | 0 | 87.6 | 174.29 | 86588.1 | 14.78
0 | 0 | 86.4 | 192.85 | 95808.8 | 16.94
0 | 0 | 85.2 | 242.95 | 120698.8 | 20.11
0 | 0 | 84 | 264.2 | 131259.5 | 22.45
0 | 0 | 82.8 | 297.7 | 147901.8 | 25.8
0 | 0 | 81.6 | 344.01 | 170910.3 | 29.18
0 | 0 | 80.4 | 371.34 | 184484.4 | 31.94
0 | 0 | 79.1 | 442.25 | 219716.4 | 36.23
0 | 0 | 77.9 | 455.78 | 226438.1 | 39.34
0 | 0 | 76.7 | 506.65 | 251708.5 | 42.74
0 | 0 | 75.5 | 562.44 | 279429.4 | 46.78
0 | 0 | 74.3 | 575.21 | 285769.2 | 49.53
0 | 0 | 73.1 | 668.74 | 332238.7 | 53.84
0 | 0 | 71.9 | 657.14 | 326477.8 | 56.93
0 | 0 | 70.7 | 708.82 | 352149.6 | 59.84
0 | 0 | 69.5 | 768.94 | 382020.3 | 63.7
0 | 0 | 68.3 | 772.12 | 383599.2 | 66.01
0 | 0 | 67.1 | 859.66 | 427091.8 | 69.34
0 | 0 | 65.9 | 857.03 | 425783.8 | 71.97
0 | 0 | 64.7 | 866.65 | 430562.2 | 74.32
0 | 0 | 63.4 | 943.96 | 468973.1 | 76.99
0 | 0 | 62.3 | 918.81 | 456478.5 | 79.17
0 | 0 | 61.1 | 1025.74 | 509602.2 | 81.1
0 | 0 | 59.9 | 991.18 | 492429.5 | 83.29
0 | 0 | 58.7 | 998.73 | 496182.7 | 84.87
0 | 0 | 57.4 | 1068.34 | 530765.6 | 86.47
0 | 0 | 56.2 | 1047.09 | 520205 | 88.04
0 | 0 | 55 | 1128.55 | 560675.2 | 89.19
0 | 0 | 53.8 | 1083.97 | 538530.6 | 90.74
0 | 0 | 52.6 | 1120.91 | 556882.9 | 91.51
0 | 0 | 51.4 | 1151 | 571832.9 | 92.53
0 | 0 | 50.2 | 1130.19 | 561491.7 | 93.45
0 | 0 | 49 | 1213.4 | 602831.9 | 94.38
0 | 0 | 47.8 | 1142.07 | 567393.3 | 94.95
0 | 0 | 46.6 | 1175.02 | 583764.4 | 95.58
0 | 0 | 45.4 | 1210.32 | 601301.4 | 96.1
0 | 0 | 44.2 | 1170.63 | 581585.2 | 96.6
0 | 0 | 42.9 | 1249.76 | 620895.3 | 97.04
0 | 0 | 41.8 | 1184.48 | 588462.2 | 97.55
0 | 0 | 40.6 | 1184.28 | 588366.4 | 97.52
0 | 0 | 39.3 | 1254.49 | 623246.6 | 98.1
0 | 0 | 38.1 | 1197.07 | 594717.8 | 98.35
0 | 0 | 36.9 | 1274.71 | 633291.9 | 98.67
0 | 0 | 35.7 | 1216.55 | 604395.1 | 98.85
0 | 0 | 34.5 | 1202.54 | 597435.2 | 98.97
0 | 0 | 33.3 | 1280.2 | 636021.1 | 99.04
0 | 0 | 32.1 | 1215.47 | 603860 | 99.19
0 | 0 | 30.9 | 1285.95 | 638877.6 | 99.47
0 | 0 | 29.7 | 1253.61 | 622809.6 | 99.55
0 | 0 | 28.5 | 1220.65 | 606432.4 | 99.63
0 | 0 | 27.3 | 1304.92 | 648301.5 | 99.72
0 | 0 | 26.1 | 1230.49 | 611323.4 | 99.73
0 | 0 | 24.9 | 1293.39 | 642570.2 | 99.69
0 | 0 | 23.7 | 1257 | 624494.6 | 99.92
0 | 0 | 22.5 | 1244.67 | 618369.4 | 99.92
0 | 0 | 21.3 | 1315.92 | 653763.4 | 100
0 | 0 | 20.1 | 1223.69 | 607943.1 | 99.93
0 | 0 | 18.9 | 1306.75 | 649207.9 | 99.79
0 | 0 | 17.7 | 1263.6 | 627769.9 | 100
0 | 0 | 16.4 | 1236.88 | 614497.3 | 99.94
0 | 0 | 15.2 | 1305.26 | 648469.9 | 99.64
0 | 0 | 14 | 1245.2 | 618629 | 99.65
0 | 0 | 12.8 | 1323.31 | 657439.3 | 99.5
0 | 0 | 11.6 | 1227.53 | 609851.7 | 99.14
0 | 0 | 10.4 | 1237.2 | 614657.3 | 99.2
0 | 0 | 9.2 | 1304.34 | 648013 | 99.1
0 | 0 | 8 | 1229.73 | 610944.1 | 98.6
0 | 0 | 6.8 | 1316.84 | 654222.8 | 98.36
0 | 0 | 5.6 | 1217.03 | 604637.2 | 97.75
0 | 0 | 4.4 | 1219.79 | 606009.2 | 97.14
0 | 0 | 3.2 | 1284.5 | 638156.7 | 96.5
0 | 0 | 2 | 1194.22 | 593305.5 | 95.34
0 | 0 | 0.8 | 1282.73 | 637276.5 | 95.3
0 | 0 | -0.5 | 1225.71 | 608948.7 | 95.66