## Diamant
0,5x0,5 cm² | [5 empty] | 1x1 cm² | [6 empty] | 2x2cm² | [5 empty] | 3x3 cm²
Measurement time: | 27/05/2023 10:11:08 |  |  |  |  | Measurement time: | 27/05/2023 10:34:11 | [5 empty] | Measurement time: | 27/05/2023 10:55:28 |  |  |  |  | Measurement time: | 27/05/2023 11:21:09
Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Controller: | CCU |  |  |  |  | Controller: | CCU | [5 empty] | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant | [5 empty] | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant
Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission | [5 empty] | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | SSD: | 900 mm | [5 empty] | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 5 x 5 mm |  |  |  |  | Field size: | 10 x 10 mm | [5 empty] | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 | [5 empty] | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water | [5 empty] | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline | [5 empty] | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Wedge: | none | [5 empty] | Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]: | [6 empty] | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 54.94 | 713 | 0.95 | -20 | 0 | 100 | 75.1 | 634.4 | 0.69 |  | -30 | 0 | 100 | 130.51 | 1102.5 | 1.08 | -30 | 0 | 100 | 324.34 | 2739.7 | 2.56
-9 | 0 | 100 | 68.32 | 886.7 | 1.18 | -19 | 0 | 100 | 71.86 | 607 | 0.66 |  | -29 | 0 | 100 | 153.93 | 1300.2 | 1.27 | -29 | 0 | 100 | 353.13 | 2982.9 | 2.79
-8 | 0 | 100 | 96.98 | 1258.7 | 1.68 | -18 | 0 | 100 | 82.56 | 697.4 | 0.75 |  | -28 | 0 | 100 | 155.47 | 1313.3 | 1.29 | -28 | 0 | 100 | 384.49 | 3247.8 | 3.04
-7 | 0 | 100 | 134.42 | 1744.6 | 2.33 | -17 | 0 | 100 | 86.03 | 726.7 | 0.79 |  | -27 | 0 | 100 | 168.92 | 1426.9 | 1.4 | -27 | 0 | 100 | 406.39 | 3432.8 | 3.21
-6 | 0 | 100 | 209.18 | 2715 | 3.63 | -16 | 0 | 100 | 92.84 | 784.2 | 0.85 |  | -26 | 0 | 100 | 182.93 | 1545.3 | 1.51 | -26 | 0 | 100 | 433.36 | 3660.6 | 3.43
-5 | 0 | 100 | 362.15 | 4700.4 | 6.29 | -15 | 0 | 100 | 109.32 | 923.4 | 1 |  | -25 | 0 | 100 | 197.03 | 1664.4 | 1.63 | -25 | 0 | 100 | 466.95 | 3944.4 | 3.69
-4 | 0 | 100 | 785.65 | 10196.9 | 13.61 | -14 | 0 | 100 | 115.79 | 978.1 | 1.06 |  | -24 | 0 | 100 | 208.13 | 1758.1 | 1.72 | -24 | 0 | 100 | 511.31 | 4319.1 | 4.04
-3 | 0 | 100 | 1949.96 | 25308.4 | 33.82 | -13 | 0 | 100 | 139.83 | 1181.2 | 1.27 |  | -23 | 0 | 100 | 217.63 | 1838.3 | 1.8 | -23 | 0 | 100 | 558.57 | 4718.3 | 4.42
-2 | 0 | 100 | 3937.31 | 51102 | 68.27 | -12 | 0 | 100 | 164.45 | 1389.1 | 1.5 |  | -22 | 0 | 100 | 237.81 | 2008.8 | 1.97 | -22 | 0 | 100 | 624.56 | 5275.7 | 4.94
-1 | 0 | 100 | 5451.24 | 70751.2 | 94.03 | -11 | 0 | 100 | 211.72 | 1788.4 | 1.93 |  | -21 | 0 | 100 | 255.65 | 2159.5 | 2.11 | -21 | 0 | 100 | 703.96 | 5946.5 | 5.57
0 | 0 | 100 | 5799.8 | 75275.1 | 100 | -10 | 0 | 100 | 285.08 | 2408.1 | 2.6 |  | -20 | 0 | 100 | 287.92 | 2432.1 | 2.38 | -20 | 0 | 100 | 817.65 | 6906.8 | 6.47
1 | 0 | 100 | 5215.41 | 67690.3 | 90.34 | -9 | 0 | 100 | 409.92 | 3462.7 | 3.75 |  | -19 | 0 | 100 | 309.6 | 2615.2 | 2.56 | -19 | 0 | 100 | 996.1 | 8414.1 | 7.88
2 | 0 | 100 | 3421.05 | 44401.5 | 59.08 | -8 | 0 | 100 | 635.04 | 5364.3 | 5.81 |  | -18 | 0 | 100 | 349.37 | 2951.1 | 2.89 | -18 | 0 | 100 | 1280.57 | 10817.1 | 10.14
3 | 0 | 100 | 1469.84 | 19076.9 | 25.4 | -7 | 0 | 100 | 1095.42 | 9253.1 | 10.01 |  | -17 | 0 | 100 | 396.29 | 3347.5 | 3.28 | -17 | 0 | 100 | 1837.55 | 15522 | 14.54
4 | 0 | 100 | 572.21 | 7426.7 | 9.91 | -6 | 0 | 100 | 2344.59 | 19804.9 | 21.41 |  | -16 | 0 | 100 | 472.3 | 3989.6 | 3.91 | -16 | 0 | 100 | 3296.43 | 27845.2 | 26.09
5 | 0 | 100 | 284.62 | 3694.1 | 4.91 | -5 | 0 | 100 | 5310.66 | 44859.6 | 48.46 |  | -15 | 0 | 100 | 565.48 | 4776.6 | 4.68 | -15 | 0 | 100 | 6398.06 | 54045 | 50.62
6 | 0 | 100 | 163.66 | 2124.2 | 2.83 | -4 | 0 | 100 | 8531.68 | 72067.8 | 77.97 |  | -14 | 0 | 100 | 719.51 | 6077.7 | 5.95 | -14 | 0 | 100 | 9612.72 | 81199.5 | 76.05
7 | 0 | 100 | 103.43 | 1342.4 | 1.79 | -3 | 0 | 100 | 10156.15 | 85789.9 | 92.76 |  | -13 | 0 | 100 | 976.3 | 8246.9 | 8.08 | -13 | 0 | 100 | 11185.93 | 94488.6 | 88.47
8 | 0 | 100 | 71.71 | 930.8 | 1.24 | -2 | 0 | 100 | 10699.5 | 90379.6 | 97.73 |  | -12 | 0 | 100 | 1486.51 | 12556.7 | 12.29 | -12 | 0 | 100 | 11745.68 | 99216.8 | 92.98
9 | 0 | 100 | 54.55 | 708 | 0.94 | -1 | 0 | 100 | 10902.16 | 92091.5 | 99.57 |  | -11 | 0 | 100 | 2786.29 | 23536.1 | 23.04 | -11 | 0 | 100 | 12066.54 | 101927.2 | 95.51
10 | 0 | 100 | 41.03 | 532.6 | 0.71 | 0 | 0 | 100 | 10947.74 | 92476.5 | 100 |  | -10 | 0 | 100 | 5843.77 | 49362.8 | 48.34 | -10 | 0 | 100 | 12249.13 | 103469.5 | 96.96
[6 empty] | 1 | 0 | 100 | 10869.45 | 91815.2 | 99.25 |  | -9 | 0 | 100 | 9145.2 | 77250.3 | 75.68 | -9 | 0 | 100 | 12345.99 | 104287.7 | 97.71
[6 empty] | 2 | 0 | 100 | 10542.16 | 89050.6 | 96.42 |  | -8 | 0 | 100 | 10813.56 | 91343.1 | 89.51 | -8 | 0 | 100 | 12447.67 | 105146.5 | 98.52
Measurement time: | 27/05/2023 10:12:41 |  |  |  |  | 3 | 0 | 100 | 9623.12 | 81287.3 | 87.93 |  | -7 | 0 | 100 | 11397.53 | 96276 | 94.35 | -7 | 0 | 100 | 12510.58 | 105678 | 99.05
Radiation device: | Novalis TB Accelerator |  |  |  |  | 4 | 0 | 100 | 7238.96 | 61148.1 | 66.17 |  | -6 | 0 | 100 | 11673.46 | 98606.8 | 96.62 | -6 | 0 | 100 | 12565.55 | 106142.3 | 99.41
Energy: | 6 MV (FFF) |  |  |  |  | 5 | 0 | 100 | 3831.67 | 32366.5 | 35 |  | -5 | 0 | 100 | 11847.9 | 100080.2 | 98.07 | -5 | 0 | 100 | 12582.41 | 106284.8 | 99.55
Controller: | CCU |  |  |  |  | 6 | 0 | 100 | 1671.22 | 14117 | 15.28 |  | -4 | 0 | 100 | 11984.99 | 101238.3 | 99.15 | -4 | 0 | 100 | 12609.12 | 106510.4 | 99.78
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 7 | 0 | 100 | 866.89 | 7322.7 | 7.92 |  | -3 | 0 | 100 | 12035.27 | 101663 | 99.59 | -3 | 0 | 100 | 12625.13 | 106645.6 | 99.94
Field detector: | PTW60019 Diamant |  |  |  |  | 8 | 0 | 100 | 531.72 | 4491.5 | 4.86 |  | -2 | 0 | 100 | 12065.42 | 101917.7 | 99.88 | -2 | 0 | 100 | 12641.31 | 106782.3 | 100.05
Reference detector: | Stealth CI transmission |  |  |  |  | 9 | 0 | 100 | 358.14 | 3025.3 | 3.27 |  | -1 | 0 | 100 | 12103.99 | 102243.5 | 100.17 | -1 | 0 | 100 | 12636.29 | 106739.9 | 100
SAD: | 1000 mm |  |  |  |  | 10 | 0 | 100 | 256.66 | 2168 | 2.35 |  | 0 | 0 | 100 | 12087.82 | 102106.9 | 100 | 0 | 0 | 100 | 12635.16 | 106730.3 | 100
SSD: | 900 mm |  |  |  |  | 11 | 0 | 100 | 192.05 | 1622.3 | 1.76 |  | 1 | 0 | 100 | 12064.21 | 101907.5 | 99.84 | 1 | 0 | 100 | 12619.16 | 106595.2 | 99.93
Field size: | 5 x 5 mm |  |  |  |  | 12 | 0 | 100 | 160.53 | 1356 | 1.46 |  | 2 | 0 | 100 | 12062.93 | 101896.6 | 99.79 | 2 | 0 | 100 | 12606.64 | 106489.4 | 99.77
Gantry angle: | 0 |  |  |  |  | 13 | 0 | 100 | 125.8 | 1062.6 | 1.15 |  | 3 | 0 | 100 | 11977.9 | 101178.4 | 99.16 | 3 | 0 | 100 | 12586.79 | 106321.8 | 99.67
Measurement medium: | Water |  |  |  |  | 14 | 0 | 100 | 107.59 | 908.8 | 0.98 |  | 4 | 0 | 100 | 11917 | 100664 | 98.67 | 4 | 0 | 100 | 12561.34 | 106106.7 | 99.46
Scan type: | Crossline |  |  |  |  | 15 | 0 | 100 | 99.07 | 836.9 | 0.91 |  | 5 | 0 | 100 | 11794.03 | 99625.2 | 97.58 | 5 | 0 | 100 | 12528.09 | 105825.9 | 99.2
Scan mode: | StepByStep |  |  |  |  | 16 | 0 | 100 | 89 | 751.8 | 0.81 |  | 6 | 0 | 100 | 11565.9 | 97698.2 | 95.7 | 6 | 0 | 100 | 12501.11 | 105598 | 98.96
Wedge: | none |  |  |  |  | 17 | 0 | 100 | 82.68 | 698.4 | 0.76 |  | 7 | 0 | 100 | 11113.65 | 93878 | 91.97 | 7 | 0 | 100 | 12437.03 | 105056.7 | 98.53
[6 empty] | 18 | 0 | 100 | 78.14 | 660.1 | 0.71 |  | 8 | 0 | 100 | 10029.39 | 84719.1 | 82.98 | 8 | 0 | 100 | 12375.35 | 104535.7 | 98
Points [mm]: | [5 empty] | 19 | 0 | 100 | 70.93 | 599.2 | 0.65 |  | 9 | 0 | 100 | 7402.99 | 62533.7 | 61.27 | 9 | 0 | 100 | 12249.77 | 103474.9 | 97.01
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 20 | 0 | 100 | 63.26 | 534.4 | 0.58 |  | 10 | 0 | 100 | 3990.77 | 33710.4 | 33.03 | 10 | 0 | 100 | 12084.65 | 102080.1 | 95.64
0 | -10 | 100 | 47.36 | 614.6 | 0.81 | [7 empty] | 11 | 0 | 100 | 1952.78 | 16495.3 | 16.16 | 11 | 0 | 100 | 11830.99 | 99937.4 | 93.7
0 | -9 | 100 | 62.56 | 811.9 | 1.07 | [7 empty] | 12 | 0 | 100 | 1151.99 | 9730.9 | 9.53 | 12 | 0 | 100 | 11383.85 | 96160.4 | 90.16
0 | -8 | 100 | 84.83 | 1101 | 1.45 | Measurement time: | 27/05/2023 10:37:20 | [5 empty] | 13 | 0 | 100 | 820.6 | 6931.6 | 6.79 | 13 | 0 | 100 | 10177.88 | 85973.4 | 80.6
0 | -7 | 100 | 139.1 | 1805.4 | 2.38 | Radiation device: | Novalis TB Accelerator | [5 empty] | 14 | 0 | 100 | 611.7 | 5167.1 | 5.06 | 14 | 0 | 100 | 7573.61 | 63974.9 | 59.95
0 | -6 | 100 | 236.57 | 3070.4 | 4.06 | Energy: | 6 MV (FFF) | [5 empty] | 15 | 0 | 100 | 493.66 | 4170 | 4.09 | 15 | 0 | 100 | 4106.23 | 34685.7 | 32.52
0 | -5 | 100 | 468.06 | 6074.9 | 8.01 | Controller: | CCU | [5 empty] | 16 | 0 | 100 | 413.56 | 3493.4 | 3.42 | 16 | 0 | 100 | 2171.29 | 18341.1 | 17.21
0 | -4 | 100 | 1020.94 | 13250.7 | 17.56 | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | 17 | 0 | 100 | 355.05 | 2999.1 | 2.94 | 17 | 0 | 100 | 1412.36 | 11930.3 | 11.18
0 | -3 | 100 | 2288.84 | 29706.7 | 39.32 | Field detector: | PTW60019 Diamant | [5 empty] | 18 | 0 | 100 | 306.88 | 2592.3 | 2.54 | 18 | 0 | 100 | 1051.06 | 8878.4 | 8.32
0 | -2 | 100 | 4199.68 | 54507.3 | 72.07 | Reference detector: | Stealth CI transmission | [5 empty] | 19 | 0 | 100 | 286.36 | 2418.9 | 2.37 | 19 | 0 | 100 | 842.49 | 7116.6 | 6.67
0 | -1 | 100 | 5433.23 | 70517.4 | 93.37 | SAD: | 1000 mm | [5 empty] | 20 | 0 | 100 | 259.36 | 2190.8 | 2.15 | 20 | 0 | 100 | 720.31 | 6084.5 | 5.7
0 | 0 | 100 | 5798.97 | 75264.4 | 100 | SSD: | 900 mm | [5 empty] | 21 | 0 | 100 | 232.3 | 1962.3 | 1.92 | 21 | 0 | 100 | 625.47 | 5283.4 | 4.95
0 | 1 | 100 | 5593.83 | 72601.9 | 95.85 | Field size: | 10 x 10 mm | [5 empty] | 22 | 0 | 100 | 219.48 | 1854 | 1.82 | 22 | 0 | 100 | 557.11 | 4706 | 4.41
0 | 2 | 100 | 4882.67 | 63371.8 | 83.9 | Gantry angle: | 0 | [5 empty] | 23 | 0 | 100 | 209.18 | 1766.9 | 1.73 | 23 | 0 | 100 | 512.45 | 4328.7 | 4.06
0 | 3 | 100 | 3231 | 41934.9 | 55.55 | Measurement medium: | Water | [5 empty] | 24 | 0 | 100 | 194.79 | 1645.4 | 1.61 | 24 | 0 | 100 | 469.35 | 3964.6 | 3.72
0 | 4 | 100 | 1482.62 | 19242.9 | 25.39 | Scan type: | Crossline | [5 empty] | 25 | 0 | 100 | 182.02 | 1537.5 | 1.51 | 25 | 0 | 100 | 426.06 | 3598.9 | 3.37
0 | 5 | 100 | 596.81 | 7745.9 | 10.23 | Scan mode: | StepByStep | [5 empty] | 26 | 0 | 100 | 173.98 | 1469.6 | 1.44 | 26 | 0 | 100 | 399.44 | 3374.1 | 3.16
0 | 6 | 100 | 269.16 | 3493.4 | 4.63 | Wedge: | none | [5 empty] | 27 | 0 | 100 | 165.45 | 1397.6 | 1.37 | 27 | 0 | 100 | 382.23 | 3228.7 | 3.03
0 | 7 | 100 | 142.04 | 1843.5 | 2.43 | [7 empty] | 28 | 0 | 100 | 162.58 | 1373.3 | 1.35 | 28 | 0 | 100 | 361.89 | 3056.9 | 2.87
0 | 8 | 100 | 94.25 | 1223.3 | 1.61 | Points [mm]: | [6 empty] | 29 | 0 | 100 | 165.87 | 1401.1 | 1.37 | 29 | 0 | 100 | 331.67 | 2801.7 | 2.63
0 | 9 | 100 | 65.73 | 853.1 | 1.12 | Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | 30 | 0 | 100 | 150.42 | 1270.6 | 1.25 | 30 | 0 | 100 | 318.4 | 2689.6 | 2.52
0 | 10 | 100 | 54.14 | 702.7 | 0.93 | 0 | -20 | 100 | 65.87 | 556.4 | 0.6
[6 empty] | 0 | -19 | 100 | 77.73 | 656.6 | 0.71
[6 empty] | 0 | -18 | 100 | 72.55 | 612.9 | 0.66 |  | Measurement time: | 27/05/2023 10:58:20 |  |  |  |  | Measurement time: | 27/05/2023 11:23:54
Measurement time: | 27/05/2023 10:14:20 |  |  |  |  | 0 | -17 | 100 | 82.75 | 699 | 0.76 |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -16 | 100 | 95.48 | 806.5 | 0.87 |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Energy: | 6 MV (FFF) |  |  |  |  | 0 | -15 | 100 | 98.43 | 831.5 | 0.9 |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Controller: | CCU |  |  |  |  | 0 | -14 | 100 | 121.62 | 1027.4 | 1.11 |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -13 | 100 | 139.14 | 1175.3 | 1.27 |  | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant
Field detector: | PTW60019 Diamant |  |  |  |  | 0 | -12 | 100 | 159.43 | 1346.7 | 1.45 |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -11 | 100 | 223.4 | 1887.1 | 2.04 |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SAD: | 1000 mm |  |  |  |  | 0 | -10 | 100 | 310.88 | 2626 | 2.84 |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
SSD: | 900 mm |  |  |  |  | 0 | -9 | 100 | 445.78 | 3765.5 | 4.06 |  | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Field size: | 5 x 5 mm |  |  |  |  | 0 | -8 | 100 | 752.2 | 6353.9 | 6.85 |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Gantry angle: | 0 |  |  |  |  | 0 | -7 | 100 | 1371.42 | 11584.5 | 12.49 |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Measurement medium: | Water |  |  |  |  | 0 | -6 | 100 | 2906.41 | 24550.7 | 26.5 |  | Scan type: | Crossline |  |  |  |  | Scan type: | Crossline
Scan type: | Beam |  |  |  |  | 0 | -5 | 100 | 5664.33 | 47847.1 | 51.61 |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Scan mode: | StepByStep |  |  |  |  | 0 | -4 | 100 | 8717.5 | 73637.5 | 79.48 |  | Wedge: | none |  |  |  |  | Wedge: | none
Wedge: | none |  |  |  |  | 0 | -3 | 100 | 10252.18 | 86601.1 | 93.42
[6 empty] | 0 | -2 | 100 | 10726.99 | 90611.9 | 97.84 |  | Points [mm]: | [5 empty] | Points [mm]:
Points [mm]: | [5 empty] | 0 | -1 | 100 | 10892.32 | 92008.4 | 99.33 |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 0 | 100 | 10968.84 | 92654.7 | 100 |  | 0 | -30 | 100 | 141.02 | 1191.2 | 1.17 | 0 | -30 | 100 | 321.71 | 2717.5 | 2.55
0 | 0 | -0.5 | 7277.92 | 94459.5 | 63.53 | 0 | 1 | 100 | 10899.46 | 92068.8 | 99.43 |  | 0 | -29 | 100 | 147.52 | 1246.1 | 1.22 | 0 | -29 | 100 | 342.85 | 2896 | 2.72
0 | 0 | 0.5 | 7294.87 | 94679.6 | 63.55 | 0 | 2 | 100 | 10745.51 | 90768.3 | 98.06 |  | 0 | -28 | 100 | 162.24 | 1370.5 | 1.34 | 0 | -28 | 100 | 367.15 | 3101.4 | 2.91
0 | 0 | 1.5 | 8012.87 | 103998.4 | 69.89 | 0 | 3 | 100 | 10292.45 | 86941.3 | 93.96 |  | 0 | -27 | 100 | 173.19 | 1462.9 | 1.43 | 0 | -27 | 100 | 393.62 | 3324.9 | 3.12
0 | 0 | 2.5 | 9332.86 | 121130.4 | 81.57 | 0 | 4 | 100 | 9378.12 | 79217.9 | 85.51 |  | 0 | -26 | 100 | 180.71 | 1526.5 | 1.5 | 0 | -26 | 100 | 413.33 | 3491.4 | 3.27
0 | 0 | 3.5 | 10292.12 | 133580.6 | 89.81 | 0 | 5 | 100 | 7184.98 | 60692.2 | 65.5 |  | 0 | -25 | 100 | 196.42 | 1659.2 | 1.63 | 0 | -25 | 100 | 451.24 | 3811.7 | 3.57
0 | 0 | 4.5 | 10854.07 | 140874.1 | 94.65 | 0 | 6 | 100 | 4015.73 | 33921.2 | 36.61 |  | 0 | -24 | 100 | 204.26 | 1725.4 | 1.69 | 0 | -24 | 100 | 487.75 | 4120.1 | 3.86
0 | 0 | 5.5 | 11176.23 | 145055.4 | 97.45 | 0 | 7 | 100 | 1766.01 | 14917.6 | 16.1 |  | 0 | -23 | 100 | 216.02 | 1824.7 | 1.79 | 0 | -23 | 100 | 536.6 | 4532.7 | 4.25
0 | 0 | 6.5 | 11342.78 | 147217 | 98.77 | 0 | 8 | 100 | 853.18 | 7206.9 | 7.78 |  | 0 | -22 | 100 | 235.39 | 1988.4 | 1.95 | 0 | -22 | 100 | 602.52 | 5089.5 | 4.78
0 | 0 | 7.5 | 11422.71 | 148254.4 | 99.6 | 0 | 9 | 100 | 503.36 | 4252 | 4.59 |  | 0 | -21 | 100 | 251.74 | 2126.5 | 2.09 | 0 | -21 | 100 | 682.02 | 5761 | 5.41
0 | 0 | 8.5 | 11457.65 | 148707.8 | 99.86 | 0 | 10 | 100 | 350.36 | 2959.5 | 3.19 |  | 0 | -20 | 100 | 281.39 | 2376.9 | 2.33 | 0 | -20 | 100 | 810.32 | 6844.9 | 6.42
0 | 0 | 9.5 | 11462.66 | 148772.9 | 99.99 | 0 | 11 | 100 | 245.5 | 2073.8 | 2.24 |  | 0 | -19 | 100 | 313.78 | 2650.6 | 2.6 | 0 | -19 | 100 | 1009.26 | 8525.3 | 7.99
0 | 0 | 10.5 | 11430.99 | 148361.8 | 100 | 0 | 12 | 100 | 188.46 | 1591.9 | 1.72 |  | 0 | -18 | 100 | 348.22 | 2941.5 | 2.88 | 0 | -18 | 100 | 1360.9 | 11495.7 | 10.78
0 | 0 | 11.5 | 11377.3 | 147665.1 | 99.33 | 0 | 13 | 100 | 153.61 | 1297.6 | 1.4 |  | 0 | -17 | 100 | 395.73 | 3342.8 | 3.28 | 0 | -17 | 100 | 2053.53 | 17346.3 | 16.28
0 | 0 | 12.5 | 11319.44 | 146914.1 | 98.88 | 0 | 14 | 100 | 136 | 1148.8 | 1.24 |  | 0 | -16 | 100 | 473.51 | 3999.8 | 3.92 | 0 | -16 | 100 | 3623.43 | 30607.4 | 28.72
0 | 0 | 13.5 | 11240.58 | 145890.5 | 98.31 | 0 | 15 | 100 | 113.49 | 958.7 | 1.04 |  | 0 | -15 | 100 | 585.6 | 4946.6 | 4.85 | 0 | -15 | 100 | 6568.98 | 55488.8 | 52.03
0 | 0 | 14.5 | 11153.09 | 144755.1 | 97.43 | 0 | 16 | 100 | 106.03 | 895.7 | 0.97 |  | 0 | -14 | 100 | 759.3 | 6413.9 | 6.28 | 0 | -14 | 100 | 9511.51 | 80344.6 | 75.33
0 | 0 | 15.5 | 11084.52 | 143865 | 97 | 0 | 17 | 100 | 94.39 | 797.3 | 0.86 |  | 0 | -13 | 100 | 1101.9 | 9307.9 | 9.12 | 0 | -13 | 100 | 11106.28 | 93815.7 | 88.01
0 | 0 | 16.5 | 11004.28 | 142823.6 | 96.21 | 0 | 18 | 100 | 83.92 | 708.9 | 0.77 |  | 0 | -12 | 100 | 1803.41 | 15233.6 | 14.93 | 0 | -12 | 100 | 11753.93 | 99286.5 | 93.14
0 | 0 | 17.5 | 10917.81 | 141701.3 | 95.45 | 0 | 19 | 100 | 81.57 | 689.1 | 0.74 |  | 0 | -11 | 100 | 3406.61 | 28775.9 | 28.2 | 0 | -11 | 100 | 12040.87 | 101710.3 | 95.35
0 | 0 | 18.5 | 10836.02 | 140639.8 | 94.9 | 0 | 20 | 100 | 72.57 | 613 | 0.66 |  | 0 | -10 | 100 | 6359.9 | 53722.7 | 52.61 | 0 | -10 | 100 | 12206.73 | 103111.3 | 96.75
0 | 0 | 19.5 | 10747.81 | 139494.9 | 93.93 | [7 empty] | 0 | -9 | 100 | 9400.17 | 79404 | 77.81 | 0 | -9 | 100 | 12322.18 | 104086.5 | 97.67
0 | 0 | 20.5 | 10658.97 | 138341.9 | 93.16 | [7 empty] | 0 | -8 | 100 | 10858.64 | 91723.9 | 89.86 | 0 | -8 | 100 | 12405.63 | 104791.4 | 98.28
0 | 0 | 21.5 | 10571.2 | 137202.7 | 92.28 | Measurement time: | 27/05/2023 10:40:15 | [5 empty] | 0 | -7 | 100 | 11434.69 | 96589.8 | 94.63 | 0 | -7 | 100 | 12459.35 | 105245.3 | 98.72
0 | 0 | 22.5 | 10497.94 | 136251.8 | 91.64 | Radiation device: | Novalis TB Accelerator | [5 empty] | 0 | -6 | 100 | 11700.62 | 98836.2 | 96.88 | 0 | -6 | 100 | 12524.98 | 105799.6 | 99.18
0 | 0 | 23.5 | 10420.14 | 135242.2 | 91.22 | Energy: | 6 MV (FFF) | [5 empty] | 0 | -5 | 100 | 11870.55 | 100271.6 | 98.3 | 0 | -5 | 100 | 12554.12 | 106045.8 | 99.42
0 | 0 | 24.5 | 10334.42 | 134129.6 | 90.6 | Controller: | CCU | [5 empty] | 0 | -4 | 100 | 11946.98 | 100917.2 | 98.87 | 0 | -4 | 100 | 12566.71 | 106152.1 | 99.55
0 | 0 | 25.5 | 10265.47 | 133234.7 | 89.99 | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | 0 | -3 | 100 | 12024.02 | 101567.9 | 99.48 | 0 | -3 | 100 | 12618.73 | 106591.5 | 99.92
0 | 0 | 26.5 | 10181.49 | 132144.7 | 89.17 | Field detector: | PTW60019 Diamant | [5 empty] | 0 | -2 | 100 | 12081.25 | 102051.4 | 99.93 | 0 | -2 | 100 | 12628.16 | 106671.2 | 99.97
0 | 0 | 27.5 | 10097.51 | 131054.8 | 88.35 | Reference detector: | Stealth CI transmission | [5 empty] | 0 | -1 | 100 | 12081.43 | 102052.9 | 99.95 | 0 | -1 | 100 | 12624.98 | 106644.3 | 99.99
0 | 0 | 28.5 | 10013.53 | 129964.8 | 87.53 | SAD: | 1000 mm | [5 empty] | 0 | 0 | 100 | 12075.92 | 102006.4 | 100 | 0 | 0 | 100 | 12629.36 | 106681.3 | 100
0 | 0 | 29.5 | 9929.55 | 128874.8 | 86.71 | SSD: | 900 mm | [5 empty] | 0 | 1 | 100 | 12081.39 | 102052.6 | 99.99 | 0 | 1 | 100 | 12633.68 | 106717.8 | 100.01
0 | 0 | 30.5 | 9851.36 | 127859.9 | 86.04 | Field size: | 10 x 10 mm | [5 empty] | 0 | 2 | 100 | 12072.75 | 101979.6 | 99.94 | 0 | 2 | 100 | 12646.26 | 106824.1 | 100.14
0 | 0 | 31.5 | 9776.82 | 126892.5 | 85.47 | Gantry angle: | 0 | [5 empty] | 0 | 3 | 100 | 12047.99 | 101770.4 | 99.75 | 0 | 3 | 100 | 12614.5 | 106555.8 | 99.91
0 | 0 | 32.5 | 9702.27 | 125925 | 84.89 | Measurement medium: | Water | [5 empty] | 0 | 4 | 100 | 11995.34 | 101325.7 | 99.29 | 0 | 4 | 100 | 12606.39 | 106487.3 | 99.79
0 | 0 | 33.5 | 9627.73 | 124957.6 | 84.32 | Scan type: | Beam | [5 empty] | 0 | 5 | 100 | 11898.65 | 100509 | 98.5 | 0 | 5 | 100 | 12573.06 | 106205.7 | 99.57
0 | 0 | 34.5 | 9553.19 | 123990.1 | 83.74 | Scan mode: | StepByStep | [5 empty] | 0 | 6 | 100 | 11762.82 | 99361.6 | 97.41 | 0 | 6 | 100 | 12555.69 | 106059.1 | 99.42
0 | 0 | 35.5 | 9477.58 | 123008.7 | 83.05 | Wedge: | none | [5 empty] | 0 | 7 | 100 | 11519.26 | 97304.2 | 95.4 | 0 | 7 | 100 | 12517.99 | 105740.6 | 99.09
0 | 0 | 36.5 | 9401.29 | 122018.5 | 82.29 | [7 empty] | 0 | 8 | 100 | 11049.81 | 93338.7 | 91.47 | 0 | 8 | 100 | 12460.71 | 105256.8 | 98.67
0 | 0 | 37.5 | 9324.99 | 121028.3 | 81.52 | Points [mm]: | [6 empty] | 0 | 9 | 100 | 10001.41 | 84482.8 | 82.79 | 0 | 9 | 100 | 12391.49 | 104672 | 98.12
0 | 0 | 38.5 | 9248.7 | 120038.1 | 80.76 | Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | 0 | 10 | 100 | 7849.08 | 66301.9 | 64.98 | 0 | 10 | 100 | 12295.01 | 103857.1 | 97.34
0 | 0 | 39.5 | 9172.41 | 119047.9 | 80 | 0 | 0 | -0.5 | 11919.65 | 100686.4 | 11942.23 | 57.0251 | 0 | 11 | 100 | 4661.24 | 39373.9 | 38.57 | 0 | 11 | 100 | 12121.2 | 102388.8 | 95.96
0 | 0 | 40.5 | 9099.62 | 118103.2 | 79.31 | 0 | 0 | 0.5 | 11969.9 | 101110.8 | 11993.35 | 57.2692 | 0 | 12 | 100 | 2226.93 | 18811.1 | 18.43 | 0 | 12 | 100 | 11869.88 | 100265.9 | 94.01
0 | 0 | 41.5 | 9029.21 | 117189.3 | 78.67 | 0 | 0 | 1.5 | 13033.58 | 110095.8 | 13045.31 | 62.2924 | 0 | 13 | 100 | 1288.36 | 10882.9 | 10.66 | 0 | 13 | 100 | 11374.61 | 96082.4 | 90.06
0 | 0 | 42.5 | 8958.8 | 116275.5 | 78.04 | 0 | 0 | 2.5 | 15276.84 | 129044.8 | 15302.54 | 73.0709 | 0 | 14 | 100 | 854.73 | 7220 | 7.07 | 0 | 14 | 100 | 10367.27 | 87573.2 | 82.08
0 | 0 | 43.5 | 8888.39 | 115361.7 | 77.4 | 0 | 0 | 3.5 | 17193.06 | 145231.3 | 17238.89 | 82.3171 | 0 | 15 | 100 | 637.16 | 5382.1 | 5.27 | 0 | 15 | 100 | 8070.13 | 68169.1 | 63.89
0 | 0 | 44.5 | 8817.99 | 114447.9 | 76.77 | 0 | 0 | 4.5 | 18488.02 | 156170 | 18537.22 | 88.5167 | 0 | 16 | 100 | 507.1 | 4283.5 | 4.2 | 0 | 16 | 100 | 4951.51 | 41825.8 | 39.2
0 | 0 | 45.5 | 8750.51 | 113572.2 | 76.16 | 0 | 0 | 5.5 | 19425.73 | 164090.9 | 19477.1 | 93.0047 | 0 | 17 | 100 | 430.73 | 3638.4 | 3.56 | 0 | 17 | 100 | 2626.21 | 22183.8 | 20.79
0 | 0 | 46.5 | 8685.03 | 112722.2 | 75.57 | 0 | 0 | 6.5 | 20057.63 | 169428.6 | 20115.21 | 96.0518 | 0 | 18 | 100 | 372.3 | 3144.9 | 3.08 | 0 | 18 | 100 | 1533.13 | 12950.5 | 12.13
0 | 0 | 47.5 | 8619.54 | 111872.3 | 74.97 | 0 | 0 | 7.5 | 20439.7 | 172656 | 20492.06 | 97.8513 | 0 | 19 | 100 | 329.91 | 2786.8 | 2.73 | 0 | 19 | 100 | 1091.03 | 9216 | 8.64
0 | 0 | 48.5 | 8554.05 | 111022.3 | 74.38 | 0 | 0 | 8.5 | 20672.56 | 174623 | 20727.19 | 98.974 | 0 | 20 | 100 | 296.8 | 2507.1 | 2.46 | 0 | 20 | 100 | 877.87 | 7415.5 | 6.95
0 | 0 | 49.5 | 8488.56 | 110172.3 | 73.79 | 0 | 0 | 9.5 | 20826.29 | 175921.5 | 20889.53 | 99.7492 | 0 | 21 | 100 | 274.93 | 2322.4 | 2.27 | 0 | 21 | 100 | 734.22 | 6202 | 5.81
0 | 0 | 50.5 | 8422.7 | 109317.5 | 73.21 | 0 | 0 | 10.5 | 20891.83 | 176475.2 | 20942.05 | 100 | 0 | 22 | 100 | 246.42 | 2081.5 | 2.04 | 0 | 22 | 100 | 636.75 | 5378.6 | 5.04
0 | 0 | 51.5 | 8356.55 | 108458.9 | 72.65 | 0 | 0 | 11.5 | 20893.47 | 176489 | 20937.94 | 99.9804 | 0 | 23 | 100 | 228.41 | 1929.4 | 1.89 | 0 | 23 | 100 | 578.24 | 4884.5 | 4.58
0 | 0 | 52.5 | 8290.4 | 107600.4 | 72.08 | 0 | 0 | 12.5 | 20860.7 | 176212.2 | 20914.19 | 99.867 | 0 | 24 | 100 | 211.98 | 1790.6 | 1.75 | 0 | 24 | 100 | 520.74 | 4398.8 | 4.12
0 | 0 | 53.5 | 8224.25 | 106741.8 | 71.51 | 0 | 0 | 13.5 | 20786.19 | 175582.8 | 20843.99 | 99.5318 | 0 | 25 | 100 | 213.28 | 1801.6 | 1.77 | 0 | 25 | 100 | 477.1 | 4030.1 | 3.78
0 | 0 | 54.5 | 8158.1 | 105883.3 | 70.95 | 0 | 0 | 14.5 | 20684.3 | 174722.1 | 20733.82 | 99.0057 | 0 | 26 | 100 | 190.38 | 1608.2 | 1.57 | 0 | 26 | 100 | 437.09 | 3692.1 | 3.46
0 | 0 | 55.5 | 8097.07 | 105091.3 | 70.46 | 0 | 0 | 15.5 | 20564.04 | 173706.3 | 20611.08 | 98.4196 | 0 | 27 | 100 | 180.28 | 1522.8 | 1.49 | 0 | 27 | 100 | 408.83 | 3453.4 | 3.24
0 | 0 | 56.5 | 8040.51 | 104357.1 | 70.05 | 0 | 0 | 16.5 | 20430.58 | 172578.9 | 20483.08 | 97.8084 | 0 | 28 | 100 | 166.26 | 1404.4 | 1.37 | 0 | 28 | 100 | 376.24 | 3178.1 | 2.98
0 | 0 | 57.5 | 7983.94 | 103622.9 | 69.64 | 0 | 0 | 17.5 | 20312.25 | 171579.4 | 20349.03 | 97.1683 | 0 | 29 | 100 | 159.79 | 1349.8 | 1.32 | 0 | 29 | 100 | 359.28 | 3034.9 | 2.84
0 | 0 | 58.5 | 7927.37 | 102888.7 | 69.22 | 0 | 0 | 18.5 | 20183.54 | 170492.2 | 20207.44 | 96.4922 | 0 | 30 | 100 | 151.88 | 1283 | 1.26 | 0 | 30 | 100 | 338.42 | 2858.7 | 2.68
0 | 0 | 59.5 | 7870.8 | 102154.5 | 68.81 | 0 | 0 | 19.5 | 20052.5 | 169385.3 | 20076.99 | 95.8693
0 | 0 | 60.5 | 7811.35 | 101382.9 | 68.29 | 0 | 0 | 20.5 | 19914.07 | 168215.9 | 19928.79 | 95.1616
0 | 0 | 61.5 | 7749.16 | 100575.7 | 67.67 | 0 | 0 | 21.5 | 19748.93 | 166821 | 19770.33 | 94.4049 | Measurement time: | 27/05/2023 11:00:52 |  |  |  |  | Measurement time: | 27/05/2023 11:26:24
0 | 0 | 62.5 | 7686.96 | 99768.4 | 67.04 | 0 | 0 | 22.5 | 19605.02 | 165605.4 | 19631.81 | 93.7435 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
0 | 0 | 63.5 | 7624.76 | 98961.2 | 66.42 | 0 | 0 | 23.5 | 19461.86 | 164396 | 19487.31 | 93.0535 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
0 | 0 | 64.5 | 7562.57 | 98153.9 | 65.8 | 0 | 0 | 24.5 | 19326.2 | 163250.2 | 19359.5 | 92.4432 | Controller: | CCU |  |  |  |  | Controller: | CCU
0 | 0 | 65.5 | 7503.53 | 97387.7 | 65.24 | 0 | 0 | 25.5 | 19184.95 | 162057 | 19220.01 | 91.7771 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | 66.5 | 7447.62 | 96662.1 | 64.74 | 0 | 0 | 26.5 | 19029.15 | 160741 | 19062.6 | 91.0255 | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant
0 | 0 | 67.5 | 7391.72 | 95936.5 | 64.24 | 0 | 0 | 27.5 | 18873.36 | 159424.9 | 18905.19 | 90.2738 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
0 | 0 | 68.5 | 7335.81 | 95210.9 | 63.74 | 0 | 0 | 28.5 | 18717.56 | 158108.9 | 18747.78 | 89.5222 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
0 | 0 | 69.5 | 7279.9 | 94485.3 | 63.24 | 0 | 0 | 29.5 | 18561.77 | 156792.9 | 18590.37 | 88.7705 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
0 | 0 | 70.5 | 7224.38 | 93764.6 | 62.77 | 0 | 0 | 30.5 | 18420.15 | 155596.7 | 18447.37 | 88.0877 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
0 | 0 | 71.5 | 7169.25 | 93049 | 62.33 | 0 | 0 | 31.5 | 18285.41 | 154458.5 | 18311.36 | 87.4382 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
0 | 0 | 72.5 | 7114.12 | 92333.5 | 61.9 | 0 | 0 | 32.5 | 18150.66 | 153320.2 | 18175.34 | 86.7887 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
0 | 0 | 73.5 | 7058.99 | 91618 | 61.46 | 0 | 0 | 33.5 | 18015.91 | 152182 | 18039.33 | 86.1393 | Scan type: | Beam |  |  |  |  | Scan type: | Beam
0 | 0 | 74.5 | 7003.86 | 90902.5 | 61.03 | 0 | 0 | 34.5 | 17881.17 | 151043.8 | 17903.31 | 85.4898 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
0 | 0 | 75.5 | 6949.8 | 90200.9 | 60.56 | 0 | 0 | 35.5 | 17747.41 | 149914 | 17768.93 | 84.8481 | Wedge: | none |  |  |  |  | Wedge: | none
0 | 0 | 76.5 | 6896.93 | 89514.6 | 60.06 | 0 | 0 | 36.5 | 17614.23 | 148789 | 17635.5 | 84.211
0 | 0 | 77.5 | 6844.05 | 88828.4 | 59.56 | 0 | 0 | 37.5 | 17481.06 | 147664.1 | 17502.07 | 83.5738 | Points [mm]: | [5 empty] | Points [mm]:
0 | 0 | 78.5 | 6791.18 | 88142.1 | 59.06 | 0 | 0 | 38.5 | 17347.88 | 146539.1 | 17368.64 | 82.9367 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 79.5 | 6738.3 | 87455.8 | 58.56 | 0 | 0 | 39.5 | 17214.7 | 145414.1 | 17235.21 | 82.2995 | 0 | 0 | -0.5 | 12354.2 | 104357 | 55.47 | 0 | 0 | -0.5 | 12697.36 | 107255.7 | 55.94
0 | 0 | 80.5 | 6687.77 | 86800 | 58.12 | 0 | 0 | 40.5 | 17082.93 | 144301 | 17102.62 | 81.6664 | 0 | 0 | 0.5 | 12388.19 | 104644.2 | 55.62 | 0 | 0 | 0.5 | 12738.08 | 107599.7 | 56.15
0 | 0 | 81.5 | 6639.91 | 86178.9 | 57.75 | 0 | 0 | 41.5 | 16952.06 | 143195.5 | 16970.57 | 81.0359 | 0 | 0 | 1.5 | 13436.01 | 113495.2 | 60.33 | 0 | 0 | 1.5 | 13653.26 | 115330.3 | 60.18
0 | 0 | 82.5 | 6592.06 | 85557.8 | 57.38 | 0 | 0 | 42.5 | 16821.19 | 142090.1 | 16838.52 | 80.4053 | 0 | 0 | 2.5 | 15588.57 | 131678 | 70.02 | 0 | 0 | 2.5 | 15792.51 | 133400.7 | 69.59
0 | 0 | 83.5 | 6544.2 | 84936.7 | 57.01 | 0 | 0 | 43.5 | 16690.32 | 140984.6 | 16706.47 | 79.7748 | 0 | 0 | 3.5 | 17566.16 | 148382.9 | 78.92 | 0 | 0 | 3.5 | 17856.88 | 150838.7 | 78.67
0 | 0 | 84.5 | 6496.35 | 84315.6 | 56.65 | 0 | 0 | 44.5 | 16559.45 | 139879.2 | 16574.42 | 79.1442 | 0 | 0 | 4.5 | 18988.01 | 160393.4 | 85.29 | 0 | 0 | 4.5 | 19292.82 | 162968.1 | 84.99
0 | 0 | 85.5 | 6447.58 | 83682.6 | 56.26 | 0 | 0 | 45.5 | 16432.47 | 138806.5 | 16446.47 | 78.5332 | 0 | 0 | 5.5 | 20029.46 | 169190.6 | 89.95 | 0 | 0 | 5.5 | 20324.44 | 171682.4 | 89.54
0 | 0 | 86.5 | 6397.6 | 83034 | 55.85 | 0 | 0 | 46.5 | 16308.33 | 137757.9 | 16321.53 | 77.9366 | 0 | 0 | 6.5 | 20797.23 | 175676.1 | 93.38 | 0 | 0 | 6.5 | 21075.06 | 178022.9 | 92.87
0 | 0 | 87.5 | 6347.63 | 82385.3 | 55.44 | 0 | 0 | 47.5 | 16184.19 | 136709.3 | 16196.59 | 77.34 | 0 | 0 | 7.5 | 21338.76 | 180250.5 | 95.76 | 0 | 0 | 7.5 | 21632.81 | 182734.3 | 95.31
0 | 0 | 88.5 | 6297.65 | 81736.7 | 55.04 | 0 | 0 | 48.5 | 16060.06 | 135660.7 | 16071.64 | 76.7434 | 0 | 0 | 8.5 | 21721.9 | 183486.8 | 97.48 | 0 | 0 | 8.5 | 22024.87 | 186046.1 | 97.03
0 | 0 | 89.5 | 6247.68 | 81088.1 | 54.63 | 0 | 0 | 49.5 | 15935.92 | 134612.1 | 15946.7 | 76.1468 | 0 | 0 | 9.5 | 21957.93 | 185480.6 | 98.59 | 0 | 0 | 9.5 | 22280.68 | 188206.9 | 98.15
0 | 0 | 90.5 | 6200.08 | 80470.4 | 54.22 | 0 | 0 | 50.5 | 15816.21 | 133601 | 15825.05 | 75.5659 | 0 | 0 | 10.5 | 22119.41 | 186844.7 | 99.34 | 0 | 0 | 10.5 | 22478.2 | 189875.3 | 99.02
0 | 0 | 91.5 | 6155.36 | 79889.9 | 53.82 | 0 | 0 | 51.5 | 15700 | 132619.3 | 15706 | 74.9974 | 0 | 0 | 11.5 | 22239.98 | 187863.1 | 99.85 | 0 | 0 | 11.5 | 22594.2 | 190855.2 | 99.58
0 | 0 | 92.5 | 6110.64 | 79309.5 | 53.41 | 0 | 0 | 52.5 | 15583.79 | 131637.7 | 15586.95 | 74.429 | 0 | 0 | 12.5 | 22269.4 | 188111.6 | 100 | 0 | 0 | 12.5 | 22660.77 | 191417.5 | 99.86
0 | 0 | 93.5 | 6065.91 | 78729 | 53.01 | 0 | 0 | 53.5 | 15467.57 | 130656 | 15467.9 | 73.8605 | 0 | 0 | 13.5 | 22242.22 | 187882 | 99.92 | 0 | 0 | 13.5 | 22701.64 | 191762.8 | 100
0 | 0 | 94.5 | 6021.19 | 78148.5 | 52.61 | 0 | 0 | 54.5 | 15351.36 | 129674.3 | 15348.85 | 73.292 | 0 | 0 | 14.5 | 22186.64 | 187412.5 | 99.66 | 0 | 0 | 14.5 | 22684.05 | 191614.2 | 99.93
0 | 0 | 95.5 | 5975.98 | 77561.8 | 52.19 | 0 | 0 | 55.5 | 15237.48 | 128712.3 | 15234.93 | 72.748 | 0 | 0 | 15.5 | 22111.47 | 186777.6 | 99.34 | 0 | 0 | 15.5 | 22597.19 | 190880.5 | 99.57
0 | 0 | 96.5 | 5930.16 | 76967.1 | 51.77 | 0 | 0 | 56.5 | 15125.63 | 127767.6 | 15125.51 | 72.2255 | 0 | 0 | 16.5 | 22020.17 | 186006.3 | 98.92 | 0 | 0 | 16.5 | 22517.13 | 190204.2 | 99.17
0 | 0 | 97.5 | 5884.34 | 76372.4 | 51.34 | 0 | 0 | 57.5 | 15013.79 | 126822.8 | 15016.08 | 71.703 | 0 | 0 | 17.5 | 21929.53 | 185240.7 | 98.46 | 0 | 0 | 17.5 | 22445.31 | 189597.5 | 98.86
0 | 0 | 98.5 | 5838.52 | 75777.7 | 50.91 | 0 | 0 | 58.5 | 14901.94 | 125878.1 | 14906.65 | 71.1805 | 0 | 0 | 18.5 | 21833.06 | 184425.8 | 98.03 | 0 | 0 | 18.5 | 22348.7 | 188781.4 | 98.47
0 | 0 | 99.5 | 5792.7 | 75183 | 50.49 | 0 | 0 | 59.5 | 14790.1 | 124933.3 | 14797.23 | 70.658 | 0 | 0 | 19.5 | 21711.76 | 183401.2 | 97.51 | 0 | 0 | 19.5 | 22237.65 | 187843.4 | 97.94
0 | 0 | 100.5 | 5749.38 | 74620.8 | 50.09 | 0 | 0 | 60.5 | 14676.58 | 123974.4 | 14684.3 | 70.1187 | 0 | 0 | 20.5 | 21572.79 | 182227.3 | 96.88 | 0 | 0 | 20.5 | 22109.46 | 186760.6 | 97.37
0 | 0 | 101.5 | 5709.12 | 74098.2 | 49.72 | 0 | 0 | 61.5 | 14561.53 | 123002.5 | 14568.19 | 69.5643 | 0 | 0 | 21.5 | 21414.72 | 180892.1 | 96.17 | 0 | 0 | 21.5 | 21970.04 | 185582.9 | 96.79
0 | 0 | 102.5 | 5668.86 | 73575.6 | 49.36 | 0 | 0 | 62.5 | 14446.48 | 122030.7 | 14452.09 | 69.0099 | 0 | 0 | 22.5 | 21273.03 | 179695.2 | 95.53 | 0 | 0 | 22.5 | 21808.11 | 184215 | 96.1
0 | 0 | 103.5 | 5628.59 | 73053 | 49 | 0 | 0 | 63.5 | 14331.43 | 121058.9 | 14335.98 | 68.4555 | 0 | 0 | 23.5 | 21145.98 | 178622 | 94.88 | 0 | 0 | 23.5 | 21693.25 | 183244.8 | 95.56
0 | 0 | 104.5 | 5588.33 | 72530.4 | 48.63 | 0 | 0 | 64.5 | 14216.38 | 120087.1 | 14219.88 | 67.9011 | 0 | 0 | 24.5 | 20995.42 | 177350.2 | 94.23 | 0 | 0 | 24.5 | 21560.06 | 182119.8 | 94.98
0 | 0 | 105.5 | 5547.83 | 72004.9 | 48.28 | 0 | 0 | 65.5 | 14110.53 | 119192.9 | 14112.79 | 67.3897 | 0 | 0 | 25.5 | 20842.03 | 176054.5 | 93.6 | 0 | 0 | 25.5 | 21391.93 | 180699.6 | 94.23
0 | 0 | 106.5 | 5507.05 | 71475.6 | 47.94 | 0 | 0 | 66.5 | 14013.42 | 118372.6 | 14014.31 | 66.9195 | 0 | 0 | 26.5 | 20695.28 | 174814.9 | 92.93 | 0 | 0 | 26.5 | 21260.44 | 179588.8 | 93.64
0 | 0 | 107.5 | 5466.27 | 70946.3 | 47.61 | 0 | 0 | 67.5 | 13916.32 | 117552.4 | 13915.82 | 66.4492 | 0 | 0 | 27.5 | 20548.53 | 173575.3 | 92.27 | 0 | 0 | 27.5 | 21128.94 | 178478.1 | 93.05
0 | 0 | 108.5 | 5425.49 | 70417 | 47.27 | 0 | 0 | 68.5 | 13819.22 | 116732.1 | 13817.33 | 65.9789 | 0 | 0 | 28.5 | 20401.78 | 172335.7 | 91.6 | 0 | 0 | 28.5 | 20997.45 | 177367.3 | 92.46
0 | 0 | 109.5 | 5384.71 | 69887.7 | 46.93 | 0 | 0 | 69.5 | 13722.11 | 115911.9 | 13718.85 | 65.5086 | 0 | 0 | 29.5 | 20255.03 | 171096.1 | 90.94 | 0 | 0 | 29.5 | 20865.95 | 176256.6 | 91.87
0 | 0 | 110.5 | 5345.87 | 69383.6 | 46.61 | 0 | 0 | 70.5 | 13622.77 | 115072.8 | 13618.76 | 65.0307 | 0 | 0 | 30.5 | 20111.54 | 169884 | 90.29 | 0 | 0 | 30.5 | 20727.66 | 175088.4 | 91.27
0 | 0 | 111.5 | 5309.54 | 68912.1 | 46.31 | 0 | 0 | 71.5 | 13520.98 | 114213 | 13516.91 | 64.5443 | 0 | 0 | 31.5 | 19970.1 | 168689.2 | 89.65 | 0 | 0 | 31.5 | 20585.33 | 173886.1 | 90.65
0 | 0 | 112.5 | 5273.21 | 68440.6 | 46 | 0 | 0 | 72.5 | 13419.2 | 113353.2 | 13415.07 | 64.0581 | 0 | 0 | 32.5 | 19828.66 | 167494.5 | 89.01 | 0 | 0 | 32.5 | 20443 | 172683.8 | 90.03
0 | 0 | 113.5 | 5236.88 | 67969 | 45.7 | 0 | 0 | 73.5 | 13317.41 | 112493.4 | 13313.23 | 63.5718 | 0 | 0 | 33.5 | 19687.22 | 166299.7 | 88.36 | 0 | 0 | 33.5 | 20300.66 | 171481.5 | 89.41
0 | 0 | 114.5 | 5200.55 | 67497.5 | 45.4 | 0 | 0 | 74.5 | 13215.63 | 111633.6 | 13211.38 | 63.0854 | 0 | 0 | 34.5 | 19545.78 | 165105 | 87.72 | 0 | 0 | 34.5 | 20158.33 | 170279.2 | 88.79
0 | 0 | 115.5 | 5163.25 | 67013.4 | 45.07 | 0 | 0 | 75.5 | 13116.42 | 110795.6 | 13113.65 | 62.6188 | 0 | 0 | 35.5 | 19407.06 | 163933.2 | 87.1 | 0 | 0 | 35.5 | 20019.51 | 169106.6 | 88.18
0 | 0 | 116.5 | 5124.64 | 66512.2 | 44.72 | 0 | 0 | 76.5 | 13019.93 | 109980.5 | 13020.23 | 62.1727 | 0 | 0 | 36.5 | 19269.98 | 162775.3 | 86.49 | 0 | 0 | 36.5 | 19882.9 | 167952.7 | 87.57
0 | 0 | 117.5 | 5086.02 | 66011 | 44.37 | 0 | 0 | 77.5 | 12923.44 | 109165.4 | 12926.81 | 61.7266 | 0 | 0 | 37.5 | 19132.9 | 161617.3 | 85.88 | 0 | 0 | 37.5 | 19746.3 | 166798.7 | 86.96
0 | 0 | 118.5 | 5047.4 | 65509.8 | 44.01 | 0 | 0 | 78.5 | 12826.95 | 108350.4 | 12833.39 | 61.2805 | 0 | 0 | 38.5 | 18995.82 | 160459.4 | 85.26 | 0 | 0 | 38.5 | 19609.69 | 165644.8 | 86.35
0 | 0 | 119.5 | 5008.79 | 65008.7 | 43.66 | 0 | 0 | 79.5 | 12730.46 | 107535.3 | 12739.98 | 60.8344 | 0 | 0 | 39.5 | 18858.74 | 159301.5 | 84.65 | 0 | 0 | 39.5 | 19473.08 | 164490.9 | 85.74
0 | 0 | 120.5 | 4971.47 | 64524.3 | 43.33 | 0 | 0 | 80.5 | 12635.99 | 106737.4 | 12646.53 | 60.3882 | 0 | 0 | 40.5 | 18720.38 | 158132.7 | 84.03 | 0 | 0 | 40.5 | 19339.98 | 163366.5 | 85.16
0 | 0 | 121.5 | 4935.74 | 64060.6 | 43.03 | 0 | 0 | 81.5 | 12543.94 | 105959.8 | 12553.04 | 59.9418 | 0 | 0 | 41.5 | 18581.14 | 156956.6 | 83.41 | 0 | 0 | 41.5 | 19209.36 | 162263.1 | 84.59
0 | 0 | 122.5 | 4900.01 | 63596.9 | 42.72 | 0 | 0 | 82.5 | 12451.89 | 105182.3 | 12459.56 | 59.4954 | 0 | 0 | 42.5 | 18441.9 | 155780.4 | 82.79 | 0 | 0 | 42.5 | 19078.73 | 161159.8 | 84.02
0 | 0 | 123.5 | 4864.28 | 63133.1 | 42.42 | 0 | 0 | 83.5 | 12359.84 | 104404.7 | 12366.07 | 59.049 | 0 | 0 | 43.5 | 18302.67 | 154604.3 | 82.16 | 0 | 0 | 43.5 | 18948.11 | 160056.4 | 83.45
0 | 0 | 124.5 | 4828.56 | 62669.4 | 42.12 | 0 | 0 | 84.5 | 12267.8 | 103627.2 | 12272.59 | 58.6026 | 0 | 0 | 44.5 | 18163.43 | 153428.1 | 81.54 | 0 | 0 | 44.5 | 18817.49 | 158953.1 | 82.88
0 | 0 | 125.5 | 4792.66 | 62203.6 | 41.81 | 0 | 0 | 85.5 | 12177.58 | 102865.1 | 12182.87 | 58.1742 | 0 | 0 | 45.5 | 18030.28 | 152303.4 | 80.94 | 0 | 0 | 45.5 | 18683.32 | 157819.7 | 82.29
0 | 0 | 126.5 | 4756.55 | 61734.8 | 41.49 | 0 | 0 | 86.5 | 12089.19 | 102118.4 | 12096.9 | 57.7637 | 0 | 0 | 46.5 | 17901.42 | 151214.9 | 80.36 | 0 | 0 | 46.5 | 18546.1 | 156660.5 | 81.69
0 | 0 | 127.5 | 4720.43 | 61266.1 | 41.17 | 0 | 0 | 87.5 | 12000.79 | 101371.8 | 12010.92 | 57.3531 | 0 | 0 | 47.5 | 17772.56 | 150126.4 | 79.78 | 0 | 0 | 47.5 | 18408.87 | 155501.4 | 81.09
0 | 0 | 128.5 | 4684.31 | 60797.3 | 40.85 | 0 | 0 | 88.5 | 11912.4 | 100625.1 | 11924.95 | 56.9426 | 0 | 0 | 48.5 | 17643.7 | 149037.9 | 79.21 | 0 | 0 | 48.5 | 18271.65 | 154342.3 | 80.49
0 | 0 | 129.5 | 4648.2 | 60328.6 | 40.53 | 0 | 0 | 89.5 | 11824.01 | 99878.4 | 11838.97 | 56.532 | 0 | 0 | 49.5 | 17514.84 | 147949.4 | 78.63 | 0 | 0 | 49.5 | 18134.43 | 153183.1 | 79.89
0 | 0 | 130.5 | 4614.38 | 59889.7 | 40.23 | 0 | 0 | 90.5 | 11736.1 | 99135.9 | 11752.16 | 56.1175 | 0 | 0 | 50.5 | 17383.16 | 146837.1 | 78.04 | 0 | 0 | 50.5 | 18003.8 | 152079.7 | 79.32
0 | 0 | 131.5 | 4583.31 | 59486.4 | 39.97 | 0 | 0 | 91.5 | 11648.76 | 98398.1 | 11664.37 | 55.6983 | 0 | 0 | 51.5 | 17249.18 | 145705.4 | 77.44 | 0 | 0 | 51.5 | 17878.15 | 151018.3 | 78.75
0 | 0 | 132.5 | 4552.23 | 59083.1 | 39.72 | 0 | 0 | 92.5 | 11561.42 | 97660.4 | 11576.57 | 55.2791 | 0 | 0 | 52.5 | 17115.2 | 144573.6 | 76.84 | 0 | 0 | 52.5 | 17752.5 | 149956.9 | 78.19
0 | 0 | 133.5 | 4521.16 | 58679.7 | 39.46 | 0 | 0 | 93.5 | 11474.09 | 96922.6 | 11488.77 | 54.8598 | 0 | 0 | 53.5 | 16981.22 | 143441.9 | 76.24 | 0 | 0 | 53.5 | 17626.85 | 148895.6 | 77.63
0 | 0 | 134.5 | 4490.08 | 58276.4 | 39.2 | 0 | 0 | 94.5 | 11386.75 | 96184.9 | 11400.97 | 54.4406 | 0 | 0 | 54.5 | 16847.24 | 142310.1 | 75.64 | 0 | 0 | 54.5 | 17501.2 | 147834.2 | 77.06
0 | 0 | 135.5 | 4458.11 | 57861.4 | 38.91 | 0 | 0 | 95.5 | 11302.62 | 95474.3 | 11315.97 | 54.0347 | 0 | 0 | 55.5 | 16721.45 | 141247.6 | 75.07 | 0 | 0 | 55.5 | 17374.62 | 146765 | 76.5
0 | 0 | 136.5 | 4424.98 | 57431.5 | 38.59 | 0 | 0 | 96.5 | 11222.67 | 94798.9 | 11234.59 | 53.6461 | 0 | 0 | 56.5 | 16603.45 | 140250.8 | 74.54 | 0 | 0 | 56.5 | 17247.22 | 145688.8 | 75.95
0 | 0 | 137.5 | 4391.86 | 57001.6 | 38.27 | 0 | 0 | 97.5 | 11142.71 | 94123.5 | 11153.21 | 53.2575 | 0 | 0 | 57.5 | 16485.44 | 139254 | 74.01 | 0 | 0 | 57.5 | 17119.82 | 144612.6 | 75.39
0 | 0 | 138.5 | 4358.73 | 56571.7 | 37.95 | 0 | 0 | 98.5 | 11062.76 | 93448.1 | 11071.84 | 52.8689 | 0 | 0 | 58.5 | 16367.44 | 138257.2 | 73.48 | 0 | 0 | 58.5 | 16992.41 | 143536.4 | 74.84
0 | 0 | 139.5 | 4325.61 | 56141.8 | 37.63 | 0 | 0 | 99.5 | 10982.81 | 92772.7 | 10990.46 | 52.4803 | 0 | 0 | 59.5 | 16249.44 | 137260.4 | 72.95 | 0 | 0 | 59.5 | 16865.01 | 142460.2 | 74.28
0 | 0 | 140.5 | 4294.54 | 55738.4 | 37.34 | 0 | 0 | 100.5 | 10905.35 | 92118.5 | 10911.69 | 52.1042 | 0 | 0 | 60.5 | 16132.23 | 136270.4 | 72.42 | 0 | 0 | 60.5 | 16745.29 | 141448.9 | 73.75
0 | 0 | 141.5 | 4266.02 | 55368.4 | 37.08 | 0 | 0 | 101.5 | 10831.34 | 91493.3 | 10836.52 | 51.7453 | 0 | 0 | 61.5 | 16015.76 | 135286.6 | 71.91 | 0 | 0 | 61.5 | 16632.75 | 140498.3 | 73.24
0 | 0 | 142.5 | 4237.51 | 54998.3 | 36.82 | 0 | 0 | 102.5 | 10757.32 | 90868.1 | 10761.34 | 51.3863 | 0 | 0 | 62.5 | 15899.3 | 134302.8 | 71.39 | 0 | 0 | 62.5 | 16520.21 | 139547.7 | 72.73
0 | 0 | 143.5 | 4209 | 54628.2 | 36.57 | 0 | 0 | 103.5 | 10683.31 | 90242.8 | 10686.17 | 51.0273 | 0 | 0 | 63.5 | 15782.83 | 133319 | 70.87 | 0 | 0 | 63.5 | 16407.67 | 138597.1 | 72.22
0 | 0 | 144.5 | 4180.49 | 54258.2 | 36.31 | 0 | 0 | 104.5 | 10609.29 | 89617.6 | 10610.99 | 50.6683 | 0 | 0 | 64.5 | 15666.36 | 132335.2 | 70.36 | 0 | 0 | 64.5 | 16295.14 | 137646.5 | 71.71
0 | 0 | 145.5 | 4150.91 | 53874.3 | 36.06 | 0 | 0 | 105.5 | 10535.04 | 88990.4 | 10536.73 | 50.3137 | 0 | 0 | 65.5 | 15552.03 | 131369.4 | 69.85 | 0 | 0 | 65.5 | 16181.15 | 136683.6 | 71.21
0 | 0 | 146.5 | 4120.03 | 53473.5 | 35.82 | 0 | 0 | 106.5 | 10460.48 | 88360.6 | 10463.65 | 49.9648 | 0 | 0 | 66.5 | 15439.8 | 130421.4 | 69.34 | 0 | 0 | 66.5 | 16065.66 | 135708 | 70.71
0 | 0 | 147.5 | 4089.15 | 53072.8 | 35.58 | 0 | 0 | 107.5 | 10385.92 | 87730.8 | 10390.56 | 49.6158 | 0 | 0 | 67.5 | 15327.58 | 129473.4 | 68.84 | 0 | 0 | 67.5 | 15950.17 | 134732.5 | 70.21
0 | 0 | 148.5 | 4058.28 | 52672 | 35.34 | 0 | 0 | 108.5 | 10311.36 | 87101 | 10317.48 | 49.2668 | 0 | 0 | 68.5 | 15215.36 | 128525.5 | 68.33 | 0 | 0 | 68.5 | 15834.67 | 133756.9 | 69.71
0 | 0 | 149.5 | 4027.4 | 52271.3 | 35.1 | 0 | 0 | 109.5 | 10236.8 | 86471.2 | 10244.4 | 48.9178 | 0 | 0 | 69.5 | 15103.13 | 127577.5 | 67.83 | 0 | 0 | 69.5 | 15719.18 | 132781.3 | 69.21
0 | 0 | 150.5 | 3998.34 | 51894.2 | 34.85 | 0 | 0 | 110.5 | 10162.48 | 85843.4 | 10170.29 | 48.564 | 0 | 0 | 70.5 | 14994.13 | 126656.7 | 67.33 | 0 | 0 | 70.5 | 15604.57 | 131813.2 | 68.72
0 | 0 | 151.5 | 3971.28 | 51543 | 34.6 | 0 | 0 | 111.5 | 10088.41 | 85217.7 | 10095.06 | 48.2047 | 0 | 0 | 71.5 | 14888.34 | 125763.1 | 66.84 | 0 | 0 | 71.5 | 15490.85 | 130852.6 | 68.22
0 | 0 | 152.5 | 3944.22 | 51191.7 | 34.34 | 0 | 0 | 112.5 | 10014.35 | 84592.1 | 10019.83 | 47.8455 | 0 | 0 | 72.5 | 14782.55 | 124869.6 | 66.36 | 0 | 0 | 72.5 | 15377.12 | 129891.9 | 67.73
0 | 0 | 153.5 | 3917.16 | 50840.5 | 34.09 | 0 | 0 | 113.5 | 9940.29 | 83966.5 | 9944.6 | 47.4863 | 0 | 0 | 73.5 | 14676.77 | 123976 | 65.87 | 0 | 0 | 73.5 | 15263.39 | 128931.2 | 67.24
0 | 0 | 154.5 | 3890.1 | 50489.3 | 33.84 | 0 | 0 | 114.5 | 9866.22 | 83340.9 | 9869.37 | 47.127 | 0 | 0 | 74.5 | 14570.98 | 123082.4 | 65.39 | 0 | 0 | 74.5 | 15149.66 | 127970.5 | 66.74
0 | 0 | 155.5 | 3862.63 | 50132.8 | 33.59 | 0 | 0 | 115.5 | 9792.91 | 82721.6 | 9795.77 | 46.7756 | 0 | 0 | 75.5 | 14465.3 | 122189.7 | 64.91 | 0 | 0 | 75.5 | 15039.14 | 127037 | 66.26
0 | 0 | 156.5 | 3834.61 | 49769.1 | 33.35 | 0 | 0 | 116.5 | 9720.67 | 82111.4 | 9724.48 | 46.4352 | 0 | 0 | 76.5 | 14359.74 | 121298 | 64.45 | 0 | 0 | 76.5 | 14932.01 | 126132 | 65.78
0 | 0 | 157.5 | 3806.59 | 49405.4 | 33.12 | 0 | 0 | 117.5 | 9648.44 | 81501.2 | 9653.2 | 46.0948 | 0 | 0 | 77.5 | 14254.18 | 120406.3 | 63.99 | 0 | 0 | 77.5 | 14824.87 | 125227 | 65.31
0 | 0 | 158.5 | 3778.57 | 49041.7 | 32.88 | 0 | 0 | 118.5 | 9576.2 | 80891 | 9581.92 | 45.7545 | 0 | 0 | 78.5 | 14148.62 | 119514.6 | 63.53 | 0 | 0 | 78.5 | 14717.74 | 124322.1 | 64.83
0 | 0 | 159.5 | 3750.55 | 48678 | 32.65 | 0 | 0 | 119.5 | 9503.96 | 80280.8 | 9510.63 | 45.414 | 0 | 0 | 79.5 | 14043.06 | 118623 | 63.07 | 0 | 0 | 79.5 | 14610.6 | 123417.1 | 64.36
0 | 0 | 160.5 | 3723.99 | 48333.4 | 32.43 | 0 | 0 | 120.5 | 9436.32 | 79709.4 | 9443.63 | 45.0941 | 0 | 0 | 80.5 | 13939.59 | 117748.9 | 62.61 | 0 | 0 | 80.5 | 14506.01 | 122533.6 | 63.89
0 | 0 | 161.5 | 3699.24 | 48012.1 | 32.24 | 0 | 0 | 121.5 | 9375.07 | 79192.1 | 9382.59 | 44.8026 | 0 | 0 | 81.5 | 13838.6 | 116895.9 | 62.14 | 0 | 0 | 81.5 | 14404.63 | 121677.2 | 63.44
0 | 0 | 162.5 | 3674.49 | 47690.9 | 32.05 | 0 | 0 | 122.5 | 9313.83 | 78674.7 | 9321.55 | 44.5112 | 0 | 0 | 82.5 | 13737.62 | 116042.9 | 61.68 | 0 | 0 | 82.5 | 14303.24 | 120820.8 | 62.99
0 | 0 | 163.5 | 3649.74 | 47369.6 | 31.85 | 0 | 0 | 123.5 | 9252.58 | 78157.4 | 9260.52 | 44.2197 | 0 | 0 | 83.5 | 13636.64 | 115189.9 | 61.22 | 0 | 0 | 83.5 | 14201.86 | 119964.4 | 62.54
0 | 0 | 164.5 | 3624.99 | 47048.4 | 31.66 | 0 | 0 | 124.5 | 9191.34 | 77640.1 | 9199.48 | 43.9283 | 0 | 0 | 84.5 | 13535.66 | 114336.9 | 60.76 | 0 | 0 | 84.5 | 14100.47 | 119108 | 62.09
0 | 0 | 165.5 | 3599.69 | 46720.1 | 31.46 | 0 | 0 | 125.5 | 9127.12 | 77097.6 | 9136.17 | 43.626 | 0 | 0 | 85.5 | 13434.38 | 113481.4 | 60.3 | 0 | 0 | 85.5 | 14001.96 | 118275.8 | 61.65
0 | 0 | 166.5 | 3573.62 | 46381.8 | 31.25 | 0 | 0 | 126.5 | 9058.93 | 76521.6 | 9069.82 | 43.3091 | 0 | 0 | 86.5 | 13332.78 | 112623.2 | 59.84 | 0 | 0 | 86.5 | 13906.6 | 117470.3 | 61.22
0 | 0 | 167.5 | 3547.55 | 46043.4 | 31.03 | 0 | 0 | 127.5 | 8990.74 | 75945.6 | 9003.48 | 42.9924 | 0 | 0 | 87.5 | 13231.17 | 111764.9 | 59.38 | 0 | 0 | 87.5 | 13811.24 | 116664.8 | 60.79
0 | 0 | 168.5 | 3521.49 | 45705.1 | 30.82 | 0 | 0 | 128.5 | 8922.55 | 75369.6 | 8937.13 | 42.6755 | 0 | 0 | 88.5 | 13129.57 | 110906.7 | 58.93 | 0 | 0 | 88.5 | 13715.88 | 115859.3 | 60.36
0 | 0 | 169.5 | 3495.42 | 45366.8 | 30.6 | 0 | 0 | 129.5 | 8854.36 | 74793.6 | 8870.79 | 42.3587 | 0 | 0 | 89.5 | 13027.97 | 110048.4 | 58.47 | 0 | 0 | 89.5 | 13620.52 | 115053.8 | 59.93
0 | 0 | 170.5 | 3470.85 | 45047.8 | 30.39 | 0 | 0 | 130.5 | 8788.49 | 74237.1 | 8804.98 | 42.0445 | 0 | 0 | 90.5 | 12931.4 | 109232.7 | 58.04 | 0 | 0 | 90.5 | 13525.44 | 114250.6 | 59.51
0 | 0 | 171.5 | 3448.12 | 44752.8 | 30.19 | 0 | 0 | 131.5 | 8725.38 | 73704.1 | 8739.81 | 41.7333 | 0 | 0 | 91.5 | 12841.66 | 108474.6 | 57.63 | 0 | 0 | 91.5 | 13430.73 | 113450.6 | 59.09
0 | 0 | 172.5 | 3425.39 | 44457.9 | 29.99 | 0 | 0 | 132.5 | 8662.27 | 73171 | 8674.65 | 41.4222 | 0 | 0 | 92.5 | 12751.91 | 107716.6 | 57.22 | 0 | 0 | 92.5 | 13336.02 | 112650.6 | 58.68
0 | 0 | 173.5 | 3402.66 | 44162.9 | 29.78 | 0 | 0 | 133.5 | 8599.16 | 72637.9 | 8609.48 | 41.111 | 0 | 0 | 93.5 | 12662.17 | 106958.5 | 56.81 | 0 | 0 | 93.5 | 13241.32 | 111850.6 | 58.27
0 | 0 | 174.5 | 3379.93 | 43867.9 | 29.58 | 0 | 0 | 134.5 | 8536.05 | 72104.8 | 8544.32 | 40.7998 | 0 | 0 | 94.5 | 12572.43 | 106200.5 | 56.4 | 0 | 0 | 94.5 | 13146.61 | 111050.6 | 57.85
0 | 0 | 175.5 | 3356.26 | 43560.6 | 29.37 | 0 | 0 | 135.5 | 8477.3 | 71608.5 | 8484.25 | 40.513 | 0 | 0 | 95.5 | 12482.49 | 105440.7 | 56 | 0 | 0 | 95.5 | 13051.3 | 110245.5 | 57.43
0 | 0 | 176.5 | 3331.45 | 43238.6 | 29.15 | 0 | 0 | 136.5 | 8424.23 | 71160.2 | 8430.85 | 40.258 | 0 | 0 | 96.5 | 12392.3 | 104678.8 | 55.6 | 0 | 0 | 96.5 | 12955.15 | 109433.3 | 57.01
0 | 0 | 177.5 | 3306.63 | 42916.5 | 28.93 | 0 | 0 | 137.5 | 8371.17 | 70712 | 8377.45 | 40.003 | 0 | 0 | 97.5 | 12302.1 | 103916.9 | 55.21 | 0 | 0 | 97.5 | 12859 | 108621.1 | 56.59
0 | 0 | 178.5 | 3281.82 | 42594.4 | 28.71 | 0 | 0 | 138.5 | 8318.1 | 70263.8 | 8324.05 | 39.748 | 0 | 0 | 98.5 | 12211.9 | 103155 | 54.81 | 0 | 0 | 98.5 | 12762.85 | 107808.9 | 56.16
0 | 0 | 179.5 | 3257 | 42272.4 | 28.49 | 0 | 0 | 139.5 | 8265.04 | 69815.5 | 8270.65 | 39.493 | 0 | 0 | 99.5 | 12121.7 | 102393.1 | 54.41 | 0 | 0 | 99.5 | 12666.7 | 106996.8 | 55.74
0 | 0 | 180.5 | 3233.13 | 41962.5 | 28.27 | 0 | 0 | 140.5 | 8207.24 | 69327.3 | 8213.1 | 39.2182 | 0 | 0 | 100.5 | 12037.32 | 101680.3 | 54.04 | 0 | 0 | 100.5 | 12575.83 | 106229.2 | 55.34
0 | 0 | 181.5 | 3210.42 | 41667.7 | 28.06 | 0 | 0 | 141.5 | 8143.75 | 68791 | 8150.56 | 38.9196 | 0 | 0 | 101.5 | 11961.28 | 101038 | 53.69 | 0 | 0 | 101.5 | 12491.93 | 105520.5 | 54.97
0 | 0 | 182.5 | 3187.71 | 41373 | 27.85 | 0 | 0 | 142.5 | 8080.27 | 68254.7 | 8088.03 | 38.621 | 0 | 0 | 102.5 | 11885.24 | 100395.7 | 53.34 | 0 | 0 | 102.5 | 12408.03 | 104811.7 | 54.59
0 | 0 | 183.5 | 3165 | 41078.2 | 27.64 | 0 | 0 | 143.5 | 8016.78 | 67718.5 | 8025.49 | 38.3224 | 0 | 0 | 103.5 | 11809.2 | 99753.4 | 53 | 0 | 0 | 103.5 | 12324.13 | 104103 | 54.22
0 | 0 | 184.5 | 3142.29 | 40783.5 | 27.43 | 0 | 0 | 144.5 | 7953.29 | 67182.2 | 7962.95 | 38.0237 | 0 | 0 | 104.5 | 11733.16 | 99111 | 52.65 | 0 | 0 | 104.5 | 12240.23 | 103394.3 | 53.85
0 | 0 | 185.5 | 3121.81 | 40517.8 | 27.25 | 0 | 0 | 145.5 | 7894.38 | 66684.6 | 7905.12 | 37.7476 | 0 | 0 | 105.5 | 11650.18 | 98410.1 | 52.28 | 0 | 0 | 105.5 | 12153.58 | 102662.3 | 53.47
0 | 0 | 186.5 | 3103.71 | 40282.8 | 27.11 | 0 | 0 | 146.5 | 7841.3 | 66236.2 | 7853.29 | 37.5001 | 0 | 0 | 106.5 | 11558.57 | 97636.3 | 51.88 | 0 | 0 | 106.5 | 12063.37 | 101900.4 | 53.08
0 | 0 | 187.5 | 3085.6 | 40047.8 | 26.97 | 0 | 0 | 147.5 | 7788.22 | 65787.8 | 7801.45 | 37.2526 | 0 | 0 | 107.5 | 11466.97 | 96862.5 | 51.47 | 0 | 0 | 107.5 | 11973.17 | 101138.4 | 52.7
0 | 0 | 188.5 | 3067.49 | 39812.8 | 26.83 | 0 | 0 | 148.5 | 7735.14 | 65339.4 | 7749.62 | 37.0051 | 0 | 0 | 108.5 | 11375.36 | 96088.7 | 51.07 | 0 | 0 | 108.5 | 11882.96 | 100376.4 | 52.31
0 | 0 | 189.5 | 3049.39 | 39577.8 | 26.69 | 0 | 0 | 149.5 | 7682.06 | 64891.1 | 7697.78 | 36.7575 | 0 | 0 | 109.5 | 11283.76 | 95314.9 | 50.67 | 0 | 0 | 109.5 | 11792.76 | 99614.5 | 51.92
0 | 0 | 190.5 | 3030.31 | 39330.1 | 26.53 | 0 | 0 | 150.5 | 7629.55 | 64447.5 | 7645.8 | 36.5093 | 0 | 0 | 110.5 | 11199.21 | 94600.7 | 50.29 | 0 | 0 | 110.5 | 11708.06 | 98899 | 51.55
0 | 0 | 191.5 | 3009.94 | 39065.8 | 26.34 | 0 | 0 | 151.5 | 7577.77 | 64010.1 | 7593.63 | 36.2602 | 0 | 0 | 111.5 | 11123.27 | 93959.3 | 49.94 | 0 | 0 | 111.5 | 11630.47 | 98243.6 | 51.21
0 | 0 | 192.5 | 2989.58 | 38801.5 | 26.15 | 0 | 0 | 152.5 | 7525.98 | 63572.6 | 7541.47 | 36.0111 | 0 | 0 | 112.5 | 11047.33 | 93317.8 | 49.59 | 0 | 0 | 112.5 | 11552.88 | 97588.2 | 50.87
0 | 0 | 193.5 | 2969.21 | 38537.2 | 25.96 | 0 | 0 | 153.5 | 7474.19 | 63135.2 | 7489.3 | 35.762 | 0 | 0 | 113.5 | 10971.39 | 92676.3 | 49.24 | 0 | 0 | 113.5 | 11475.3 | 96932.9 | 50.53
0 | 0 | 194.5 | 2948.85 | 38272.9 | 25.77 | 0 | 0 | 154.5 | 7422.41 | 62697.7 | 7437.13 | 35.5129 | 0 | 0 | 114.5 | 10895.45 | 92034.8 | 48.89 | 0 | 0 | 114.5 | 11397.71 | 96277.5 | 50.19
0 | 0 | 195.5 | 2928.24 | 38005.4 | 25.58 | 0 | 0 | 155.5 | 7370.48 | 62259.1 | 7383.59 | 35.2572 | 0 | 0 | 115.5 | 10820.82 | 91404.5 | 48.55 | 0 | 0 | 115.5 | 11318.57 | 95608.9 | 49.84
0 | 0 | 196.5 | 2907.37 | 37734.6 | 25.38 | 0 | 0 | 156.5 | 7318.37 | 61819 | 7328.2 | 34.9928 | 0 | 0 | 116.5 | 10747.76 | 90787.3 | 48.22 | 0 | 0 | 116.5 | 11237.63 | 94925.3 | 49.47
0 | 0 | 197.5 | 2886.5 | 37463.7 | 25.19 | 0 | 0 | 157.5 | 7266.26 | 61378.8 | 7272.8 | 34.7282 | 0 | 0 | 117.5 | 10674.71 | 90170.2 | 47.9 | 0 | 0 | 117.5 | 11156.69 | 94241.6 | 49.11
0 | 0 | 198.5 | 2865.63 | 37192.8 | 25 | 0 | 0 | 158.5 | 7214.15 | 60938.6 | 7217.41 | 34.4637 | 0 | 0 | 118.5 | 10601.65 | 89553.1 | 47.57 | 0 | 0 | 118.5 | 11075.76 | 93557.9 | 48.75
0 | 0 | 199.5 | 2844.76 | 36922 | 24.8 | 0 | 0 | 159.5 | 7162.04 | 60498.4 | 7162.01 | 34.1992 | 0 | 0 | 119.5 | 10528.59 | 88935.9 | 47.25 | 0 | 0 | 119.5 | 10994.82 | 92874.2 | 48.38
0 | 0 | 200.5 | 2825.64 | 36673.7 | 24.63 | 0 | 0 | 160.5 | 7112.3 | 60078.2 | 7111.87 | 33.9598 | 0 | 0 | 120.5 | 10453.88 | 88304.9 | 46.91 | 0 | 0 | 120.5 | 10917.53 | 92221.3 | 48.04
0 | 0 | 201.5 | 2808.51 | 36451.5 | 24.5 | 0 | 0 | 161.5 | 7065.52 | 59683.1 | 7068.33 | 33.7519 | 0 | 0 | 121.5 | 10376.81 | 87653.8 | 46.57 | 0 | 0 | 121.5 | 10844.91 | 91607.9 | 47.73
0 | 0 | 202.5 | 2791.39 | 36229.2 | 24.37 | 0 | 0 | 162.5 | 7018.75 | 59288 | 7024.8 | 33.544 | 0 | 0 | 122.5 | 10299.73 | 87002.8 | 46.23 | 0 | 0 | 122.5 | 10772.3 | 90994.5 | 47.41
0 | 0 | 203.5 | 2774.26 | 36006.9 | 24.23 | 0 | 0 | 163.5 | 6971.97 | 58892.8 | 6981.26 | 33.3361 | 0 | 0 | 123.5 | 10222.66 | 86351.7 | 45.89 | 0 | 0 | 123.5 | 10699.68 | 90381.1 | 47.09
0 | 0 | 204.5 | 2757.13 | 35784.6 | 24.1 | 0 | 0 | 164.5 | 6925.19 | 58497.7 | 6937.72 | 33.1282 | 0 | 0 | 124.5 | 10145.58 | 85700.6 | 45.54 | 0 | 0 | 124.5 | 10627.06 | 89767.7 | 46.78
0 | 0 | 205.5 | 2739.01 | 35549.4 | 23.95 | 0 | 0 | 165.5 | 6879.86 | 58114.8 | 6893.73 | 32.9181 | 0 | 0 | 125.5 | 10072.62 | 85084.3 | 45.22 | 0 | 0 | 125.5 | 10553.5 | 89146.3 | 46.46
0 | 0 | 206.5 | 2719.64 | 35298 | 23.78 | 0 | 0 | 166.5 | 6836.31 | 57746.9 | 6849.17 | 32.7053 | 0 | 0 | 126.5 | 10004.94 | 84512.6 | 44.91 | 0 | 0 | 126.5 | 10478.72 | 88514.7 | 46.13
0 | 0 | 207.5 | 2700.28 | 35046.7 | 23.62 | 0 | 0 | 167.5 | 6792.75 | 57379 | 6804.62 | 32.4926 | 0 | 0 | 127.5 | 9937.25 | 83940.9 | 44.61 | 0 | 0 | 127.5 | 10403.94 | 87883 | 45.8
0 | 0 | 208.5 | 2680.91 | 34795.3 | 23.45 | 0 | 0 | 168.5 | 6749.19 | 57011 | 6760.06 | 32.2798 | 0 | 0 | 128.5 | 9869.57 | 83369.2 | 44.3 | 0 | 0 | 128.5 | 10329.16 | 87251.3 | 45.47
0 | 0 | 209.5 | 2661.54 | 34543.9 | 23.29 | 0 | 0 | 169.5 | 6705.64 | 56643.1 | 6715.5 | 32.0671 | 0 | 0 | 129.5 | 9801.89 | 82797.4 | 44 | 0 | 0 | 129.5 | 10254.37 | 86619.6 | 45.14
0 | 0 | 210.5 | 2643.26 | 34306.7 | 23.13 | 0 | 0 | 170.5 | 6660.84 | 56264.7 | 6670.59 | 31.8526 | 0 | 0 | 130.5 | 9734.32 | 82226.6 | 43.69 | 0 | 0 | 130.5 | 10180.91 | 85999 | 44.81
0 | 0 | 211.5 | 2626.36 | 34087.3 | 22.98 | 0 | 0 | 171.5 | 6614.43 | 55872.6 | 6625.2 | 31.6359 | 0 | 0 | 131.5 | 9666.89 | 81657.1 | 43.39 | 0 | 0 | 131.5 | 10109.24 | 85393.7 | 44.5
0 | 0 | 212.5 | 2609.46 | 33867.9 | 22.84 | 0 | 0 | 172.5 | 6568.01 | 55480.6 | 6579.82 | 31.4192 | 0 | 0 | 132.5 | 9599.47 | 81087.6 | 43.09 | 0 | 0 | 132.5 | 10037.57 | 84788.3 | 44.18
0 | 0 | 213.5 | 2592.56 | 33648.6 | 22.69 | 0 | 0 | 173.5 | 6521.6 | 55088.5 | 6534.43 | 31.2024 | 0 | 0 | 133.5 | 9532.04 | 80518 | 42.78 | 0 | 0 | 133.5 | 9965.9 | 84182.9 | 43.86
0 | 0 | 214.5 | 2575.65 | 33429.2 | 22.55 | 0 | 0 | 174.5 | 6475.18 | 54696.4 | 6489.05 | 30.9857 | 0 | 0 | 134.5 | 9464.62 | 79948.5 | 42.48 | 0 | 0 | 134.5 | 9894.23 | 83577.5 | 43.55
0 | 0 | 215.5 | 2558.72 | 33209.4 | 22.4 | 0 | 0 | 175.5 | 6429.98 | 54314.6 | 6444.63 | 30.7736 | 0 | 0 | 135.5 | 9398.95 | 79393.8 | 42.18 | 0 | 0 | 135.5 | 9825.91 | 83000.4 | 43.24
0 | 0 | 216.5 | 2541.74 | 32989.1 | 22.25 | 0 | 0 | 176.5 | 6386.33 | 53945.9 | 6401.46 | 30.5675 | 0 | 0 | 136.5 | 9335.33 | 78856.4 | 41.89 | 0 | 0 | 136.5 | 9762 | 82460.5 | 42.96
0 | 0 | 217.5 | 2524.76 | 32768.7 | 22.1 | 0 | 0 | 177.5 | 6342.68 | 53577.2 | 6358.28 | 30.3613 | 0 | 0 | 137.5 | 9271.72 | 78319 | 41.61 | 0 | 0 | 137.5 | 9698.09 | 81920.7 | 42.67
0 | 0 | 218.5 | 2507.79 | 32548.4 | 21.96 | 0 | 0 | 178.5 | 6299.02 | 53208.4 | 6315.1 | 30.1551 | 0 | 0 | 138.5 | 9208.1 | 77781.6 | 41.32 | 0 | 0 | 138.5 | 9634.19 | 81380.8 | 42.39
0 | 0 | 219.5 | 2490.81 | 32328 | 21.81 | 0 | 0 | 179.5 | 6255.37 | 52839.7 | 6271.92 | 29.9489 | 0 | 0 | 139.5 | 9144.48 | 77244.3 | 41.03 | 0 | 0 | 139.5 | 9570.28 | 80841 | 42.1
0 | 0 | 220.5 | 2474.24 | 32113 | 21.67 | 0 | 0 | 180.5 | 6212.16 | 52474.6 | 6229.06 | 29.7443 | 0 | 0 | 140.5 | 9079.42 | 76694.7 | 40.74 | 0 | 0 | 140.5 | 9505.54 | 80294.1 | 41.81
0 | 0 | 221.5 | 2458.22 | 31905.1 | 21.53 | 0 | 0 | 181.5 | 6169.47 | 52114.1 | 6186.58 | 29.5414 | 0 | 0 | 141.5 | 9012.41 | 76128.6 | 40.44 | 0 | 0 | 141.5 | 9439.79 | 79738.8 | 41.52
0 | 0 | 222.5 | 2442.2 | 31697.1 | 21.4 | 0 | 0 | 182.5 | 6126.78 | 51753.5 | 6144.09 | 29.3385 | 0 | 0 | 142.5 | 8945.39 | 75562.5 | 40.14 | 0 | 0 | 142.5 | 9374.05 | 79183.4 | 41.23
0 | 0 | 223.5 | 2426.18 | 31489.2 | 21.27 | 0 | 0 | 183.5 | 6084.1 | 51392.9 | 6101.61 | 29.1357 | 0 | 0 | 143.5 | 8878.37 | 74996.4 | 39.84 | 0 | 0 | 143.5 | 9308.31 | 78628.1 | 40.94
0 | 0 | 224.5 | 2410.16 | 31281.3 | 21.14 | 0 | 0 | 184.5 | 6041.41 | 51032.3 | 6059.12 | 28.9328 | 0 | 0 | 144.5 | 8811.36 | 74430.3 | 39.54 | 0 | 0 | 144.5 | 9242.56 | 78072.8 | 40.65
0 | 0 | 225.5 | 2393.86 | 31069.8 | 21 | 0 | 0 | 185.5 | 6000.31 | 50685.1 | 6017.38 | 28.7335 | 0 | 0 | 145.5 | 8748.88 | 73902.6 | 39.26 | 0 | 0 | 145.5 | 9177.34 | 77521.8 | 40.37
0 | 0 | 226.5 | 2377.25 | 30854.2 | 20.84 | 0 | 0 | 186.5 | 5961.12 | 50354.1 | 5976.54 | 28.5385 | 0 | 0 | 146.5 | 8692.44 | 73425.8 | 39.01 | 0 | 0 | 146.5 | 9112.82 | 76976.8 | 40.09
0 | 0 | 227.5 | 2360.64 | 30638.6 | 20.69 | 0 | 0 | 187.5 | 5921.93 | 50023.1 | 5935.69 | 28.3434 | 0 | 0 | 147.5 | 8635.99 | 72949 | 38.75 | 0 | 0 | 147.5 | 9048.3 | 76431.8 | 39.81
0 | 0 | 228.5 | 2344.03 | 30423 | 20.54 | 0 | 0 | 188.5 | 5882.74 | 49692.1 | 5894.85 | 28.1484 | 0 | 0 | 148.5 | 8579.55 | 72472.2 | 38.5 | 0 | 0 | 148.5 | 8983.79 | 75886.9 | 39.53
0 | 0 | 229.5 | 2327.42 | 30207.4 | 20.38 | 0 | 0 | 189.5 | 5843.56 | 49361 | 5854.01 | 27.9534 | 0 | 0 | 149.5 | 8523.1 | 71995.4 | 38.24 | 0 | 0 | 149.5 | 8919.27 | 75341.9 | 39.25
0 | 0 | 230.5 | 2311.76 | 30004.1 | 20.24 | 0 | 0 | 190.5 | 5805.14 | 49036.5 | 5814.37 | 27.7641 | 0 | 0 | 150.5 | 8466.39 | 71516.3 | 37.99 | 0 | 0 | 150.5 | 8857.78 | 74822.5 | 38.98
0 | 0 | 231.5 | 2297.24 | 29815.6 | 20.12 | 0 | 0 | 191.5 | 5767.72 | 48720.5 | 5776.28 | 27.5822 | 0 | 0 | 151.5 | 8409.33 | 71034.4 | 37.73 | 0 | 0 | 151.5 | 8800.3 | 74336.9 | 38.73
0 | 0 | 232.5 | 2282.72 | 29627.2 | 19.99 | 0 | 0 | 192.5 | 5730.31 | 48404.4 | 5738.2 | 27.4004 | 0 | 0 | 152.5 | 8352.27 | 70552.4 | 37.48 | 0 | 0 | 152.5 | 8742.81 | 73851.3 | 38.47
0 | 0 | 233.5 | 2268.2 | 29438.7 | 19.87 | 0 | 0 | 193.5 | 5692.89 | 48088.4 | 5700.11 | 27.2185 | 0 | 0 | 153.5 | 8295.22 | 70070.4 | 37.22 | 0 | 0 | 153.5 | 8685.33 | 73365.8 | 38.22
0 | 0 | 234.5 | 2253.67 | 29250.3 | 19.74 | 0 | 0 | 194.5 | 5655.47 | 47772.3 | 5662.02 | 27.0366 | 0 | 0 | 154.5 | 8238.16 | 69588.5 | 36.97 | 0 | 0 | 154.5 | 8627.85 | 72880.2 | 37.97
0 | 0 | 235.5 | 2239 | 29059.8 | 19.61 | 0 | 0 | 195.5 | 5618.51 | 47460.1 | 5624.19 | 26.856 | 0 | 0 | 155.5 | 8181.86 | 69112.9 | 36.72 | 0 | 0 | 155.5 | 8568.71 | 72380.6 | 37.7
0 | 0 | 236.5 | 2224.14 | 28866.9 | 19.48 | 0 | 0 | 196.5 | 5582.08 | 47152.3 | 5586.66 | 26.6768 | 0 | 0 | 156.5 | 8126.51 | 68645.4 | 36.47 | 0 | 0 | 156.5 | 8507.11 | 71860.3 | 37.43
0 | 0 | 237.5 | 2209.28 | 28674 | 19.36 | 0 | 0 | 197.5 | 5545.64 | 46844.5 | 5549.12 | 26.4975 | 0 | 0 | 157.5 | 8071.17 | 68177.9 | 36.22 | 0 | 0 | 157.5 | 8445.52 | 71340 | 37.16
0 | 0 | 238.5 | 2194.41 | 28481.1 | 19.23 | 0 | 0 | 198.5 | 5509.21 | 46536.8 | 5511.59 | 26.3183 | 0 | 0 | 158.5 | 8015.83 | 67710.5 | 35.97 | 0 | 0 | 158.5 | 8383.92 | 70819.8 | 36.89
0 | 0 | 239.5 | 2179.55 | 28288.2 | 19.1 | 0 | 0 | 199.5 | 5472.77 | 46229 | 5474.05 | 26.139 | 0 | 0 | 159.5 | 7960.49 | 67243 | 35.72 | 0 | 0 | 159.5 | 8322.33 | 70299.5 | 36.62
0 | 0 | 240.5 | 2165.24 | 28102.4 | 18.96 | 0 | 0 | 200.5 | 5436.47 | 45922.3 | 5437.96 | 25.9667 | 0 | 0 | 160.5 | 7905.55 | 66778.9 | 35.48 | 0 | 0 | 160.5 | 8264.62 | 69811.9 | 36.37
0 | 0 | 241.5 | 2151.49 | 27924 | 18.82 | 0 | 0 | 201.5 | 5400.32 | 45617 | 5403.61 | 25.8027 | 0 | 0 | 161.5 | 7851.16 | 66319.4 | 35.23 | 0 | 0 | 161.5 | 8212.39 | 69370.8 | 36.14
0 | 0 | 242.5 | 2137.75 | 27745.7 | 18.68 | 0 | 0 | 202.5 | 5364.17 | 45311.6 | 5369.25 | 25.6386 | 0 | 0 | 162.5 | 7796.77 | 65860 | 34.99 | 0 | 0 | 162.5 | 8160.17 | 68929.7 | 35.91
0 | 0 | 243.5 | 2124.01 | 27567.3 | 18.55 | 0 | 0 | 203.5 | 5328.02 | 45006.3 | 5334.9 | 25.4746 | 0 | 0 | 163.5 | 7742.37 | 65400.5 | 34.74 | 0 | 0 | 163.5 | 8107.95 | 68488.6 | 35.68
0 | 0 | 244.5 | 2110.26 | 27388.9 | 18.41 | 0 | 0 | 204.5 | 5291.87 | 44700.9 | 5300.55 | 25.3106 | 0 | 0 | 164.5 | 7687.98 | 64941 | 34.5 | 0 | 0 | 164.5 | 8055.73 | 68047.4 | 35.45
0 | 0 | 245.5 | 2097.19 | 27219.3 | 18.28 | 0 | 0 | 205.5 | 5257.02 | 44406.5 | 5266.29 | 25.147 | 0 | 0 | 165.5 | 7634.29 | 64487.5 | 34.26 | 0 | 0 | 165.5 | 8002.69 | 67599.4 | 35.22
0 | 0 | 246.5 | 2084.87 | 27059.4 | 18.16 | 0 | 0 | 206.5 | 5223.67 | 44124.8 | 5232.16 | 24.984 | 0 | 0 | 166.5 | 7581.59 | 64042.4 | 34.02 | 0 | 0 | 166.5 | 7948.56 | 67142.2 | 34.98
0 | 0 | 247.5 | 2072.56 | 26899.6 | 18.04 | 0 | 0 | 207.5 | 5190.32 | 43843.1 | 5198.03 | 24.821 | 0 | 0 | 167.5 | 7528.89 | 63597.2 | 33.78 | 0 | 0 | 167.5 | 7894.44 | 66685.1 | 34.74
0 | 0 | 248.5 | 2060.25 | 26739.7 | 17.93 | 0 | 0 | 208.5 | 5156.98 | 43561.4 | 5163.9 | 24.658 | 0 | 0 | 168.5 | 7476.19 | 63152 | 33.54 | 0 | 0 | 168.5 | 7840.32 | 66227.9 | 34.5
0 | 0 | 249.5 | 2047.93 | 26579.9 | 17.81 | 0 | 0 | 209.5 | 5123.63 | 43279.7 | 5129.77 | 24.4951 | 0 | 0 | 169.5 | 7423.48 | 62706.8 | 33.3 | 0 | 0 | 169.5 | 7786.2 | 65770.7 | 34.26
[6 empty] | 0 | 0 | 210.5 | 5089.38 | 42990.4 | 5095.05 | 24.3293 | 0 | 0 | 170.5 | 7371.6 | 62268.6 | 33.07 | 0 | 0 | 170.5 | 7732.26 | 65315.1 | 34.02
[6 empty] | 0 | 0 | 211.5 | 5053.93 | 42691 | 5059.56 | 24.1598 | 0 | 0 | 171.5 | 7320.79 | 61839.4 | 32.84 | 0 | 0 | 171.5 | 7678.58 | 64861.6 | 33.79
[6 empty] | 0 | 0 | 212.5 | 5018.49 | 42391.6 | 5024.07 | 23.9903 | 0 | 0 | 172.5 | 7269.97 | 61410.1 | 32.62 | 0 | 0 | 172.5 | 7624.89 | 64408.1 | 33.55
[6 empty] | 0 | 0 | 213.5 | 4983.05 | 42092.3 | 4988.58 | 23.8209 | 0 | 0 | 173.5 | 7219.16 | 60980.9 | 32.39 | 0 | 0 | 173.5 | 7571.2 | 63954.6 | 33.32
[6 empty] | 0 | 0 | 214.5 | 4947.6 | 41792.9 | 4953.08 | 23.6514 | 0 | 0 | 174.5 | 7168.35 | 60551.7 | 32.17 | 0 | 0 | 174.5 | 7517.51 | 63501.1 | 33.09
[6 empty] | 0 | 0 | 215.5 | 4912.94 | 41500 | 4918.76 | 23.4875 | 0 | 0 | 175.5 | 7119.82 | 60141.7 | 31.95 | 0 | 0 | 175.5 | 7465.99 | 63065.9 | 32.86
[6 empty] | 0 | 0 | 216.5 | 4879.15 | 41214.6 | 4885.73 | 23.3298 | 0 | 0 | 176.5 | 7074.02 | 59754.9 | 31.75 | 0 | 0 | 176.5 | 7417.5 | 62656.3 | 32.65
[6 empty] | 0 | 0 | 217.5 | 4845.35 | 40929.1 | 4852.7 | 23.172 | 0 | 0 | 177.5 | 7028.23 | 59368.1 | 31.54 | 0 | 0 | 177.5 | 7369.02 | 62246.8 | 32.43
[6 empty] | 0 | 0 | 218.5 | 4811.56 | 40643.7 | 4819.68 | 23.0144 | 0 | 0 | 178.5 | 6982.44 | 58981.3 | 31.34 | 0 | 0 | 178.5 | 7320.53 | 61837.2 | 32.22
[6 empty] | 0 | 0 | 219.5 | 4777.77 | 40358.2 | 4786.65 | 22.8566 | 0 | 0 | 179.5 | 6936.65 | 58594.5 | 31.13 | 0 | 0 | 179.5 | 7272.05 | 61427.6 | 32
[6 empty] | 0 | 0 | 220.5 | 4747.34 | 40101.2 | 4756.3 | 22.7117 | 0 | 0 | 180.5 | 6890.87 | 58207.8 | 30.93 | 0 | 0 | 180.5 | 7222.28 | 61007.3 | 31.78
[6 empty] | 0 | 0 | 221.5 | 4720.98 | 39878.6 | 4729.21 | 22.5824 | 0 | 0 | 181.5 | 6845.1 | 57821.2 | 30.72 | 0 | 0 | 181.5 | 7170.83 | 60572.6 | 31.55
[6 empty] | 0 | 0 | 222.5 | 4694.63 | 39656 | 4702.12 | 22.453 | 0 | 0 | 182.5 | 6799.33 | 57434.6 | 30.51 | 0 | 0 | 182.5 | 7119.37 | 60138 | 31.33
[6 empty] | 0 | 0 | 223.5 | 4668.28 | 39433.4 | 4675.03 | 22.3237 | 0 | 0 | 183.5 | 6753.56 | 57047.9 | 30.3 | 0 | 0 | 183.5 | 7067.92 | 59703.3 | 31.1
[6 empty] | 0 | 0 | 224.5 | 4641.93 | 39210.8 | 4647.95 | 22.1943 | 0 | 0 | 184.5 | 6707.79 | 56661.3 | 30.1 | 0 | 0 | 184.5 | 7016.46 | 59268.7 | 30.87
[6 empty] | 0 | 0 | 225.5 | 4611.81 | 38956.4 | 4617.96 | 22.0511 | 0 | 0 | 185.5 | 6662.72 | 56280.6 | 29.89 | 0 | 0 | 185.5 | 6968.82 | 58866.3 | 30.66
[6 empty] | 0 | 0 | 226.5 | 4577.44 | 38666.1 | 4584.72 | 21.8924 | 0 | 0 | 186.5 | 6618.43 | 55906.5 | 29.69 | 0 | 0 | 186.5 | 6925.88 | 58503.5 | 30.47
[6 empty] | 0 | 0 | 227.5 | 4543.07 | 38375.8 | 4551.48 | 21.7337 | 0 | 0 | 187.5 | 6574.15 | 55532.4 | 29.49 | 0 | 0 | 187.5 | 6882.94 | 58140.8 | 30.28
[6 empty] | 0 | 0 | 228.5 | 4508.71 | 38085.4 | 4518.23 | 21.5749 | 0 | 0 | 188.5 | 6529.86 | 55158.3 | 29.29 | 0 | 0 | 188.5 | 6840 | 57778.1 | 30.09
[6 empty] | 0 | 0 | 229.5 | 4474.34 | 37795.1 | 4484.99 | 21.4162 | 0 | 0 | 189.5 | 6485.57 | 54784.2 | 29.09 | 0 | 0 | 189.5 | 6797.06 | 57415.4 | 29.9
[6 empty] | 0 | 0 | 230.5 | 4444.71 | 37544.8 | 4455.3 | 21.2744 | 0 | 0 | 190.5 | 6442.56 | 54420.9 | 28.9 | 0 | 0 | 190.5 | 6754.03 | 57051.9 | 29.72
[6 empty] | 0 | 0 | 231.5 | 4420.34 | 37339 | 4429.56 | 21.1515 | 0 | 0 | 191.5 | 6401.15 | 54071.1 | 28.71 | 0 | 0 | 191.5 | 6710.88 | 56687.4 | 29.53
[6 empty] | 0 | 0 | 232.5 | 4395.97 | 37133.2 | 4403.82 | 21.0286 | 0 | 0 | 192.5 | 6359.73 | 53721.2 | 28.53 | 0 | 0 | 192.5 | 6667.73 | 56322.9 | 29.34
[6 empty] | 0 | 0 | 233.5 | 4371.61 | 36927.4 | 4378.08 | 20.9057 | 0 | 0 | 193.5 | 6318.31 | 53371.3 | 28.35 | 0 | 0 | 193.5 | 6624.58 | 55958.4 | 29.15
[6 empty] | 0 | 0 | 234.5 | 4347.24 | 36721.5 | 4352.34 | 20.7828 | 0 | 0 | 194.5 | 6276.89 | 53021.5 | 28.16 | 0 | 0 | 194.5 | 6581.43 | 55593.9 | 28.96
[6 empty] | 0 | 0 | 235.5 | 4320.41 | 36494.9 | 4324.74 | 20.651 | 0 | 0 | 195.5 | 6233.53 | 52655.2 | 27.97 | 0 | 0 | 195.5 | 6537.46 | 55222.5 | 28.76
[6 empty] | 0 | 0 | 236.5 | 4290.67 | 36243.7 | 4294.96 | 20.5088 | 0 | 0 | 196.5 | 6187.98 | 52270.4 | 27.77 | 0 | 0 | 196.5 | 6492.54 | 54843.1 | 28.57
[6 empty] | 0 | 0 | 237.5 | 4260.93 | 35992.4 | 4265.18 | 20.3666 | 0 | 0 | 197.5 | 6142.43 | 51885.6 | 27.56 | 0 | 0 | 197.5 | 6447.63 | 54463.7 | 28.37
[6 empty] | 0 | 0 | 238.5 | 4231.18 | 35741.2 | 4235.4 | 20.2244 | 0 | 0 | 198.5 | 6096.88 | 51500.9 | 27.36 | 0 | 0 | 198.5 | 6402.72 | 54084.4 | 28.17
[6 empty] | 0 | 0 | 239.5 | 4201.44 | 35489.9 | 4205.62 | 20.0822 | 0 | 0 | 199.5 | 6051.33 | 51116.1 | 27.15 | 0 | 0 | 199.5 | 6357.81 | 53705 | 27.97
[6 empty] | 0 | 0 | 240.5 | 4173.82 | 35256.6 | 4177.88 | 19.9497 | 0 | 0 | 200.5 | 6009.31 | 50761.2 | 26.96 | 0 | 0 | 200.5 | 6314.91 | 53342.6 | 27.78
[6 empty] | 0 | 0 | 241.5 | 4148.37 | 35041.7 | 4152.22 | 19.8272 | 0 | 0 | 201.5 | 5971.93 | 50445.4 | 26.79 | 0 | 0 | 201.5 | 6274.3 | 52999.6 | 27.61
[6 empty] | 0 | 0 | 242.5 | 4122.92 | 34826.7 | 4126.56 | 19.7047 | 0 | 0 | 202.5 | 5934.55 | 50129.6 | 26.62 | 0 | 0 | 202.5 | 6233.7 | 52656.6 | 27.43
[6 empty] | 0 | 0 | 243.5 | 4097.47 | 34611.7 | 4100.9 | 19.5821 | 0 | 0 | 203.5 | 5897.16 | 49813.9 | 26.45 | 0 | 0 | 203.5 | 6193.09 | 52313.6 | 27.25
[6 empty] | 0 | 0 | 244.5 | 4072.02 | 34396.8 | 4075.24 | 19.4596 | 0 | 0 | 204.5 | 5859.78 | 49498.1 | 26.28 | 0 | 0 | 204.5 | 6152.48 | 51970.6 | 27.08
[6 empty] | 0 | 0 | 245.5 | 4047.36 | 34188.4 | 4050.58 | 19.3419 | 0 | 0 | 205.5 | 5822.63 | 49184.2 | 26.12 | 0 | 0 | 205.5 | 6111.02 | 51620.3 | 26.9
[6 empty] | 0 | 0 | 246.5 | 4023.54 | 33987.2 | 4027.02 | 19.2293 | 0 | 0 | 206.5 | 5785.77 | 48872.9 | 25.95 | 0 | 0 | 206.5 | 6068.29 | 51259.4 | 26.7
[6 empty] | 0 | 0 | 247.5 | 3999.72 | 33786 | 4003.46 | 19.1168 | 0 | 0 | 207.5 | 5748.92 | 48561.6 | 25.79 | 0 | 0 | 207.5 | 6025.56 | 50898.5 | 26.51
[6 empty] | 0 | 0 | 248.5 | 3975.89 | 33584.7 | 3979.89 | 19.0043 | 0 | 0 | 208.5 | 5712.06 | 48250.3 | 25.62 | 0 | 0 | 208.5 | 5982.83 | 50537.5 | 26.32
[6 empty] | 0 | 0 | 249.5 | 3952.07 | 33383.5 | 3956.33 | 18.8918 | 0 | 0 | 209.5 | 5675.2 | 47939 | 25.45 | 0 | 0 | 209.5 | 5940.1 | 50176.6 | 26.13
[13 empty] | 0 | 0 | 210.5 | 5637.93 | 47624.1 | 25.29 | 0 | 0 | 210.5 | 5901.04 | 49846.7 | 25.96
[13 empty] | 0 | 0 | 211.5 | 5600.13 | 47304.8 | 25.12 | 0 | 0 | 211.5 | 5866.24 | 49552.7 | 25.81
[13 empty] | 0 | 0 | 212.5 | 5562.33 | 46985.5 | 24.95 | 0 | 0 | 212.5 | 5831.44 | 49258.7 | 25.65
[13 empty] | 0 | 0 | 213.5 | 5524.53 | 46666.2 | 24.79 | 0 | 0 | 213.5 | 5796.64 | 48964.7 | 25.5
[13 empty] | 0 | 0 | 214.5 | 5486.73 | 46346.9 | 24.62 | 0 | 0 | 214.5 | 5761.83 | 48670.7 | 25.35
[13 empty] | 0 | 0 | 215.5 | 5450.93 | 46044.5 | 24.46 | 0 | 0 | 215.5 | 5725.57 | 48364.4 | 25.19
[13 empty] | 0 | 0 | 216.5 | 5417.7 | 45763.8 | 24.31 | 0 | 0 | 216.5 | 5687.53 | 48043.1 | 25.02
[13 empty] | 0 | 0 | 217.5 | 5384.46 | 45483 | 24.15 | 0 | 0 | 217.5 | 5649.49 | 47721.8 | 24.85
[13 empty] | 0 | 0 | 218.5 | 5351.22 | 45202.2 | 24 | 0 | 0 | 218.5 | 5611.45 | 47400.4 | 24.68
[13 empty] | 0 | 0 | 219.5 | 5317.98 | 44921.5 | 23.85 | 0 | 0 | 219.5 | 5573.41 | 47079.1 | 24.51
[13 empty] | 0 | 0 | 220.5 | 5282.34 | 44620.4 | 23.69 | 0 | 0 | 220.5 | 5536.04 | 46763.4 | 24.34
[13 empty] | 0 | 0 | 221.5 | 5243.07 | 44288.7 | 23.51 | 0 | 0 | 221.5 | 5499.53 | 46455 | 24.19
[13 empty] | 0 | 0 | 222.5 | 5203.79 | 43956.9 | 23.34 | 0 | 0 | 222.5 | 5463.01 | 46146.6 | 24.03
[13 empty] | 0 | 0 | 223.5 | 5164.52 | 43625.1 | 23.16 | 0 | 0 | 223.5 | 5426.5 | 45838.1 | 23.87
[13 empty] | 0 | 0 | 224.5 | 5125.24 | 43293.4 | 22.99 | 0 | 0 | 224.5 | 5389.99 | 45529.7 | 23.72
[13 empty] | 0 | 0 | 225.5 | 5089.85 | 42994.4 | 22.83 | 0 | 0 | 225.5 | 5355.04 | 45234.5 | 23.57
[13 empty] | 0 | 0 | 226.5 | 5058.92 | 42733.1 | 22.69 | 0 | 0 | 226.5 | 5322.08 | 44956.1 | 23.42
[13 empty] | 0 | 0 | 227.5 | 5027.98 | 42471.8 | 22.55 | 0 | 0 | 227.5 | 5289.12 | 44677.7 | 23.27
[13 empty] | 0 | 0 | 228.5 | 4997.05 | 42210.5 | 22.42 | 0 | 0 | 228.5 | 5256.16 | 44399.3 | 23.13
[13 empty] | 0 | 0 | 229.5 | 4966.11 | 41949.2 | 22.28 | 0 | 0 | 229.5 | 5223.2 | 44120.8 | 22.98
[13 empty] | 0 | 0 | 230.5 | 4935.64 | 41691.8 | 22.14 | 0 | 0 | 230.5 | 5188.64 | 43828.9 | 22.83
[13 empty] | 0 | 0 | 231.5 | 4905.75 | 41439.3 | 22.01 | 0 | 0 | 231.5 | 5152.24 | 43521.4 | 22.67
[13 empty] | 0 | 0 | 232.5 | 4875.86 | 41186.8 | 21.87 | 0 | 0 | 232.5 | 5115.84 | 43214 | 22.51
[13 empty] | 0 | 0 | 233.5 | 4845.96 | 40934.3 | 21.74 | 0 | 0 | 233.5 | 5079.44 | 42906.5 | 22.35
[13 empty] | 0 | 0 | 234.5 | 4816.07 | 40681.7 | 21.61 | 0 | 0 | 234.5 | 5043.04 | 42599 | 22.18
[13 empty] | 0 | 0 | 235.5 | 4785.75 | 40425.6 | 21.47 | 0 | 0 | 235.5 | 5009.29 | 42313.9 | 22.04
[13 empty] | 0 | 0 | 236.5 | 4754.92 | 40165.2 | 21.33 | 0 | 0 | 236.5 | 4978.94 | 42057.6 | 21.9
[13 empty] | 0 | 0 | 237.5 | 4724.1 | 39904.9 | 21.19 | 0 | 0 | 237.5 | 4948.59 | 41801.2 | 21.77
[13 empty] | 0 | 0 | 238.5 | 4693.27 | 39644.5 | 21.05 | 0 | 0 | 238.5 | 4918.24 | 41544.8 | 21.63
[13 empty] | 0 | 0 | 239.5 | 4662.45 | 39384.1 | 20.91 | 0 | 0 | 239.5 | 4887.88 | 41288.4 | 21.5
[13 empty] | 0 | 0 | 240.5 | 4632.16 | 39128.3 | 20.78 | 0 | 0 | 240.5 | 4857.17 | 41028.9 | 21.36
[13 empty] | 0 | 0 | 241.5 | 4602.5 | 38877.7 | 20.64 | 0 | 0 | 241.5 | 4826.03 | 40765.9 | 21.22
[13 empty] | 0 | 0 | 242.5 | 4572.83 | 38627.1 | 20.51 | 0 | 0 | 242.5 | 4794.89 | 40502.9 | 21.09
[13 empty] | 0 | 0 | 243.5 | 4543.17 | 38376.5 | 20.38 | 0 | 0 | 243.5 | 4763.75 | 40239.8 | 20.95
[13 empty] | 0 | 0 | 244.5 | 4513.5 | 38125.9 | 20.25 | 0 | 0 | 244.5 | 4732.61 | 39976.8 | 20.81
[13 empty] | 0 | 0 | 245.5 | 4484.02 | 37876.9 | 20.11 | 0 | 0 | 245.5 | 4702.01 | 39718.3 | 20.68
[13 empty] | 0 | 0 | 246.5 | 4454.75 | 37629.7 | 19.98 | 0 | 0 | 246.5 | 4672.05 | 39465.2 | 20.55
[13 empty] | 0 | 0 | 247.5 | 4425.48 | 37382.5 | 19.85 | 0 | 0 | 247.5 | 4642.1 | 39212.2 | 20.42
[13 empty] | 0 | 0 | 248.5 | 4396.22 | 37135.2 | 19.72 | 0 | 0 | 248.5 | 4612.14 | 38959.2 | 20.29
[13 empty] | 0 | 0 | 249.5 | 4366.95 | 36888 | 19.59 | 0 | 0 | 249.5 | 4582.19 | 38706.2 | 20.16

## Diode SFD
0,5x0,5 cm² | [5 empty] | 1x1 cm² | [5 empty] | 2x2 cm² | [5 empty] | 3x3 cm²
Measurement time: | 12/04/2023 18:06:08 |  |  |  |  | Controller: | CCU |  |  |  |  | Measurement time: | 12/04/2023 18:56:46 |  |  |  |  | Measurement time: | 12/04/2023 19:15:44
Radiation device: | Novalis TB Accelerator |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Controller: | CCU |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | SAD: | 1000 mm |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  |  |  |  | SSD: | 900 mm |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele
Reference detector: | Stealth CI transmission |  |  |  |  | Field size: | 10 x 10 mm |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | Gantry angle: | 0 |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | Measurement medium: | Water |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 5 x 5 mm |  |  |  |  | Scan type: | Inline |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Gantry angle: | 0 |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Wedge: | none |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline | [10 empty] | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Points [mm]: | [5 empty] | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Wedge: | none |  |  |  |  | Wedge: | none
[6 empty] | -10 | 0 | 100 | 2.24 | 23220.6 | 2.24
Points [mm]: | [5 empty] | -9 | 0 | 100 | 3.47 | 35958.7 | 3.47 | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | -8 | 0 | 100 | 5.6 | 58041.1 | 5.6 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-5 | 0 | 100 | 122.57 | 42830.6 | 4.99 | -7 | 0 | 100 | 9.59 | 99440.6 | 9.6 | -20 | 0 | 100 | 2.36 | 24477.1 | 2.16 | -30 | 0 | 100 | 2.67 | 31606.2 | 2.68
-4 | 0 | 100 | 265.42 | 92752.8 | 10.8 | -6 | 0 | 100 | 19.65 | 203612.8 | 19.66 | -19 | 0 | 100 | 2.66 | 27555.4 | 2.43 | -29 | 0 | 100 | 2.84 | 33637.9 | 2.85
-3 | 0 | 100 | 666.49 | 232904.3 | 27.12 | -5 | 0 | 100 | 48.91 | 506879.6 | 48.93 | -18 | 0 | 100 | 3.01 | 31181.1 | 2.74 | -28 | 0 | 100 | 3.06 | 36338.1 | 3.08
-2 | 0 | 100 | 1674.49 | 585152 | 68.05 | -4 | 0 | 100 | 83.4 | 864432.6 | 83.4 | -17 | 0 | 100 | 3.51 | 36353.9 | 3.2 | -27 | 0 | 100 | 3.27 | 38835.4 | 3.29
-1 | 0 | 100 | 2342.77 | 818683.6 | 95.24 | -3 | 0 | 100 | 94.18 | 976137.9 | 94.26 | -16 | 0 | 100 | 4.25 | 44083.9 | 3.88 | -26 | 0 | 100 | 3.57 | 42366.8 | 3.59
0 | 0 | 100 | 2458.43 | 859099.9 | 100 | -2 | 0 | 100 | 98.23 | 1018120.9 | 98.28 | -15 | 0 | 100 | 5.25 | 54447.3 | 4.79 | -25 | 0 | 100 | 3.84 | 45558.3 | 3.86
1 | 0 | 100 | 2282.07 | 797470.3 | 92.67 | -1 | 0 | 100 | 99.7 | 1033311.2 | 99.77 | -14 | 0 | 100 | 6.9 | 71499.9 | 6.29 | -24 | 0 | 100 | 4.23 | 50191.9 | 4.25
2 | 0 | 100 | 1477.46 | 516299.4 | 60.07 | 0 | 0 | 100 | 99.94 | 1035779.3 | 100 | -13 | 0 | 100 | 9.42 | 97684.1 | 8.6 | -23 | 0 | 100 | 4.66 | 55242.1 | 4.68
3 | 0 | 100 | 510.89 | 178531.3 | 20.74 | 1 | 0 | 100 | 99.07 | 1026765.2 | 99.13 | -12 | 0 | 100 | 14.24 | 147604.1 | 13 | -22 | 0 | 100 | 5.21 | 61823.5 | 5.23
4 | 0 | 100 | 202.75 | 70851 | 8.24 | 2 | 0 | 100 | 96.22 | 997276.3 | 96.3 | -11 | 0 | 100 | 25.63 | 265644.6 | 23.38 | -21 | 0 | 100 | 5.89 | 69845.8 | 5.91
5 | 0 | 100 | 97.71 | 34144.8 | 3.97 | 3 | 0 | 100 | 89.22 | 924697.5 | 89.28 | -10 | 0 | 100 | 57.04 | 591210.3 | 52.02 | -20 | 0 | 100 | 6.95 | 82356.5 | 6.97
[6 empty] | 4 | 0 | 100 | 69.35 | 718782.5 | 69.4 | -9 | 0 | 100 | 89.75 | 930229.9 | 81.86 | -19 | 0 | 100 | 8.44 | 100062.6 | 8.47
[6 empty] | 5 | 0 | 100 | 31.77 | 329316.4 | 31.8 | -8 | 0 | 100 | 99.67 | 1033076.7 | 90.91 | -18 | 0 | 100 | 10.81 | 128181.9 | 10.85
Measurement time: | 12/04/2023 18:06:57 |  |  |  |  | 6 | 0 | 100 | 13.18 | 136578.5 | 13.19 | -7 | 0 | 100 | 103.81 | 1075932.6 | 94.68 | -17 | 0 | 100 | 15.48 | 183510.9 | 15.54
Radiation device: | Novalis TB Accelerator |  |  |  |  | 7 | 0 | 100 | 7.01 | 72634.7 | 7.01 | -6 | 0 | 100 | 106.26 | 1101296.7 | 96.94 | -16 | 0 | 100 | 25.93 | 307521.1 | 26.03
Energy: | 6 MV (FFF) |  |  |  |  | 8 | 0 | 100 | 4.27 | 44279.6 | 4.28 | -5 | 0 | 100 | 107.8 | 1117305.2 | 98.29 | -15 | 0 | 100 | 55.05 | 652839.2 | 55.27
Controller: | CCU |  |  |  |  | 9 | 0 | 100 | 2.7 | 27954.5 | 2.7 | -4 | 0 | 100 | 108.8 | 1127612.1 | 99.21 | -14 | 0 | 100 | 80.52 | 954815.3 | 80.83
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 10 | 0 | 100 | 1.76 | 18254.1 | 1.76 | -3 | 0 | 100 | 109.11 | 1130919.5 | 99.53 | -13 | 0 | 100 | 89.28 | 1058769.8 | 89.63
Field detector: | SFD parallele | [10 empty] | -2 | 0 | 100 | 109.61 | 1136058.1 | 99.97 | -12 | 0 | 100 | 93.24 | 1105719 | 93.61
Reference detector: | Stealth CI transmission | [10 empty] | -1 | 0 | 100 | 109.53 | 1135252.9 | 99.91 | -11 | 0 | 100 | 95.16 | 1128439.2 | 95.53
SAD: | 1000 mm |  |  |  |  | Measurement time: | 12/04/2023 18:30:13 |  |  |  |  | 0 | 0 | 100 | 109.63 | 1136255 | 100 | -10 | 0 | 100 | 96.76 | 1147419.8 | 97.14
SSD: | 900 mm |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | 1 | 0 | 100 | 109.62 | 1136130.7 | 100 | -9 | 0 | 100 | 97.58 | 1157106.6 | 97.96
Field size: | 5 x 5 mm |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | 2 | 0 | 100 | 109.22 | 1132001.9 | 99.61 | -8 | 0 | 100 | 98.36 | 1166365.8 | 98.74
Gantry angle: | 0 |  |  |  |  | Controller: | CCU |  |  |  |  | 3 | 0 | 100 | 108.65 | 1126078.8 | 99.1 | -7 | 0 | 100 | 98.7 | 1170439.1 | 99.09
Measurement medium: | Water |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 4 | 0 | 100 | 107.94 | 1118781.9 | 98.4 | -6 | 0 | 100 | 99.14 | 1175603.4 | 99.52
Scan type: | Crossline |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | 5 | 0 | 100 | 106.49 | 1103696.4 | 97.07 | -5 | 0 | 100 | 99.33 | 1177845.5 | 99.71
Scan mode: | StepByStep |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | 6 | 0 | 100 | 104.52 | 1083347 | 95.25 | -4 | 0 | 100 | 99.63 | 1181484.4 | 100.02
Wedge: | none |  |  |  |  | SAD: | 1000 mm |  |  |  |  | 7 | 0 | 100 | 100.16 | 1038109.3 | 91.31 | -3 | 0 | 100 | 99.89 | 1184478 | 100.27
[6 empty] | SSD: | 900 mm |  |  |  |  | 8 | 0 | 100 | 90.5 | 937950.6 | 82.51 | -2 | 0 | 100 | 99.71 | 1182411.4 | 100.1
Points [mm]: | [5 empty] | Field size: | 10 x 10 mm |  |  |  |  | 9 | 0 | 100 | 65.58 | 679754.2 | 59.81 | -1 | 0 | 100 | 100.04 | 1186330.8 | 100.43
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Gantry angle: | 0 |  |  |  |  | 10 | 0 | 100 | 29.8 | 308812.3 | 27.17 | 0 | 0 | 100 | 99.61 | 1181242.5 | 100
0 | -5 | 100 | 210.22 | 73462 | 8.55 | Measurement medium: | Water |  |  |  |  | 11 | 0 | 100 | 15.5 | 160633.3 | 14.13 | 1 | 0 | 100 | 99.57 | 1180738.9 | 99.96
0 | -4 | 100 | 492.45 | 172086.7 | 20.03 | Scan type: | Crossline |  |  |  |  | 12 | 0 | 100 | 9.84 | 101982.8 | 8.97 | 2 | 0 | 100 | 99.7 | 1182254.9 | 100.09
0 | -3 | 100 | 1169.55 | 408700.3 | 47.58 | Scan mode: | StepByStep |  |  |  |  | 13 | 0 | 100 | 7.08 | 73364.9 | 6.46 | 3 | 0 | 100 | 99.88 | 1184428.1 | 100.27
0 | -2 | 100 | 2065.33 | 721731.3 | 83.99 | Wedge: | none |  |  |  |  | 14 | 0 | 100 | 5.44 | 56379.2 | 4.96 | 4 | 0 | 100 | 99.31 | 1177655.1 | 99.7
0 | -1 | 100 | 2419.53 | 845505.3 | 98.32 | [6 empty] | 15 | 0 | 100 | 4.41 | 45707 | 4.02 | 5 | 0 | 100 | 99.01 | 1174152.9 | 99.4
0 | 0 | 100 | 2459.88 | 859606.4 | 100 | Points [mm]: | [5 empty] | 16 | 0 | 100 | 3.54 | 36689.6 | 3.23 | 6 | 0 | 100 | 98.79 | 1171494.8 | 99.17
0 | 1 | 100 | 2360.54 | 824893.3 | 95.83 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 17 | 0 | 100 | 3.1 | 32107.4 | 2.82 | 7 | 0 | 100 | 98.31 | 1165783.1 | 98.69
0 | 2 | 100 | 1983.12 | 693001.8 | 80.67 | 0 | -10 | 100 | 2.55 | 26458.6 | 2.56 | 18 | 0 | 100 | 2.7 | 27943.9 | 2.46 | 8 | 0 | 100 | 97.55 | 1156728.5 | 97.92
0 | 3 | 100 | 1081.9 | 378070.8 | 44 | 0 | -9 | 100 | 4.04 | 41907.8 | 4.05 | 19 | 0 | 100 | 2.34 | 24280 | 2.13 | 9 | 0 | 100 | 96.42 | 1143374.6 | 96.79
0 | 4 | 100 | 406.76 | 142143.4 | 16.55 | 0 | -8 | 100 | 7.01 | 72666 | 7.03 | 20 | 0 | 100 | 2.13 | 22053.6 | 1.94 | 10 | 0 | 100 | 95.38 | 1131032.4 | 95.75
0 | 5 | 100 | 155.02 | 54172.8 | 6.3 | 0 | -7 | 100 | 13.64 | 141398.9 | 13.68 | [6 empty] | 11 | 0 | 100 | 93.21 | 1105299.3 | 93.57
[6 empty] | 0 | -6 | 100 | 28.74 | 297842.4 | 28.82 | [6 empty] | 12 | 0 | 100 | 89.27 | 1058550.9 | 89.61
[6 empty] | 0 | -5 | 100 | 61.55 | 637973.4 | 61.73 | Measurement time: | 12/04/2023 18:59:35 |  |  |  |  | 13 | 0 | 100 | 80.57 | 955447.7 | 80.88
Measurement time: | 12/04/2023 18:07:44 |  |  |  |  | 0 | -4 | 100 | 87.12 | 902943.3 | 87.3 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 14 | 0 | 100 | 57.45 | 681255.3 | 57.67
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -3 | 100 | 95.17 | 986388.5 | 95.43 | Energy: | 6 MV (FFF) |  |  |  |  | 15 | 0 | 100 | 27.28 | 323529 | 27.39
Energy: | 6 MV (FFF) |  |  |  |  | 0 | -2 | 100 | 98.07 | 1016475.8 | 98.36 | Controller: | CCU |  |  |  |  | 16 | 0 | 100 | 15.89 | 188386.9 | 15.95
Controller: | CCU |  |  |  |  | 0 | -1 | 100 | 99.32 | 1029355 | 99.57 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 17 | 0 | 100 | 10.91 | 129421 | 10.96
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | 0 | 100 | 99.72 | 1033580.3 | 100 | Field detector: | SFD parallele |  |  |  |  | 18 | 0 | 100 | 8.56 | 101537.2 | 8.6
Field detector: | SFD parallele |  |  |  |  | 0 | 1 | 100 | 99.33 | 1029555.2 | 99.6 | Reference detector: | Stealth CI transmission |  |  |  |  | 19 | 0 | 100 | 6.97 | 82645 | 7
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | 2 | 100 | 97.52 | 1010800.3 | 97.76 | SAD: | 1000 mm |  |  |  |  | 20 | 0 | 100 | 5.93 | 70321.7 | 5.95
SAD: | 1000 mm |  |  |  |  | 0 | 3 | 100 | 93.13 | 965282 | 93.4 | SSD: | 900 mm |  |  |  |  | 21 | 0 | 100 | 5.19 | 61543.4 | 5.21
SSD: | 900 mm |  |  |  |  | 0 | 4 | 100 | 83.83 | 868845.2 | 84.1 | Field size: | 20 x 20 mm |  |  |  |  | 22 | 0 | 100 | 4.69 | 55584.8 | 4.71
Field size: | 5 x 5 mm |  |  |  |  | 0 | 5 | 100 | 61.89 | 641414.4 | 62.08 | Gantry angle: | 0 |  |  |  |  | 23 | 0 | 100 | 4.21 | 49889.4 | 4.22
Gantry angle: | 0 |  |  |  |  | 0 | 6 | 100 | 27.78 | 287883 | 27.86 | Measurement medium: | Water |  |  |  |  | 24 | 0 | 100 | 3.86 | 45758.8 | 3.87
Measurement medium: | Water |  |  |  |  | 0 | 7 | 100 | 11.97 | 124067.6 | 12 | Scan type: | Crossline |  |  |  |  | 25 | 0 | 100 | 3.57 | 42305.8 | 3.58
Scan type: | Beam |  |  |  |  | 0 | 8 | 100 | 6.32 | 65489.1 | 6.34 | Scan mode: | StepByStep |  |  |  |  | 26 | 0 | 100 | 3.32 | 39356.6 | 3.33
Scan mode: | StepByStep |  |  |  |  | 0 | 9 | 100 | 3.76 | 39000.6 | 3.77 | Wedge: | none |  |  |  |  | 27 | 0 | 100 | 3.07 | 36423.1 | 3.08
Wedge: | none |  |  |  |  | 0 | 10 | 100 | 2.43 | 25216.2 | 2.44 | [6 empty] | 28 | 0 | 100 | 2.87 | 34019.6 | 2.88
[12 empty] | Points [mm]: | [5 empty] | 29 | 0 | 100 | 2.7 | 31996 | 2.71
Points [mm]: | [11 empty] | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 30 | 0 | 100 | 2.53 | 29960.2 | 2.54
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Measurement time: | 12/04/2023 18:31:48 |  |  |  |  | 0 | -20 | 100 | 8.37 | 25242.5 | 2.22
0 | 0 | -0.5 | 2773.72 | 969277.8 | 54.6 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -19 | 100 | 9.49 | 28633.4 | 2.52
0 | 0 | 0.5 | 2956.49 | 1033147 | 58.31 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | -18 | 100 | 10.91 | 32890.9 | 2.9 | Measurement time: | 12/04/2023 19:18:35
0 | 0 | 1.5 | 3583.71 | 1252329.3 | 70.68 | Controller: | CCU |  |  |  |  | 0 | -17 | 100 | 12.92 | 38977.4 | 3.43 | Radiation device: | Novalis TB Accelerator
0 | 0 | 2.5 | 4270.33 | 1492267.7 | 84.2 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -16 | 100 | 15.8 | 47654.1 | 4.19 | Energy: | 6 MV (FFF)
0 | 0 | 3.5 | 4635.25 | 1619791.2 | 91.39 | Field detector: | SFD parallele |  |  |  |  | 0 | -15 | 100 | 20.2 | 60926.4 | 5.37 | Controller: | CCU
0 | 0 | 4.5 | 4850.47 | 1695000.4 | 95.61 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -14 | 100 | 26.93 | 81220.6 | 7.16 | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | 5.5 | 4971.74 | 1737377 | 98.04 | SAD: | 1000 mm |  |  |  |  | 0 | -13 | 100 | 39.23 | 118305.7 | 10.41 | Field detector: | SFD parallele
0 | 0 | 6.5 | 5035.81 | 1759766.9 | 99.33 | SSD: | 900 mm |  |  |  |  | 0 | -12 | 100 | 63.81 | 192447.7 | 16.95 | Reference detector: | Stealth CI transmission
0 | 0 | 7.5 | 5070.14 | 1771762.9 | 99.9 | Field size: | 10 x 10 mm |  |  |  |  | 0 | -11 | 100 | 123.14 | 371368.7 | 32.71 | SAD: | 1000 mm
0 | 0 | 8.5 | 5075.24 | 1773546.5 | 100 | Gantry angle: | 0 |  |  |  |  | 0 | -10 | 100 | 235.2 | 709343.1 | 62.46 | SSD: | 900 mm
0 | 0 | 9.5 | 5061.53 | 1768753.1 | 99.75 | Measurement medium: | Water |  |  |  |  | 0 | -9 | 100 | 320.34 | 966104.2 | 85.1 | Field size: | 30 x 30 mm
0 | 0 | 10.5 | 5042.96 | 1762264.1 | 99.32 | Scan type: | Beam |  |  |  |  | 0 | -8 | 100 | 347.96 | 1049400.9 | 92.42 | Gantry angle: | 0
0 | 0 | 11.5 | 5016.95 | 1753175.5 | 98.78 | Scan mode: | StepByStep |  |  |  |  | 0 | -7 | 100 | 358.52 | 1081248.6 | 95.19 | Measurement medium: | Water
0 | 0 | 12.5 | 4977.53 | 1739400.4 | 98.04 | Wedge: | none |  |  |  |  | 0 | -6 | 100 | 366.29 | 1104686.4 | 97.26 | Scan type: | Crossline
0 | 0 | 13.5 | 4947.59 | 1728935.7 | 97.57 | [6 empty] | 0 | -5 | 100 | 369.84 | 1115383.6 | 98.22 | Scan mode: | StepByStep
0 | 0 | 14.5 | 4915.58 | 1717751.9 | 96.85 | Points [mm]: | [5 empty] | 0 | -4 | 100 | 373.16 | 1125407.3 | 99.08 | Wedge: | none
0 | 0 | 15.5 | 4881.78 | 1705939.5 | 96.11 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | -3 | 100 | 374.7 | 1130066.3 | 99.56
0 | 0 | 16.5 | 4839.37 | 1691118.2 | 95.37 | 0 | 0 | -0.5 | 97.67 | 1012313 | 50.15 | 0 | -2 | 100 | 375.61 | 1132800.8 | 99.77 | Points [mm]:
0 | 0 | 17.5 | 4797.78 | 1676587.2 | 94.59 | 0 | 0 | 0.5 | 102.98 | 1067305.5 | 52.87 | 0 | -1 | 100 | 376.73 | 1136161.4 | 100.09 | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 18.5 | 4762.59 | 1664289.4 | 93.9 | 0 | 0 | 1.5 | 124.85 | 1294028.4 | 64.09 | 0 | 0 | 100 | 376.62 | 1135834.9 | 100 | 0 | -30 | 100 | 2.65 | 31448 | 2.66
0 | 0 | 19.5 | 4719.47 | 1649221.8 | 93.03 | 0 | 0 | 2.5 | 150.25 | 1557258.4 | 77.12 | 0 | 1 | 100 | 376.14 | 1134390.5 | 99.93 | 0 | -29 | 100 | 2.85 | 33800.8 | 2.86
0 | 0 | 20.5 | 4680.21 | 1635500.5 | 92.19 | 0 | 0 | 3.5 | 166.04 | 1720912.2 | 85.22 | 0 | 2 | 100 | 375.95 | 1133810.6 | 99.82 | 0 | -28 | 100 | 3.09 | 36622.2 | 3.1
0 | 0 | 21.5 | 4641.84 | 1622093 | 91.46 | 0 | 0 | 4.5 | 176.67 | 1831100.8 | 90.7 | 0 | 3 | 100 | 374.39 | 1129107 | 99.44 | 0 | -27 | 100 | 3.29 | 38972.1 | 3.3
0 | 0 | 22.5 | 4603.47 | 1608685.6 | 90.72 | 0 | 0 | 5.5 | 183.95 | 1906557.8 | 94.48 | 0 | 4 | 100 | 373 | 1124927.6 | 99.02 | 0 | -26 | 100 | 3.57 | 42304.5 | 3.58
0 | 0 | 23.5 | 4565.11 | 1595278.1 | 89.98 | 0 | 0 | 6.5 | 188.65 | 1955235.4 | 96.89 | 0 | 5 | 100 | 369.38 | 1114004.6 | 98.02 | 0 | -25 | 100 | 3.87 | 45912.9 | 3.89
0 | 0 | 24.5 | 4526.74 | 1581870.6 | 89.24 | 0 | 0 | 7.5 | 191.62 | 1986025.4 | 98.38 | 0 | 6 | 100 | 364.3 | 1098676.7 | 96.74 | 0 | -24 | 100 | 4.23 | 50122.2 | 4.24
0 | 0 | 25.5 | 4490.27 | 1569125.5 | 88.52 | 0 | 0 | 8.5 | 193.42 | 2004727.9 | 99.29 | 0 | 7 | 100 | 356.17 | 1074162.9 | 94.54 | 0 | -23 | 100 | 4.75 | 56302 | 4.77
0 | 0 | 26.5 | 4454.84 | 1556746.4 | 87.81 | 0 | 0 | 9.5 | 194.54 | 2016300.7 | 99.87 | 0 | 8 | 100 | 337.77 | 1018665.2 | 89.67 | 0 | -22 | 100 | 5.33 | 63230.4 | 5.35
0 | 0 | 27.5 | 4419.42 | 1544367.3 | 87.1 | 0 | 0 | 10.5 | 194.79 | 2018891.4 | 100 | 0 | 9 | 100 | 301.06 | 907950.2 | 79.95 | 0 | -21 | 100 | 6.09 | 72216.2 | 6.12
0 | 0 | 28.5 | 4383.99 | 1531988.2 | 86.39 | 0 | 0 | 11.5 | 194.6 | 2016927.3 | 99.89 | 0 | 10 | 100 | 214.11 | 645718.7 | 56.84 | 0 | -20 | 100 | 7.36 | 87297.9 | 7.39
0 | 0 | 29.5 | 4348.57 | 1519609 | 85.68 | 0 | 0 | 12.5 | 194.19 | 2012680.7 | 99.7 | 0 | 11 | 100 | 107.01 | 322733.8 | 28.41 | 0 | -19 | 100 | 9.18 | 108881 | 9.22
0 | 0 | 30.5 | 4313.78 | 1507453.8 | 85 | 0 | 0 | 13.5 | 193.39 | 2004386 | 99.3 | 0 | 12 | 100 | 54.45 | 164216.5 | 14.46 | 0 | -18 | 100 | 12.46 | 147710 | 12.51
0 | 0 | 31.5 | 4279.42 | 1495446.2 | 84.33 | 0 | 0 | 14.5 | 192.4 | 1994129.1 | 98.78 | 0 | 13 | 100 | 33.96 | 102408.8 | 9.02 | 0 | -17 | 100 | 19.36 | 229548.3 | 19.44
0 | 0 | 32.5 | 4245.06 | 1483438.7 | 83.66 | 0 | 0 | 15.5 | 191.28 | 1982578.4 | 98.17 | 0 | 14 | 100 | 24.6 | 74177.9 | 6.53 | 0 | -16 | 100 | 36.03 | 427311.4 | 36.18
0 | 0 | 33.5 | 4210.7 | 1471431.1 | 82.99 | 0 | 0 | 16.5 | 190.02 | 1969448.2 | 97.52 | 0 | 15 | 100 | 18.3 | 55196.8 | 4.86 | 0 | -15 | 100 | 63.46 | 752514.2 | 63.72
0 | 0 | 34.5 | 4176.34 | 1459423.6 | 82.32 | 0 | 0 | 17.5 | 188.82 | 1957006.8 | 96.92 | 0 | 16 | 100 | 14.96 | 45126.1 | 3.97 | 0 | -14 | 100 | 83.17 | 986223.9 | 83.51
0 | 0 | 35.5 | 4141.96 | 1447410.4 | 81.64 | 0 | 0 | 18.5 | 187.5 | 1943336.8 | 96.25 | 0 | 17 | 100 | 12.23 | 36881.8 | 3.25 | 0 | -13 | 100 | 90.27 | 1070404.3 | 90.64
0 | 0 | 36.5 | 4107.57 | 1435393.1 | 80.96 | 0 | 0 | 19.5 | 186.04 | 1928206.1 | 95.49 | 0 | 18 | 100 | 10.61 | 31991.5 | 2.82 | 0 | -12 | 100 | 93.68 | 1110862.7 | 94.07
0 | 0 | 37.5 | 4073.18 | 1423375.9 | 80.28 | 0 | 0 | 20.5 | 184.75 | 1914866.8 | 94.81 | 0 | 19 | 100 | 9.16 | 27622.8 | 2.43 | 0 | -11 | 100 | 95.41 | 1131369.5 | 95.8
0 | 0 | 38.5 | 4038.8 | 1411358.7 | 79.59 | 0 | 0 | 21.5 | 183.36 | 1900422.1 | 94.09 | 0 | 20 | 100 | 8.04 | 24252.4 | 2.14 | 0 | -10 | 100 | 96.36 | 1142626.3 | 96.76
0 | 0 | 39.5 | 4004.41 | 1399341.5 | 78.91 | 0 | 0 | 22.5 | 181.96 | 1885977.4 | 93.37 | [6 empty] | 0 | -9 | 100 | 97.26 | 1153388.3 | 97.67
0 | 0 | 40.5 | 3971.14 | 1387716.9 | 78.24 | 0 | 0 | 23.5 | 180.57 | 1871532.7 | 92.66 | [6 empty] | 0 | -8 | 100 | 98.05 | 1162710.6 | 98.46
0 | 0 | 41.5 | 3938.72 | 1376388 | 77.59 | 0 | 0 | 24.5 | 179.18 | 1857088 | 91.94 | Measurement time: | 12/04/2023 19:01:28 |  |  |  |  | 0 | -7 | 100 | 98.46 | 1167568.5 | 98.87
0 | 0 | 42.5 | 3906.3 | 1365059 | 76.94 | 0 | 0 | 25.5 | 177.77 | 1842507.5 | 91.22 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -6 | 100 | 98.82 | 1171876 | 99.23
0 | 0 | 43.5 | 3873.88 | 1353730 | 76.29 | 0 | 0 | 26.5 | 176.36 | 1827852.1 | 90.5 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | -5 | 100 | 99.18 | 1176169.8 | 99.6
0 | 0 | 44.5 | 3841.46 | 1342401.1 | 75.64 | 0 | 0 | 27.5 | 174.94 | 1813196.7 | 89.78 | Controller: | CCU |  |  |  |  | 0 | -4 | 100 | 99.39 | 1178551.5 | 99.8
0 | 0 | 45.5 | 3809.4 | 1331196.6 | 75.01 | 0 | 0 | 28.5 | 173.53 | 1798541.2 | 89.07 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -3 | 100 | 99.28 | 1177354.1 | 99.7
0 | 0 | 46.5 | 3777.62 | 1320090.1 | 74.4 | 0 | 0 | 29.5 | 172.11 | 1783885.8 | 88.35 | Field detector: | SFD parallele |  |  |  |  | 0 | -2 | 100 | 99.68 | 1182008.5 | 100.09
0 | 0 | 47.5 | 3745.83 | 1308983.5 | 73.79 | 0 | 0 | 30.5 | 170.77 | 1769922.2 | 87.65 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -1 | 100 | 99.67 | 1181930.6 | 100.08
0 | 0 | 48.5 | 3714.05 | 1297877 | 73.18 | 0 | 0 | 31.5 | 169.46 | 1756396.1 | 86.98 | SAD: | 1000 mm |  |  |  |  | 0 | 0 | 100 | 99.59 | 1180942 | 100
0 | 0 | 49.5 | 3682.27 | 1286770.4 | 72.57 | 0 | 0 | 32.5 | 168.16 | 1742870 | 86.3 | SSD: | 900 mm |  |  |  |  | 0 | 1 | 100 | 99.77 | 1183083.1 | 100.18
0 | 0 | 50.5 | 3652.28 | 1276289.3 | 71.99 | 0 | 0 | 33.5 | 166.85 | 1729343.8 | 85.62 | Field size: | 20 x 20 mm |  |  |  |  | 0 | 2 | 100 | 99.49 | 1179812.9 | 99.9
0 | 0 | 51.5 | 3623.75 | 1266320.6 | 71.43 | 0 | 0 | 34.5 | 165.55 | 1715817.7 | 84.95 | Gantry angle: | 0 |  |  |  |  | 0 | 3 | 100 | 99.49 | 1179830.5 | 99.91
0 | 0 | 52.5 | 3595.22 | 1256351.9 | 70.87 | 0 | 0 | 35.5 | 164.23 | 1702152.3 | 84.27 | Measurement medium: | Water |  |  |  |  | 0 | 4 | 100 | 99.5 | 1179898.8 | 99.91
0 | 0 | 53.5 | 3566.7 | 1246383.3 | 70.32 | 0 | 0 | 36.5 | 162.9 | 1688384.7 | 83.6 | Scan type: | Beam |  |  |  |  | 0 | 5 | 100 | 99.56 | 1180648.6 | 99.98
0 | 0 | 54.5 | 3538.17 | 1236414.6 | 69.76 | 0 | 0 | 37.5 | 161.57 | 1674617.1 | 82.92 | Scan mode: | StepByStep |  |  |  |  | 0 | 6 | 100 | 99.09 | 1174986.8 | 99.5
0 | 0 | 55.5 | 3509.06 | 1226242 | 69.18 | 0 | 0 | 38.5 | 160.24 | 1660849.5 | 82.25 | Wedge: | none |  |  |  |  | 0 | 7 | 100 | 98.58 | 1168959.5 | 98.99
0 | 0 | 56.5 | 3479.44 | 1215892.7 | 68.59 | 0 | 0 | 39.5 | 158.92 | 1647081.9 | 81.57 | [6 empty] | 0 | 8 | 100 | 98.02 | 1162381.4 | 98.43
0 | 0 | 57.5 | 3449.83 | 1205543.3 | 68 | 0 | 0 | 40.5 | 157.68 | 1634262.5 | 80.93 | Points [mm]: | [5 empty] | 0 | 9 | 100 | 97.27 | 1153466.8 | 97.67
0 | 0 | 58.5 | 3420.21 | 1195194 | 67.4 | 0 | 0 | 41.5 | 156.5 | 1622079 | 80.32 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 10 | 100 | 96.5 | 1144335.9 | 96.9
0 | 0 | 59.5 | 3390.59 | 1184844.6 | 66.81 | 0 | 0 | 42.5 | 155.33 | 1609895.5 | 79.72 | 0 | 0 | -0.5 | 343.52 | 1036029.9 | 48.6 | 0 | 11 | 100 | 94.97 | 1126220.4 | 95.37
0 | 0 | 60.5 | 3363.03 | 1175210.9 | 66.26 | 0 | 0 | 43.5 | 154.15 | 1597712 | 79.11 | 0 | 0 | 0.5 | 356.35 | 1074711.8 | 50.42 | 0 | 12 | 100 | 92.68 | 1099043.3 | 93.06
0 | 0 | 61.5 | 3337.38 | 1166249.5 | 65.75 | 0 | 0 | 44.5 | 152.98 | 1585528.5 | 78.5 | 0 | 0 | 1.5 | 426.14 | 1285187.9 | 60.3 | 0 | 13 | 100 | 88.09 | 1044564.9 | 88.45
0 | 0 | 62.5 | 3311.74 | 1157288.2 | 65.24 | 0 | 0 | 45.5 | 151.81 | 1573465.9 | 77.9 | 0 | 0 | 2.5 | 519.39 | 1566424.7 | 73.5 | 0 | 14 | 100 | 78.48 | 930602.9 | 78.8
0 | 0 | 63.5 | 3286.09 | 1148326.8 | 64.73 | 0 | 0 | 46.5 | 150.66 | 1561494.5 | 77.3 | 0 | 0 | 3.5 | 579.38 | 1747329.8 | 81.98 | 0 | 15 | 100 | 56.29 | 667565.5 | 56.53
0 | 0 | 64.5 | 3260.45 | 1139365.4 | 64.22 | 0 | 0 | 47.5 | 149.5 | 1549523.1 | 76.71 | 0 | 0 | 4.5 | 620.34 | 1870882.4 | 87.77 | 0 | 16 | 100 | 28.87 | 342307.2 | 28.99
0 | 0 | 65.5 | 3235.05 | 1130489.4 | 63.72 | 0 | 0 | 48.5 | 148.35 | 1537551.6 | 76.12 | 0 | 0 | 5.5 | 648.47 | 1955697 | 91.72 | 0 | 17 | 100 | 15.9 | 188604.4 | 15.97
0 | 0 | 66.5 | 3209.89 | 1121696.5 | 63.23 | 0 | 0 | 49.5 | 147.19 | 1525580.2 | 75.52 | 0 | 0 | 6.5 | 668.44 | 2015951.6 | 94.55 | 0 | 18 | 100 | 11.09 | 131551.3 | 11.14
0 | 0 | 67.5 | 3184.73 | 1112903.6 | 62.73 | 0 | 0 | 50.5 | 146.02 | 1513418.2 | 74.92 | 0 | 0 | 7.5 | 682.51 | 2058378.1 | 96.56 | 0 | 19 | 100 | 8.5 | 100738.3 | 8.53
0 | 0 | 68.5 | 3159.56 | 1104110.8 | 62.24 | 0 | 0 | 51.5 | 144.83 | 1501106.7 | 74.31 | 0 | 0 | 8.5 | 693.47 | 2091417.7 | 98.09 | 0 | 20 | 100 | 6.99 | 82840.6 | 7.01
0 | 0 | 69.5 | 3134.4 | 1095317.9 | 61.75 | 0 | 0 | 52.5 | 143.64 | 1488795.2 | 73.7 | 0 | 0 | 9.5 | 701.13 | 2114538.5 | 99.13 | 0 | 21 | 100 | 5.84 | 69291.8 | 5.87
0 | 0 | 70.5 | 3109.66 | 1086672 | 61.28 | 0 | 0 | 53.5 | 142.46 | 1476483.7 | 73.09 | 0 | 0 | 10.5 | 705.4 | 2127399.3 | 99.73 | 0 | 22 | 100 | 5.15 | 61095.1 | 5.17
0 | 0 | 71.5 | 3085.35 | 1078175.5 | 60.82 | 0 | 0 | 54.5 | 141.27 | 1464172.2 | 72.48 | 0 | 0 | 11.5 | 706.61 | 2131051.1 | 99.92 | 0 | 23 | 100 | 4.61 | 54614.9 | 4.62
0 | 0 | 72.5 | 3061.03 | 1069679 | 60.37 | 0 | 0 | 55.5 | 140.12 | 1452306.7 | 71.89 | 0 | 0 | 12.5 | 707.06 | 2132425.1 | 100 | 0 | 24 | 100 | 4.15 | 49170.5 | 4.16
0 | 0 | 73.5 | 3036.72 | 1061182.5 | 59.91 | 0 | 0 | 56.5 | 139.02 | 1440836.8 | 71.33 | 0 | 0 | 13.5 | 706.47 | 2130643.7 | 99.93 | 0 | 25 | 100 | 3.82 | 45315 | 3.84
0 | 0 | 74.5 | 3012.4 | 1052686 | 59.45 | 0 | 0 | 57.5 | 137.91 | 1429366.8 | 70.76 | 0 | 0 | 14.5 | 704.82 | 2125664.5 | 99.67 | 0 | 26 | 100 | 3.5 | 41535.7 | 3.52
0 | 0 | 75.5 | 2988.39 | 1044294.4 | 58.98 | 0 | 0 | 58.5 | 136.8 | 1417896.9 | 70.19 | 0 | 0 | 15.5 | 702.3 | 2118060.8 | 99.3 | 0 | 27 | 100 | 3.27 | 38734.6 | 3.28
0 | 0 | 76.5 | 2964.72 | 1036021.8 | 58.5 | 0 | 0 | 59.5 | 135.7 | 1406427 | 69.63 | 0 | 0 | 16.5 | 698.91 | 2107841.9 | 98.86 | 0 | 28 | 100 | 3.04 | 36084.5 | 3.06
0 | 0 | 77.5 | 2941.04 | 1027749.3 | 58.01 | 0 | 0 | 60.5 | 134.66 | 1395668.8 | 69.1 | 0 | 0 | 17.5 | 695.22 | 2096715.4 | 98.36 | 0 | 29 | 100 | 2.83 | 33601.2 | 2.85
0 | 0 | 78.5 | 2917.37 | 1019476.8 | 57.52 | 0 | 0 | 61.5 | 133.68 | 1385527.5 | 68.6 | 0 | 0 | 18.5 | 691.77 | 2086302.9 | 97.82 | 0 | 30 | 100 | 2.66 | 31505.5 | 2.67
0 | 0 | 79.5 | 2893.7 | 1011204.2 | 57.04 | 0 | 0 | 62.5 | 132.7 | 1375386.2 | 68.1 | 0 | 0 | 19.5 | 687.63 | 2073814.7 | 97.24
0 | 0 | 80.5 | 2870.67 | 1003157.9 | 56.58 | 0 | 0 | 63.5 | 131.72 | 1365244.9 | 67.6 | 0 | 0 | 20.5 | 682.77 | 2059158.6 | 96.58
0 | 0 | 81.5 | 2848.28 | 995331.8 | 56.14 | 0 | 0 | 64.5 | 130.74 | 1355103.6 | 67.1 | 0 | 0 | 21.5 | 677.99 | 2044726.2 | 95.91 | Measurement time: | 12/04/2023 19:22:22
0 | 0 | 82.5 | 2825.88 | 987505.6 | 55.7 | 0 | 0 | 65.5 | 129.74 | 1344697.2 | 66.58 | 0 | 0 | 22.5 | 673.2 | 2030293.8 | 95.23 | Radiation device: | Novalis TB Accelerator
0 | 0 | 83.5 | 2803.49 | 979679.4 | 55.27 | 0 | 0 | 66.5 | 128.71 | 1334065.1 | 66.05 | 0 | 0 | 23.5 | 668.41 | 2015861.3 | 94.55 | Energy: | 6 MV (FFF)
0 | 0 | 84.5 | 2781.09 | 971853.2 | 54.83 | 0 | 0 | 67.5 | 127.69 | 1323433.1 | 65.52 | 0 | 0 | 24.5 | 663.63 | 2001428.9 | 93.87 | Controller: | CCU
0 | 0 | 85.5 | 2758.4 | 963923.3 | 54.39 | 0 | 0 | 68.5 | 126.66 | 1312801 | 64.99 | 0 | 0 | 25.5 | 658.84 | 1986989.3 | 93.19 | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | 86.5 | 2735.44 | 955899.1 | 53.94 | 0 | 0 | 69.5 | 125.64 | 1302168.9 | 64.45 | 0 | 0 | 26.5 | 654.05 | 1972544.9 | 92.5 | Field detector: | SFD parallele
0 | 0 | 87.5 | 2712.47 | 947874.9 | 53.5 | 0 | 0 | 70.5 | 124.68 | 1292230.5 | 63.96 | 0 | 0 | 27.5 | 649.26 | 1958100.6 | 91.82 | Reference detector: | Stealth CI transmission
0 | 0 | 88.5 | 2689.51 | 939850.7 | 53.05 | 0 | 0 | 71.5 | 123.78 | 1282879.3 | 63.5 | 0 | 0 | 28.5 | 644.47 | 1943656.2 | 91.13 | SAD: | 1000 mm
0 | 0 | 89.5 | 2666.55 | 931826.5 | 52.6 | 0 | 0 | 72.5 | 122.87 | 1273528.2 | 63.04 | 0 | 0 | 29.5 | 639.68 | 1929211.8 | 90.45 | SSD: | 900 mm
0 | 0 | 90.5 | 2645.19 | 924361.3 | 52.17 | 0 | 0 | 73.5 | 121.97 | 1264177 | 62.59 | 0 | 0 | 30.5 | 634.98 | 1915018.6 | 89.78 | Field size: | 30 x 30 mm
0 | 0 | 91.5 | 2625.86 | 917607.1 | 51.77 | 0 | 0 | 74.5 | 121.07 | 1254825.8 | 62.13 | 0 | 0 | 31.5 | 630.32 | 1900980.7 | 89.13 | Gantry angle: | 0
0 | 0 | 92.5 | 2606.53 | 910852.9 | 51.37 | 0 | 0 | 75.5 | 120.12 | 1245002.5 | 61.64 | 0 | 0 | 32.5 | 625.67 | 1886942.7 | 88.47 | Measurement medium: | Water
0 | 0 | 93.5 | 2587.2 | 904098.7 | 50.97 | 0 | 0 | 76.5 | 119.12 | 1234664 | 61.13 | 0 | 0 | 33.5 | 621.01 | 1872904.7 | 87.82 | Scan type: | Beam
0 | 0 | 94.5 | 2567.87 | 897344.5 | 50.57 | 0 | 0 | 77.5 | 118.13 | 1224325.4 | 60.61 | 0 | 0 | 34.5 | 616.36 | 1858866.7 | 87.16 | Scan mode: | StepByStep
0 | 0 | 95.5 | 2548.32 | 890510.3 | 50.18 | 0 | 0 | 78.5 | 117.13 | 1213986.9 | 60.1 | 0 | 0 | 35.5 | 611.73 | 1844904.4 | 86.51 | Wedge: | none
0 | 0 | 96.5 | 2528.46 | 883572.3 | 49.79 | 0 | 0 | 79.5 | 116.13 | 1203648.4 | 59.59 | 0 | 0 | 36.5 | 607.12 | 1830997.5 | 85.86
0 | 0 | 97.5 | 2508.61 | 876634.3 | 49.4 | 0 | 0 | 80.5 | 115.21 | 1194124.9 | 59.12 | 0 | 0 | 37.5 | 602.51 | 1817090.7 | 85.21 | Points [mm]:
0 | 0 | 98.5 | 2488.75 | 869696.4 | 49.02 | 0 | 0 | 81.5 | 114.38 | 1185496.1 | 58.69 | 0 | 0 | 38.5 | 597.9 | 1803183.8 | 84.55 | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 99.5 | 2468.9 | 862758.4 | 48.63 | 0 | 0 | 82.5 | 113.55 | 1176867.3 | 58.27 | 0 | 0 | 39.5 | 593.28 | 1789277 | 83.9 | 0 | 0 | -0.5 | 89.98 | 1066990.7 | 49.53
0 | 0 | 100.5 | 2449.39 | 855941.7 | 48.25 | 0 | 0 | 83.5 | 112.72 | 1168238.5 | 57.85 | 0 | 0 | 40.5 | 588.88 | 1775994.3 | 83.28 | 0 | 0 | 0.5 | 92.47 | 1096583.5 | 50.9
0 | 0 | 101.5 | 2430.39 | 849299.6 | 47.89 | 0 | 0 | 84.5 | 111.88 | 1159609.7 | 57.42 | 0 | 0 | 41.5 | 584.62 | 1763135.6 | 82.67 | 0 | 0 | 1.5 | 109.64 | 1300158.8 | 60.35
0 | 0 | 102.5 | 2411.38 | 842657.4 | 47.52 | 0 | 0 | 85.5 | 111.05 | 1151015.8 | 57 | 0 | 0 | 42.5 | 580.35 | 1750277 | 82.07 | 0 | 0 | 2.5 | 133.55 | 1583681.2 | 73.51
0 | 0 | 103.5 | 2392.37 | 836015.3 | 47.16 | 0 | 0 | 86.5 | 110.23 | 1142456.6 | 56.57 | 0 | 0 | 43.5 | 576.09 | 1737418.4 | 81.46 | 0 | 0 | 3.5 | 148.36 | 1759262.4 | 81.66
0 | 0 | 104.5 | 2373.36 | 829373.2 | 46.79 | 0 | 0 | 87.5 | 109.4 | 1133897.4 | 56.14 | 0 | 0 | 44.5 | 571.83 | 1724559.7 | 80.85 | 0 | 0 | 4.5 | 158.42 | 1878591 | 87.2
0 | 0 | 105.5 | 2355.08 | 822983.6 | 46.43 | 0 | 0 | 88.5 | 108.58 | 1125338.3 | 55.71 | 0 | 0 | 45.5 | 567.58 | 1711764.1 | 80.25 | 0 | 0 | 5.5 | 165.57 | 1963433.9 | 91.14
0 | 0 | 106.5 | 2337.74 | 816925.8 | 46.07 | 0 | 0 | 89.5 | 107.75 | 1116779.1 | 55.28 | 0 | 0 | 46.5 | 563.36 | 1699016 | 79.65 | 0 | 0 | 6.5 | 171.08 | 2028716.4 | 94.17
0 | 0 | 107.5 | 2320.41 | 810868 | 45.71 | 0 | 0 | 90.5 | 106.94 | 1108398.2 | 54.87 | 0 | 0 | 47.5 | 559.13 | 1686268 | 79.06 | 0 | 0 | 7.5 | 175.28 | 2078566.2 | 96.48
0 | 0 | 108.5 | 2303.07 | 804810.2 | 45.36 | 0 | 0 | 91.5 | 106.15 | 1100244.3 | 54.47 | 0 | 0 | 48.5 | 554.9 | 1673519.9 | 78.46 | 0 | 0 | 8.5 | 177.89 | 2109478.1 | 97.92
0 | 0 | 109.5 | 2285.74 | 798752.4 | 45 | 0 | 0 | 92.5 | 105.37 | 1092090.3 | 54.07 | 0 | 0 | 49.5 | 550.67 | 1660771.8 | 77.86 | 0 | 0 | 9.5 | 179.65 | 2130389.5 | 98.89
0 | 0 | 110.5 | 2268.26 | 792645.6 | 44.65 | 0 | 0 | 93.5 | 104.58 | 1083936.4 | 53.67 | 0 | 0 | 50.5 | 546.42 | 1647928.3 | 77.26 | 0 | 0 | 10.5 | 180.92 | 2145371.7 | 99.58
0 | 0 | 111.5 | 2250.6 | 786473.3 | 44.3 | 0 | 0 | 94.5 | 103.79 | 1075782.5 | 53.26 | 0 | 0 | 51.5 | 542.13 | 1635000.1 | 76.66 | 0 | 0 | 11.5 | 181.4 | 2151133 | 99.85
0 | 0 | 112.5 | 2232.94 | 780301.1 | 43.96 | 0 | 0 | 95.5 | 103.01 | 1067700.3 | 52.87 | 0 | 0 | 52.5 | 537.84 | 1622071.8 | 76.06 | 0 | 0 | 12.5 | 181.67 | 2154313.7 | 100
0 | 0 | 113.5 | 2215.28 | 774128.9 | 43.61 | 0 | 0 | 96.5 | 102.24 | 1059705.6 | 52.47 | 0 | 0 | 53.5 | 533.56 | 1609143.5 | 75.45 | 0 | 0 | 13.5 | 181.63 | 2153793.1 | 99.98
0 | 0 | 114.5 | 2197.61 | 767956.7 | 43.26 | 0 | 0 | 97.5 | 101.47 | 1051711 | 52.07 | 0 | 0 | 54.5 | 529.27 | 1596215.2 | 74.85 | 0 | 0 | 14.5 | 181.38 | 2150876.6 | 99.84
0 | 0 | 115.5 | 2180.26 | 761891.6 | 42.93 | 0 | 0 | 98.5 | 100.7 | 1043716.3 | 51.68 | 0 | 0 | 55.5 | 525.3 | 1584246.1 | 74.29 | 0 | 0 | 15.5 | 180.96 | 2145891.2 | 99.61
0 | 0 | 116.5 | 2163.32 | 755973.2 | 42.6 | 0 | 0 | 99.5 | 99.93 | 1035721.6 | 51.28 | 0 | 0 | 56.5 | 521.6 | 1573094.1 | 73.77 | 0 | 0 | 16.5 | 180.15 | 2136319.6 | 99.16
0 | 0 | 117.5 | 2146.38 | 750054.8 | 42.28 | 0 | 0 | 100.5 | 99.18 | 1027977.7 | 50.9 | 0 | 0 | 57.5 | 517.9 | 1561942 | 73.25 | 0 | 0 | 17.5 | 179.54 | 2129015.2 | 98.83
0 | 0 | 118.5 | 2129.45 | 744136.3 | 41.95 | 0 | 0 | 101.5 | 98.46 | 1020532.7 | 50.53 | 0 | 0 | 58.5 | 514.21 | 1550790 | 72.73 | 0 | 0 | 18.5 | 178.73 | 2119446.7 | 98.38
0 | 0 | 119.5 | 2112.51 | 738217.9 | 41.63 | 0 | 0 | 102.5 | 97.75 | 1013087.6 | 50.16 | 0 | 0 | 59.5 | 510.51 | 1539638 | 72.21 | 0 | 0 | 19.5 | 177.78 | 2108121.9 | 97.86
0 | 0 | 120.5 | 2096.41 | 732590.2 | 41.32 | 0 | 0 | 103.5 | 97.03 | 1005642.6 | 49.79 | 0 | 0 | 60.5 | 506.75 | 1528292.1 | 71.68 | 0 | 0 | 20.5 | 176.83 | 2096949.5 | 97.34
0 | 0 | 121.5 | 2081.34 | 727323.5 | 41.02 | 0 | 0 | 104.5 | 96.31 | 998197.5 | 49.43 | 0 | 0 | 61.5 | 502.92 | 1516756.5 | 71.13 | 0 | 0 | 21.5 | 175.68 | 2083271.9 | 96.7
0 | 0 | 122.5 | 2066.26 | 722056.8 | 40.73 | 0 | 0 | 105.5 | 95.61 | 990909.5 | 49.07 | 0 | 0 | 62.5 | 499.1 | 1505220.9 | 70.59 | 0 | 0 | 22.5 | 174.53 | 2069594.4 | 96.07
0 | 0 | 123.5 | 2051.19 | 716790 | 40.43 | 0 | 0 | 106.5 | 94.92 | 983809.9 | 48.71 | 0 | 0 | 63.5 | 495.27 | 1493685.4 | 70.04 | 0 | 0 | 23.5 | 173.37 | 2055916.8 | 95.43
0 | 0 | 124.5 | 2036.12 | 711523.3 | 40.14 | 0 | 0 | 107.5 | 94.24 | 976710.4 | 48.36 | 0 | 0 | 64.5 | 491.45 | 1482149.8 | 69.5 | 0 | 0 | 24.5 | 172.22 | 2042239.3 | 94.8
0 | 0 | 125.5 | 2021.34 | 706356.8 | 39.85 | 0 | 0 | 108.5 | 93.55 | 969610.8 | 48.01 | 0 | 0 | 65.5 | 487.66 | 1470722.5 | 68.96 | 0 | 0 | 25.5 | 171.05 | 2028395.8 | 94.16
0 | 0 | 126.5 | 2006.87 | 701302.6 | 39.56 | 0 | 0 | 109.5 | 92.87 | 962511.3 | 47.66 | 0 | 0 | 66.5 | 483.91 | 1459409.6 | 68.43 | 0 | 0 | 26.5 | 169.88 | 2014453.3 | 93.51
0 | 0 | 127.5 | 1992.41 | 696248.3 | 39.28 | 0 | 0 | 110.5 | 92.18 | 955377.3 | 47.31 | 0 | 0 | 67.5 | 480.16 | 1448096.7 | 67.9 | 0 | 0 | 27.5 | 168.7 | 2000510.8 | 92.86
0 | 0 | 128.5 | 1977.95 | 691194 | 38.99 | 0 | 0 | 111.5 | 91.49 | 948199.5 | 46.95 | 0 | 0 | 68.5 | 476.41 | 1436783.8 | 67.37 | 0 | 0 | 28.5 | 167.52 | 1986568.2 | 92.21
0 | 0 | 129.5 | 1963.48 | 686139.7 | 38.71 | 0 | 0 | 112.5 | 90.79 | 941021.7 | 46.59 | 0 | 0 | 69.5 | 472.65 | 1425470.8 | 66.83 | 0 | 0 | 29.5 | 166.35 | 1972625.7 | 91.57
0 | 0 | 130.5 | 1948.83 | 681019.1 | 38.42 | 0 | 0 | 113.5 | 90.1 | 933843.9 | 46.24 | 0 | 0 | 70.5 | 469.06 | 1414640.4 | 66.33 | 0 | 0 | 30.5 | 165.23 | 1959346.7 | 90.95
0 | 0 | 131.5 | 1933.9 | 675800.8 | 38.13 | 0 | 0 | 114.5 | 89.41 | 926666.2 | 45.88 | 0 | 0 | 71.5 | 465.64 | 1404325.2 | 65.84 | 0 | 0 | 31.5 | 164.15 | 1946527.4 | 90.35
0 | 0 | 132.5 | 1918.96 | 670582.5 | 37.84 | 0 | 0 | 115.5 | 88.73 | 919619.8 | 45.53 | 0 | 0 | 72.5 | 462.22 | 1394010 | 65.35 | 0 | 0 | 32.5 | 163.07 | 1933708.1 | 89.76
0 | 0 | 133.5 | 1904.03 | 665364.1 | 37.55 | 0 | 0 | 116.5 | 88.07 | 912753.9 | 45.19 | 0 | 0 | 73.5 | 458.8 | 1383694.8 | 64.87 | 0 | 0 | 33.5 | 161.99 | 1920888.8 | 89.16
0 | 0 | 134.5 | 1889.1 | 660145.8 | 37.25 | 0 | 0 | 117.5 | 87.4 | 905888 | 44.84 | 0 | 0 | 74.5 | 455.38 | 1373379.7 | 64.38 | 0 | 0 | 34.5 | 160.91 | 1908069.5 | 88.57
0 | 0 | 135.5 | 1874.45 | 655026 | 36.97 | 0 | 0 | 118.5 | 86.74 | 899022.1 | 44.5 | 0 | 0 | 75.5 | 451.97 | 1363075.5 | 63.9 | 0 | 0 | 35.5 | 159.79 | 1894873.1 | 87.96
0 | 0 | 136.5 | 1860.16 | 650033.5 | 36.7 | 0 | 0 | 119.5 | 86.08 | 892156.2 | 44.16 | 0 | 0 | 76.5 | 448.55 | 1352785.6 | 63.42 | 0 | 0 | 36.5 | 158.66 | 1881447.1 | 87.33
0 | 0 | 137.5 | 1845.87 | 645041.1 | 36.43 | 0 | 0 | 120.5 | 85.43 | 885417.7 | 43.83 | 0 | 0 | 77.5 | 445.14 | 1342495.7 | 62.94 | 0 | 0 | 37.5 | 157.53 | 1868021.1 | 86.71
0 | 0 | 138.5 | 1831.59 | 640048.7 | 36.16 | 0 | 0 | 121.5 | 84.79 | 878853.5 | 43.5 | 0 | 0 | 78.5 | 441.73 | 1332205.8 | 62.46 | 0 | 0 | 38.5 | 156.4 | 1854595 | 86.09
0 | 0 | 139.5 | 1817.3 | 635056.2 | 35.89 | 0 | 0 | 122.5 | 84.16 | 872289.4 | 43.18 | 0 | 0 | 79.5 | 438.32 | 1321915.9 | 61.98 | 0 | 0 | 39.5 | 155.26 | 1841169 | 85.46
0 | 0 | 140.5 | 1803.45 | 630217.8 | 35.61 | 0 | 0 | 123.5 | 83.53 | 865725.3 | 42.85 | 0 | 0 | 80.5 | 435.08 | 1312162.7 | 61.52 | 0 | 0 | 40.5 | 154.15 | 1827958.3 | 84.85
0 | 0 | 141.5 | 1790.14 | 625565.8 | 35.33 | 0 | 0 | 124.5 | 82.89 | 859161.2 | 42.53 | 0 | 0 | 81.5 | 432.02 | 1302931.7 | 61.08 | 0 | 0 | 41.5 | 153.05 | 1814919.7 | 84.25
0 | 0 | 142.5 | 1776.83 | 620913.9 | 35.04 | 0 | 0 | 125.5 | 82.27 | 852700.2 | 42.21 | 0 | 0 | 82.5 | 428.96 | 1293700.8 | 60.64 | 0 | 0 | 42.5 | 151.95 | 1801881 | 83.64
0 | 0 | 143.5 | 1763.52 | 616262 | 34.76 | 0 | 0 | 126.5 | 81.66 | 846365.9 | 41.9 | 0 | 0 | 83.5 | 425.9 | 1284469.8 | 60.21 | 0 | 0 | 43.5 | 150.85 | 1788842.4 | 83.04
0 | 0 | 144.5 | 1750.21 | 611610.1 | 34.48 | 0 | 0 | 127.5 | 81.05 | 840031.5 | 41.59 | 0 | 0 | 84.5 | 422.84 | 1275238.8 | 59.77 | 0 | 0 | 44.5 | 149.75 | 1775803.7 | 82.43
0 | 0 | 145.5 | 1737.22 | 607072 | 34.21 | 0 | 0 | 128.5 | 80.44 | 833697.2 | 41.28 | 0 | 0 | 85.5 | 419.71 | 1265793.2 | 59.32 | 0 | 0 | 45.5 | 148.66 | 1762911 | 81.83
0 | 0 | 146.5 | 1724.67 | 602687.9 | 33.96 | 0 | 0 | 129.5 | 79.83 | 827362.8 | 40.97 | 0 | 0 | 86.5 | 416.5 | 1256113 | 58.87 | 0 | 0 | 46.5 | 147.59 | 1750123.9 | 81.24
0 | 0 | 147.5 | 1712.13 | 598303.8 | 33.71 | 0 | 0 | 130.5 | 79.24 | 821242.5 | 40.66 | 0 | 0 | 87.5 | 413.29 | 1246432.7 | 58.42 | 0 | 0 | 47.5 | 146.51 | 1737336.9 | 80.64
0 | 0 | 148.5 | 1699.58 | 593919.7 | 33.46 | 0 | 0 | 131.5 | 78.67 | 815385.9 | 40.37 | 0 | 0 | 88.5 | 410.08 | 1236752.5 | 57.97 | 0 | 0 | 48.5 | 145.43 | 1724549.8 | 80.05
0 | 0 | 149.5 | 1687.04 | 589535.5 | 33.21 | 0 | 0 | 132.5 | 78.11 | 809529.2 | 40.08 | 0 | 0 | 89.5 | 406.87 | 1227072.3 | 57.52 | 0 | 0 | 49.5 | 144.35 | 1711762.8 | 79.46
0 | 0 | 150.5 | 1674.47 | 585144.3 | 32.96 | 0 | 0 | 133.5 | 77.54 | 803672.5 | 39.78 | 0 | 0 | 90.5 | 403.79 | 1217773.8 | 57.09 | 0 | 0 | 50.5 | 143.3 | 1699336.7 | 78.88
0 | 0 | 151.5 | 1661.88 | 580744.6 | 32.72 | 0 | 0 | 134.5 | 76.98 | 797815.9 | 39.49 | 0 | 0 | 91.5 | 400.87 | 1208970.1 | 56.67 | 0 | 0 | 51.5 | 142.28 | 1687197.2 | 78.32
0 | 0 | 152.5 | 1649.29 | 576344.8 | 32.48 | 0 | 0 | 135.5 | 76.41 | 792003 | 39.2 | 0 | 0 | 92.5 | 397.95 | 1200166.4 | 56.26 | 0 | 0 | 52.5 | 141.26 | 1675057.7 | 77.75
0 | 0 | 153.5 | 1636.7 | 571945 | 32.23 | 0 | 0 | 136.5 | 75.86 | 786241.4 | 38.91 | 0 | 0 | 93.5 | 395.03 | 1191362.6 | 55.85 | 0 | 0 | 53.5 | 140.23 | 1662918.2 | 77.19
0 | 0 | 154.5 | 1624.11 | 567545.3 | 31.99 | 0 | 0 | 137.5 | 75.3 | 780479.9 | 38.63 | 0 | 0 | 94.5 | 392.11 | 1182558.9 | 55.43 | 0 | 0 | 54.5 | 139.21 | 1650778.7 | 76.63
0 | 0 | 155.5 | 1611.83 | 563256.4 | 31.75 | 0 | 0 | 138.5 | 74.75 | 774718.4 | 38.34 | 0 | 0 | 95.5 | 389.22 | 1173852 | 55.02 | 0 | 0 | 55.5 | 138.18 | 1638610.5 | 76.06
0 | 0 | 156.5 | 1599.97 | 559110.9 | 31.52 | 0 | 0 | 139.5 | 74.19 | 768956.9 | 38.06 | 0 | 0 | 96.5 | 386.38 | 1165274.3 | 54.62 | 0 | 0 | 56.5 | 137.15 | 1626416.9 | 75.5
0 | 0 | 157.5 | 1588.11 | 554965.5 | 31.3 | 0 | 0 | 140.5 | 73.66 | 763410.1 | 37.78 | 0 | 0 | 97.5 | 383.53 | 1156696.6 | 54.22 | 0 | 0 | 57.5 | 136.13 | 1614223.3 | 74.93
0 | 0 | 158.5 | 1576.25 | 550820 | 31.07 | 0 | 0 | 141.5 | 73.14 | 758094.5 | 37.52 | 0 | 0 | 98.5 | 380.69 | 1148118.9 | 53.81 | 0 | 0 | 58.5 | 135.1 | 1602029.7 | 74.36
0 | 0 | 159.5 | 1564.38 | 546674.5 | 30.84 | 0 | 0 | 142.5 | 72.63 | 752778.9 | 37.26 | 0 | 0 | 99.5 | 377.85 | 1139541.2 | 53.41 | 0 | 0 | 59.5 | 134.07 | 1589836.1 | 73.8
0 | 0 | 160.5 | 1553.08 | 542725.6 | 30.62 | 0 | 0 | 143.5 | 72.12 | 747463.2 | 37 | 0 | 0 | 100.5 | 374.96 | 1130843.7 | 53 | 0 | 0 | 60.5 | 133.08 | 1578121.9 | 73.25
0 | 0 | 161.5 | 1542.62 | 539068.2 | 30.41 | 0 | 0 | 144.5 | 71.6 | 742147.6 | 36.73 | 0 | 0 | 101.5 | 372.02 | 1121984.8 | 52.59 | 0 | 0 | 61.5 | 132.13 | 1566858 | 72.73
0 | 0 | 162.5 | 1532.15 | 535410.8 | 30.2 | 0 | 0 | 145.5 | 71.08 | 736712.7 | 36.47 | 0 | 0 | 102.5 | 369.09 | 1113125.9 | 52.17 | 0 | 0 | 62.5 | 131.18 | 1555594 | 72.21
0 | 0 | 163.5 | 1521.68 | 531753.3 | 29.99 | 0 | 0 | 146.5 | 70.54 | 731123 | 36.19 | 0 | 0 | 103.5 | 366.15 | 1104267 | 51.76 | 0 | 0 | 63.5 | 130.23 | 1544330.1 | 71.69
0 | 0 | 164.5 | 1511.22 | 528095.9 | 29.77 | 0 | 0 | 147.5 | 70 | 725533.2 | 35.92 | 0 | 0 | 104.5 | 363.21 | 1095408.1 | 51.34 | 0 | 0 | 64.5 | 129.28 | 1533066.2 | 71.16
0 | 0 | 165.5 | 1500.35 | 524296.8 | 29.56 | 0 | 0 | 148.5 | 69.46 | 719943.5 | 35.64 | 0 | 0 | 105.5 | 360.46 | 1087118.5 | 50.95 | 0 | 0 | 65.5 | 128.35 | 1521969 | 70.65
0 | 0 | 166.5 | 1488.98 | 520324.2 | 29.33 | 0 | 0 | 149.5 | 68.92 | 714353.8 | 35.36 | 0 | 0 | 106.5 | 358 | 1079691.1 | 50.6 | 0 | 0 | 66.5 | 127.42 | 1511028.2 | 70.14
0 | 0 | 167.5 | 1477.61 | 516351.5 | 29.11 | 0 | 0 | 150.5 | 68.42 | 709103 | 35.1 | 0 | 0 | 107.5 | 355.54 | 1072263.7 | 50.25 | 0 | 0 | 67.5 | 126.5 | 1500087.3 | 69.63
0 | 0 | 168.5 | 1466.24 | 512378.8 | 28.88 | 0 | 0 | 151.5 | 67.95 | 704262.9 | 34.86 | 0 | 0 | 108.5 | 353.08 | 1064836.3 | 49.9 | 0 | 0 | 68.5 | 125.58 | 1489146.4 | 69.12
0 | 0 | 169.5 | 1454.87 | 508406.1 | 28.66 | 0 | 0 | 152.5 | 67.48 | 699422.8 | 34.62 | 0 | 0 | 109.5 | 350.61 | 1057409 | 49.55 | 0 | 0 | 69.5 | 124.66 | 1478205.6 | 68.62
0 | 0 | 170.5 | 1444.09 | 504637 | 28.45 | 0 | 0 | 153.5 | 67.02 | 694582.7 | 34.37 | 0 | 0 | 110.5 | 348.09 | 1049803.4 | 49.2 | 0 | 0 | 70.5 | 123.74 | 1467311.9 | 68.11
0 | 0 | 171.5 | 1434.03 | 501121 | 28.25 | 0 | 0 | 154.5 | 66.55 | 689742.6 | 34.13 | 0 | 0 | 111.5 | 345.49 | 1041952.8 | 48.83 | 0 | 0 | 71.5 | 122.82 | 1456464.7 | 67.61
0 | 0 | 172.5 | 1423.96 | 497604.9 | 28.05 | 0 | 0 | 155.5 | 66.1 | 685067 | 33.89 | 0 | 0 | 112.5 | 342.89 | 1034102.2 | 48.46 | 0 | 0 | 72.5 | 121.91 | 1445617.5 | 67.1
0 | 0 | 173.5 | 1413.9 | 494088.8 | 27.86 | 0 | 0 | 156.5 | 65.67 | 680614.9 | 33.67 | 0 | 0 | 113.5 | 340.28 | 1026251.6 | 48.1 | 0 | 0 | 73.5 | 120.99 | 1434770.3 | 66.6
0 | 0 | 174.5 | 1403.84 | 490572.8 | 27.66 | 0 | 0 | 157.5 | 65.24 | 676162.8 | 33.45 | 0 | 0 | 114.5 | 337.68 | 1018401 | 47.73 | 0 | 0 | 74.5 | 120.08 | 1423923.1 | 66.1
0 | 0 | 175.5 | 1393.23 | 486866.1 | 27.46 | 0 | 0 | 158.5 | 64.81 | 671710.7 | 33.23 | 0 | 0 | 115.5 | 335.21 | 1010969.7 | 47.38 | 0 | 0 | 75.5 | 119.18 | 1413243.1 | 65.6
0 | 0 | 176.5 | 1381.9 | 482907 | 27.24 | 0 | 0 | 159.5 | 64.38 | 667258.6 | 33 | 0 | 0 | 116.5 | 332.92 | 1004043 | 47.06 | 0 | 0 | 76.5 | 118.29 | 1402753.1 | 65.11
0 | 0 | 177.5 | 1370.57 | 478947.8 | 27.02 | 0 | 0 | 160.5 | 63.92 | 662546.5 | 32.77 | 0 | 0 | 117.5 | 330.62 | 997116.3 | 46.74 | 0 | 0 | 77.5 | 117.41 | 1392263.2 | 64.63
0 | 0 | 178.5 | 1359.24 | 474988.6 | 26.81 | 0 | 0 | 161.5 | 63.44 | 657502 | 32.52 | 0 | 0 | 118.5 | 328.32 | 990189.7 | 46.41 | 0 | 0 | 78.5 | 116.52 | 1381773.3 | 64.14
0 | 0 | 179.5 | 1347.91 | 471029.4 | 26.59 | 0 | 0 | 162.5 | 62.95 | 652457.5 | 32.28 | 0 | 0 | 119.5 | 326.03 | 983263 | 46.09 | 0 | 0 | 79.5 | 115.64 | 1371283.3 | 63.65
0 | 0 | 180.5 | 1337.4 | 467354.5 | 26.39 | 0 | 0 | 163.5 | 62.46 | 647413 | 32.03 | 0 | 0 | 120.5 | 323.67 | 976140.5 | 45.75 | 0 | 0 | 80.5 | 114.79 | 1361208.1 | 63.19
0 | 0 | 181.5 | 1327.93 | 464047.1 | 26.2 | 0 | 0 | 164.5 | 61.98 | 642368.5 | 31.78 | 0 | 0 | 121.5 | 321.23 | 968778.6 | 45.4 | 0 | 0 | 81.5 | 113.98 | 1351591.6 | 62.74
0 | 0 | 182.5 | 1318.47 | 460739.7 | 26.02 | 0 | 0 | 165.5 | 61.52 | 637671.2 | 31.55 | 0 | 0 | 122.5 | 318.78 | 961416.6 | 45.06 | 0 | 0 | 82.5 | 113.17 | 1341975.1 | 62.29
0 | 0 | 183.5 | 1309 | 457432.3 | 25.83 | 0 | 0 | 166.5 | 61.11 | 633370.4 | 31.34 | 0 | 0 | 123.5 | 316.34 | 954054.7 | 44.71 | 0 | 0 | 83.5 | 112.36 | 1332358.6 | 61.85
0 | 0 | 184.5 | 1299.54 | 454124.9 | 25.64 | 0 | 0 | 167.5 | 60.69 | 629069.7 | 31.12 | 0 | 0 | 124.5 | 313.9 | 946692.8 | 44.36 | 0 | 0 | 84.5 | 111.55 | 1322742.1 | 61.4
0 | 0 | 185.5 | 1290.12 | 450832.7 | 25.45 | 0 | 0 | 168.5 | 60.28 | 624768.9 | 30.91 | 0 | 0 | 125.5 | 311.53 | 939548.9 | 44.03 | 0 | 0 | 85.5 | 110.73 | 1313080.8 | 60.95
0 | 0 | 186.5 | 1280.75 | 447560.2 | 25.26 | 0 | 0 | 169.5 | 59.86 | 620468.1 | 30.7 | 0 | 0 | 126.5 | 309.27 | 932719.9 | 43.71 | 0 | 0 | 86.5 | 109.91 | 1303371.4 | 60.5
0 | 0 | 187.5 | 1271.39 | 444287.7 | 25.07 | 0 | 0 | 170.5 | 59.46 | 616226.5 | 30.49 | 0 | 0 | 127.5 | 307 | 925890.9 | 43.39 | 0 | 0 | 87.5 | 109.09 | 1293662 | 60.05
0 | 0 | 188.5 | 1262.03 | 441015.2 | 24.87 | 0 | 0 | 171.5 | 59.05 | 612069.7 | 30.28 | 0 | 0 | 128.5 | 304.74 | 919061.9 | 43.07 | 0 | 0 | 88.5 | 108.27 | 1283952.6 | 59.6
0 | 0 | 189.5 | 1252.66 | 437742.7 | 24.68 | 0 | 0 | 172.5 | 58.65 | 607912.8 | 30.08 | 0 | 0 | 129.5 | 302.48 | 912233 | 42.75 | 0 | 0 | 89.5 | 107.46 | 1274243.1 | 59.15
0 | 0 | 190.5 | 1243.48 | 434533.2 | 24.49 | 0 | 0 | 173.5 | 58.25 | 603756 | 29.87 | 0 | 0 | 130.5 | 300.31 | 905706.6 | 42.45 | 0 | 0 | 90.5 | 106.67 | 1264870.5 | 58.71
0 | 0 | 191.5 | 1234.52 | 431404 | 24.32 | 0 | 0 | 174.5 | 57.85 | 599599.1 | 29.67 | 0 | 0 | 131.5 | 298.29 | 899619.3 | 42.16 | 0 | 0 | 91.5 | 105.91 | 1255971.5 | 58.3
0 | 0 | 192.5 | 1225.57 | 428274.7 | 24.15 | 0 | 0 | 175.5 | 57.44 | 595331.9 | 29.46 | 0 | 0 | 132.5 | 296.28 | 893531.9 | 41.87 | 0 | 0 | 92.5 | 105.16 | 1247072.5 | 57.89
0 | 0 | 193.5 | 1216.61 | 425145.5 | 23.98 | 0 | 0 | 176.5 | 57.01 | 590924.2 | 29.24 | 0 | 0 | 133.5 | 294.26 | 887444.6 | 41.58 | 0 | 0 | 93.5 | 104.41 | 1238173.5 | 57.47
0 | 0 | 194.5 | 1207.66 | 422016.3 | 23.81 | 0 | 0 | 177.5 | 56.59 | 586516.6 | 29.03 | 0 | 0 | 134.5 | 292.24 | 881357.2 | 41.29 | 0 | 0 | 94.5 | 103.66 | 1229274.5 | 57.06
0 | 0 | 195.5 | 1199.49 | 419163 | 23.65 | 0 | 0 | 178.5 | 56.16 | 582108.9 | 28.81 | 0 | 0 | 135.5 | 290.11 | 874952.4 | 40.98 | 0 | 0 | 95.5 | 102.91 | 1220292.6 | 56.64
0 | 0 | 196.5 | 1192.21 | 416616.9 | 23.5 | 0 | 0 | 179.5 | 55.74 | 577701.3 | 28.6 | 0 | 0 | 136.5 | 287.84 | 868096.8 | 40.67 | 0 | 0 | 96.5 | 102.14 | 1211201.2 | 56.22
0 | 0 | 197.5 | 1184.92 | 414070.7 | 23.36 | 0 | 0 | 180.5 | 55.32 | 573352.9 | 28.38 | 0 | 0 | 137.5 | 285.57 | 861241.2 | 40.35 | 0 | 0 | 97.5 | 101.37 | 1202109.8 | 55.8
0 | 0 | 198.5 | 1177.63 | 411524.5 | 23.21 | 0 | 0 | 181.5 | 54.91 | 569071.3 | 28.17 | 0 | 0 | 138.5 | 283.3 | 854385.7 | 40.04 | 0 | 0 | 98.5 | 100.61 | 1193018.4 | 55.38
0 | 0 | 199.5 | 1170.35 | 408978.3 | 23.07 | 0 | 0 | 182.5 | 54.49 | 564789.7 | 27.95 | 0 | 0 | 139.5 | 281.02 | 847530.1 | 39.72 | 0 | 0 | 99.5 | 99.84 | 1183927 | 54.96
0 | 0 | 200.5 | 1162.63 | 406283 | 22.92 | 0 | 0 | 183.5 | 54.08 | 560508.1 | 27.74 | 0 | 0 | 140.5 | 278.88 | 841080.3 | 39.42 | 0 | 0 | 100.5 | 99.11 | 1175265.3 | 54.55
0 | 0 | 201.5 | 1154.39 | 403400.7 | 22.75 | 0 | 0 | 184.5 | 53.67 | 556226.5 | 27.53 | 0 | 0 | 141.5 | 276.91 | 835122.3 | 39.14 | 0 | 0 | 101.5 | 98.42 | 1167105.9 | 54.18
0 | 0 | 202.5 | 1146.14 | 400518.5 | 22.59 | 0 | 0 | 185.5 | 53.3 | 552394.7 | 27.34 | 0 | 0 | 142.5 | 274.93 | 829164.3 | 38.86 | 0 | 0 | 102.5 | 97.73 | 1158946.5 | 53.8
0 | 0 | 203.5 | 1137.89 | 397636.2 | 22.43 | 0 | 0 | 186.5 | 52.98 | 549140.3 | 27.18 | 0 | 0 | 143.5 | 272.96 | 823206.3 | 38.58 | 0 | 0 | 103.5 | 97.04 | 1150787.2 | 53.42
0 | 0 | 204.5 | 1129.64 | 394754 | 22.26 | 0 | 0 | 187.5 | 52.67 | 545886 | 27.02 | 0 | 0 | 144.5 | 270.98 | 817248.3 | 38.3 | 0 | 0 | 104.5 | 96.36 | 1142627.8 | 53.04
0 | 0 | 205.5 | 1121.23 | 391813 | 22.1 | 0 | 0 | 188.5 | 52.35 | 542631.6 | 26.86 | 0 | 0 | 145.5 | 269.03 | 811365.9 | 38.02 | 0 | 0 | 105.5 | 95.66 | 1134319.9 | 52.65
0 | 0 | 206.5 | 1112.59 | 388795.7 | 21.93 | 0 | 0 | 189.5 | 52.04 | 539377.2 | 26.7 | 0 | 0 | 146.5 | 267.11 | 805577 | 37.75 | 0 | 0 | 106.5 | 94.94 | 1125816.9 | 52.26
0 | 0 | 207.5 | 1103.96 | 385778.5 | 21.76 | 0 | 0 | 190.5 | 51.69 | 535765.5 | 26.52 | 0 | 0 | 147.5 | 265.19 | 799788.2 | 37.48 | 0 | 0 | 107.5 | 94.22 | 1117314 | 51.86
0 | 0 | 208.5 | 1095.32 | 382761.2 | 21.59 | 0 | 0 | 191.5 | 51.3 | 531724.4 | 26.32 | 0 | 0 | 148.5 | 263.27 | 793999.3 | 37.21 | 0 | 0 | 108.5 | 93.5 | 1108811 | 51.47
0 | 0 | 209.5 | 1086.69 | 379744 | 21.43 | 0 | 0 | 192.5 | 50.91 | 527683.3 | 26.12 | 0 | 0 | 149.5 | 261.35 | 788210.4 | 36.94 | 0 | 0 | 109.5 | 92.79 | 1100308 | 51.07
0 | 0 | 210.5 | 1078.69 | 376948.6 | 21.27 | 0 | 0 | 193.5 | 50.52 | 523642.1 | 25.92 | 0 | 0 | 150.5 | 259.48 | 782574.2 | 36.67 | 0 | 0 | 110.5 | 92.11 | 1092300.4 | 50.7
0 | 0 | 211.5 | 1071.46 | 374421.1 | 21.13 | 0 | 0 | 194.5 | 50.13 | 519601 | 25.71 | 0 | 0 | 151.5 | 257.68 | 777146.9 | 36.42 | 0 | 0 | 111.5 | 91.49 | 1084920.2 | 50.36
0 | 0 | 212.5 | 1064.22 | 371893.6 | 20.99 | 0 | 0 | 195.5 | 49.78 | 515947.9 | 25.53 | 0 | 0 | 152.5 | 255.88 | 771719.6 | 36.16 | 0 | 0 | 112.5 | 90.87 | 1077540.1 | 50.02
0 | 0 | 213.5 | 1056.99 | 369366.1 | 20.85 | 0 | 0 | 196.5 | 49.47 | 512750.6 | 25.37 | 0 | 0 | 153.5 | 254.09 | 766292.3 | 35.91 | 0 | 0 | 113.5 | 90.25 | 1070159.9 | 49.68
0 | 0 | 214.5 | 1049.76 | 366838.5 | 20.71 | 0 | 0 | 197.5 | 49.16 | 509553.2 | 25.21 | 0 | 0 | 154.5 | 252.29 | 760865 | 35.65 | 0 | 0 | 114.5 | 89.62 | 1062779.8 | 49.33
0 | 0 | 215.5 | 1042.67 | 364359.9 | 20.57 | 0 | 0 | 198.5 | 48.85 | 506355.8 | 25.05 | 0 | 0 | 155.5 | 250.52 | 755551.9 | 35.4 | 0 | 0 | 115.5 | 88.98 | 1055200.3 | 48.98
0 | 0 | 216.5 | 1035.74 | 361941.2 | 20.43 | 0 | 0 | 199.5 | 48.55 | 503158.4 | 24.89 | 0 | 0 | 156.5 | 248.81 | 750376.3 | 35.16 | 0 | 0 | 116.5 | 88.32 | 1047347.9 | 48.62
0 | 0 | 217.5 | 1028.82 | 359522.4 | 20.29 | 0 | 0 | 200.5 | 48.22 | 499776.3 | 24.72 | 0 | 0 | 157.5 | 247.09 | 745200.6 | 34.92 | 0 | 0 | 117.5 | 87.66 | 1039495.6 | 48.25
0 | 0 | 218.5 | 1021.9 | 357103.7 | 20.15 | 0 | 0 | 201.5 | 47.87 | 496172.1 | 24.54 | 0 | 0 | 158.5 | 245.38 | 740025 | 34.68 | 0 | 0 | 118.5 | 87 | 1031643.3 | 47.89
0 | 0 | 219.5 | 1014.98 | 354685 | 20.01 | 0 | 0 | 202.5 | 47.52 | 492567.9 | 24.37 | 0 | 0 | 159.5 | 243.66 | 734849.3 | 34.43 | 0 | 0 | 119.5 | 86.34 | 1023791 | 47.52
0 | 0 | 220.5 | 1008.08 | 352275.4 | 19.87 | 0 | 0 | 203.5 | 47.18 | 488963.7 | 24.19 | 0 | 0 | 160.5 | 241.96 | 729711.1 | 34.19 | 0 | 0 | 120.5 | 85.69 | 1016170.1 | 47.17
0 | 0 | 221.5 | 1001.22 | 349876.6 | 19.73 | 0 | 0 | 204.5 | 46.83 | 485359.5 | 24.01 | 0 | 0 | 161.5 | 240.27 | 724624.8 | 33.95 | 0 | 0 | 121.5 | 85.07 | 1008832.4 | 46.83
0 | 0 | 222.5 | 994.36 | 347477.9 | 19.58 | 0 | 0 | 205.5 | 46.48 | 481790.7 | 23.83 | 0 | 0 | 162.5 | 238.58 | 719538.4 | 33.71 | 0 | 0 | 122.5 | 84.45 | 1001494.8 | 46.49
0 | 0 | 223.5 | 987.49 | 345079.1 | 19.44 | 0 | 0 | 206.5 | 46.14 | 478268.4 | 23.66 | 0 | 0 | 163.5 | 236.9 | 714452.1 | 33.47 | 0 | 0 | 123.5 | 83.84 | 994157.1 | 46.15
0 | 0 | 224.5 | 980.63 | 342680.4 | 19.3 | 0 | 0 | 207.5 | 45.8 | 474746.1 | 23.49 | 0 | 0 | 164.5 | 235.21 | 709365.7 | 33.23 | 0 | 0 | 124.5 | 83.22 | 986819.5 | 45.81
0 | 0 | 225.5 | 973.54 | 340203.7 | 19.16 | 0 | 0 | 208.5 | 45.47 | 471223.8 | 23.32 | 0 | 0 | 165.5 | 233.59 | 704483.3 | 33 | 0 | 0 | 125.5 | 82.61 | 979673.6 | 45.47
0 | 0 | 226.5 | 966.2 | 337639.7 | 19.02 | 0 | 0 | 209.5 | 45.13 | 467701.5 | 23.14 | 0 | 0 | 166.5 | 232.06 | 699863.4 | 32.79 | 0 | 0 | 126.5 | 82.03 | 972719.3 | 45.15
0 | 0 | 227.5 | 958.86 | 335075.7 | 18.88 | 0 | 0 | 210.5 | 44.82 | 464583.7 | 22.99 | 0 | 0 | 167.5 | 230.53 | 695243.5 | 32.57 | 0 | 0 | 127.5 | 81.44 | 965765.1 | 44.83
0 | 0 | 228.5 | 951.53 | 332511.7 | 18.74 | 0 | 0 | 211.5 | 44.56 | 461891.6 | 22.86 | 0 | 0 | 168.5 | 229 | 690623.6 | 32.36 | 0 | 0 | 128.5 | 80.86 | 958810.8 | 44.51
0 | 0 | 229.5 | 944.19 | 329947.7 | 18.61 | 0 | 0 | 212.5 | 44.3 | 459199.4 | 22.72 | 0 | 0 | 169.5 | 227.46 | 686003.7 | 32.15 | 0 | 0 | 129.5 | 80.27 | 951856.6 | 44.18
0 | 0 | 230.5 | 937.22 | 327513.1 | 18.47 | 0 | 0 | 213.5 | 44.05 | 456507.2 | 22.59 | 0 | 0 | 170.5 | 225.92 | 681335.1 | 31.93 | 0 | 0 | 130.5 | 79.7 | 945121.6 | 43.87
0 | 0 | 231.5 | 930.68 | 325226.9 | 18.34 | 0 | 0 | 214.5 | 43.79 | 453815.1 | 22.46 | 0 | 0 | 171.5 | 224.35 | 676603.6 | 31.7 | 0 | 0 | 131.5 | 79.16 | 938656.6 | 43.57
0 | 0 | 232.5 | 924.14 | 322940.7 | 18.2 | 0 | 0 | 215.5 | 43.51 | 450924.9 | 22.32 | 0 | 0 | 172.5 | 222.78 | 671872.1 | 31.48 | 0 | 0 | 132.5 | 78.61 | 932191.6 | 43.27
0 | 0 | 233.5 | 917.6 | 320654.5 | 18.07 | 0 | 0 | 216.5 | 43.2 | 447775.3 | 22.16 | 0 | 0 | 173.5 | 221.21 | 667140.5 | 31.26 | 0 | 0 | 133.5 | 78.07 | 925726.6 | 42.97
0 | 0 | 234.5 | 911.05 | 318368.3 | 17.93 | 0 | 0 | 217.5 | 42.9 | 444625.7 | 22 | 0 | 0 | 174.5 | 219.64 | 662409 | 31.04 | 0 | 0 | 134.5 | 77.52 | 919261.6 | 42.67
0 | 0 | 235.5 | 904.84 | 316196.7 | 17.81 | 0 | 0 | 218.5 | 42.59 | 441476.1 | 21.85 | 0 | 0 | 175.5 | 218.1 | 657755.2 | 30.82 | 0 | 0 | 135.5 | 76.96 | 912598.7 | 42.36
0 | 0 | 236.5 | 899.02 | 314161.6 | 17.7 | 0 | 0 | 219.5 | 42.29 | 438326.4 | 21.69 | 0 | 0 | 176.5 | 216.59 | 653207.1 | 30.61 | 0 | 0 | 136.5 | 76.37 | 905669.1 | 42.04
0 | 0 | 237.5 | 893.19 | 312126.5 | 17.59 | 0 | 0 | 220.5 | 41.99 | 435228.8 | 21.54 | 0 | 0 | 177.5 | 215.08 | 648658.9 | 30.39 | 0 | 0 | 137.5 | 75.79 | 898739.4 | 41.72
0 | 0 | 238.5 | 887.37 | 310091.4 | 17.48 | 0 | 0 | 221.5 | 41.7 | 432191.7 | 21.39 | 0 | 0 | 178.5 | 213.57 | 644110.7 | 30.18 | 0 | 0 | 138.5 | 75.21 | 891809.8 | 41.4
0 | 0 | 239.5 | 881.55 | 308056.4 | 17.37 | 0 | 0 | 222.5 | 41.41 | 429154.6 | 21.23 | 0 | 0 | 179.5 | 212.06 | 639562.6 | 29.97 | 0 | 0 | 139.5 | 74.62 | 884880.1 | 41.07
0 | 0 | 240.5 | 875.38 | 305901.5 | 17.25 | 0 | 0 | 223.5 | 41.11 | 426117.5 | 21.08 | 0 | 0 | 180.5 | 210.56 | 635030 | 29.75 | 0 | 0 | 140.5 | 74.08 | 878428.1 | 40.78
0 | 0 | 241.5 | 868.8 | 303602.8 | 17.12 | 0 | 0 | 224.5 | 40.82 | 423080.4 | 20.93 | 0 | 0 | 181.5 | 209.07 | 630516.5 | 29.54 | 0 | 0 | 141.5 | 73.58 | 872559.8 | 40.5
0 | 0 | 242.5 | 862.22 | 301304.2 | 17 | 0 | 0 | 225.5 | 40.54 | 420150.2 | 20.79 | 0 | 0 | 182.5 | 207.57 | 626003.1 | 29.32 | 0 | 0 | 142.5 | 73.09 | 866691.6 | 40.23
0 | 0 | 243.5 | 855.64 | 299005.5 | 16.87 | 0 | 0 | 226.5 | 40.27 | 417344.4 | 20.65 | 0 | 0 | 183.5 | 206.07 | 621489.6 | 29.11 | 0 | 0 | 143.5 | 72.59 | 860823.3 | 39.96
0 | 0 | 244.5 | 849.07 | 296706.8 | 16.74 | 0 | 0 | 227.5 | 40 | 414538.6 | 20.51 | 0 | 0 | 184.5 | 204.58 | 616976.2 | 28.9 | 0 | 0 | 144.5 | 72.1 | 854955.1 | 39.69
0 | 0 | 245.5 | 843.14 | 294635.6 | 16.62 | 0 | 0 | 228.5 | 39.73 | 411732.9 | 20.37 | 0 | 0 | 185.5 | 203.15 | 612670.9 | 28.7 | 0 | 0 | 145.5 | 71.6 | 849037.1 | 39.41
0 | 0 | 246.5 | 837.93 | 292814.8 | 16.52 | 0 | 0 | 229.5 | 39.45 | 408927.1 | 20.23 | 0 | 0 | 186.5 | 201.81 | 608632.6 | 28.51 | 0 | 0 | 146.5 | 71.09 | 843050.4 | 39.13
0 | 0 | 247.5 | 832.72 | 290994 | 16.41 | 0 | 0 | 230.5 | 39.18 | 406071.2 | 20.09 | 0 | 0 | 187.5 | 200.47 | 604594.4 | 28.32 | 0 | 0 | 147.5 | 70.59 | 837063.7 | 38.86
0 | 0 | 248.5 | 827.51 | 289173.2 | 16.3 | 0 | 0 | 231.5 | 38.9 | 403155.4 | 19.95 | 0 | 0 | 188.5 | 199.13 | 600556.2 | 28.13 | 0 | 0 | 148.5 | 70.08 | 831077 | 38.58
0 | 0 | 249.5 | 822.3 | 287352.4 | 16.19 | 0 | 0 | 232.5 | 38.62 | 400239.6 | 19.8 | 0 | 0 | 189.5 | 197.79 | 596517.9 | 27.95 | 0 | 0 | 149.5 | 69.58 | 825090.3 | 38.3
[6 empty] | 0 | 0 | 233.5 | 38.33 | 397323.8 | 19.66 | 0 | 0 | 190.5 | 196.39 | 592304.3 | 27.75 | 0 | 0 | 150.5 | 69.07 | 819107.7 | 38.02
[6 empty] | 0 | 0 | 234.5 | 38.05 | 394408 | 19.51 | 0 | 0 | 191.5 | 194.93 | 587887.7 | 27.54 | 0 | 0 | 151.5 | 68.57 | 813130.2 | 37.74
[6 empty] | 0 | 0 | 235.5 | 37.79 | 391713.2 | 19.38 | 0 | 0 | 192.5 | 193.47 | 583471.1 | 27.33 | 0 | 0 | 152.5 | 68.07 | 807152.7 | 37.47
[6 empty] | 0 | 0 | 236.5 | 37.56 | 389285.7 | 19.26 | 0 | 0 | 193.5 | 192 | 579054.4 | 27.13 | 0 | 0 | 153.5 | 67.56 | 801175.2 | 37.19
[6 empty] | 0 | 0 | 237.5 | 37.33 | 386858.3 | 19.14 | 0 | 0 | 194.5 | 190.54 | 574637.8 | 26.92 | 0 | 0 | 154.5 | 67.06 | 795197.7 | 36.91
[6 empty] | 0 | 0 | 238.5 | 37.09 | 384430.8 | 19.03 | 0 | 0 | 195.5 | 189.2 | 570614.6 | 26.73 | 0 | 0 | 155.5 | 66.59 | 789631.9 | 36.65
[6 empty] | 0 | 0 | 239.5 | 36.86 | 382003.4 | 18.91 | 0 | 0 | 196.5 | 188.01 | 567014.6 | 26.56 | 0 | 0 | 156.5 | 66.16 | 784586.3 | 36.42
[6 empty] | 0 | 0 | 240.5 | 36.61 | 379426.4 | 18.78 | 0 | 0 | 197.5 | 186.82 | 563414.5 | 26.39 | 0 | 0 | 157.5 | 65.74 | 779540.8 | 36.19
[6 empty] | 0 | 0 | 241.5 | 36.34 | 376693.1 | 18.65 | 0 | 0 | 198.5 | 185.62 | 559814.4 | 26.22 | 0 | 0 | 158.5 | 65.31 | 774495.2 | 35.95
[6 empty] | 0 | 0 | 242.5 | 36.08 | 373959.8 | 18.51 | 0 | 0 | 199.5 | 184.43 | 556214.4 | 26.05 | 0 | 0 | 159.5 | 64.89 | 769449.6 | 35.72
[6 empty] | 0 | 0 | 243.5 | 35.82 | 371226.5 | 18.37 | 0 | 0 | 200.5 | 183.18 | 552461.6 | 25.87 | 0 | 0 | 160.5 | 64.46 | 764353.2 | 35.48
[6 empty] | 0 | 0 | 244.5 | 35.55 | 368493.2 | 18.24 | 0 | 0 | 201.5 | 181.87 | 548490.2 | 25.69 | 0 | 0 | 161.5 | 64.02 | 759186.4 | 35.24
[6 empty] | 0 | 0 | 245.5 | 35.31 | 365956.8 | 18.11 | 0 | 0 | 202.5 | 180.55 | 544518.8 | 25.5 | 0 | 0 | 162.5 | 63.59 | 754019.6 | 35
[6 empty] | 0 | 0 | 246.5 | 35.08 | 363637.2 | 17.99 | 0 | 0 | 203.5 | 179.23 | 540547.4 | 25.32 | 0 | 0 | 163.5 | 63.15 | 748852.7 | 34.76
[6 empty] | 0 | 0 | 247.5 | 34.86 | 361317.6 | 17.87 | 0 | 0 | 204.5 | 177.92 | 536576 | 25.14 | 0 | 0 | 164.5 | 62.71 | 743685.9 | 34.52
[6 empty] | 0 | 0 | 248.5 | 34.64 | 358998 | 17.76 | 0 | 0 | 205.5 | 176.71 | 532929.1 | 24.97 | 0 | 0 | 165.5 | 62.28 | 738492.2 | 34.28
[6 empty] | 0 | 0 | 249.5 | 34.41 | 356678.4 | 17.64 | 0 | 0 | 206.5 | 175.63 | 529687.2 | 24.81 | 0 | 0 | 166.5 | 61.84 | 733265.4 | 34.04
[12 empty] | 0 | 0 | 207.5 | 174.56 | 526445.3 | 24.66 | 0 | 0 | 167.5 | 61.39 | 728038.6 | 33.79
[12 empty] | 0 | 0 | 208.5 | 173.48 | 523203.4 | 24.51 | 0 | 0 | 168.5 | 60.95 | 722811.8 | 33.55
[12 empty] | 0 | 0 | 209.5 | 172.41 | 519961.5 | 24.35 | 0 | 0 | 169.5 | 60.51 | 717585 | 33.31
[12 empty] | 0 | 0 | 210.5 | 171.3 | 516620.5 | 24.19 | 0 | 0 | 170.5 | 60.06 | 712229.4 | 33.06
[12 empty] | 0 | 0 | 211.5 | 170.15 | 513147.4 | 24.03 | 0 | 0 | 171.5 | 59.6 | 706700.1 | 32.8
[12 empty] | 0 | 0 | 212.5 | 169 | 509674.3 | 23.87 | 0 | 0 | 172.5 | 59.13 | 701170.8 | 32.55
[12 empty] | 0 | 0 | 213.5 | 167.84 | 506201.1 | 23.71 | 0 | 0 | 173.5 | 58.66 | 695641.5 | 32.29
[12 empty] | 0 | 0 | 214.5 | 166.69 | 502728 | 23.54 | 0 | 0 | 174.5 | 58.2 | 690112.3 | 32.03
[12 empty] | 0 | 0 | 215.5 | 165.51 | 499173.7 | 23.38 | 0 | 0 | 175.5 | 57.78 | 685118.1 | 31.8
[12 empty] | 0 | 0 | 216.5 | 164.3 | 495515.7 | 23.2 | 0 | 0 | 176.5 | 57.42 | 680862.6 | 31.6
[12 empty] | 0 | 0 | 217.5 | 163.09 | 491857.6 | 23.03 | 0 | 0 | 177.5 | 57.06 | 676607.1 | 31.41
[12 empty] | 0 | 0 | 218.5 | 161.88 | 488199.6 | 22.86 | 0 | 0 | 178.5 | 56.7 | 672351.6 | 31.21
[12 empty] | 0 | 0 | 219.5 | 160.66 | 484541.6 | 22.69 | 0 | 0 | 179.5 | 56.34 | 668096.2 | 31.01
[12 empty] | 0 | 0 | 220.5 | 159.55 | 481172.5 | 22.53 | 0 | 0 | 180.5 | 55.96 | 663561.4 | 30.8
[12 empty] | 0 | 0 | 221.5 | 158.55 | 478168 | 22.39 | 0 | 0 | 181.5 | 55.55 | 658682.9 | 30.58
[12 empty] | 0 | 0 | 222.5 | 157.55 | 475163.5 | 22.25 | 0 | 0 | 182.5 | 55.13 | 653804.4 | 30.35
[12 empty] | 0 | 0 | 223.5 | 156.56 | 472159 | 22.11 | 0 | 0 | 183.5 | 54.72 | 648925.8 | 30.12
[12 empty] | 0 | 0 | 224.5 | 155.56 | 469154.5 | 21.97 | 0 | 0 | 184.5 | 54.31 | 644047.3 | 29.9
[12 empty] | 0 | 0 | 225.5 | 154.53 | 466035 | 21.83 | 0 | 0 | 185.5 | 53.93 | 639547 | 29.69
[12 empty] | 0 | 0 | 226.5 | 153.45 | 462791.7 | 21.67 | 0 | 0 | 186.5 | 53.59 | 635512.4 | 29.5
[12 empty] | 0 | 0 | 227.5 | 152.38 | 459548.5 | 21.52 | 0 | 0 | 187.5 | 53.25 | 631477.8 | 29.31
[12 empty] | 0 | 0 | 228.5 | 151.3 | 456305.3 | 21.37 | 0 | 0 | 188.5 | 52.91 | 627443.3 | 29.12
[12 empty] | 0 | 0 | 229.5 | 150.23 | 453062.1 | 21.22 | 0 | 0 | 189.5 | 52.57 | 623408.7 | 28.94
[12 empty] | 0 | 0 | 230.5 | 149.22 | 450036.2 | 21.08 | 0 | 0 | 190.5 | 52.23 | 619392.6 | 28.75
[12 empty] | 0 | 0 | 231.5 | 148.3 | 447257.5 | 20.95 | 0 | 0 | 191.5 | 51.9 | 615403.2 | 28.57
[12 empty] | 0 | 0 | 232.5 | 147.38 | 444478.8 | 20.82 | 0 | 0 | 192.5 | 51.56 | 611413.8 | 28.38
[12 empty] | 0 | 0 | 233.5 | 146.46 | 441700.1 | 20.69 | 0 | 0 | 193.5 | 51.22 | 607424.4 | 28.2
[12 empty] | 0 | 0 | 234.5 | 145.54 | 438921.4 | 20.56 | 0 | 0 | 194.5 | 50.89 | 603434.9 | 28.01
[12 empty] | 0 | 0 | 235.5 | 144.52 | 435842.4 | 20.42 | 0 | 0 | 195.5 | 50.54 | 599374.2 | 27.82
[12 empty] | 0 | 0 | 236.5 | 143.38 | 432413.5 | 20.25 | 0 | 0 | 196.5 | 50.19 | 595226.3 | 27.63
[12 empty] | 0 | 0 | 237.5 | 142.24 | 428984.6 | 20.09 | 0 | 0 | 197.5 | 49.84 | 591078.5 | 27.44
[12 empty] | 0 | 0 | 238.5 | 141.1 | 425555.7 | 19.93 | 0 | 0 | 198.5 | 49.5 | 586930.6 | 27.24
[12 empty] | 0 | 0 | 239.5 | 139.97 | 422126.8 | 19.77 | 0 | 0 | 199.5 | 49.15 | 582782.7 | 27.05
[12 empty] | 0 | 0 | 240.5 | 139.04 | 419324.8 | 19.64 | 0 | 0 | 200.5 | 48.81 | 578753.1 | 26.86
[12 empty] | 0 | 0 | 241.5 | 138.35 | 417261.8 | 19.54 | 0 | 0 | 201.5 | 48.48 | 574881.4 | 26.69
[12 empty] | 0 | 0 | 242.5 | 137.67 | 415198.8 | 19.45 | 0 | 0 | 202.5 | 48.15 | 571009.8 | 26.51
[12 empty] | 0 | 0 | 243.5 | 136.99 | 413135.8 | 19.35 | 0 | 0 | 203.5 | 47.83 | 567138.1 | 26.33
[12 empty] | 0 | 0 | 244.5 | 136.3 | 411072.8 | 19.25 | 0 | 0 | 204.5 | 47.5 | 563266.4 | 26.15
[12 empty] | 0 | 0 | 245.5 | 135.46 | 408528.7 | 19.13 | 0 | 0 | 205.5 | 47.18 | 559421 | 25.97
[12 empty] | 0 | 0 | 246.5 | 134.43 | 405410.6 | 18.99 | 0 | 0 | 206.5 | 46.85 | 555609.8 | 25.79
[12 empty] | 0 | 0 | 247.5 | 133.39 | 402292.4 | 18.84 | 0 | 0 | 207.5 | 46.53 | 551798.6 | 25.61
[12 empty] | 0 | 0 | 248.5 | 132.36 | 399174.3 | 18.7 | 0 | 0 | 208.5 | 46.21 | 547987.4 | 25.44
[12 empty] | 0 | 0 | 249.5 | 131.32 | 396056.2 | 18.55 | 0 | 0 | 209.5 | 45.89 | 544176.2 | 25.26
[18 empty] | 0 | 0 | 210.5 | 45.57 | 540345 | 25.08
[18 empty] | 0 | 0 | 211.5 | 45.24 | 536491.4 | 24.9
[18 empty] | 0 | 0 | 212.5 | 44.92 | 532637.9 | 24.72
[18 empty] | 0 | 0 | 213.5 | 44.59 | 528784.3 | 24.55
[18 empty] | 0 | 0 | 214.5 | 44.27 | 524930.7 | 24.37
[18 empty] | 0 | 0 | 215.5 | 43.94 | 521080 | 24.19
[18 empty] | 0 | 0 | 216.5 | 43.62 | 517232.8 | 24.01
[18 empty] | 0 | 0 | 217.5 | 43.29 | 513385.7 | 23.83
[18 empty] | 0 | 0 | 218.5 | 42.97 | 509538.5 | 23.65
[18 empty] | 0 | 0 | 219.5 | 42.64 | 505691.4 | 23.47
[18 empty] | 0 | 0 | 220.5 | 42.36 | 502307.8 | 23.32
[18 empty] | 0 | 0 | 221.5 | 42.12 | 499509.1 | 23.19
[18 empty] | 0 | 0 | 222.5 | 41.89 | 496710.5 | 23.06
[18 empty] | 0 | 0 | 223.5 | 41.65 | 493911.8 | 22.93
[18 empty] | 0 | 0 | 224.5 | 41.41 | 491113.2 | 22.8
[18 empty] | 0 | 0 | 225.5 | 41.15 | 487963.1 | 22.65
[18 empty] | 0 | 0 | 226.5 | 40.85 | 484403.7 | 22.49
[18 empty] | 0 | 0 | 227.5 | 40.55 | 480844.3 | 22.32
[18 empty] | 0 | 0 | 228.5 | 40.25 | 477284.9 | 22.15
[18 empty] | 0 | 0 | 229.5 | 39.95 | 473725.5 | 21.99
[18 empty] | 0 | 0 | 230.5 | 39.67 | 470433.3 | 21.84
[18 empty] | 0 | 0 | 231.5 | 39.42 | 467445.1 | 21.7
[18 empty] | 0 | 0 | 232.5 | 39.17 | 464457 | 21.56
[18 empty] | 0 | 0 | 233.5 | 38.92 | 461468.9 | 21.42
[18 empty] | 0 | 0 | 234.5 | 38.66 | 458480.8 | 21.28
[18 empty] | 0 | 0 | 235.5 | 38.41 | 455430.8 | 21.14
[18 empty] | 0 | 0 | 236.5 | 38.14 | 452314 | 21
[18 empty] | 0 | 0 | 237.5 | 37.88 | 449197.1 | 20.85
[18 empty] | 0 | 0 | 238.5 | 37.62 | 446080.3 | 20.71
[18 empty] | 0 | 0 | 239.5 | 37.35 | 442963.5 | 20.56
[18 empty] | 0 | 0 | 240.5 | 37.11 | 440043.4 | 20.43
[18 empty] | 0 | 0 | 241.5 | 36.88 | 437317.6 | 20.3
[18 empty] | 0 | 0 | 242.5 | 36.65 | 434591.8 | 20.17
[18 empty] | 0 | 0 | 243.5 | 36.42 | 431866.1 | 20.05
[18 empty] | 0 | 0 | 244.5 | 36.19 | 429140.3 | 19.92
[18 empty] | 0 | 0 | 245.5 | 35.95 | 426364.4 | 19.79
[18 empty] | 0 | 0 | 246.5 | 35.72 | 423534.9 | 19.66
[18 empty] | 0 | 0 | 247.5 | 35.48 | 420705.4 | 19.53
[18 empty] | 0 | 0 | 248.5 | 35.24 | 417875.9 | 19.4
[18 empty] | 0 | 0 | 249.5 | 35 | 415046.5 | 19.27

## Diode SFD pas et 6 MV
Pas de mesure | [5 empty] | 6 MV
Measurement time: | 12/04/2023 18:28:42 |  |  |  |  | Measurement time: | 12/04/2023 18:39:08
Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV
Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele
Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 10 x 10 mm |  |  |  |  | Field size: | 10 x 10 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 2.24 | 23220.6 | 2.24 | -10 | 0 | 100 | 0.34 | 3551.6 | 1.24
-9.5 | 0 | 100 | 2.81 | 29134.9 | 2.81 | -9 | 0 | 100 | 0.79 | 8179.5 | 2.85
-8.9 | 0 | 100 | 3.55 | 36812.3 | 3.55 | -8 | 0 | 100 | 1.56 | 16135.2 | 5.63
-8.5 | 0 | 100 | 4.44 | 46032.2 | 4.44 | -7 | 0 | 100 | 3.04 | 31498.4 | 10.98
-8 | 0 | 100 | 5.67 | 58807.9 | 5.68 | -6 | 0 | 100 | 6.45 | 66865.9 | 23.35
-7.5 | 0 | 100 | 7.28 | 75471.7 | 7.29 | -5 | 0 | 100 | 15.48 | 160437.3 | 55.89
-7 | 0 | 100 | 9.78 | 101327.9 | 9.78 | -4 | 0 | 100 | 23.67 | 245295.3 | 85.61
-6.5 | 0 | 100 | 13.43 | 139196.1 | 13.44 | -3 | 0 | 100 | 26.27 | 272322.1 | 94.97
-5.9 | 0 | 100 | 20.42 | 211640.3 | 20.44 | -2 | 0 | 100 | 27.37 | 283646.4 | 98.89
-5.5 | 0 | 100 | 30.94 | 320728.2 | 30.98 | -1 | 0 | 100 | 27.75 | 287644.5 | 100.29
-5 | 0 | 100 | 50.52 | 523631.7 | 50.58 | 0 | 0 | 100 | 27.65 | 286558.4 | 100
-4.5 | 0 | 100 | 70.45 | 730144.2 | 70.51 | 1 | 0 | 100 | 27.42 | 284147.2 | 99.11
-4 | 0 | 100 | 84.37 | 874466.9 | 84.41 | 2 | 0 | 100 | 26.34 | 272961 | 95.22
-3.5 | 0 | 100 | 90.47 | 937651 | 90.61 | 3 | 0 | 100 | 23.7 | 245681.8 | 85.71
-3 | 0 | 100 | 94.39 | 978284.7 | 94.52 | 4 | 0 | 100 | 17.35 | 179846.6 | 62.72
-2.5 | 0 | 100 | 96.56 | 1000806.1 | 96.67 | 5 | 0 | 100 | 7.56 | 78310.1 | 27.29
-2 | 0 | 100 | 98.26 | 1018395.3 | 98.35 | 6 | 0 | 100 | 3.26 | 33814.2 | 11.8
-1.5 | 0 | 100 | 99.17 | 1027821.6 | 99.26 | 7 | 0 | 100 | 1.64 | 17022.9 | 5.94
-1 | 0 | 100 | 99.69 | 1033220 | 99.81 | 8 | 0 | 100 | 0.83 | 8626.4 | 3.01
-0.4 | 0 | 100 | 100.08 | 1037245.6 | 100.17 | 9 | 0 | 100 | 0.37 | 3848.1 | 1.34
0 | 0 | 100 | 99.94 | 1035803.1 | 100.05 | 10 | 0 | 100 | 0.08 | 815.2 | 0.28
0.5 | 0 | 100 | 99.41 | 1030325.9 | 99.5
1 | 0 | 100 | 99.05 | 1026603.6 | 99.16
1.5 | 0 | 100 | 97.91 | 1014791.9 | 98.11 | Measurement time: | 12/04/2023 18:40:43
2 | 0 | 100 | 96.17 | 996789.1 | 96.3 | Radiation device: | Novalis TB Accelerator
2.5 | 0 | 100 | 93.43 | 968307.9 | 93.5 | Energy: | 6 MV
3 | 0 | 100 | 89.17 | 924157.1 | 89.28 | Controller: | CCU
3.5 | 0 | 100 | 82.73 | 857417.4 | 82.82 | Measurement device: | Blue Phantom 2 (48*48*41)
4 | 0 | 100 | 70.87 | 734548.8 | 70.96 | Field detector: | SFD parallele
4.5 | 0 | 100 | 52.02 | 539187.9 | 52.1 | Reference detector: | Stealth CI transmission
5 | 0 | 100 | 32.11 | 332774.7 | 32.15 | SAD: | 1000 mm
5.5 | 0 | 100 | 20.66 | 214162.9 | 20.69 | SSD: | 900 mm
6 | 0 | 100 | 13.29 | 137747.3 | 13.31 | Field size: | 10 x 10 mm
6.5 | 0 | 100 | 9.52 | 98680.8 | 9.53 | Gantry angle: | 0
7 | 0 | 100 | 7.04 | 72998.1 | 7.05 | Measurement medium: | Water
7.4 | 0 | 100 | 5.49 | 56952.5 | 5.5 | Scan type: | Crossline
8 | 0 | 100 | 4.29 | 44483.7 | 4.3 | Scan mode: | StepByStep
8.5 | 0 | 100 | 3.44 | 35604.4 | 3.44 | Wedge: | none
9 | 0 | 100 | 2.71 | 28130.3 | 2.72
9.5 | 0 | 100 | 2.18 | 22588.5 | 2.18 | Points [mm]:
10 | 0 | 100 | 1.76 | 18254.1 | 1.76 | Inline | Crossline | Depth | Normalized field | Current field | Ratio
[6 empty] | 0 | -10 | 100 | 0.37 | 3843.1 | 1.34
[6 empty] | 0 | -9 | 100 | 0.85 | 8844.1 | 3.09
Measurement time: | 12/04/2023 18:30:13 |  |  |  |  | 0 | -8 | 100 | 1.8 | 18654.2 | 6.52
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -7 | 100 | 3.83 | 39744.2 | 13.87
Energy: | 6 MV (FFF) |  |  |  |  | 0 | -6 | 100 | 8.33 | 86366.7 | 30.12
Controller: | CCU |  |  |  |  | 0 | -5 | 100 | 16.21 | 168032.4 | 58.68
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -4 | 100 | 23.33 | 241838.2 | 84.38
Field detector: | SFD parallele |  |  |  |  | 0 | -3 | 100 | 26.3 | 272604.3 | 95.07
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -2 | 100 | 27.09 | 280738 | 97.97
SAD: | 1000 mm |  |  |  |  | 0 | -1 | 100 | 27.56 | 285618.9 | 99.61
SSD: | 900 mm |  |  |  |  | 0 | 0 | 100 | 27.65 | 286554.8 | 100
Field size: | 10 x 10 mm |  |  |  |  | 0 | 1 | 100 | 27.64 | 286473.4 | 100
Gantry angle: | 0 |  |  |  |  | 0 | 2 | 100 | 26.96 | 279385.3 | 97.48
Measurement medium: | Water |  |  |  |  | 0 | 3 | 100 | 25.68 | 266178.3 | 92.86
Scan type: | Crossline |  |  |  |  | 0 | 4 | 100 | 23.09 | 239282.9 | 83.51
Scan mode: | StepByStep |  |  |  |  | 0 | 5 | 100 | 17.37 | 180067.4 | 62.87
Wedge: | none |  |  |  |  | 0 | 6 | 100 | 8.35 | 86565.3 | 30.22
[6 empty] | 0 | 7 | 100 | 3.56 | 36891.5 | 12.88
Points [mm]: | [5 empty] | 0 | 8 | 100 | 1.69 | 17484.3 | 6.11
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 9 | 100 | 0.85 | 8810.9 | 3.08
0 | -10 | 100 | 2.55 | 26458.6 | 2.56 | 0 | 10 | 100 | 0.37 | 3826.7 | 1.34
0 | -9.5 | 100 | 3.22 | 33400.9 | 3.23
0 | -9 | 100 | 3.97 | 41189.2 | 3.98
0 | -8.5 | 100 | 5.28 | 54759.5 | 5.29 | Measurement time: | 12/04/2023 18:43:02
0 | -8 | 100 | 7.17 | 74319 | 7.18 | Radiation device: | Novalis TB Accelerator
0 | -7.5 | 100 | 9.63 | 99804 | 9.65 | Energy: | 6 MV
0 | -7 | 100 | 13.33 | 138150.2 | 13.35 | Controller: | CCU
0 | -6.4 | 100 | 20.27 | 210139.7 | 20.31 | Measurement device: | Blue Phantom 2 (48*48*41)
0 | -6 | 100 | 29.43 | 305029.8 | 29.5 | Field detector: | SFD parallele
0 | -5.5 | 100 | 43.21 | 447869.2 | 43.28 | Reference detector: | Stealth CI transmission
0 | -5 | 100 | 61.58 | 638248.4 | 61.71 | SAD: | 1000 mm
0 | -4.5 | 100 | 75.85 | 786119.9 | 75.92 | SSD: | 900 mm
0 | -4.1 | 100 | 86.35 | 894976.4 | 86.46 | Field size: | 10 x 10 mm
0 | -3.6 | 100 | 92.01 | 953651.3 | 92.2 | Gantry angle: | 0
0 | -3 | 100 | 95.17 | 986387.3 | 95.36 | Measurement medium: | Water
0 | -2.6 | 100 | 96.71 | 1002404.2 | 96.89 | Scan type: | Beam
0 | -2 | 100 | 98.04 | 1016119.6 | 98.25 | Scan mode: | StepByStep
0 | -1.4 | 100 | 99.1 | 1027173 | 99.3 | Wedge: | none
0 | -1.1 | 100 | 99.27 | 1028878.4 | 99.45
0 | -0.6 | 100 | 99.66 | 1032900.6 | 99.87 | Points [mm]:
0 | 0 | 100 | 99.73 | 1033632.5 | 99.94 | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0.5 | 100 | 99.66 | 1032965.2 | 99.84 | 0 | 0 | -0.5 | 18.51 | 191797 | 36.72
0 | 0.9 | 100 | 99.41 | 1030332 | 99.6 | 0 | 0 | 0.5 | 19.89 | 206150.9 | 39.48
0 | 1.4 | 100 | 98.7 | 1023011.3 | 98.92 | 0 | 0 | 1.5 | 26.02 | 269698 | 51.66
0 | 2 | 100 | 97.66 | 1012214.8 | 97.83 | 0 | 0 | 2.5 | 33.68 | 349027.7 | 66.82
0 | 2.4 | 100 | 95.99 | 994904.5 | 96.2 | 0 | 0 | 3.5 | 38.91 | 403309.1 | 77.21
0 | 3 | 100 | 93.47 | 968799.1 | 93.68 | 0 | 0 | 4.5 | 42.72 | 442756.2 | 84.73
0 | 3.5 | 100 | 89.49 | 927488.6 | 89.64 | 0 | 0 | 5.5 | 45.45 | 471016.6 | 90.14
0 | 4 | 100 | 83.71 | 867614.6 | 83.93 | 0 | 0 | 6.5 | 47.28 | 490079.4 | 93.86
0 | 4.4 | 100 | 75.95 | 787173.4 | 76.05 | 0 | 0 | 7.5 | 48.55 | 503161 | 96.38
0 | 5 | 100 | 62.6 | 648801.1 | 62.75 | 0 | 0 | 8.5 | 49.41 | 512131.4 | 98.14
0 | 5.4 | 100 | 45.69 | 473582.2 | 45.81 | 0 | 0 | 9.5 | 49.91 | 517253 | 99.06
0 | 6 | 100 | 28.23 | 292540.2 | 28.29 | 0 | 0 | 10.5 | 50.2 | 520260.2 | 99.61
0 | 6.5 | 100 | 17.79 | 184337.6 | 17.82 | 0 | 0 | 11.5 | 50.38 | 522146.6 | 99.95
0 | 7 | 100 | 12.2 | 126413.8 | 12.22 | 0 | 0 | 12.5 | 50.41 | 522511.3 | 100
0 | 7.5 | 100 | 8.76 | 90768.4 | 8.78 | 0 | 0 | 13.5 | 50.32 | 521545.1 | 99.83
0 | 7.9 | 100 | 6.52 | 67526.4 | 6.53 | 0 | 0 | 14.5 | 50.17 | 519992.1 | 99.52
0 | 8.4 | 100 | 5.02 | 52025.8 | 5.03 | 0 | 0 | 15.5 | 49.87 | 516929.3 | 99.06
0 | 8.9 | 100 | 3.85 | 39936.5 | 3.86 | 0 | 0 | 16.5 | 49.63 | 514369.1 | 98.53
0 | 9.4 | 100 | 3.06 | 31711 | 3.07 | 0 | 0 | 17.5 | 49.44 | 512386.1 | 98.08
0 | 10 | 100 | 2.43 | 25216.2 | 2.44 | 0 | 0 | 18.5 | 49.08 | 508661.1 | 97.37
[6 empty] | 0 | 0 | 19.5 | 48.71 | 504876.9 | 96.66
[6 empty] | 0 | 0 | 20.5 | 48.43 | 501920.4 | 96.17
[6 empty] | 0 | 0 | 21.5 | 48.1 | 498584.3 | 95.53
[6 empty] | 0 | 0 | 22.5 | 47.78 | 495248.2 | 94.89
[6 empty] | 0 | 0 | 23.5 | 47.46 | 491912.1 | 94.26
[6 empty] | 0 | 0 | 24.5 | 47.14 | 488575.9 | 93.62
[6 empty] | 0 | 0 | 25.5 | 46.81 | 485202.3 | 92.97
[6 empty] | 0 | 0 | 26.5 | 46.49 | 481807 | 92.32
[6 empty] | 0 | 0 | 27.5 | 46.16 | 478411.7 | 91.66
[6 empty] | 0 | 0 | 28.5 | 45.83 | 475016.5 | 91
[6 empty] | 0 | 0 | 29.5 | 45.5 | 471621.2 | 90.34
[6 empty] | 0 | 0 | 30.5 | 45.2 | 468457.1 | 89.74
[6 empty] | 0 | 0 | 31.5 | 44.91 | 465435.8 | 89.17
[6 empty] | 0 | 0 | 32.5 | 44.62 | 462414.5 | 88.6
[6 empty] | 0 | 0 | 33.5 | 44.32 | 459393.2 | 88.04
[6 empty] | 0 | 0 | 34.5 | 44.03 | 456371.9 | 87.47
[6 empty] | 0 | 0 | 35.5 | 43.72 | 453118.1 | 86.84
[6 empty] | 0 | 0 | 36.5 | 43.39 | 449705.9 | 86.16
[6 empty] | 0 | 0 | 37.5 | 43.06 | 446293.8 | 85.49
[6 empty] | 0 | 0 | 38.5 | 42.73 | 442881.6 | 84.82
[6 empty] | 0 | 0 | 39.5 | 42.4 | 439469.4 | 84.14
[6 empty] | 0 | 0 | 40.5 | 42.09 | 436262.8 | 83.52
[6 empty] | 0 | 0 | 41.5 | 41.8 | 433210.9 | 82.95
[6 empty] | 0 | 0 | 42.5 | 41.5 | 430158.9 | 82.37
[6 empty] | 0 | 0 | 43.5 | 41.21 | 427107 | 81.8
[6 empty] | 0 | 0 | 44.5 | 40.91 | 424055.1 | 81.22
[6 empty] | 0 | 0 | 45.5 | 40.64 | 421177.6 | 80.67
[6 empty] | 0 | 0 | 46.5 | 40.37 | 418431.8 | 80.14
[6 empty] | 0 | 0 | 47.5 | 40.11 | 415685.9 | 79.61
[6 empty] | 0 | 0 | 48.5 | 39.84 | 412940.1 | 79.08
[6 empty] | 0 | 0 | 49.5 | 39.58 | 410194.2 | 78.55
[6 empty] | 0 | 0 | 50.5 | 39.3 | 407279.4 | 77.99
[6 empty] | 0 | 0 | 51.5 | 39 | 404214.4 | 77.43
[6 empty] | 0 | 0 | 52.5 | 38.7 | 401149.5 | 76.86
[6 empty] | 0 | 0 | 53.5 | 38.41 | 398084.5 | 76.29
[6 empty] | 0 | 0 | 54.5 | 38.11 | 395019.5 | 75.73
[6 empty] | 0 | 0 | 55.5 | 37.82 | 392010.3 | 75.14
[6 empty] | 0 | 0 | 56.5 | 37.54 | 389041.4 | 74.54
[6 empty] | 0 | 0 | 57.5 | 37.25 | 386072.5 | 73.94
[6 empty] | 0 | 0 | 58.5 | 36.96 | 383103.6 | 73.35
[6 empty] | 0 | 0 | 59.5 | 36.68 | 380134.7 | 72.75
[6 empty] | 0 | 0 | 60.5 | 36.41 | 377380.2 | 72.21
[6 empty] | 0 | 0 | 61.5 | 36.16 | 374811.6 | 71.73
[6 empty] | 0 | 0 | 62.5 | 35.92 | 372242.9 | 71.24
[6 empty] | 0 | 0 | 63.5 | 35.67 | 369674.3 | 70.76
[6 empty] | 0 | 0 | 64.5 | 35.42 | 367105.6 | 70.27
[6 empty] | 0 | 0 | 65.5 | 35.18 | 364672.6 | 69.81
[6 empty] | 0 | 0 | 66.5 | 34.96 | 362366.7 | 69.38
[6 empty] | 0 | 0 | 67.5 | 34.74 | 360060.8 | 68.95
[6 empty] | 0 | 0 | 68.5 | 34.52 | 357754.9 | 68.51
[6 empty] | 0 | 0 | 69.5 | 34.29 | 355449 | 68.08
[6 empty] | 0 | 0 | 70.5 | 34.07 | 353099.4 | 67.63
[6 empty] | 0 | 0 | 71.5 | 33.84 | 350706.1 | 67.16
[6 empty] | 0 | 0 | 72.5 | 33.61 | 348312.7 | 66.7
[6 empty] | 0 | 0 | 73.5 | 33.38 | 345919.3 | 66.24
[6 empty] | 0 | 0 | 74.5 | 33.14 | 343526 | 65.77
[6 empty] | 0 | 0 | 75.5 | 32.9 | 341023.7 | 65.29
[6 empty] | 0 | 0 | 76.5 | 32.65 | 338402.8 | 64.79
[6 empty] | 0 | 0 | 77.5 | 32.4 | 335781.8 | 64.29
[6 empty] | 0 | 0 | 78.5 | 32.14 | 333160.8 | 63.79
[6 empty] | 0 | 0 | 79.5 | 31.89 | 330539.9 | 63.29
[6 empty] | 0 | 0 | 80.5 | 31.66 | 328190.4 | 62.84
[6 empty] | 0 | 0 | 81.5 | 31.46 | 326080.9 | 62.43
[6 empty] | 0 | 0 | 82.5 | 31.26 | 323971.3 | 62.03
[6 empty] | 0 | 0 | 83.5 | 31.05 | 321861.8 | 61.62
[6 empty] | 0 | 0 | 84.5 | 30.85 | 319752.3 | 61.21
[6 empty] | 0 | 0 | 85.5 | 30.63 | 317464.1 | 60.77
[6 empty] | 0 | 0 | 86.5 | 30.39 | 314998.8 | 60.29
[6 empty] | 0 | 0 | 87.5 | 30.15 | 312533.5 | 59.81
[6 empty] | 0 | 0 | 88.5 | 29.92 | 310068.2 | 59.32
[6 empty] | 0 | 0 | 89.5 | 29.68 | 307602.9 | 58.84
[6 empty] | 0 | 0 | 90.5 | 29.46 | 305341.7 | 58.41
[6 empty] | 0 | 0 | 91.5 | 29.27 | 303345 | 58.04
[6 empty] | 0 | 0 | 92.5 | 29.07 | 301348.3 | 57.68
[6 empty] | 0 | 0 | 93.5 | 28.88 | 299351.6 | 57.31
[6 empty] | 0 | 0 | 94.5 | 28.69 | 297354.9 | 56.95
[6 empty] | 0 | 0 | 95.5 | 28.51 | 295446.3 | 56.59
[6 empty] | 0 | 0 | 96.5 | 28.33 | 293655.1 | 56.26
[6 empty] | 0 | 0 | 97.5 | 28.16 | 291864 | 55.93
[6 empty] | 0 | 0 | 98.5 | 27.99 | 290072.8 | 55.6
[6 empty] | 0 | 0 | 99.5 | 27.81 | 288281.7 | 55.27
[6 empty] | 0 | 0 | 100.5 | 27.63 | 286381.5 | 54.9
[6 empty] | 0 | 0 | 101.5 | 27.43 | 284334.5 | 54.5
[6 empty] | 0 | 0 | 102.5 | 27.24 | 282287.4 | 54.09
[6 empty] | 0 | 0 | 103.5 | 27.04 | 280240.4 | 53.69
[6 empty] | 0 | 0 | 104.5 | 26.84 | 278193.4 | 53.28
[6 empty] | 0 | 0 | 105.5 | 26.65 | 276212 | 52.9
[6 empty] | 0 | 0 | 106.5 | 26.47 | 274312.4 | 52.55
[6 empty] | 0 | 0 | 107.5 | 26.28 | 272412.7 | 52.2
[6 empty] | 0 | 0 | 108.5 | 26.1 | 270513 | 51.84
[6 empty] | 0 | 0 | 109.5 | 25.92 | 268613.4 | 51.49
[6 empty] | 0 | 0 | 110.5 | 25.73 | 266698.6 | 51.13
[6 empty] | 0 | 0 | 111.5 | 25.55 | 264762.8 | 50.75
[6 empty] | 0 | 0 | 112.5 | 25.36 | 262827.1 | 50.37
[6 empty] | 0 | 0 | 113.5 | 25.17 | 260891.4 | 49.99
[6 empty] | 0 | 0 | 114.5 | 24.98 | 258955.7 | 49.61
[6 empty] | 0 | 0 | 115.5 | 24.82 | 257209.7 | 49.27
[6 empty] | 0 | 0 | 116.5 | 24.67 | 255712.9 | 48.99
[6 empty] | 0 | 0 | 117.5 | 24.53 | 254216.1 | 48.72
[6 empty] | 0 | 0 | 118.5 | 24.38 | 252719.3 | 48.44
[6 empty] | 0 | 0 | 119.5 | 24.24 | 251222.5 | 48.17
[6 empty] | 0 | 0 | 120.5 | 24.07 | 249506.9 | 47.84
[6 empty] | 0 | 0 | 121.5 | 23.88 | 247513.3 | 47.45
[6 empty] | 0 | 0 | 122.5 | 23.69 | 245519.7 | 47.06
[6 empty] | 0 | 0 | 123.5 | 23.5 | 243526.2 | 46.67
[6 empty] | 0 | 0 | 124.5 | 23.3 | 241532.6 | 46.28
[6 empty] | 0 | 0 | 125.5 | 23.13 | 239716.2 | 45.93
[6 empty] | 0 | 0 | 126.5 | 22.98 | 238145.4 | 45.64
[6 empty] | 0 | 0 | 127.5 | 22.83 | 236574.6 | 45.36
[6 empty] | 0 | 0 | 128.5 | 22.67 | 235003.8 | 45.07
[6 empty] | 0 | 0 | 129.5 | 22.52 | 233433 | 44.78
[6 empty] | 0 | 0 | 130.5 | 22.38 | 231979.6 | 44.5
[6 empty] | 0 | 0 | 131.5 | 22.25 | 230652.6 | 44.22
[6 empty] | 0 | 0 | 132.5 | 22.13 | 229325.6 | 43.94
[6 empty] | 0 | 0 | 133.5 | 22 | 227998.6 | 43.66
[6 empty] | 0 | 0 | 134.5 | 21.87 | 226671.6 | 43.38
[6 empty] | 0 | 0 | 135.5 | 21.72 | 225158 | 43.08
[6 empty] | 0 | 0 | 136.5 | 21.55 | 223399.2 | 42.77
[6 empty] | 0 | 0 | 137.5 | 21.38 | 221640.3 | 42.45
[6 empty] | 0 | 0 | 138.5 | 21.21 | 219881.5 | 42.14
[6 empty] | 0 | 0 | 139.5 | 21.05 | 218122.6 | 41.82
[6 empty] | 0 | 0 | 140.5 | 20.89 | 216544.6 | 41.53
[6 empty] | 0 | 0 | 141.5 | 20.76 | 215168.6 | 41.25
[6 empty] | 0 | 0 | 142.5 | 20.63 | 213792.7 | 40.98
[6 empty] | 0 | 0 | 143.5 | 20.49 | 212416.7 | 40.7
[6 empty] | 0 | 0 | 144.5 | 20.36 | 211040.8 | 40.43
[6 empty] | 0 | 0 | 145.5 | 20.23 | 209657.9 | 40.16
[6 empty] | 0 | 0 | 146.5 | 20.09 | 208266 | 39.9
[6 empty] | 0 | 0 | 147.5 | 19.96 | 206874 | 39.63
[6 empty] | 0 | 0 | 148.5 | 19.83 | 205482.1 | 39.37
[6 empty] | 0 | 0 | 149.5 | 19.69 | 204090.2 | 39.11
[6 empty] | 0 | 0 | 150.5 | 19.56 | 202746.8 | 38.85
[6 empty] | 0 | 0 | 151.5 | 19.44 | 201467.3 | 38.6
[6 empty] | 0 | 0 | 152.5 | 19.31 | 200187.7 | 38.35
[6 empty] | 0 | 0 | 153.5 | 19.19 | 198908.2 | 38.1
[6 empty] | 0 | 0 | 154.5 | 19.07 | 197628.7 | 37.85
[6 empty] | 0 | 0 | 155.5 | 18.94 | 196335.3 | 37.6
[6 empty] | 0 | 0 | 156.5 | 18.82 | 195024.4 | 37.35
[6 empty] | 0 | 0 | 157.5 | 18.69 | 193713.6 | 37.1
[6 empty] | 0 | 0 | 158.5 | 18.56 | 192402.7 | 36.85
[6 empty] | 0 | 0 | 159.5 | 18.44 | 191091.9 | 36.6
[6 empty] | 0 | 0 | 160.5 | 18.32 | 189850.9 | 36.36
[6 empty] | 0 | 0 | 161.5 | 18.21 | 188695.6 | 36.15
[6 empty] | 0 | 0 | 162.5 | 18.09 | 187540.3 | 35.93
[6 empty] | 0 | 0 | 163.5 | 17.98 | 186385.1 | 35.72
[6 empty] | 0 | 0 | 164.5 | 17.87 | 185229.8 | 35.5
[6 empty] | 0 | 0 | 165.5 | 17.75 | 184016.9 | 35.27
[6 empty] | 0 | 0 | 166.5 | 17.63 | 182732.6 | 35.02
[6 empty] | 0 | 0 | 167.5 | 17.51 | 181448.4 | 34.78
[6 empty] | 0 | 0 | 168.5 | 17.38 | 180164.1 | 34.53
[6 empty] | 0 | 0 | 169.5 | 17.26 | 178879.8 | 34.28
[6 empty] | 0 | 0 | 170.5 | 17.14 | 177678.7 | 34.05
[6 empty] | 0 | 0 | 171.5 | 17.04 | 176579 | 33.83
[6 empty] | 0 | 0 | 172.5 | 16.93 | 175479.3 | 33.62
[6 empty] | 0 | 0 | 173.5 | 16.82 | 174379.6 | 33.4
[6 empty] | 0 | 0 | 174.5 | 16.72 | 173279.9 | 33.18
[6 empty] | 0 | 0 | 175.5 | 16.61 | 172136.8 | 32.97
[6 empty] | 0 | 0 | 176.5 | 16.49 | 170938.5 | 32.75
[6 empty] | 0 | 0 | 177.5 | 16.38 | 169740.2 | 32.54
[6 empty] | 0 | 0 | 178.5 | 16.26 | 168541.9 | 32.32
[6 empty] | 0 | 0 | 179.5 | 16.15 | 167343.6 | 32.11
[6 empty] | 0 | 0 | 180.5 | 16.04 | 166230.4 | 31.89
[6 empty] | 0 | 0 | 181.5 | 15.94 | 165218.6 | 31.68
[6 empty] | 0 | 0 | 182.5 | 15.84 | 164206.8 | 31.46
[6 empty] | 0 | 0 | 183.5 | 15.75 | 163194.9 | 31.24
[6 empty] | 0 | 0 | 184.5 | 15.65 | 162183.1 | 31.03
[6 empty] | 0 | 0 | 185.5 | 15.54 | 161071.1 | 30.81
[6 empty] | 0 | 0 | 186.5 | 15.42 | 159825.3 | 30.58
[6 empty] | 0 | 0 | 187.5 | 15.3 | 158579.5 | 30.35
[6 empty] | 0 | 0 | 188.5 | 15.18 | 157333.7 | 30.12
[6 empty] | 0 | 0 | 189.5 | 15.06 | 156087.9 | 29.89
[6 empty] | 0 | 0 | 190.5 | 14.95 | 154978.3 | 29.68
[6 empty] | 0 | 0 | 191.5 | 14.86 | 154039.3 | 29.5
[6 empty] | 0 | 0 | 192.5 | 14.77 | 153100.4 | 29.32
[6 empty] | 0 | 0 | 193.5 | 14.68 | 152161.4 | 29.14
[6 empty] | 0 | 0 | 194.5 | 14.59 | 151222.4 | 28.97
[6 empty] | 0 | 0 | 195.5 | 14.5 | 150290.4 | 28.79
[6 empty] | 0 | 0 | 196.5 | 14.41 | 149366.2 | 28.61
[6 empty] | 0 | 0 | 197.5 | 14.32 | 148442 | 28.44
[6 empty] | 0 | 0 | 198.5 | 14.23 | 147517.8 | 28.27
[6 empty] | 0 | 0 | 199.5 | 14.14 | 146593.6 | 28.09
[6 empty] | 0 | 0 | 200.5 | 14.05 | 145628.2 | 27.91
[6 empty] | 0 | 0 | 201.5 | 13.95 | 144611.3 | 27.72
[6 empty] | 0 | 0 | 202.5 | 13.85 | 143594.3 | 27.53
[6 empty] | 0 | 0 | 203.5 | 13.76 | 142577.4 | 27.34
[6 empty] | 0 | 0 | 204.5 | 13.66 | 141560.4 | 27.14
[6 empty] | 0 | 0 | 205.5 | 13.57 | 140640 | 26.97
[6 empty] | 0 | 0 | 206.5 | 13.49 | 139844.9 | 26.82
[6 empty] | 0 | 0 | 207.5 | 13.42 | 139049.8 | 26.67
[6 empty] | 0 | 0 | 208.5 | 13.34 | 138254.8 | 26.53
[6 empty] | 0 | 0 | 209.5 | 13.26 | 137459.7 | 26.38
[6 empty] | 0 | 0 | 210.5 | 13.18 | 136554.8 | 26.2
[6 empty] | 0 | 0 | 211.5 | 13.07 | 135511.9 | 25.99
[6 empty] | 0 | 0 | 212.5 | 12.97 | 134469 | 25.79
[6 empty] | 0 | 0 | 213.5 | 12.87 | 133426.1 | 25.58
[6 empty] | 0 | 0 | 214.5 | 12.77 | 132383.2 | 25.37
[6 empty] | 0 | 0 | 215.5 | 12.68 | 131448 | 25.19
[6 empty] | 0 | 0 | 216.5 | 12.61 | 130645.1 | 25.03
[6 empty] | 0 | 0 | 217.5 | 12.53 | 129842.2 | 24.87
[6 empty] | 0 | 0 | 218.5 | 12.45 | 129039.2 | 24.71
[6 empty] | 0 | 0 | 219.5 | 12.37 | 128236.3 | 24.55
[6 empty] | 0 | 0 | 220.5 | 12.29 | 127417.3 | 24.4
[6 empty] | 0 | 0 | 221.5 | 12.21 | 126578.2 | 24.23
[6 empty] | 0 | 0 | 222.5 | 12.13 | 125739 | 24.07
[6 empty] | 0 | 0 | 223.5 | 12.05 | 124899.8 | 23.91
[6 empty] | 0 | 0 | 224.5 | 11.97 | 124060.6 | 23.75
[6 empty] | 0 | 0 | 225.5 | 11.89 | 123273.4 | 23.6
[6 empty] | 0 | 0 | 226.5 | 11.82 | 122549.3 | 23.47
[6 empty] | 0 | 0 | 227.5 | 11.75 | 121825.2 | 23.33
[6 empty] | 0 | 0 | 228.5 | 11.68 | 121101 | 23.2
[6 empty] | 0 | 0 | 229.5 | 11.61 | 120376.9 | 23.06
[6 empty] | 0 | 0 | 230.5 | 11.54 | 119613.8 | 22.92
[6 empty] | 0 | 0 | 231.5 | 11.46 | 118802.4 | 22.76
[6 empty] | 0 | 0 | 232.5 | 11.38 | 117991 | 22.61
[6 empty] | 0 | 0 | 233.5 | 11.31 | 117179.5 | 22.45
[6 empty] | 0 | 0 | 234.5 | 11.23 | 116368.1 | 22.29
[6 empty] | 0 | 0 | 235.5 | 11.15 | 115558 | 22.13
[6 empty] | 0 | 0 | 236.5 | 11.07 | 114749.6 | 21.98
[6 empty] | 0 | 0 | 237.5 | 10.99 | 113941.1 | 21.82
[6 empty] | 0 | 0 | 238.5 | 10.92 | 113132.7 | 21.67
[6 empty] | 0 | 0 | 239.5 | 10.84 | 112324.3 | 21.51
[6 empty] | 0 | 0 | 240.5 | 10.77 | 111641 | 21.38
[6 empty] | 0 | 0 | 241.5 | 10.72 | 111091.8 | 21.28
[6 empty] | 0 | 0 | 242.5 | 10.67 | 110542.7 | 21.18
[6 empty] | 0 | 0 | 243.5 | 10.61 | 109993.5 | 21.08
[6 empty] | 0 | 0 | 244.5 | 10.56 | 109444.4 | 20.97
[6 empty] | 0 | 0 | 245.5 | 10.49 | 108750.5 | 20.84
[6 empty] | 0 | 0 | 246.5 | 10.41 | 107876.7 | 20.67
[6 empty] | 0 | 0 | 247.5 | 10.32 | 107002.9 | 20.51
[6 empty] | 0 | 0 | 248.5 | 10.24 | 106129.2 | 20.34
[6 empty] | 0 | 0 | 249.5 | 10.16 | 105255.4 | 20.17

## Diode SFD + cône
1x1 cm² - cône
Measurement time: | 12/04/2023 19:58:03
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 2.45 | 24823.8 | 0
-9 | 0 | 100 | 3.32 | 33612.9 | 0
-8 | 0 | 100 | 4.89 | 49482.1 | 0
-7 | 0 | 100 | 8.02 | 81172.1 | 0
-6 | 0 | 100 | 15.78 | 159648.2 | 0
-5 | 0 | 100 | 51.46 | 520663.8 | 0
-4 | 0 | 100 | 85.76 | 867669.5 | 0
-3 | 0 | 100 | 94.35 | 954641.2 | 0
-2 | 0 | 100 | 98.12 | 992759.6 | 0
-1 | 0 | 100 | 99.54 | 1007162.3 | 0
0 | 0 | 100 | 100 | 1011771.1 | 0
1 | 0 | 100 | 99.03 | 1002006.3 | 0
2 | 0 | 100 | 96.86 | 979972.6 | 0
3 | 0 | 100 | 92.34 | 934225.4 | 0
4 | 0 | 100 | 79.06 | 799865.2 | 0
5 | 0 | 100 | 33.12 | 335071.7 | 0
6 | 0 | 100 | 11.68 | 118167.2 | 0
7 | 0 | 100 | 6.3 | 63785.2 | 0
8 | 0 | 100 | 4.1 | 41504.1 | 0
9 | 0 | 100 | 2.85 | 28820.4 | 0
10 | 0 | 100 | 2.08 | 21021.6 | 0
Measurement time: | 12/04/2023 19:59:40
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -10 | 100 | 2.24 | 22637.2 | 0
0 | -9 | 100 | 3.13 | 31662.8 | 0
0 | -8 | 100 | 4.64 | 46935 | 0
0 | -7 | 100 | 7.13 | 72101 | 0
0 | -6 | 100 | 13.52 | 136750.9 | 0
0 | -5 | 100 | 40.22 | 406869.4 | 0
0 | -4 | 100 | 81.57 | 825182.8 | 0
0 | -3 | 100 | 92.65 | 937296.8 | 0
0 | -2 | 100 | 97.17 | 982992.6 | 0
0 | -1 | 100 | 99.16 | 1003093.8 | 0
0 | 0 | 100 | 100 | 1011616.1 | 0
0 | 1 | 100 | 99.45 | 1006057.6 | 0
0 | 2 | 100 | 97.65 | 987884.4 | 0
0 | 3 | 100 | 93.16 | 942409.8 | 0
0 | 4 | 100 | 83.85 | 848226.2 | 0
0 | 5 | 100 | 42.23 | 427224.5 | 0
0 | 6 | 100 | 13.18 | 133338.8 | 0
0 | 7 | 100 | 7.06 | 71453.6 | 0
0 | 8 | 100 | 4.5 | 45477.9 | 0
0 | 9 | 100 | 3.03 | 30635.4 | 0
0 | 10 | 100 | 2.19 | 22175.6 | 0
Measurement time: | 12/04/2023 20:01:27
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | -0.5 | 51.59 | 1026298.4 | 0
0 | 0 | 0.5 | 52.77 | 1049961.4 | 0
0 | 0 | 1.5 | 62.38 | 1240978.9 | 0
0 | 0 | 2.5 | 76.58 | 1523649.1 | 0
0 | 0 | 3.5 | 85.38 | 1698625.2 | 0
0 | 0 | 4.5 | 91.01 | 1810578.2 | 0
0 | 0 | 5.5 | 94.63 | 1882755.5 | 0
0 | 0 | 6.5 | 96.98 | 1929484.6 | 0
0 | 0 | 7.5 | 98.56 | 1960797.1 | 0
0 | 0 | 8.5 | 99.46 | 1978817.6 | 0
0 | 0 | 9.5 | 99.87 | 1986846.5 | 0
0 | 0 | 10.5 | 100 | 1989507.6 | 0
0 | 0 | 11.5 | 99.82 | 1985830.3 | 0
0 | 0 | 12.5 | 99.54 | 1980283.8 | 0
0 | 0 | 13.5 | 99.14 | 1972369.4 | 0
0 | 0 | 14.5 | 98.64 | 1962384.1 | 0
0 | 0 | 15.5 | 97.98 | 1949307.8 | 0
0 | 0 | 16.5 | 97.23 | 1934477 | 0
0 | 0 | 17.5 | 96.56 | 1921039.5 | 0
0 | 0 | 18.5 | 95.82 | 1906268 | 0
0 | 0 | 19.5 | 95.13 | 1892558.1 | 0
0 | 0 | 20.5 | 94.48 | 1879751.1 | 0
0 | 0 | 21.5 | 93.77 | 1865482.3 | 0
0 | 0 | 22.5 | 93.05 | 1851213.5 | 0
0 | 0 | 23.5 | 92.33 | 1836944.7 | 0
0 | 0 | 24.5 | 91.61 | 1822675.9 | 0
0 | 0 | 25.5 | 90.89 | 1808263.9 | 0
0 | 0 | 26.5 | 90.16 | 1793758 | 0
0 | 0 | 27.5 | 89.43 | 1779252.2 | 0
0 | 0 | 28.5 | 88.7 | 1764746.3 | 0
0 | 0 | 29.5 | 87.97 | 1750240.4 | 0
0 | 0 | 30.5 | 87.26 | 1735972.1 | 0
0 | 0 | 31.5 | 86.55 | 1721854 | 0
0 | 0 | 32.5 | 85.84 | 1707735.8 | 0
0 | 0 | 33.5 | 85.13 | 1693617.7 | 0
0 | 0 | 34.5 | 84.42 | 1679499.5 | 0
0 | 0 | 35.5 | 83.75 | 1666232.7 | 0
0 | 0 | 36.5 | 83.11 | 1653540.4 | 0
0 | 0 | 37.5 | 82.48 | 1640848.1 | 0
0 | 0 | 38.5 | 81.84 | 1628155.7 | 0
0 | 0 | 39.5 | 81.2 | 1615463.4 | 0
0 | 0 | 40.5 | 80.58 | 1603164.9 | 0
0 | 0 | 41.5 | 79.98 | 1591151.1 | 0
0 | 0 | 42.5 | 79.37 | 1579137.3 | 0
0 | 0 | 43.5 | 78.77 | 1567123.6 | 0
0 | 0 | 44.5 | 78.17 | 1555109.8 | 0
0 | 0 | 45.5 | 77.55 | 1542820.1 | 0
0 | 0 | 46.5 | 76.92 | 1530330.6 | 0
0 | 0 | 47.5 | 76.29 | 1517841 | 0
0 | 0 | 48.5 | 75.66 | 1505351.5 | 0
0 | 0 | 49.5 | 75.04 | 1492862 | 0
0 | 0 | 50.5 | 74.42 | 1480560.4 | 0
0 | 0 | 51.5 | 73.81 | 1468425.7 | 0
0 | 0 | 52.5 | 73.2 | 1456291 | 0
0 | 0 | 53.5 | 72.59 | 1444156.4 | 0
0 | 0 | 54.5 | 71.98 | 1432021.7 | 0
0 | 0 | 55.5 | 71.42 | 1421004.7 | 0
0 | 0 | 56.5 | 70.92 | 1410901.6 | 0
0 | 0 | 57.5 | 70.41 | 1400798.5 | 0
0 | 0 | 58.5 | 69.9 | 1390695.4 | 0
0 | 0 | 59.5 | 69.39 | 1380592.3 | 0
0 | 0 | 60.5 | 68.86 | 1370010.2 | 0
0 | 0 | 61.5 | 68.31 | 1358959.7 | 0
0 | 0 | 62.5 | 67.75 | 1347909.1 | 0
0 | 0 | 63.5 | 67.2 | 1336858.6 | 0
0 | 0 | 64.5 | 66.64 | 1325808 | 0
0 | 0 | 65.5 | 66.11 | 1315222.3 | 0
0 | 0 | 66.5 | 65.6 | 1305031.6 | 0
0 | 0 | 67.5 | 65.08 | 1294840.8 | 0
0 | 0 | 68.5 | 64.57 | 1284650.1 | 0
0 | 0 | 69.5 | 64.06 | 1274459.4 | 0
0 | 0 | 70.5 | 63.56 | 1264542.9 | 0
0 | 0 | 71.5 | 63.08 | 1254894.9 | 0
0 | 0 | 72.5 | 62.59 | 1245246.8 | 0
0 | 0 | 73.5 | 62.11 | 1235598.8 | 0
0 | 0 | 74.5 | 61.62 | 1225950.8 | 0
0 | 0 | 75.5 | 61.15 | 1216641 | 0
0 | 0 | 76.5 | 60.7 | 1207715.9 | 0
0 | 0 | 77.5 | 60.26 | 1198790.8 | 0
0 | 0 | 78.5 | 59.81 | 1189865.6 | 0
0 | 0 | 79.5 | 59.36 | 1180940.5 | 0
0 | 0 | 80.5 | 58.91 | 1171991.2 | 0
0 | 0 | 81.5 | 58.46 | 1163013.3 | 0
0 | 0 | 82.5 | 58.01 | 1154035.3 | 0
0 | 0 | 83.5 | 57.55 | 1145057.3 | 0
0 | 0 | 84.5 | 57.1 | 1136079.4 | 0
0 | 0 | 85.5 | 56.66 | 1127336.6 | 0
0 | 0 | 86.5 | 56.24 | 1118846.7 | 0
0 | 0 | 87.5 | 55.81 | 1110356.7 | 0
0 | 0 | 88.5 | 55.38 | 1101866.7 | 0
0 | 0 | 89.5 | 54.96 | 1093376.8 | 0
0 | 0 | 90.5 | 54.54 | 1085033.2 | 0
0 | 0 | 91.5 | 54.13 | 1076887.4 | 0
0 | 0 | 92.5 | 53.72 | 1068741.5 | 0
0 | 0 | 93.5 | 53.31 | 1060595.6 | 0
0 | 0 | 94.5 | 52.9 | 1052449.8 | 0
0 | 0 | 95.5 | 52.5 | 1044534.8 | 0
0 | 0 | 96.5 | 52.12 | 1036915.9 | 0
0 | 0 | 97.5 | 51.74 | 1029297 | 0
0 | 0 | 98.5 | 51.35 | 1021678.1 | 0
0 | 0 | 99.5 | 50.97 | 1014059.2 | 0
0 | 0 | 100.5 | 50.59 | 1006490.2 | 0
0 | 0 | 101.5 | 50.21 | 998989.7 | 0
0 | 0 | 102.5 | 49.84 | 991489.1 | 0
0 | 0 | 103.5 | 49.46 | 983988.5 | 0
0 | 0 | 104.5 | 49.08 | 976487.9 | 0
0 | 0 | 105.5 | 48.71 | 969083.7 | 0
0 | 0 | 106.5 | 48.34 | 961802.4 | 0
0 | 0 | 107.5 | 47.98 | 954521.2 | 0
0 | 0 | 108.5 | 47.61 | 947240 | 0
0 | 0 | 109.5 | 47.25 | 939958.7 | 0
0 | 0 | 110.5 | 46.89 | 932963.4 | 0
0 | 0 | 111.5 | 46.56 | 926350.3 | 0
0 | 0 | 112.5 | 46.23 | 919737.2 | 0
0 | 0 | 113.5 | 45.9 | 913124.1 | 0
0 | 0 | 114.5 | 45.56 | 906511 | 0
0 | 0 | 115.5 | 45.24 | 899983.9 | 0
0 | 0 | 116.5 | 44.91 | 893565.2 | 0
0 | 0 | 117.5 | 44.59 | 887146.5 | 0
0 | 0 | 118.5 | 44.27 | 880727.8 | 0
0 | 0 | 119.5 | 43.95 | 874309.1 | 0
0 | 0 | 120.5 | 43.62 | 867897 | 0
0 | 0 | 121.5 | 43.3 | 861492.9 | 0
0 | 0 | 122.5 | 42.98 | 855088.8 | 0
0 | 0 | 123.5 | 42.66 | 848684.7 | 0
0 | 0 | 124.5 | 42.34 | 842280.6 | 0
0 | 0 | 125.5 | 42.01 | 835807.7 | 0
0 | 0 | 126.5 | 41.68 | 829253.8 | 0
0 | 0 | 127.5 | 41.35 | 822699.9 | 0
0 | 0 | 128.5 | 41.02 | 816146 | 0
0 | 0 | 129.5 | 40.69 | 809592.1 | 0
0 | 0 | 130.5 | 40.39 | 803507.9 | 0
0 | 0 | 131.5 | 40.11 | 798077.4 | 0
0 | 0 | 132.5 | 39.84 | 792646.9 | 0
0 | 0 | 133.5 | 39.57 | 787216.5 | 0
0 | 0 | 134.5 | 39.3 | 781786 | 0
0 | 0 | 135.5 | 39.03 | 776517.5 | 0
0 | 0 | 136.5 | 38.78 | 771450.4 | 0
0 | 0 | 137.5 | 38.52 | 766383.2 | 0
0 | 0 | 138.5 | 38.27 | 761316.1 | 0
0 | 0 | 139.5 | 38.01 | 756248.9 | 0
0 | 0 | 140.5 | 37.74 | 750781.6 | 0
0 | 0 | 141.5 | 37.44 | 744844.5 | 0
0 | 0 | 142.5 | 37.14 | 738907.3 | 0
0 | 0 | 143.5 | 36.84 | 732970.2 | 0
0 | 0 | 144.5 | 36.54 | 727033.1 | 0
0 | 0 | 145.5 | 36.26 | 721474.6 | 0
0 | 0 | 146.5 | 36.01 | 716458.9 | 0
0 | 0 | 147.5 | 35.76 | 711443.3 | 0
0 | 0 | 148.5 | 35.51 | 706427.6 | 0
0 | 0 | 149.5 | 35.26 | 701412 | 0
0 | 0 | 150.5 | 35.01 | 696453 | 0
0 | 0 | 151.5 | 34.76 | 691568.2 | 0
0 | 0 | 152.5 | 34.52 | 686683.4 | 0
0 | 0 | 153.5 | 34.27 | 681798.6 | 0
0 | 0 | 154.5 | 34.02 | 676913.9 | 0
0 | 0 | 155.5 | 33.78 | 672076.1 | 0
0 | 0 | 156.5 | 33.54 | 667302.2 | 0
0 | 0 | 157.5 | 33.3 | 662528.3 | 0
0 | 0 | 158.5 | 33.06 | 657754.4 | 0
0 | 0 | 159.5 | 32.82 | 652980.4 | 0
0 | 0 | 160.5 | 32.6 | 648530.5 | 0
0 | 0 | 161.5 | 32.39 | 644486.9 | 0
0 | 0 | 162.5 | 32.19 | 640443.3 | 0
0 | 0 | 163.5 | 31.99 | 636399.7 | 0
0 | 0 | 164.5 | 31.78 | 632356.1 | 0
0 | 0 | 165.5 | 31.56 | 627952.3 | 0
0 | 0 | 166.5 | 31.32 | 623044.6 | 0
0 | 0 | 167.5 | 31.07 | 618137 | 0
0 | 0 | 168.5 | 30.82 | 613229.4 | 0
0 | 0 | 169.5 | 30.58 | 608321.7 | 0
0 | 0 | 170.5 | 30.35 | 603776.5 | 0
0 | 0 | 171.5 | 30.15 | 599746.3 | 0
0 | 0 | 172.5 | 29.94 | 595716.1 | 0
0 | 0 | 173.5 | 29.74 | 591685.9 | 0
0 | 0 | 174.5 | 29.54 | 587655.7 | 0
0 | 0 | 175.5 | 29.33 | 583553.4 | 0
0 | 0 | 176.5 | 29.12 | 579355.5 | 0
0 | 0 | 177.5 | 28.91 | 575157.6 | 0
0 | 0 | 178.5 | 28.7 | 570959.7 | 0
0 | 0 | 179.5 | 28.49 | 566761.8 | 0
0 | 0 | 180.5 | 28.29 | 562800.5 | 0
0 | 0 | 181.5 | 28.1 | 559112.6 | 0
0 | 0 | 182.5 | 27.92 | 555424.7 | 0
0 | 0 | 183.5 | 27.73 | 551736.8 | 0
0 | 0 | 184.5 | 27.55 | 548048.9 | 0
0 | 0 | 185.5 | 27.36 | 544269.2 | 0
0 | 0 | 186.5 | 27.16 | 540360.7 | 0
0 | 0 | 187.5 | 26.96 | 536452.2 | 0
0 | 0 | 188.5 | 26.77 | 532543.7 | 0
0 | 0 | 189.5 | 26.57 | 528635.1 | 0
0 | 0 | 190.5 | 26.38 | 524895.6 | 0
0 | 0 | 191.5 | 26.21 | 521373.5 | 0
0 | 0 | 192.5 | 26.03 | 517851.3 | 0
0 | 0 | 193.5 | 25.85 | 514329.2 | 0
0 | 0 | 194.5 | 25.68 | 510807.1 | 0
0 | 0 | 195.5 | 25.5 | 507304.3 | 0
0 | 0 | 196.5 | 25.32 | 503827.4 | 0
0 | 0 | 197.5 | 25.15 | 500350.4 | 0
0 | 0 | 198.5 | 24.97 | 496873.5 | 0
0 | 0 | 199.5 | 24.8 | 493396.5 | 0
0 | 0 | 200.5 | 24.63 | 489934.1 | 0
0 | 0 | 201.5 | 24.45 | 486491.6 | 0
0 | 0 | 202.5 | 24.28 | 483049.2 | 0
0 | 0 | 203.5 | 24.11 | 479606.7 | 0
0 | 0 | 204.5 | 23.93 | 476164.2 | 0
0 | 0 | 205.5 | 23.77 | 472920 | 0
0 | 0 | 206.5 | 23.62 | 469925.6 | 0
0 | 0 | 207.5 | 23.47 | 466931.1 | 0
0 | 0 | 208.5 | 23.32 | 463936.7 | 0
0 | 0 | 209.5 | 23.17 | 460942.2 | 0
0 | 0 | 210.5 | 23.02 | 457956.5 | 0
0 | 0 | 211.5 | 22.87 | 454982.2 | 0
0 | 0 | 212.5 | 22.72 | 452008 | 0
0 | 0 | 213.5 | 22.57 | 449033.7 | 0
0 | 0 | 214.5 | 22.42 | 446059.5 | 0
0 | 0 | 215.5 | 22.27 | 443117.1 | 0
0 | 0 | 216.5 | 22.13 | 440213.9 | 0
0 | 0 | 217.5 | 21.98 | 437310.7 | 0
0 | 0 | 218.5 | 21.83 | 434407.6 | 0
0 | 0 | 219.5 | 21.69 | 431504.4 | 0
0 | 0 | 220.5 | 21.54 | 428589.3 | 0
0 | 0 | 221.5 | 21.4 | 425659.2 | 0
0 | 0 | 222.5 | 21.25 | 422729.1 | 0
0 | 0 | 223.5 | 21.1 | 419799.1 | 0
0 | 0 | 224.5 | 20.95 | 416869 | 0
0 | 0 | 225.5 | 20.81 | 414012 | 0
0 | 0 | 226.5 | 20.67 | 411240.4 | 0
0 | 0 | 227.5 | 20.53 | 408468.7 | 0
0 | 0 | 228.5 | 20.39 | 405697 | 0
0 | 0 | 229.5 | 20.25 | 402925.3 | 0
0 | 0 | 230.5 | 20.11 | 400160.4 | 0
0 | 0 | 231.5 | 19.97 | 397403.7 | 0
0 | 0 | 232.5 | 19.84 | 394647 | 0
0 | 0 | 233.5 | 19.7 | 391890.3 | 0
0 | 0 | 234.5 | 19.56 | 389133.6 | 0
0 | 0 | 235.5 | 19.43 | 386493.1 | 0
0 | 0 | 236.5 | 19.3 | 383988.2 | 0
0 | 0 | 237.5 | 19.17 | 381483.2 | 0
0 | 0 | 238.5 | 19.05 | 378978.3 | 0
0 | 0 | 239.5 | 18.92 | 376473.3 | 0
0 | 0 | 240.5 | 18.79 | 373848.4 | 0
0 | 0 | 241.5 | 18.65 | 371087.6 | 0
0 | 0 | 242.5 | 18.51 | 368326.8 | 0
0 | 0 | 243.5 | 18.37 | 365565.9 | 0
0 | 0 | 244.5 | 18.24 | 362805.1 | 0
0 | 0 | 245.5 | 18.11 | 360370.6 | 0
0 | 0 | 246.5 | 18.01 | 358341.5 | 0
0 | 0 | 247.5 | 17.91 | 356312.4 | 0
0 | 0 | 248.5 | 17.81 | 354283.3 | 0
0 | 0 | 249.5 | 17.71 | 352254.1 | 0

## NanoRazor
0,5x0,5 cm² | [5 empty] | 1x1 cm² | [5 empty] | 2x2 cm² | [5 empty] | 3x3 cm²
Measurement time: | 27/05/2023 12:25:45 |  |  |  |  | Measurement time: | 27/05/2023 12:42:01 |  |  |  |  | Measurement time: | 27/05/2023 13:03:36 |  |  |  |  | Measurement time: | 27/05/2023 13:34:18
Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 5 x 5 mm |  |  |  |  | Field size: | 10 x 10 mm |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 23.3 | 135.8 | 1.18 | -20 | 0 | 100 | 68.08 | 396.8 | 2.71 | -30 | 0 | 100 | 97.06 | 565.7 | 3.47 | -30 | 0 | 100 | 134.13 | 781.7 | 4.62
-9 | 0 | 100 | 37.29 | 217.3 | 1.9 | -19 | 0 | 100 | 84.58 | 493 | 3.37 | -29 | 0 | 100 | 97.72 | 569.5 | 3.49 | -29 | 0 | 100 | 136.35 | 794.6 | 4.69
-8 | 0 | 100 | 36.1 | 210.4 | 1.84 | -18 | 0 | 100 | 97.55 | 568.5 | 3.88 | -28 | 0 | 100 | 93.28 | 543.6 | 3.33 | -28 | 0 | 100 | 155.9 | 908.6 | 5.36
-7 | 0 | 100 | 79.61 | 464 | 4.06 | -17 | 0 | 100 | 92.5 | 539.1 | 3.68 | -27 | 0 | 100 | 98.22 | 572.4 | 3.51 | -27 | 0 | 100 | 171.64 | 1000.3 | 5.9
-6 | 0 | 100 | 132.41 | 771.7 | 6.74 | -16 | 0 | 100 | 98.59 | 574.6 | 3.93 | -26 | 0 | 100 | 118.27 | 689.3 | 4.23 | -26 | 0 | 100 | 163.19 | 951.1 | 5.61
-5 | 0 | 100 | 205.03 | 1194.9 | 10.49 | -15 | 0 | 100 | 72.38 | 421.8 | 2.88 | -25 | 0 | 100 | 117.17 | 682.9 | 4.19 | -25 | 0 | 100 | 153.61 | 895.2 | 5.29
-4 | 0 | 100 | 340.81 | 1986.3 | 17.44 | -14 | 0 | 100 | 65.85 | 383.8 | 2.62 | -24 | 0 | 100 | 127.82 | 744.9 | 4.57 | -24 | 0 | 100 | 151.01 | 880.1 | 5.19
-3 | 0 | 100 | 668.69 | 3897.1 | 34.15 | -13 | 0 | 100 | 74.89 | 436.5 | 2.98 | -23 | 0 | 100 | 136.36 | 794.7 | 4.87 | -23 | 0 | 100 | 164.43 | 958.3 | 5.66
-2 | 0 | 100 | 1239.57 | 7224.2 | 62.99 | -12 | 0 | 100 | 71.18 | 414.8 | 2.83 | -22 | 0 | 100 | 145.43 | 847.6 | 5.2 | -22 | 0 | 100 | 175.91 | 1025.2 | 6.05
-1 | 0 | 100 | 1784.63 | 10400.8 | 90.94 | -11 | 0 | 100 | 69.64 | 405.9 | 2.77 | -21 | 0 | 100 | 137.78 | 803 | 4.93 | -21 | 0 | 100 | 190.88 | 1112.5 | 6.57
0 | 0 | 100 | 1964.63 | 11449.9 | 100 | -10 | 0 | 100 | 85.55 | 498.6 | 3.41 | -20 | 0 | 100 | 127.45 | 742.8 | 4.55 | -20 | 0 | 100 | 229.95 | 1340.2 | 7.92
1 | 0 | 100 | 1853.7 | 10803.4 | 94.01 | -9 | 0 | 100 | 116.6 | 679.5 | 4.64 | -19 | 0 | 100 | 133.01 | 775.2 | 4.75 | -19 | 0 | 100 | 281.37 | 1639.8 | 9.68
2 | 0 | 100 | 1365.47 | 7958 | 69.55 | -8 | 0 | 100 | 177.29 | 1033.2 | 7.06 | -18 | 0 | 100 | 135.24 | 788.2 | 4.84 | -18 | 0 | 100 | 351.68 | 2049.6 | 12.1
3 | 0 | 100 | 708.03 | 4126.4 | 36.11 | -7 | 0 | 100 | 293.1 | 1708.2 | 11.68 | -17 | 0 | 100 | 134.79 | 785.6 | 4.82 | -17 | 0 | 100 | 484.87 | 2825.8 | 16.68
4 | 0 | 100 | 339.24 | 1977.1 | 17.31 | -6 | 0 | 100 | 541.44 | 3155.5 | 21.56 | -16 | 0 | 100 | 145.94 | 850.6 | 5.21 | -16 | 0 | 100 | 763.87 | 4451.9 | 26.28
5 | 0 | 100 | 198.05 | 1154.3 | 10.1 | -5 | 0 | 100 | 1075.94 | 6270.6 | 42.81 | -15 | 0 | 100 | 174.49 | 1016.9 | 6.24 | -15 | 0 | 100 | 1328.08 | 7740.1 | 45.7
6 | 0 | 100 | 145.55 | 848.3 | 7.42 | -4 | 0 | 100 | 1799.5 | 10487.5 | 71.66 | -14 | 0 | 100 | 236.15 | 1376.3 | 8.45 | -14 | 0 | 100 | 2036.3 | 11867.6 | 70.08
7 | 0 | 100 | 108.38 | 631.6 | 5.53 | -3 | 0 | 100 | 2258.7 | 13163.7 | 89.91 | -13 | 0 | 100 | 298.69 | 1740.7 | 10.68 | -13 | 0 | 100 | 2474.23 | 14419.8 | 85.14
8 | 0 | 100 | 89.6 | 522.2 | 4.59 | -2 | 0 | 100 | 2427.15 | 14145.5 | 96.66 | -12 | 0 | 100 | 421.18 | 2454.6 | 15.07 | -12 | 0 | 100 | 2658.48 | 15493.7 | 91.48
9 | 0 | 100 | 92.33 | 538.1 | 4.71 | -1 | 0 | 100 | 2485.37 | 14484.8 | 99 | -11 | 0 | 100 | 679.65 | 3961 | 24.3 | -11 | 0 | 100 | 2754.3 | 16052.1 | 94.81
10 | 0 | 100 | 78.07 | 455 | 3.97 | 0 | 0 | 100 | 2510.69 | 14632.3 | 100 | -10 | 0 | 100 | 1232.83 | 7185 | 44.07 | -10 | 0 | 100 | 2799.49 | 16315.5 | 96.37
[6 empty] | 1 | 0 | 100 | 2502.65 | 14585.5 | 99.6 | -9 | 0 | 100 | 1940.14 | 11307.2 | 69.38 | -9 | 0 | 100 | 2832.07 | 16505.3 | 97.49
[6 empty] | 2 | 0 | 100 | 2429.37 | 14158.4 | 96.84 | -8 | 0 | 100 | 2421.21 | 14110.9 | 86.56 | -8 | 0 | 100 | 2850.02 | 16610 | 98.1
Measurement time: | 27/05/2023 12:27:28 |  |  |  |  | 3 | 0 | 100 | 2255.05 | 13142.5 | 89.83 | -7 | 0 | 100 | 2584.09 | 15060.1 | 92.42 | -7 | 0 | 100 | 2866.57 | 16706.4 | 98.59
Radiation device: | Novalis TB Accelerator |  |  |  |  | 4 | 0 | 100 | 1860.4 | 10842.4 | 74.12 | -6 | 0 | 100 | 2671.06 | 15567 | 95.52 | -6 | 0 | 100 | 2885.69 | 16817.9 | 99.28
Energy: | 6 MV (FFF) |  |  |  |  | 5 | 0 | 100 | 1201.76 | 7003.9 | 47.86 | -5 | 0 | 100 | 2717.37 | 15836.9 | 97.14 | -5 | 0 | 100 | 2905.67 | 16934.3 | 99.99
Controller: | CCU |  |  |  |  | 6 | 0 | 100 | 587.22 | 3422.3 | 23.41 | -4 | 0 | 100 | 2732.93 | 15927.6 | 97.7 | -4 | 0 | 100 | 2910.42 | 16962 | 100.12
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 7 | 0 | 100 | 340.33 | 1983.4 | 13.56 | -3 | 0 | 100 | 2738.67 | 15961 | 97.95 | -3 | 0 | 100 | 2909.54 | 16956.8 | 100.07
Field detector: | Razor Nano |  |  |  |  | 8 | 0 | 100 | 224.49 | 1308.3 | 8.95 | -2 | 0 | 100 | 2754.16 | 16051.3 | 98.46 | -2 | 0 | 100 | 2905.79 | 16935 | 99.99
Reference detector: | Stealth CI transmission |  |  |  |  | 9 | 0 | 100 | 174.29 | 1015.7 | 6.94 | -1 | 0 | 100 | 2779.27 | 16197.6 | 99.38 | -1 | 0 | 100 | 2895.9 | 16877.4 | 99.68
SAD: | 1000 mm |  |  |  |  | 10 | 0 | 100 | 129.62 | 755.4 | 5.16 | 0 | 0 | 100 | 2795.25 | 16290.8 | 100 | 0 | 0 | 100 | 2904.67 | 16928.5 | 100
SSD: | 900 mm |  |  |  |  | 11 | 0 | 100 | 116.88 | 681.2 | 4.66 | 1 | 0 | 100 | 2810.66 | 16380.6 | 100.51 | 1 | 0 | 100 | 2882.59 | 16799.8 | 99.21
Field size: | 5 x 5 mm |  |  |  |  | 12 | 0 | 100 | 104.3 | 607.9 | 4.16 | 2 | 0 | 100 | 2800.58 | 16321.8 | 100.12 | 2 | 0 | 100 | 2880.5 | 16787.6 | 99.18
Gantry angle: | 0 |  |  |  |  | 13 | 0 | 100 | 106.57 | 621.1 | 4.24 | 3 | 0 | 100 | 2793.03 | 16277.9 | 99.9 | 3 | 0 | 100 | 2888.94 | 16836.8 | 99.42
Measurement medium: | Water |  |  |  |  | 14 | 0 | 100 | 101.3 | 590.4 | 4.04 | 4 | 0 | 100 | 2777.5 | 16187.3 | 99.36 | 4 | 0 | 100 | 2891.15 | 16849.7 | 99.49
Scan type: | Crossline |  |  |  |  | 15 | 0 | 100 | 88.47 | 515.6 | 3.53 | 5 | 0 | 100 | 2739.22 | 15964.3 | 97.94 | 5 | 0 | 100 | 2899.12 | 16896.1 | 99.79
Scan mode: | StepByStep |  |  |  |  | 16 | 0 | 100 | 97.56 | 568.6 | 3.89 | 6 | 0 | 100 | 2684.42 | 15644.9 | 96.04 | 6 | 0 | 100 | 2887.15 | 16826.4 | 99.37
Wedge: | none |  |  |  |  | 17 | 0 | 100 | 83.15 | 484.6 | 3.31 | 7 | 0 | 100 | 2585.69 | 15069.5 | 92.45 | 7 | 0 | 100 | 2892.16 | 16855.6 | 99.58
[6 empty] | 18 | 0 | 100 | 83.85 | 488.7 | 3.34 | 8 | 0 | 100 | 2379.48 | 13867.7 | 85.09 | 8 | 0 | 100 | 2872.1 | 16738.6 | 98.86
Points [mm]: | [5 empty] | 19 | 0 | 100 | 81.12 | 472.8 | 3.24 | 9 | 0 | 100 | 1889.94 | 11014.6 | 67.59 | 9 | 0 | 100 | 2841.72 | 16561.6 | 97.81
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 20 | 0 | 100 | 90.01 | 524.6 | 3.59 | 10 | 0 | 100 | 1203.47 | 7013.8 | 43.06 | 10 | 0 | 100 | 2800.16 | 16319.4 | 96.39
0 | -10 | 100 | 79.93 | 465.8 | 4.07 | [6 empty] | 11 | 0 | 100 | 650.92 | 3793.6 | 23.28 | 11 | 0 | 100 | 2724.38 | 15877.7 | 93.8
0 | -9 | 100 | 85.52 | 498.4 | 4.35 | [6 empty] | 12 | 0 | 100 | 393.69 | 2294.4 | 14.08 | 12 | 0 | 100 | 2631.54 | 15336.6 | 90.53
0 | -8 | 100 | 105.48 | 614.8 | 5.36 | Measurement time: | 27/05/2023 12:45:21 |  |  |  |  | 13 | 0 | 100 | 270.56 | 1576.8 | 9.67 | 13 | 0 | 100 | 2413.62 | 14066.6 | 83.06
0 | -7 | 100 | 119.48 | 696.3 | 6.09 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 14 | 0 | 100 | 224.43 | 1308 | 8.02 | 14 | 0 | 100 | 1901.53 | 11082.1 | 65.45
0 | -6 | 100 | 162.45 | 946.8 | 8.28 | Energy: | 6 MV (FFF) |  |  |  |  | 15 | 0 | 100 | 178.85 | 1042.4 | 6.39 | 15 | 0 | 100 | 1189.92 | 6934.8 | 40.95
0 | -5 | 100 | 264.23 | 1539.9 | 13.48 | Controller: | CCU |  |  |  |  | 16 | 0 | 100 | 158.83 | 925.7 | 5.68 | 16 | 0 | 100 | 671.68 | 3914.5 | 23.12
0 | -4 | 100 | 505.73 | 2947.4 | 25.75 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 17 | 0 | 100 | 140.03 | 816.1 | 5.01 | 17 | 0 | 100 | 440.4 | 2566.7 | 15.16
0 | -3 | 100 | 966.53 | 5632.9 | 49.02 | Field detector: | Razor Nano |  |  |  |  | 18 | 0 | 100 | 124.98 | 728.4 | 4.47 | 18 | 0 | 100 | 332.84 | 1939.8 | 11.46
0 | -2 | 100 | 1553.16 | 9051.9 | 78.54 | Reference detector: | Stealth CI transmission |  |  |  |  | 19 | 0 | 100 | 125.53 | 731.6 | 4.49 | 19 | 0 | 100 | 273.34 | 1593 | 9.41
0 | -1 | 100 | 1913.43 | 11151.5 | 97.29 | SAD: | 1000 mm |  |  |  |  | 20 | 0 | 100 | 118.85 | 692.7 | 4.25 | 20 | 0 | 100 | 217.51 | 1267.7 | 7.49
0 | 0 | 100 | 1968.72 | 11473.7 | 100 | SSD: | 900 mm |  |  |  |  | 21 | 0 | 100 | 113.21 | 659.8 | 4.05 | 21 | 0 | 100 | 205.43 | 1197.2 | 7.07
0 | 1 | 100 | 1887.12 | 10998.2 | 95.92 | Field size: | 10 x 10 mm |  |  |  |  | 22 | 0 | 100 | 110.76 | 645.5 | 3.96 | 22 | 0 | 100 | 189.72 | 1105.7 | 6.53
0 | 2 | 100 | 1552.16 | 9046 | 78.9 | Gantry angle: | 0 |  |  |  |  | 23 | 0 | 100 | 94.75 | 552.2 | 3.39 | 23 | 0 | 100 | 166 | 967.4 | 5.72
0 | 3 | 100 | 953.16 | 5555 | 48.62 | Measurement medium: | Water |  |  |  |  | 24 | 0 | 100 | 91.03 | 530.5 | 3.26 | 24 | 0 | 100 | 154.64 | 901.3 | 5.32
0 | 4 | 100 | 473.13 | 2757.4 | 24.05 | Scan type: | Crossline |  |  |  |  | 25 | 0 | 100 | 92.88 | 541.3 | 3.32 | 25 | 0 | 100 | 161.07 | 938.7 | 5.54
0 | 5 | 100 | 254.11 | 1481 | 12.95 | Scan mode: | StepByStep |  |  |  |  | 26 | 0 | 100 | 93.22 | 543.3 | 3.33 | 26 | 0 | 100 | 141.68 | 825.7 | 4.88
0 | 6 | 100 | 147.65 | 860.5 | 7.5 | Wedge: | none |  |  |  |  | 27 | 0 | 100 | 100.77 | 587.3 | 3.6 | 27 | 0 | 100 | 142.38 | 829.8 | 4.9
0 | 7 | 100 | 113.06 | 658.9 | 5.75 | [6 empty] | 28 | 0 | 100 | 93.06 | 542.3 | 3.33 | 28 | 0 | 100 | 132.72 | 773.5 | 4.57
0 | 8 | 100 | 96.04 | 559.7 | 4.88 | Points [mm]: | [5 empty] | 29 | 0 | 100 | 93.51 | 544.9 | 3.35 | 29 | 0 | 100 | 126.17 | 735.3 | 4.34
0 | 9 | 100 | 77.82 | 453.5 | 3.98 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 30 | 0 | 100 | 89.9 | 524 | 3.22 | 30 | 0 | 100 | 120.82 | 704.2 | 4.16
0 | 10 | 100 | 75.67 | 441 | 3.84 | 0 | -20 | 100 | 89.11 | 519.3 | 3.56
[6 empty] | 0 | -19 | 100 | 81.67 | 476 | 3.26
[6 empty] | 0 | -18 | 100 | 85.06 | 495.7 | 3.4 | Measurement time: | 27/05/2023 13:06:18 |  |  |  |  | Measurement time: | 27/05/2023 13:37:05
Measurement time: | 27/05/2023 12:28:57 |  |  |  |  | 0 | -17 | 100 | 89.63 | 522.4 | 3.58 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -16 | 100 | 84.28 | 491.2 | 3.37 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Energy: | 6 MV (FFF) |  |  |  |  | 0 | -15 | 100 | 90.18 | 525.6 | 3.6 | Controller: | CCU |  |  |  |  | Controller: | CCU
Controller: | CCU |  |  |  |  | 0 | -14 | 100 | 89.03 | 518.9 | 3.55 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -13 | 100 | 105.95 | 617.5 | 4.23 | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano
Field detector: | Razor Nano |  |  |  |  | 0 | -12 | 100 | 113.28 | 660.2 | 4.52 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -11 | 100 | 126 | 734.4 | 5.03 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SAD: | 1000 mm |  |  |  |  | 0 | -10 | 100 | 155.41 | 905.7 | 6.21 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
SSD: | 900 mm |  |  |  |  | 0 | -9 | 100 | 195 | 1136.5 | 7.79 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Field size: | 5 x 5 mm |  |  |  |  | 0 | -8 | 100 | 261.7 | 1525.2 | 10.45 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Gantry angle: | 0 |  |  |  |  | 0 | -7 | 100 | 463.82 | 2703.1 | 18.51 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Measurement medium: | Water |  |  |  |  | 0 | -6 | 100 | 845.24 | 4926.1 | 33.79 | Scan type: | Crossline |  |  |  |  | Scan type: | Crossline
Scan type: | Beam |  |  |  |  | 0 | -5 | 100 | 1521.97 | 8870.1 | 60.9 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Scan mode: | StepByStep |  |  |  |  | 0 | -4 | 100 | 2084.98 | 12151.3 | 83.31 | Wedge: | none |  |  |  |  | Wedge: | none
Wedge: | none |  |  |  |  | 0 | -3 | 100 | 2364.15 | 13778.3 | 94.41
[6 empty] | 0 | -2 | 100 | 2453.93 | 14301.5 | 97.99 | Points [mm]: | [5 empty] | Points [mm]:
Points [mm]: | [5 empty] | 0 | -1 | 100 | 2497.46 | 14555.3 | 99.76 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 0 | 100 | 2505.53 | 14602.3 | 100 | 0 | -30 | 100 | 94.09 | 548.4 | 3.39 | 0 | -30 | 100 | 143.31 | 835.2 | 4.92
0 | 0 | -0.5 | 2373.54 | 13833 | 60.54 | 0 | 1 | 100 | 2488.23 | 14501.4 | 99.33 | 0 | -29 | 100 | 96.81 | 564.2 | 3.49 | 0 | -29 | 100 | 154.69 | 901.5 | 5.31
0 | 0 | 0.5 | 2411.72 | 14055.6 | 61.23 | 0 | 2 | 100 | 2424.52 | 14130.1 | 96.86 | 0 | -28 | 100 | 98.05 | 571.5 | 3.53 | 0 | -28 | 100 | 159.88 | 931.8 | 5.49
0 | 0 | 1.5 | 2652.9 | 15461.1 | 67.51 | 0 | 3 | 100 | 2309.51 | 13459.9 | 92.23 | 0 | -27 | 100 | 101.09 | 589.2 | 3.65 | 0 | -27 | 100 | 164.81 | 960.5 | 5.66
0 | 0 | 2.5 | 3100.67 | 18070.7 | 78.97 | 0 | 4 | 100 | 2035.95 | 11865.5 | 81.44 | 0 | -26 | 100 | 99.56 | 580.2 | 3.59 | 0 | -26 | 100 | 169.58 | 988.3 | 5.82
0 | 0 | 3.5 | 3466.9 | 20205.2 | 88.5 | 0 | 5 | 100 | 1487.72 | 8670.5 | 59.42 | 0 | -25 | 100 | 101.82 | 593.4 | 3.67 | 0 | -25 | 100 | 164.77 | 960.3 | 5.65
0 | 0 | 4.5 | 3686.17 | 21483.1 | 94.22 | 0 | 6 | 100 | 839.77 | 4894.2 | 33.55 | 0 | -24 | 100 | 114.96 | 670 | 4.14 | 0 | -24 | 100 | 188.6 | 1099.1 | 6.47
0 | 0 | 5.5 | 3813.02 | 22222.3 | 97.19 | 0 | 7 | 100 | 454.6 | 2649.4 | 18.16 | 0 | -23 | 100 | 91.96 | 535.9 | 3.32 | 0 | -23 | 100 | 203.85 | 1188.1 | 7
0 | 0 | 6.5 | 3880.7 | 22616.8 | 98.84 | 0 | 8 | 100 | 268.18 | 1563 | 10.72 | 0 | -22 | 100 | 111.37 | 649.1 | 4.02 | 0 | -22 | 100 | 222.76 | 1298.3 | 7.65
0 | 0 | 7.5 | 3914.51 | 22813.9 | 99.8 | 0 | 9 | 100 | 188.3 | 1097.4 | 7.53 | 0 | -21 | 100 | 116.11 | 676.7 | 4.19 | 0 | -21 | 100 | 236.52 | 1378.4 | 8.12
0 | 0 | 8.5 | 3920.25 | 22847.3 | 99.88 | 0 | 10 | 100 | 142.63 | 831.2 | 5.69 | 0 | -20 | 100 | 123.78 | 721.4 | 4.46 | 0 | -20 | 100 | 264.89 | 1543.8 | 9.1
0 | 0 | 9.5 | 3921.86 | 22856.7 | 100 | 0 | 11 | 100 | 121.33 | 707.1 | 4.85 | 0 | -19 | 100 | 130.8 | 762.3 | 4.72 | 0 | -19 | 100 | 353.75 | 2061.6 | 12.14
0 | 0 | 10.5 | 3909.29 | 22783.4 | 99.79 | 0 | 12 | 100 | 106.14 | 618.6 | 4.24 | 0 | -18 | 100 | 144.22 | 840.5 | 5.2 | 0 | -18 | 100 | 470.65 | 2743 | 16.15
0 | 0 | 11.5 | 3894.82 | 22699.1 | 99.31 | 0 | 13 | 100 | 90.32 | 526.4 | 3.61 | 0 | -17 | 100 | 149.18 | 869.5 | 5.38 | 0 | -17 | 100 | 714.4 | 4163.5 | 24.52
0 | 0 | 12.5 | 3879.73 | 22611.1 | 98.82 | 0 | 14 | 100 | 82.15 | 478.8 | 3.28 | 0 | -16 | 100 | 167.99 | 979 | 6.06 | 0 | -16 | 100 | 1143.71 | 6665.6 | 39.26
0 | 0 | 13.5 | 3851.13 | 22444.5 | 97.95 | 0 | 15 | 100 | 81.84 | 477 | 3.27 | 0 | -15 | 100 | 193.82 | 1129.6 | 6.99 | 0 | -15 | 100 | 1712.54 | 9980.7 | 58.81
0 | 0 | 14.5 | 3821.62 | 22272.5 | 96.99 | 0 | 16 | 100 | 81.81 | 476.8 | 3.27 | 0 | -14 | 100 | 257.23 | 1499.2 | 9.28 | 0 | -14 | 100 | 2241.24 | 13062 | 76.9
0 | 0 | 15.5 | 3789.51 | 22085.3 | 96.21 | 0 | 17 | 100 | 66.61 | 388.2 | 2.67 | 0 | -13 | 100 | 366.81 | 2137.8 | 13.24 | 0 | -13 | 100 | 2545.09 | 14832.8 | 87.33
0 | 0 | 16.5 | 3771 | 21977.4 | 95.99 | 0 | 18 | 100 | 71.19 | 414.9 | 2.85 | 0 | -12 | 100 | 579.96 | 3380 | 20.9 | 0 | -12 | 100 | 2677.68 | 15605.5 | 91.9
0 | 0 | 17.5 | 3748.25 | 21844.8 | 95.61 | 0 | 19 | 100 | 87.5 | 509.9 | 3.5 | 0 | -11 | 100 | 995.27 | 5800.5 | 35.88 | 0 | -11 | 100 | 2757.13 | 16068.6 | 94.62
0 | 0 | 18.5 | 3713.37 | 21641.6 | 94.62 | 0 | 20 | 100 | 71.69 | 417.8 | 2.87 | 0 | -10 | 100 | 1617.27 | 9425.5 | 58.37 | 0 | -10 | 100 | 2796.63 | 16298.8 | 95.99
0 | 0 | 19.5 | 3680.34 | 21449.1 | 93.77 | [6 empty] | 0 | -9 | 100 | 2187.11 | 12746.5 | 78.9 | 0 | -9 | 100 | 2834.63 | 16520.3 | 97.2
0 | 0 | 20.5 | 3644.94 | 21242.8 | 93.15 | [6 empty] | 0 | -8 | 100 | 2475.71 | 14428.5 | 89.38 | 0 | -8 | 100 | 2860.83 | 16673 | 98.18
0 | 0 | 21.5 | 3616.35 | 21076.1 | 92.29 | Measurement time: | 27/05/2023 12:48:22 |  |  |  |  | 0 | -7 | 100 | 2612.51 | 15225.8 | 94.28 | 0 | -7 | 100 | 2878.29 | 16774.7 | 98.78
0 | 0 | 22.5 | 3588.61 | 20914.5 | 91.55 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -6 | 100 | 2686.22 | 15655.4 | 96.85 | 0 | -6 | 100 | 2899.1 | 16896 | 99.48
0 | 0 | 23.5 | 3557.93 | 20735.7 | 90.55 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | -5 | 100 | 2721.92 | 15863.4 | 98.13 | 0 | -5 | 100 | 2896.16 | 16878.9 | 99.33
0 | 0 | 24.5 | 3533.07 | 20590.8 | 89.84 | Controller: | CCU |  |  |  |  | 0 | -4 | 100 | 2740.82 | 15973.6 | 98.82 | 0 | -4 | 100 | 2909.04 | 16954 | 99.8
0 | 0 | 25.5 | 3504.89 | 20426.5 | 89.41 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -3 | 100 | 2760.99 | 16091.1 | 99.57 | 0 | -3 | 100 | 2920.19 | 17018.9 | 100.16
0 | 0 | 26.5 | 3477.59 | 20267.5 | 88.74 | Field detector: | Razor Nano |  |  |  |  | 0 | -2 | 100 | 2765.66 | 16118.3 | 99.8 | 0 | -2 | 100 | 2915.6 | 16992.2 | 99.99
0 | 0 | 27.5 | 3450.29 | 20108.4 | 88.07 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -1 | 100 | 2787.43 | 16245.2 | 100.57 | 0 | -1 | 100 | 2915.39 | 16990.9 | 100.07
0 | 0 | 28.5 | 3422.99 | 19949.3 | 87.4 | SAD: | 1000 mm |  |  |  |  | 0 | 0 | 100 | 2773.89 | 16166.3 | 100 | 0 | 0 | 100 | 2914.79 | 16987.4 | 100
0 | 0 | 29.5 | 3395.7 | 19790.2 | 86.73 | SSD: | 900 mm |  |  |  |  | 0 | 1 | 100 | 2776.98 | 16184.3 | 100.16 | 0 | 1 | 100 | 2914.67 | 16986.7 | 100.03
0 | 0 | 30.5 | 3369.68 | 19638.5 | 86.09 | Field size: | 10 x 10 mm |  |  |  |  | 0 | 2 | 100 | 2776.85 | 16183.5 | 100.22 | 0 | 2 | 100 | 2920.45 | 17020.4 | 100.22
0 | 0 | 31.5 | 3344.47 | 19491.6 | 85.48 | Gantry angle: | 0 |  |  |  |  | 0 | 3 | 100 | 2764.5 | 16111.6 | 99.75 | 0 | 3 | 100 | 2905.3 | 16932.2 | 99.7
0 | 0 | 32.5 | 3319.26 | 19344.7 | 84.86 | Measurement medium: | Water |  |  |  |  | 0 | 4 | 100 | 2744.82 | 15996.8 | 99 | 0 | 4 | 100 | 2898.16 | 16890.6 | 99.46
0 | 0 | 33.5 | 3294.05 | 19197.8 | 84.25 | Scan type: | Beam |  |  |  |  | 0 | 5 | 100 | 2709.67 | 15792 | 97.72 | 0 | 5 | 100 | 2893.22 | 16861.8 | 99.25
0 | 0 | 34.5 | 3268.85 | 19050.9 | 83.63 | Scan mode: | StepByStep |  |  |  |  | 0 | 6 | 100 | 2674.98 | 15589.9 | 96.48 | 0 | 6 | 100 | 2894.91 | 16871.6 | 99.25
0 | 0 | 35.5 | 3244.99 | 18911.9 | 83 | Wedge: | none |  |  |  |  | 0 | 7 | 100 | 2588.3 | 15084.7 | 93.34 | 0 | 7 | 100 | 2871.59 | 16735.7 | 98.53
0 | 0 | 36.5 | 3222.01 | 18777.9 | 82.34 | [6 empty] | 0 | 8 | 100 | 2435.63 | 14194.9 | 87.88 | 0 | 8 | 100 | 2862.2 | 16680.9 | 98.23
0 | 0 | 37.5 | 3199.02 | 18643.9 | 81.68 | Points [mm]: | [5 empty] | 0 | 9 | 100 | 2154.16 | 12554.5 | 77.69 | 0 | 9 | 100 | 2836.49 | 16531.1 | 97.32
0 | 0 | 38.5 | 3176.03 | 18510 | 81.03 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 10 | 100 | 1585.97 | 9243.1 | 57.22 | 0 | 10 | 100 | 2809.18 | 16371.9 | 96.37
0 | 0 | 39.5 | 3153.05 | 18376 | 80.37 | 0 | 0 | -0.5 | 2467.99 | 14383.5 | 52.13 | 0 | 11 | 100 | 971.87 | 5664.1 | 35.06 | 0 | 11 | 100 | 2745.74 | 16002.2 | 94.21
0 | 0 | 40.5 | 3128.87 | 18235.1 | 79.76 | 0 | 0 | 0.5 | 2517.37 | 14671.3 | 53.18 | 0 | 12 | 100 | 554.05 | 3229 | 19.99 | 0 | 12 | 100 | 2654.8 | 15472.2 | 91.08
0 | 0 | 41.5 | 3103.75 | 18088.7 | 79.18 | 0 | 0 | 1.5 | 2829.65 | 16491.2 | 59.74 | 0 | 13 | 100 | 364.7 | 2125.5 | 13.16 | 0 | 13 | 100 | 2499.68 | 14568.2 | 85.78
0 | 0 | 42.5 | 3078.62 | 17942.3 | 78.59 | 0 | 0 | 2.5 | 3369.61 | 19638.2 | 71.13 | 0 | 14 | 100 | 264.53 | 1541.7 | 9.54 | 0 | 14 | 100 | 2174.43 | 12672.6 | 74.58
0 | 0 | 43.5 | 3053.49 | 17795.8 | 78.01 | 0 | 0 | 3.5 | 3844.35 | 22404.9 | 81.18 | 0 | 15 | 100 | 206.8 | 1205.2 | 7.46 | 0 | 15 | 100 | 1649.91 | 9615.7 | 56.61
0 | 0 | 44.5 | 3028.37 | 17649.4 | 77.43 | 0 | 0 | 4.5 | 4170.85 | 24307.8 | 88.06 | 0 | 16 | 100 | 166.35 | 969.5 | 6 | 0 | 16 | 100 | 1091.23 | 6359.7 | 37.45
0 | 0 | 45.5 | 3003.28 | 17503.2 | 76.76 | 0 | 0 | 5.5 | 4390.5 | 25587.9 | 92.67 | 0 | 17 | 100 | 151.57 | 883.4 | 5.47 | 0 | 17 | 100 | 671.82 | 3915.4 | 23.06
0 | 0 | 46.5 | 2978.22 | 17357.1 | 76.02 | 0 | 0 | 6.5 | 4533.97 | 26424.1 | 95.68 | 0 | 18 | 100 | 132.5 | 772.2 | 4.78 | 0 | 18 | 100 | 455.26 | 2653.2 | 15.62
0 | 0 | 47.5 | 2953.16 | 17211.1 | 75.28 | 0 | 0 | 7.5 | 4627.97 | 26971.9 | 97.61 | 0 | 19 | 100 | 120.87 | 704.4 | 4.36 | 0 | 19 | 100 | 341.31 | 1989.2 | 11.71
0 | 0 | 48.5 | 2928.1 | 17065 | 74.55 | 0 | 0 | 8.5 | 4695.83 | 27367.4 | 99.03 | 0 | 20 | 100 | 116.8 | 680.7 | 4.21 | 0 | 20 | 100 | 267.45 | 1558.7 | 9.17
0 | 0 | 49.5 | 2903.05 | 16919 | 73.81 | 0 | 0 | 9.5 | 4730.42 | 27569 | 99.77 | 0 | 21 | 100 | 113.85 | 663.5 | 4.11 | 0 | 21 | 100 | 228.04 | 1329 | 7.83
0 | 0 | 50.5 | 2880.42 | 16787.1 | 73.2 | 0 | 0 | 10.5 | 4741.66 | 27634.5 | 100 | 0 | 22 | 100 | 96.62 | 563.1 | 3.49 | 0 | 22 | 100 | 207.8 | 1211.1 | 7.13
0 | 0 | 51.5 | 2859.7 | 16666.4 | 72.71 | 0 | 0 | 11.5 | 4740.36 | 27626.9 | 99.94 | 0 | 23 | 100 | 95.78 | 558.2 | 3.46 | 0 | 23 | 100 | 191.81 | 1117.9 | 6.58
0 | 0 | 52.5 | 2838.99 | 16545.7 | 72.21 | 0 | 0 | 12.5 | 4728.41 | 27557.3 | 99.67 | 0 | 24 | 100 | 91.83 | 535.2 | 3.31 | 0 | 24 | 100 | 182.88 | 1065.8 | 6.27
0 | 0 | 53.5 | 2818.27 | 16424.9 | 71.71 | 0 | 0 | 13.5 | 4722.82 | 27524.7 | 99.53 | 0 | 25 | 100 | 90.06 | 524.9 | 3.25 | 0 | 25 | 100 | 170.06 | 991.1 | 5.83
0 | 0 | 54.5 | 2797.55 | 16304.2 | 71.21 | 0 | 0 | 14.5 | 4704.95 | 27420.5 | 99.14 | 0 | 26 | 100 | 84.99 | 495.3 | 3.07 | 0 | 26 | 100 | 160.94 | 937.9 | 5.52
0 | 0 | 55.5 | 2775.65 | 16176.5 | 70.66 | 0 | 0 | 15.5 | 4674.75 | 27244.5 | 98.56 | 0 | 27 | 100 | 97.8 | 570 | 3.53 | 0 | 27 | 100 | 146.73 | 855.2 | 5.03
0 | 0 | 56.5 | 2752.65 | 16042.5 | 70.04 | 0 | 0 | 16.5 | 4651.51 | 27109.1 | 98.11 | 0 | 28 | 100 | 90.87 | 529.6 | 3.28 | 0 | 28 | 100 | 145.28 | 846.7 | 4.98
0 | 0 | 57.5 | 2729.66 | 15908.5 | 69.43 | 0 | 0 | 17.5 | 4631.62 | 26993.1 | 97.62 | 0 | 29 | 100 | 74 | 431.2 | 2.67 | 0 | 29 | 100 | 139.01 | 810.2 | 4.77
0 | 0 | 58.5 | 2706.67 | 15774.5 | 68.81 | 0 | 0 | 18.5 | 4601.57 | 26818 | 96.96 | 0 | 30 | 100 | 66.15 | 385.5 | 2.39 | 0 | 30 | 100 | 128.32 | 747.8 | 4.4
0 | 0 | 59.5 | 2683.67 | 15640.5 | 68.2 | 0 | 0 | 19.5 | 4560.77 | 26580.3 | 96.15
0 | 0 | 60.5 | 2662.77 | 15518.7 | 67.69 | 0 | 0 | 20.5 | 4528.8 | 26393.9 | 95.51
0 | 0 | 61.5 | 2643.83 | 15408.3 | 67.27 | 0 | 0 | 21.5 | 4500.49 | 26228.9 | 94.95 | Measurement time: | 27/05/2023 13:09:05 |  |  |  |  | Measurement time: | 27/05/2023 13:39:45
0 | 0 | 62.5 | 2624.89 | 15297.9 | 66.85 | 0 | 0 | 22.5 | 4465.25 | 26023.6 | 94.14 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
0 | 0 | 63.5 | 2605.94 | 15187.5 | 66.43 | 0 | 0 | 23.5 | 4430.9 | 25823.4 | 93.44 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
0 | 0 | 64.5 | 2587 | 15077.1 | 66.01 | 0 | 0 | 24.5 | 4393.87 | 25607.6 | 92.74 | Controller: | CCU |  |  |  |  | Controller: | CCU
0 | 0 | 65.5 | 2568.18 | 14967.4 | 65.57 | 0 | 0 | 25.5 | 4356.83 | 25391.7 | 91.94 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | 66.5 | 2549.49 | 14858.5 | 65.1 | 0 | 0 | 26.5 | 4321.12 | 25183.6 | 91.21 | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano
0 | 0 | 67.5 | 2530.8 | 14749.5 | 64.64 | 0 | 0 | 27.5 | 4285.41 | 24975.4 | 90.47 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
0 | 0 | 68.5 | 2512.11 | 14640.6 | 64.17 | 0 | 0 | 28.5 | 4249.7 | 24767.3 | 89.73 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
0 | 0 | 69.5 | 2493.41 | 14531.7 | 63.7 | 0 | 0 | 29.5 | 4213.98 | 24559.2 | 89 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
0 | 0 | 70.5 | 2475.4 | 14426.7 | 63.25 | 0 | 0 | 30.5 | 4180.87 | 24366.2 | 88.29 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
0 | 0 | 71.5 | 2457.99 | 14325.2 | 62.82 | 0 | 0 | 31.5 | 4149.53 | 24183.5 | 87.6 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
0 | 0 | 72.5 | 2440.59 | 14223.8 | 62.38 | 0 | 0 | 32.5 | 4118.19 | 24000.9 | 86.91 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
0 | 0 | 73.5 | 2423.19 | 14122.4 | 61.95 | 0 | 0 | 33.5 | 4086.85 | 23818.3 | 86.22 | Scan type: | Beam |  |  |  |  | Scan type: | Beam
0 | 0 | 74.5 | 2405.78 | 14020.9 | 61.51 | 0 | 0 | 34.5 | 4055.51 | 23635.6 | 85.53 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
0 | 0 | 75.5 | 2389.37 | 13925.3 | 61.08 | 0 | 0 | 35.5 | 4026.99 | 23469.4 | 84.92 | Wedge: | none |  |  |  |  | Wedge: | none
0 | 0 | 76.5 | 2374.12 | 13836.4 | 60.64 | 0 | 0 | 36.5 | 4000.33 | 23314 | 84.37
0 | 0 | 77.5 | 2358.86 | 13747.5 | 60.2 | 0 | 0 | 37.5 | 3973.66 | 23158.6 | 83.81 | Points [mm]: | [5 empty] | Points [mm]:
0 | 0 | 78.5 | 2343.61 | 13658.6 | 59.76 | 0 | 0 | 38.5 | 3947 | 23003.2 | 83.25 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 79.5 | 2328.36 | 13569.7 | 59.32 | 0 | 0 | 39.5 | 3920.33 | 22847.8 | 82.69 | 0 | 0 | -0.5 | 2444.68 | 14247.6 | 48.01 | 0 | 0 | -0.5 | 2388.81 | 13922 | 46.59
0 | 0 | 80.5 | 2312 | 13474.4 | 58.9 | 0 | 0 | 40.5 | 3890.65 | 22674.8 | 82.08 | 0 | 0 | 0.5 | 2515.57 | 14660.8 | 49.41 | 0 | 0 | 0.5 | 2452.35 | 14292.3 | 47.83
0 | 0 | 81.5 | 2294.21 | 13370.7 | 58.5 | 0 | 0 | 41.5 | 3858.95 | 22490 | 81.43 | 0 | 0 | 1.5 | 2783.29 | 16221 | 54.68 | 0 | 0 | 1.5 | 2734.58 | 15937.2 | 53.33
0 | 0 | 82.5 | 2276.43 | 13267.1 | 58.11 | 0 | 0 | 42.5 | 3827.24 | 22305.2 | 80.78 | 0 | 0 | 2.5 | 3338.43 | 19456.4 | 65.57 | 0 | 0 | 2.5 | 3258.22 | 18989 | 63.53
0 | 0 | 83.5 | 2258.64 | 13163.4 | 57.71 | 0 | 0 | 43.5 | 3795.53 | 22120.4 | 80.13 | 0 | 0 | 3.5 | 3864.66 | 22523.3 | 75.93 | 0 | 0 | 3.5 | 3771.51 | 21980.4 | 73.55
0 | 0 | 84.5 | 2240.86 | 13059.7 | 57.31 | 0 | 0 | 44.5 | 3763.83 | 21935.7 | 79.48 | 0 | 0 | 4.5 | 4227.77 | 24639.5 | 83.06 | 0 | 0 | 4.5 | 4152.16 | 24198.9 | 80.96
0 | 0 | 85.5 | 2223.11 | 12956.3 | 56.87 | 0 | 0 | 45.5 | 3734.05 | 21762.1 | 78.85 | 0 | 0 | 5.5 | 4480.4 | 26111.9 | 88.01 | 0 | 0 | 5.5 | 4445.99 | 25911.3 | 86.66
0 | 0 | 86.5 | 2205.41 | 12853.2 | 56.37 | 0 | 0 | 46.5 | 3705.75 | 21597.2 | 78.23 | 0 | 0 | 6.5 | 4674.4 | 27242.5 | 91.82 | 0 | 0 | 6.5 | 4660.19 | 27159.7 | 90.85
0 | 0 | 87.5 | 2187.71 | 12750 | 55.88 | 0 | 0 | 47.5 | 3677.46 | 21432.3 | 77.61 | 0 | 0 | 7.5 | 4824.84 | 28119.2 | 94.73 | 0 | 0 | 7.5 | 4809.02 | 28027.1 | 93.76
0 | 0 | 88.5 | 2170.01 | 12646.9 | 55.38 | 0 | 0 | 48.5 | 3649.16 | 21267.4 | 77 | 0 | 0 | 8.5 | 4924.51 | 28700.2 | 96.72 | 0 | 0 | 8.5 | 4918.71 | 28666.3 | 95.88
0 | 0 | 89.5 | 2152.31 | 12543.7 | 54.88 | 0 | 0 | 49.5 | 3620.87 | 21102.5 | 76.38 | 0 | 0 | 9.5 | 4993.05 | 29099.6 | 98.08 | 0 | 0 | 9.5 | 4999.37 | 29136.5 | 97.47
0 | 0 | 90.5 | 2135.33 | 12444.7 | 54.42 | 0 | 0 | 50.5 | 3594.33 | 20947.8 | 75.82 | 0 | 0 | 10.5 | 5045.05 | 29402.7 | 99.06 | 0 | 0 | 10.5 | 5052.31 | 29445 | 98.52
0 | 0 | 91.5 | 2119.3 | 12351.3 | 54.01 | 0 | 0 | 51.5 | 3569.11 | 20800.8 | 75.29 | 0 | 0 | 11.5 | 5076.53 | 29586.1 | 99.68 | 0 | 0 | 11.5 | 5087.69 | 29651.2 | 99.18
0 | 0 | 92.5 | 2103.27 | 12257.9 | 53.6 | 0 | 0 | 52.5 | 3543.89 | 20653.9 | 74.77 | 0 | 0 | 12.5 | 5089.2 | 29660 | 99.96 | 0 | 0 | 12.5 | 5118.06 | 29828.2 | 99.76
0 | 0 | 93.5 | 2087.24 | 12164.5 | 53.19 | 0 | 0 | 53.5 | 3518.68 | 20506.9 | 74.24 | 0 | 0 | 13.5 | 5092.94 | 29681.8 | 100 | 0 | 0 | 13.5 | 5127.01 | 29880.3 | 99.95
0 | 0 | 94.5 | 2071.22 | 12071.1 | 52.78 | 0 | 0 | 54.5 | 3493.46 | 20359.9 | 73.72 | 0 | 0 | 14.5 | 5092.29 | 29678 | 99.96 | 0 | 0 | 14.5 | 5128.58 | 29889.5 | 100
0 | 0 | 95.5 | 2055.14 | 11977.4 | 52.37 | 0 | 0 | 55.5 | 3468.07 | 20212 | 73.18 | 0 | 0 | 15.5 | 5083.13 | 29624.6 | 99.8 | 0 | 0 | 15.5 | 5120.01 | 29839.5 | 99.84
0 | 0 | 96.5 | 2038.99 | 11883.3 | 51.95 | 0 | 0 | 56.5 | 3442.52 | 20063.1 | 72.62 | 0 | 0 | 16.5 | 5054.56 | 29458.1 | 99.25 | 0 | 0 | 16.5 | 5101.07 | 29729.1 | 99.45
0 | 0 | 97.5 | 2022.84 | 11789.2 | 51.53 | 0 | 0 | 57.5 | 3416.98 | 19914.2 | 72.06 | 0 | 0 | 17.5 | 5023.27 | 29275.7 | 98.65 | 0 | 0 | 17.5 | 5083.73 | 29628.1 | 99.08
0 | 0 | 98.5 | 2006.7 | 11695.1 | 51.11 | 0 | 0 | 58.5 | 3391.43 | 19765.3 | 71.5 | 0 | 0 | 18.5 | 5003.92 | 29162.9 | 98.27 | 0 | 0 | 18.5 | 5061.82 | 29500.4 | 98.65
0 | 0 | 99.5 | 1990.55 | 11601 | 50.69 | 0 | 0 | 59.5 | 3365.88 | 19616.4 | 70.95 | 0 | 0 | 19.5 | 4979.48 | 29020.5 | 97.76 | 0 | 0 | 19.5 | 5039.06 | 29367.7 | 98.23
0 | 0 | 100.5 | 1976.66 | 11520 | 50.33 | 0 | 0 | 60.5 | 3341.43 | 19473.9 | 70.42 | 0 | 0 | 20.5 | 4946.82 | 28830.2 | 97.1 | 0 | 0 | 20.5 | 5022.29 | 29270 | 97.89
0 | 0 | 101.5 | 1965.58 | 11455.5 | 50.05 | 0 | 0 | 61.5 | 3317.95 | 19337.1 | 69.93 | 0 | 0 | 21.5 | 4907.81 | 28602.8 | 96.34 | 0 | 0 | 21.5 | 5001.54 | 29149.1 | 97.49
0 | 0 | 102.5 | 1954.51 | 11390.9 | 49.76 | 0 | 0 | 62.5 | 3294.46 | 19200.2 | 69.43 | 0 | 0 | 22.5 | 4865.7 | 28357.4 | 95.54 | 0 | 0 | 22.5 | 4974.92 | 28993.9 | 96.98
0 | 0 | 103.5 | 1943.43 | 11326.3 | 49.48 | 0 | 0 | 63.5 | 3270.98 | 19063.3 | 68.93 | 0 | 0 | 23.5 | 4828 | 28137.7 | 94.82 | 0 | 0 | 23.5 | 4940.07 | 28790.8 | 96.31
0 | 0 | 104.5 | 1932.35 | 11261.8 | 49.19 | 0 | 0 | 64.5 | 3247.5 | 18926.5 | 68.44 | 0 | 0 | 24.5 | 4797.24 | 27958.4 | 94.16 | 0 | 0 | 24.5 | 4900.37 | 28559.4 | 95.54
0 | 0 | 105.5 | 1920.36 | 11191.9 | 48.9 | 0 | 0 | 65.5 | 3223.65 | 18787.5 | 67.94 | 0 | 0 | 25.5 | 4767.2 | 27783.3 | 93.57 | 0 | 0 | 25.5 | 4870.06 | 28382.8 | 94.94
0 | 0 | 106.5 | 1907.1 | 11114.6 | 48.58 | 0 | 0 | 66.5 | 3199.49 | 18646.7 | 67.45 | 0 | 0 | 26.5 | 4735.64 | 27599.4 | 92.96 | 0 | 0 | 26.5 | 4840.59 | 28211 | 94.36
0 | 0 | 107.5 | 1893.85 | 11037.4 | 48.27 | 0 | 0 | 67.5 | 3175.32 | 18505.9 | 66.96 | 0 | 0 | 27.5 | 4704.08 | 27415.4 | 92.35 | 0 | 0 | 27.5 | 4811.12 | 28039.3 | 93.79
0 | 0 | 108.5 | 1880.6 | 10960.2 | 47.95 | 0 | 0 | 68.5 | 3151.16 | 18365 | 66.47 | 0 | 0 | 28.5 | 4672.51 | 27231.5 | 91.74 | 0 | 0 | 28.5 | 4781.65 | 27867.5 | 93.21
0 | 0 | 109.5 | 1867.34 | 10882.9 | 47.64 | 0 | 0 | 69.5 | 3127 | 18224.2 | 65.98 | 0 | 0 | 29.5 | 4640.95 | 27047.6 | 91.13 | 0 | 0 | 29.5 | 4752.18 | 27695.8 | 92.63
0 | 0 | 110.5 | 1854.29 | 10806.8 | 47.31 | 0 | 0 | 70.5 | 3103.98 | 18090.1 | 65.49 | 0 | 0 | 30.5 | 4608.06 | 26855.9 | 90.49 | 0 | 0 | 30.5 | 4721.26 | 27515.6 | 92.03
0 | 0 | 111.5 | 1841.49 | 10732.2 | 46.97 | 0 | 0 | 71.5 | 3082.03 | 17962.2 | 65.01 | 0 | 0 | 31.5 | 4574.35 | 26659.4 | 89.82 | 0 | 0 | 31.5 | 4689.41 | 27330 | 91.41
0 | 0 | 112.5 | 1828.7 | 10657.7 | 46.63 | 0 | 0 | 72.5 | 3060.09 | 17834.2 | 64.52 | 0 | 0 | 32.5 | 4540.65 | 26463 | 89.16 | 0 | 0 | 32.5 | 4657.56 | 27144.4 | 90.79
0 | 0 | 113.5 | 1815.9 | 10583.1 | 46.28 | 0 | 0 | 73.5 | 3038.14 | 17706.3 | 64.04 | 0 | 0 | 33.5 | 4506.94 | 26266.6 | 88.49 | 0 | 0 | 33.5 | 4625.71 | 26958.7 | 90.16
0 | 0 | 114.5 | 1803.11 | 10508.5 | 45.94 | 0 | 0 | 74.5 | 3016.19 | 17578.4 | 63.56 | 0 | 0 | 34.5 | 4473.24 | 26070.1 | 87.83 | 0 | 0 | 34.5 | 4593.86 | 26773.1 | 89.54
0 | 0 | 115.5 | 1790.33 | 10434.1 | 45.62 | 0 | 0 | 75.5 | 2994.33 | 17451 | 63.1 | 0 | 0 | 35.5 | 4441.07 | 25882.7 | 87.19 | 0 | 0 | 35.5 | 4563.89 | 26598.4 | 88.96
0 | 0 | 116.5 | 1777.58 | 10359.7 | 45.32 | 0 | 0 | 76.5 | 2972.54 | 17324 | 62.66 | 0 | 0 | 36.5 | 4409.91 | 25701 | 86.58 | 0 | 0 | 36.5 | 4535.05 | 26430.3 | 88.39
0 | 0 | 117.5 | 1764.82 | 10285.4 | 45.02 | 0 | 0 | 77.5 | 2950.76 | 17197.1 | 62.22 | 0 | 0 | 37.5 | 4378.75 | 25519.4 | 85.96 | 0 | 0 | 37.5 | 4506.2 | 26262.2 | 87.82
0 | 0 | 118.5 | 1752.07 | 10211.1 | 44.71 | 0 | 0 | 78.5 | 2928.98 | 17070.1 | 61.78 | 0 | 0 | 38.5 | 4347.58 | 25337.8 | 85.34 | 0 | 0 | 38.5 | 4477.36 | 26094.1 | 87.26
0 | 0 | 119.5 | 1739.31 | 10136.7 | 44.41 | 0 | 0 | 79.5 | 2907.2 | 16943.2 | 61.34 | 0 | 0 | 39.5 | 4316.42 | 25156.2 | 84.72 | 0 | 0 | 39.5 | 4448.51 | 25926 | 86.69
0 | 0 | 120.5 | 1726.52 | 10062.2 | 44.11 | 0 | 0 | 80.5 | 2885.7 | 16817.9 | 60.9 | 0 | 0 | 40.5 | 4286.46 | 24981.6 | 84.13 | 0 | 0 | 40.5 | 4418.87 | 25753.3 | 86.11
0 | 0 | 121.5 | 1713.69 | 9987.4 | 43.8 | 0 | 0 | 81.5 | 2864.54 | 16694.6 | 60.45 | 0 | 0 | 41.5 | 4257.35 | 24811.9 | 83.56 | 0 | 0 | 41.5 | 4388.69 | 25577.4 | 85.53
0 | 0 | 122.5 | 1700.86 | 9912.7 | 43.49 | 0 | 0 | 82.5 | 2843.38 | 16571.3 | 60 | 0 | 0 | 42.5 | 4228.24 | 24642.2 | 82.98 | 0 | 0 | 42.5 | 4358.51 | 25401.5 | 84.94
0 | 0 | 123.5 | 1688.03 | 9837.9 | 43.18 | 0 | 0 | 83.5 | 2822.21 | 16447.9 | 59.55 | 0 | 0 | 43.5 | 4199.12 | 24472.6 | 82.41 | 0 | 0 | 43.5 | 4328.33 | 25225.6 | 84.36
0 | 0 | 124.5 | 1675.2 | 9763.1 | 42.87 | 0 | 0 | 84.5 | 2801.05 | 16324.6 | 59.1 | 0 | 0 | 44.5 | 4170.01 | 24302.9 | 81.84 | 0 | 0 | 44.5 | 4298.15 | 25049.7 | 83.77
0 | 0 | 125.5 | 1662.41 | 9688.5 | 42.54 | 0 | 0 | 85.5 | 2780.83 | 16206.7 | 58.67 | 0 | 0 | 45.5 | 4139.09 | 24122.7 | 81.23 | 0 | 0 | 45.5 | 4266.83 | 24867.2 | 83.16
0 | 0 | 126.5 | 1649.66 | 9614.2 | 42.2 | 0 | 0 | 86.5 | 2761.8 | 16095.8 | 58.27 | 0 | 0 | 46.5 | 4106.95 | 23935.4 | 80.61 | 0 | 0 | 46.5 | 4234.69 | 24679.9 | 82.53
0 | 0 | 127.5 | 1636.91 | 9539.9 | 41.85 | 0 | 0 | 87.5 | 2742.77 | 15984.9 | 57.87 | 0 | 0 | 47.5 | 4074.8 | 23748 | 79.99 | 0 | 0 | 47.5 | 4202.55 | 24492.5 | 81.9
0 | 0 | 128.5 | 1624.16 | 9465.6 | 41.51 | 0 | 0 | 88.5 | 2723.74 | 15874 | 57.47 | 0 | 0 | 48.5 | 4042.66 | 23560.7 | 79.36 | 0 | 0 | 48.5 | 4170.4 | 24305.2 | 81.27
0 | 0 | 129.5 | 1611.41 | 9391.3 | 41.16 | 0 | 0 | 89.5 | 2704.71 | 15763.1 | 57.07 | 0 | 0 | 49.5 | 4010.51 | 23373.3 | 78.74 | 0 | 0 | 49.5 | 4138.26 | 24117.9 | 80.64
0 | 0 | 130.5 | 1599.23 | 9320.4 | 40.84 | 0 | 0 | 90.5 | 2685.93 | 15653.6 | 56.67 | 0 | 0 | 50.5 | 3981.23 | 23202.7 | 78.16 | 0 | 0 | 50.5 | 4107.8 | 23940.4 | 80.05
0 | 0 | 131.5 | 1587.84 | 9254 | 40.54 | 0 | 0 | 91.5 | 2667.46 | 15546 | 56.29 | 0 | 0 | 51.5 | 3954.03 | 23044.1 | 77.62 | 0 | 0 | 51.5 | 4078.46 | 23769.3 | 79.48
0 | 0 | 132.5 | 1576.45 | 9187.6 | 40.25 | 0 | 0 | 92.5 | 2648.99 | 15438.3 | 55.91 | 0 | 0 | 52.5 | 3926.82 | 22885.6 | 77.09 | 0 | 0 | 52.5 | 4049.12 | 23598.3 | 78.91
0 | 0 | 133.5 | 1565.06 | 9121.2 | 39.95 | 0 | 0 | 93.5 | 2630.52 | 15330.7 | 55.52 | 0 | 0 | 53.5 | 3899.62 | 22727 | 76.55 | 0 | 0 | 53.5 | 4019.77 | 23427.3 | 78.34
0 | 0 | 134.5 | 1553.67 | 9054.8 | 39.66 | 0 | 0 | 94.5 | 2612.05 | 15223 | 55.14 | 0 | 0 | 54.5 | 3872.41 | 22568.5 | 76.01 | 0 | 0 | 54.5 | 3990.43 | 23256.3 | 77.77
0 | 0 | 135.5 | 1542.94 | 8992.3 | 39.38 | 0 | 0 | 95.5 | 2593.01 | 15112.1 | 54.74 | 0 | 0 | 55.5 | 3844.29 | 22404.6 | 75.45 | 0 | 0 | 55.5 | 3963.85 | 23101.4 | 77.25
0 | 0 | 136.5 | 1533.06 | 8934.7 | 39.14 | 0 | 0 | 96.5 | 2573.16 | 14996.4 | 54.31 | 0 | 0 | 56.5 | 3815.36 | 22236 | 74.88 | 0 | 0 | 56.5 | 3939.57 | 22959.9 | 76.78
0 | 0 | 137.5 | 1523.17 | 8877.1 | 38.9 | 0 | 0 | 97.5 | 2553.32 | 14880.8 | 53.89 | 0 | 0 | 57.5 | 3786.43 | 22067.4 | 74.3 | 0 | 0 | 57.5 | 3915.29 | 22818.4 | 76.31
0 | 0 | 138.5 | 1513.29 | 8819.5 | 38.66 | 0 | 0 | 98.5 | 2533.47 | 14765.1 | 53.46 | 0 | 0 | 58.5 | 3757.5 | 21898.8 | 73.73 | 0 | 0 | 58.5 | 3891.01 | 22676.9 | 75.84
0 | 0 | 139.5 | 1503.41 | 8761.9 | 38.42 | 0 | 0 | 99.5 | 2513.63 | 14649.5 | 53.03 | 0 | 0 | 59.5 | 3728.58 | 21730.2 | 73.16 | 0 | 0 | 59.5 | 3866.73 | 22535.4 | 75.37
0 | 0 | 140.5 | 1493.1 | 8701.8 | 38.16 | 0 | 0 | 100.5 | 2494.38 | 14537.3 | 52.62 | 0 | 0 | 60.5 | 3700.82 | 21568.5 | 72.61 | 0 | 0 | 60.5 | 3841.41 | 22387.8 | 74.87
0 | 0 | 141.5 | 1482.21 | 8638.4 | 37.86 | 0 | 0 | 101.5 | 2475.9 | 14429.6 | 52.24 | 0 | 0 | 61.5 | 3674.06 | 21412.5 | 72.09 | 0 | 0 | 61.5 | 3815.17 | 22234.9 | 74.36
0 | 0 | 142.5 | 1471.32 | 8574.9 | 37.56 | 0 | 0 | 102.5 | 2457.43 | 14321.9 | 51.86 | 0 | 0 | 62.5 | 3647.29 | 21256.5 | 71.57 | 0 | 0 | 62.5 | 3788.93 | 22082 | 73.84
0 | 0 | 143.5 | 1460.43 | 8511.4 | 37.26 | 0 | 0 | 103.5 | 2438.95 | 14214.2 | 51.48 | 0 | 0 | 63.5 | 3620.53 | 21100.5 | 71.05 | 0 | 0 | 63.5 | 3762.7 | 21929.1 | 73.33
0 | 0 | 144.5 | 1449.54 | 8447.9 | 36.95 | 0 | 0 | 104.5 | 2420.47 | 14106.6 | 51.1 | 0 | 0 | 64.5 | 3593.76 | 20944.5 | 70.53 | 0 | 0 | 64.5 | 3736.46 | 21776.1 | 72.81
0 | 0 | 145.5 | 1438.8 | 8385.4 | 36.68 | 0 | 0 | 105.5 | 2402.38 | 14001.1 | 50.72 | 0 | 0 | 65.5 | 3567.32 | 20790.4 | 70.02 | 0 | 0 | 65.5 | 3709.4 | 21618.4 | 72.29
0 | 0 | 146.5 | 1428.28 | 8324 | 36.45 | 0 | 0 | 106.5 | 2384.77 | 13898.5 | 50.35 | 0 | 0 | 66.5 | 3541.14 | 20637.9 | 69.5 | 0 | 0 | 66.5 | 3681.58 | 21456.3 | 71.75
0 | 0 | 147.5 | 1417.76 | 8262.7 | 36.21 | 0 | 0 | 107.5 | 2367.17 | 13795.9 | 49.98 | 0 | 0 | 67.5 | 3514.97 | 20485.3 | 68.99 | 0 | 0 | 67.5 | 3653.76 | 21294.2 | 71.21
0 | 0 | 148.5 | 1407.24 | 8201.4 | 35.98 | 0 | 0 | 108.5 | 2349.57 | 13693.4 | 49.61 | 0 | 0 | 68.5 | 3488.8 | 20332.8 | 68.48 | 0 | 0 | 68.5 | 3625.94 | 21132.1 | 70.67
0 | 0 | 149.5 | 1396.72 | 8140.1 | 35.74 | 0 | 0 | 109.5 | 2331.97 | 13590.8 | 49.24 | 0 | 0 | 69.5 | 3462.63 | 20180.3 | 67.97 | 0 | 0 | 69.5 | 3598.12 | 20969.9 | 70.13
0 | 0 | 150.5 | 1387.47 | 8086.2 | 35.51 | 0 | 0 | 110.5 | 2316.13 | 13498.5 | 48.9 | 0 | 0 | 70.5 | 3438.42 | 20039.2 | 67.5 | 0 | 0 | 70.5 | 3571.82 | 20816.6 | 69.62
0 | 0 | 151.5 | 1379.88 | 8042 | 35.29 | 0 | 0 | 111.5 | 2302.45 | 13418.7 | 48.61 | 0 | 0 | 71.5 | 3415.98 | 19908.4 | 67.05 | 0 | 0 | 71.5 | 3546.92 | 20671.5 | 69.13
0 | 0 | 152.5 | 1372.29 | 7997.7 | 35.07 | 0 | 0 | 112.5 | 2288.76 | 13338.9 | 48.31 | 0 | 0 | 72.5 | 3393.54 | 19777.6 | 66.6 | 0 | 0 | 72.5 | 3522.02 | 20526.4 | 68.64
0 | 0 | 153.5 | 1364.69 | 7953.5 | 34.85 | 0 | 0 | 113.5 | 2275.07 | 13259.2 | 48.02 | 0 | 0 | 73.5 | 3371.09 | 19646.8 | 66.15 | 0 | 0 | 73.5 | 3497.12 | 20381.3 | 68.15
0 | 0 | 154.5 | 1357.1 | 7909.2 | 34.63 | 0 | 0 | 114.5 | 2261.39 | 13179.4 | 47.73 | 0 | 0 | 74.5 | 3348.65 | 19516 | 65.7 | 0 | 0 | 74.5 | 3472.22 | 20236.2 | 67.66
0 | 0 | 155.5 | 1347.35 | 7852.4 | 34.36 | 0 | 0 | 115.5 | 2246.41 | 13092.1 | 47.4 | 0 | 0 | 75.5 | 3325.21 | 19379.4 | 65.23 | 0 | 0 | 75.5 | 3447.93 | 20094.6 | 67.18
0 | 0 | 156.5 | 1334.75 | 7778.9 | 34.04 | 0 | 0 | 116.5 | 2229.96 | 12996.2 | 47.04 | 0 | 0 | 76.5 | 3300.58 | 19235.8 | 64.76 | 0 | 0 | 76.5 | 3424.24 | 19956.5 | 66.72
0 | 0 | 157.5 | 1322.14 | 7705.5 | 33.71 | 0 | 0 | 117.5 | 2213.51 | 12900.4 | 46.68 | 0 | 0 | 77.5 | 3275.95 | 19092.3 | 64.28 | 0 | 0 | 77.5 | 3400.54 | 19818.4 | 66.26
0 | 0 | 158.5 | 1309.54 | 7632 | 33.38 | 0 | 0 | 118.5 | 2197.06 | 12804.5 | 46.31 | 0 | 0 | 78.5 | 3251.32 | 18948.8 | 63.8 | 0 | 0 | 78.5 | 3376.85 | 19680.4 | 65.8
0 | 0 | 159.5 | 1296.94 | 7558.6 | 33.05 | 0 | 0 | 119.5 | 2180.61 | 12708.6 | 45.95 | 0 | 0 | 79.5 | 3226.69 | 18805.2 | 63.32 | 0 | 0 | 79.5 | 3353.16 | 19542.3 | 65.33
0 | 0 | 160.5 | 1287.97 | 7506.3 | 32.81 | 0 | 0 | 120.5 | 2165.21 | 12618.9 | 45.63 | 0 | 0 | 80.5 | 3202.32 | 18663.2 | 62.85 | 0 | 0 | 80.5 | 3330.76 | 19411.7 | 64.9
0 | 0 | 161.5 | 1283.05 | 7477.7 | 32.67 | 0 | 0 | 121.5 | 2151.18 | 12537.1 | 45.35 | 0 | 0 | 81.5 | 3178.25 | 18522.9 | 62.38 | 0 | 0 | 81.5 | 3309.75 | 19289.3 | 64.49
0 | 0 | 162.5 | 1278.13 | 7449 | 32.53 | 0 | 0 | 122.5 | 2137.14 | 12455.3 | 45.07 | 0 | 0 | 82.5 | 3154.17 | 18382.6 | 61.91 | 0 | 0 | 82.5 | 3288.74 | 19166.8 | 64.08
0 | 0 | 163.5 | 1273.21 | 7420.3 | 32.39 | 0 | 0 | 123.5 | 2123.1 | 12373.5 | 44.79 | 0 | 0 | 83.5 | 3130.1 | 18242.3 | 61.43 | 0 | 0 | 83.5 | 3267.73 | 19044.4 | 63.68
0 | 0 | 164.5 | 1268.3 | 7391.7 | 32.25 | 0 | 0 | 124.5 | 2109.06 | 12291.7 | 44.51 | 0 | 0 | 84.5 | 3106.02 | 18101.9 | 60.96 | 0 | 0 | 84.5 | 3246.71 | 18921.9 | 63.27
0 | 0 | 165.5 | 1262.08 | 7355.4 | 32.1 | 0 | 0 | 125.5 | 2093.82 | 12202.8 | 44.19 | 0 | 0 | 85.5 | 3082.68 | 17965.9 | 60.5 | 0 | 0 | 85.5 | 3225.07 | 18795.8 | 62.85
0 | 0 | 166.5 | 1254.12 | 7309 | 31.92 | 0 | 0 | 126.5 | 2077 | 12104.8 | 43.84 | 0 | 0 | 86.5 | 3060.22 | 17835 | 60.06 | 0 | 0 | 86.5 | 3202.76 | 18665.7 | 62.4
0 | 0 | 167.5 | 1246.16 | 7262.6 | 31.74 | 0 | 0 | 127.5 | 2060.19 | 12006.8 | 43.49 | 0 | 0 | 87.5 | 3037.76 | 17704.1 | 59.62 | 0 | 0 | 87.5 | 3180.44 | 18535.7 | 61.96
0 | 0 | 168.5 | 1238.2 | 7216.3 | 31.56 | 0 | 0 | 128.5 | 2043.37 | 11908.8 | 43.13 | 0 | 0 | 88.5 | 3015.31 | 17573.3 | 59.18 | 0 | 0 | 88.5 | 3158.13 | 18405.6 | 61.52
0 | 0 | 169.5 | 1230.24 | 7169.9 | 31.38 | 0 | 0 | 129.5 | 2026.56 | 11810.8 | 42.78 | 0 | 0 | 89.5 | 2992.85 | 17442.4 | 58.73 | 0 | 0 | 89.5 | 3135.81 | 18275.6 | 61.08
0 | 0 | 170.5 | 1221.76 | 7120.5 | 31.18 | 0 | 0 | 130.5 | 2011.77 | 11724.7 | 42.46 | 0 | 0 | 90.5 | 2972.27 | 17322.5 | 58.33 | 0 | 0 | 90.5 | 3113.2 | 18143.8 | 60.64
0 | 0 | 171.5 | 1212.65 | 7067.3 | 30.95 | 0 | 0 | 131.5 | 1999.42 | 11652.6 | 42.2 | 0 | 0 | 91.5 | 2954.11 | 17216.6 | 57.97 | 0 | 0 | 91.5 | 3090.16 | 18009.5 | 60.19
0 | 0 | 172.5 | 1203.53 | 7014.2 | 30.72 | 0 | 0 | 132.5 | 1987.06 | 11580.6 | 41.93 | 0 | 0 | 92.5 | 2935.95 | 17110.8 | 57.62 | 0 | 0 | 92.5 | 3067.13 | 17875.3 | 59.75
0 | 0 | 173.5 | 1194.41 | 6961.1 | 30.49 | 0 | 0 | 133.5 | 1974.71 | 11508.6 | 41.67 | 0 | 0 | 93.5 | 2917.79 | 17004.9 | 57.27 | 0 | 0 | 93.5 | 3044.1 | 17741.1 | 59.3
0 | 0 | 174.5 | 1185.29 | 6907.9 | 30.26 | 0 | 0 | 134.5 | 1962.36 | 11436.6 | 41.4 | 0 | 0 | 94.5 | 2899.63 | 16899.1 | 56.91 | 0 | 0 | 94.5 | 3021.07 | 17606.8 | 58.86
0 | 0 | 175.5 | 1175.64 | 6851.7 | 30.02 | 0 | 0 | 135.5 | 1951.04 | 11370.7 | 41.15 | 0 | 0 | 95.5 | 2879.66 | 16782.7 | 56.52 | 0 | 0 | 95.5 | 2999.18 | 17479.3 | 58.43
0 | 0 | 176.5 | 1165.46 | 6792.3 | 29.76 | 0 | 0 | 136.5 | 1940.92 | 11311.7 | 40.93 | 0 | 0 | 96.5 | 2857.4 | 16653 | 56.08 | 0 | 0 | 96.5 | 2978.89 | 17361 | 58.03
0 | 0 | 177.5 | 1155.27 | 6732.9 | 29.5 | 0 | 0 | 137.5 | 1930.8 | 11252.7 | 40.71 | 0 | 0 | 97.5 | 2835.14 | 16523.3 | 55.64 | 0 | 0 | 97.5 | 2958.61 | 17242.8 | 57.64
0 | 0 | 178.5 | 1145.08 | 6673.5 | 29.24 | 0 | 0 | 138.5 | 1920.68 | 11193.7 | 40.49 | 0 | 0 | 98.5 | 2812.89 | 16393.6 | 55.2 | 0 | 0 | 98.5 | 2938.32 | 17124.6 | 57.24
0 | 0 | 179.5 | 1134.89 | 6614.2 | 28.99 | 0 | 0 | 139.5 | 1910.56 | 11134.8 | 40.27 | 0 | 0 | 99.5 | 2790.63 | 16263.9 | 54.76 | 0 | 0 | 99.5 | 2918.04 | 17006.4 | 56.84
0 | 0 | 180.5 | 1126.67 | 6566.3 | 28.78 | 0 | 0 | 140.5 | 1897.3 | 11057.5 | 39.98 | 0 | 0 | 100.5 | 2770.39 | 16145.9 | 54.37 | 0 | 0 | 100.5 | 2898.81 | 16894.3 | 56.47
0 | 0 | 181.5 | 1120.95 | 6532.9 | 28.62 | 0 | 0 | 141.5 | 1879.87 | 10955.9 | 39.61 | 0 | 0 | 101.5 | 2752.67 | 16042.6 | 54.03 | 0 | 0 | 101.5 | 2880.96 | 16790.3 | 56.12
0 | 0 | 182.5 | 1115.23 | 6499.6 | 28.47 | 0 | 0 | 142.5 | 1862.45 | 10854.4 | 39.25 | 0 | 0 | 102.5 | 2734.95 | 15939.4 | 53.69 | 0 | 0 | 102.5 | 2863.11 | 16686.3 | 55.77
0 | 0 | 183.5 | 1109.51 | 6466.2 | 28.32 | 0 | 0 | 143.5 | 1845.02 | 10752.8 | 38.88 | 0 | 0 | 103.5 | 2717.23 | 15836.1 | 53.34 | 0 | 0 | 103.5 | 2845.26 | 16582.2 | 55.42
0 | 0 | 184.5 | 1103.79 | 6432.9 | 28.17 | 0 | 0 | 144.5 | 1827.6 | 10651.3 | 38.51 | 0 | 0 | 104.5 | 2699.51 | 15732.8 | 53 | 0 | 0 | 104.5 | 2827.41 | 16478.2 | 55.08
0 | 0 | 185.5 | 1097.44 | 6395.9 | 28 | 0 | 0 | 145.5 | 1812.97 | 10566 | 38.21 | 0 | 0 | 105.5 | 2681.36 | 15627 | 52.65 | 0 | 0 | 105.5 | 2808.84 | 16370 | 54.72
0 | 0 | 186.5 | 1090.25 | 6354 | 27.8 | 0 | 0 | 146.5 | 1802.05 | 10502.4 | 37.99 | 0 | 0 | 106.5 | 2662.66 | 15518 | 52.28 | 0 | 0 | 106.5 | 2789.32 | 16256.2 | 54.35
0 | 0 | 187.5 | 1083.05 | 6312 | 27.61 | 0 | 0 | 147.5 | 1791.14 | 10438.8 | 37.77 | 0 | 0 | 107.5 | 2643.95 | 15409 | 51.91 | 0 | 0 | 107.5 | 2769.8 | 16142.4 | 53.97
0 | 0 | 188.5 | 1075.86 | 6270.1 | 27.42 | 0 | 0 | 148.5 | 1780.22 | 10375.2 | 37.55 | 0 | 0 | 108.5 | 2625.25 | 15300 | 51.54 | 0 | 0 | 108.5 | 2750.28 | 16028.7 | 53.6
0 | 0 | 189.5 | 1068.66 | 6228.2 | 27.23 | 0 | 0 | 149.5 | 1769.31 | 10311.6 | 37.33 | 0 | 0 | 109.5 | 2606.54 | 15191 | 51.17 | 0 | 0 | 109.5 | 2730.76 | 15914.9 | 53.23
0 | 0 | 190.5 | 1062.1 | 6190 | 27.06 | 0 | 0 | 150.5 | 1757.77 | 10244.3 | 37.09 | 0 | 0 | 110.5 | 2586.67 | 15075.1 | 50.79 | 0 | 0 | 110.5 | 2711.32 | 15801.6 | 52.85
0 | 0 | 191.5 | 1056.34 | 6156.4 | 26.92 | 0 | 0 | 151.5 | 1745.43 | 10172.4 | 36.82 | 0 | 0 | 111.5 | 2565.42 | 14951.3 | 50.37 | 0 | 0 | 111.5 | 2691.99 | 15689 | 52.47
0 | 0 | 192.5 | 1050.59 | 6122.8 | 26.79 | 0 | 0 | 152.5 | 1733.08 | 10100.4 | 36.55 | 0 | 0 | 112.5 | 2544.18 | 14827.5 | 49.96 | 0 | 0 | 112.5 | 2672.67 | 15576.4 | 52.09
0 | 0 | 193.5 | 1044.83 | 6089.3 | 26.65 | 0 | 0 | 153.5 | 1720.74 | 10028.5 | 36.28 | 0 | 0 | 113.5 | 2522.94 | 14703.7 | 49.54 | 0 | 0 | 113.5 | 2653.34 | 15463.7 | 51.71
0 | 0 | 194.5 | 1039.07 | 6055.7 | 26.52 | 0 | 0 | 154.5 | 1708.39 | 9956.5 | 36.01 | 0 | 0 | 114.5 | 2501.69 | 14579.9 | 49.13 | 0 | 0 | 114.5 | 2634.01 | 15351.1 | 51.33
0 | 0 | 195.5 | 1030.69 | 6006.9 | 26.31 | 0 | 0 | 155.5 | 1698.31 | 9897.8 | 35.8 | 0 | 0 | 115.5 | 2482.84 | 14470.1 | 48.76 | 0 | 0 | 115.5 | 2615.09 | 15240.8 | 50.96
0 | 0 | 196.5 | 1018.9 | 5938.2 | 25.99 | 0 | 0 | 156.5 | 1691.22 | 9856.5 | 35.66 | 0 | 0 | 116.5 | 2466.51 | 14374.9 | 48.43 | 0 | 0 | 116.5 | 2596.67 | 15133.5 | 50.6
0 | 0 | 197.5 | 1007.11 | 5869.5 | 25.68 | 0 | 0 | 157.5 | 1684.13 | 9815.2 | 35.52 | 0 | 0 | 117.5 | 2450.18 | 14279.7 | 48.1 | 0 | 0 | 117.5 | 2578.25 | 15026.1 | 50.24
0 | 0 | 198.5 | 995.32 | 5800.8 | 25.37 | 0 | 0 | 158.5 | 1677.04 | 9773.8 | 35.38 | 0 | 0 | 118.5 | 2433.85 | 14184.5 | 47.77 | 0 | 0 | 118.5 | 2559.83 | 14918.8 | 49.88
0 | 0 | 199.5 | 983.54 | 5732.1 | 25.06 | 0 | 0 | 159.5 | 1669.96 | 9732.5 | 35.24 | 0 | 0 | 119.5 | 2417.52 | 14089.3 | 47.44 | 0 | 0 | 119.5 | 2541.41 | 14811.4 | 49.52
0 | 0 | 200.5 | 975.41 | 5684.7 | 24.84 | 0 | 0 | 160.5 | 1661.29 | 9682 | 35.06 | 0 | 0 | 120.5 | 2401.55 | 13996.3 | 47.12 | 0 | 0 | 120.5 | 2523.85 | 14709 | 49.18
0 | 0 | 201.5 | 972.07 | 5665.2 | 24.74 | 0 | 0 | 161.5 | 1650.46 | 9618.9 | 34.82 | 0 | 0 | 121.5 | 2386 | 13905.7 | 46.81 | 0 | 0 | 121.5 | 2507.44 | 14613.4 | 48.85
0 | 0 | 202.5 | 968.73 | 5645.8 | 24.65 | 0 | 0 | 162.5 | 1639.64 | 9555.9 | 34.59 | 0 | 0 | 122.5 | 2370.46 | 13815.1 | 46.51 | 0 | 0 | 122.5 | 2491.03 | 14517.8 | 48.53
0 | 0 | 203.5 | 965.39 | 5626.3 | 24.55 | 0 | 0 | 163.5 | 1628.82 | 9492.8 | 34.35 | 0 | 0 | 123.5 | 2354.92 | 13724.5 | 46.2 | 0 | 0 | 123.5 | 2474.62 | 14422.1 | 48.21
0 | 0 | 204.5 | 962.05 | 5606.8 | 24.46 | 0 | 0 | 164.5 | 1618 | 9429.7 | 34.12 | 0 | 0 | 124.5 | 2339.38 | 13633.9 | 45.9 | 0 | 0 | 124.5 | 2458.21 | 14326.5 | 47.89
0 | 0 | 205.5 | 956.95 | 5577.1 | 24.33 | 0 | 0 | 165.5 | 1605.23 | 9355.3 | 33.85 | 0 | 0 | 125.5 | 2323.22 | 13539.8 | 45.58 | 0 | 0 | 125.5 | 2441.48 | 14229 | 47.56
0 | 0 | 206.5 | 949.62 | 5534.4 | 24.16 | 0 | 0 | 166.5 | 1590.24 | 9267.9 | 33.54 | 0 | 0 | 126.5 | 2306.28 | 13441 | 45.25 | 0 | 0 | 126.5 | 2424.34 | 14129.1 | 47.22
0 | 0 | 207.5 | 942.29 | 5491.7 | 23.99 | 0 | 0 | 167.5 | 1575.25 | 9180.6 | 33.23 | 0 | 0 | 127.5 | 2289.33 | 13342.3 | 44.92 | 0 | 0 | 127.5 | 2407.19 | 14029.2 | 46.89
0 | 0 | 208.5 | 934.96 | 5448.9 | 23.83 | 0 | 0 | 168.5 | 1560.26 | 9093.2 | 32.92 | 0 | 0 | 128.5 | 2272.39 | 13243.5 | 44.59 | 0 | 0 | 128.5 | 2390.05 | 13929.2 | 46.55
0 | 0 | 209.5 | 927.63 | 5406.2 | 23.66 | 0 | 0 | 169.5 | 1545.27 | 9005.9 | 32.61 | 0 | 0 | 129.5 | 2255.45 | 13144.8 | 44.26 | 0 | 0 | 129.5 | 2372.9 | 13829.3 | 46.21
0 | 0 | 210.5 | 921.47 | 5370.3 | 23.52 | 0 | 0 | 170.5 | 1533.68 | 8938.3 | 32.37 | 0 | 0 | 130.5 | 2240.48 | 13057.6 | 43.96 | 0 | 0 | 130.5 | 2356.74 | 13735.1 | 45.9
0 | 0 | 211.5 | 916.81 | 5343.2 | 23.41 | 0 | 0 | 171.5 | 1526.57 | 8896.9 | 32.21 | 0 | 0 | 131.5 | 2227.77 | 12983.5 | 43.71 | 0 | 0 | 131.5 | 2341.86 | 13648.4 | 45.61
0 | 0 | 212.5 | 912.16 | 5316.1 | 23.3 | 0 | 0 | 172.5 | 1519.46 | 8855.4 | 32.05 | 0 | 0 | 132.5 | 2215.06 | 12909.4 | 43.47 | 0 | 0 | 132.5 | 2326.97 | 13561.6 | 45.33
0 | 0 | 213.5 | 907.5 | 5289 | 23.19 | 0 | 0 | 173.5 | 1512.35 | 8814 | 31.9 | 0 | 0 | 133.5 | 2202.35 | 12835.3 | 43.22 | 0 | 0 | 133.5 | 2312.08 | 13474.9 | 45.05
0 | 0 | 214.5 | 902.85 | 5261.8 | 23.08 | 0 | 0 | 174.5 | 1505.24 | 8772.5 | 31.74 | 0 | 0 | 134.5 | 2189.64 | 12761.2 | 42.97 | 0 | 0 | 134.5 | 2297.19 | 13388.1 | 44.76
0 | 0 | 215.5 | 898.38 | 5235.8 | 22.96 | 0 | 0 | 175.5 | 1496.04 | 8718.9 | 31.54 | 0 | 0 | 135.5 | 2175.82 | 12680.7 | 42.7 | 0 | 0 | 135.5 | 2282.23 | 13300.9 | 44.47
0 | 0 | 216.5 | 894.15 | 5211.1 | 22.83 | 0 | 0 | 176.5 | 1483.92 | 8648.3 | 31.29 | 0 | 0 | 136.5 | 2160.54 | 12591.7 | 42.41 | 0 | 0 | 136.5 | 2267.17 | 13213.1 | 44.17
0 | 0 | 217.5 | 889.93 | 5186.5 | 22.7 | 0 | 0 | 177.5 | 1471.8 | 8577.7 | 31.03 | 0 | 0 | 137.5 | 2145.27 | 12502.7 | 42.11 | 0 | 0 | 137.5 | 2252.11 | 13125.3 | 43.88
0 | 0 | 218.5 | 885.71 | 5161.9 | 22.57 | 0 | 0 | 178.5 | 1459.68 | 8507 | 30.78 | 0 | 0 | 138.5 | 2129.99 | 12413.6 | 41.82 | 0 | 0 | 138.5 | 2237.04 | 13037.5 | 43.58
0 | 0 | 219.5 | 881.48 | 5137.3 | 22.44 | 0 | 0 | 179.5 | 1447.56 | 8436.4 | 30.52 | 0 | 0 | 139.5 | 2114.71 | 12324.6 | 41.53 | 0 | 0 | 139.5 | 2221.98 | 12949.7 | 43.28
0 | 0 | 220.5 | 877.17 | 5112.1 | 22.32 | 0 | 0 | 180.5 | 1437.12 | 8375.6 | 30.3 | 0 | 0 | 140.5 | 2099.23 | 12234.3 | 41.22 | 0 | 0 | 140.5 | 2207.22 | 12863.7 | 42.99
0 | 0 | 221.5 | 872.73 | 5086.3 | 22.21 | 0 | 0 | 181.5 | 1428.83 | 8327.3 | 30.12 | 0 | 0 | 141.5 | 2083.46 | 12142.4 | 40.91 | 0 | 0 | 141.5 | 2192.88 | 12780.1 | 42.72
0 | 0 | 222.5 | 868.29 | 5060.4 | 22.1 | 0 | 0 | 182.5 | 1420.54 | 8278.9 | 29.95 | 0 | 0 | 142.5 | 2067.69 | 12050.5 | 40.6 | 0 | 0 | 142.5 | 2178.54 | 12696.5 | 42.44
0 | 0 | 223.5 | 863.85 | 5034.5 | 22 | 0 | 0 | 183.5 | 1412.25 | 8230.6 | 29.77 | 0 | 0 | 143.5 | 2051.92 | 11958.6 | 40.28 | 0 | 0 | 143.5 | 2164.19 | 12613 | 42.16
0 | 0 | 224.5 | 859.41 | 5008.7 | 21.89 | 0 | 0 | 184.5 | 1403.96 | 8182.3 | 29.59 | 0 | 0 | 144.5 | 2036.15 | 11866.7 | 39.97 | 0 | 0 | 144.5 | 2149.85 | 12529.4 | 41.88
0 | 0 | 225.5 | 854.06 | 4977.5 | 21.77 | 0 | 0 | 185.5 | 1395.1 | 8130.6 | 29.4 | 0 | 0 | 145.5 | 2020.76 | 11777.1 | 39.67 | 0 | 0 | 145.5 | 2135.63 | 12446.5 | 41.61
0 | 0 | 226.5 | 847.64 | 4940 | 21.62 | 0 | 0 | 186.5 | 1385.52 | 8074.8 | 29.2 | 0 | 0 | 146.5 | 2005.88 | 11690.3 | 39.37 | 0 | 0 | 146.5 | 2121.56 | 12364.5 | 41.33
0 | 0 | 227.5 | 841.21 | 4902.6 | 21.48 | 0 | 0 | 187.5 | 1375.94 | 8019 | 28.99 | 0 | 0 | 147.5 | 1990.99 | 11603.5 | 39.07 | 0 | 0 | 147.5 | 2107.5 | 12282.6 | 41.06
0 | 0 | 228.5 | 834.78 | 4865.1 | 21.34 | 0 | 0 | 188.5 | 1366.36 | 7963.2 | 28.79 | 0 | 0 | 148.5 | 1976.1 | 11516.7 | 38.78 | 0 | 0 | 148.5 | 2093.44 | 12200.6 | 40.78
0 | 0 | 229.5 | 828.36 | 4827.7 | 21.2 | 0 | 0 | 189.5 | 1356.78 | 7907.3 | 28.58 | 0 | 0 | 149.5 | 1961.21 | 11430 | 38.48 | 0 | 0 | 149.5 | 2079.38 | 12118.7 | 40.51
0 | 0 | 230.5 | 823.13 | 4797.2 | 21.06 | 0 | 0 | 190.5 | 1348.44 | 7858.8 | 28.4 | 0 | 0 | 150.5 | 1947.4 | 11349.5 | 38.21 | 0 | 0 | 150.5 | 2066.52 | 12043.7 | 40.26
0 | 0 | 231.5 | 819.33 | 4775.1 | 20.94 | 0 | 0 | 191.5 | 1341.77 | 7819.9 | 28.27 | 0 | 0 | 151.5 | 1935 | 11277.2 | 37.96 | 0 | 0 | 151.5 | 2055.32 | 11978.4 | 40.04
0 | 0 | 232.5 | 815.53 | 4752.9 | 20.82 | 0 | 0 | 192.5 | 1335.1 | 7781 | 28.13 | 0 | 0 | 152.5 | 1922.59 | 11204.9 | 37.72 | 0 | 0 | 152.5 | 2044.11 | 11913.1 | 39.83
0 | 0 | 233.5 | 811.72 | 4730.7 | 20.7 | 0 | 0 | 193.5 | 1328.43 | 7742.1 | 27.99 | 0 | 0 | 153.5 | 1910.19 | 11132.6 | 37.48 | 0 | 0 | 153.5 | 2032.9 | 11847.8 | 39.61
0 | 0 | 234.5 | 807.92 | 4708.6 | 20.59 | 0 | 0 | 194.5 | 1321.77 | 7703.3 | 27.86 | 0 | 0 | 154.5 | 1897.78 | 11060.3 | 37.23 | 0 | 0 | 154.5 | 2021.7 | 11782.5 | 39.39
0 | 0 | 235.5 | 804.23 | 4687.1 | 20.48 | 0 | 0 | 195.5 | 1313.58 | 7655.6 | 27.69 | 0 | 0 | 155.5 | 1885.45 | 10988.4 | 36.99 | 0 | 0 | 155.5 | 2008.6 | 11706.1 | 39.14
0 | 0 | 236.5 | 800.71 | 4666.5 | 20.39 | 0 | 0 | 196.5 | 1303.46 | 7596.6 | 27.48 | 0 | 0 | 156.5 | 1873.21 | 10917.1 | 36.76 | 0 | 0 | 156.5 | 1992.76 | 11613.9 | 38.83
0 | 0 | 237.5 | 797.18 | 4646 | 20.3 | 0 | 0 | 197.5 | 1293.35 | 7537.7 | 27.27 | 0 | 0 | 157.5 | 1860.97 | 10845.8 | 36.52 | 0 | 0 | 157.5 | 1976.92 | 11521.6 | 38.52
0 | 0 | 238.5 | 793.65 | 4625.4 | 20.21 | 0 | 0 | 198.5 | 1283.23 | 7478.7 | 27.06 | 0 | 0 | 158.5 | 1848.73 | 10774.4 | 36.29 | 0 | 0 | 158.5 | 1961.09 | 11429.3 | 38.21
0 | 0 | 239.5 | 790.12 | 4604.8 | 20.12 | 0 | 0 | 199.5 | 1273.12 | 7419.8 | 26.86 | 0 | 0 | 159.5 | 1836.49 | 10703.1 | 36.05 | 0 | 0 | 159.5 | 1945.25 | 11337 | 37.89
0 | 0 | 240.5 | 784.7 | 4573.2 | 19.97 | 0 | 0 | 200.5 | 1263.51 | 7363.8 | 26.65 | 0 | 0 | 160.5 | 1825.71 | 10640.3 | 35.84 | 0 | 0 | 160.5 | 1931.63 | 11257.6 | 37.63
0 | 0 | 241.5 | 776.99 | 4528.3 | 19.77 | 0 | 0 | 201.5 | 1254.55 | 7311.5 | 26.46 | 0 | 0 | 161.5 | 1816.83 | 10588.5 | 35.67 | 0 | 0 | 161.5 | 1920.9 | 11195 | 37.42
0 | 0 | 242.5 | 769.28 | 4483.4 | 19.56 | 0 | 0 | 202.5 | 1245.58 | 7259.3 | 26.26 | 0 | 0 | 162.5 | 1807.95 | 10536.8 | 35.49 | 0 | 0 | 162.5 | 1910.17 | 11132.5 | 37.21
0 | 0 | 243.5 | 761.57 | 4438.5 | 19.36 | 0 | 0 | 203.5 | 1236.62 | 7207 | 26.06 | 0 | 0 | 163.5 | 1799.07 | 10485 | 35.32 | 0 | 0 | 163.5 | 1899.44 | 11070 | 36.99
0 | 0 | 244.5 | 753.86 | 4393.5 | 19.15 | 0 | 0 | 204.5 | 1227.65 | 7154.8 | 25.87 | 0 | 0 | 164.5 | 1790.19 | 10433.3 | 35.14 | 0 | 0 | 164.5 | 1888.71 | 11007.4 | 36.78
0 | 0 | 245.5 | 747.89 | 4358.7 | 18.99 | 0 | 0 | 205.5 | 1218.52 | 7101.6 | 25.68 | 0 | 0 | 165.5 | 1779.79 | 10372.6 | 34.94 | 0 | 0 | 165.5 | 1877.6 | 10942.7 | 36.57
0 | 0 | 246.5 | 744.12 | 4336.8 | 18.89 | 0 | 0 | 206.5 | 1209.19 | 7047.2 | 25.49 | 0 | 0 | 166.5 | 1767.43 | 10300.6 | 34.69 | 0 | 0 | 166.5 | 1865.94 | 10874.8 | 36.34
0 | 0 | 247.5 | 740.36 | 4314.8 | 18.8 | 0 | 0 | 207.5 | 1199.87 | 6992.8 | 25.3 | 0 | 0 | 167.5 | 1755.07 | 10228.6 | 34.45 | 0 | 0 | 167.5 | 1854.29 | 10806.8 | 36.11
0 | 0 | 248.5 | 736.6 | 4292.9 | 18.7 | 0 | 0 | 208.5 | 1190.54 | 6938.5 | 25.11 | 0 | 0 | 168.5 | 1742.7 | 10156.5 | 34.2 | 0 | 0 | 168.5 | 1842.63 | 10738.9 | 35.88
0 | 0 | 249.5 | 732.84 | 4271 | 18.6 | 0 | 0 | 209.5 | 1181.21 | 6884.1 | 24.92 | 0 | 0 | 169.5 | 1730.34 | 10084.5 | 33.96 | 0 | 0 | 169.5 | 1830.97 | 10670.9 | 35.65
[6 empty] | 0 | 0 | 210.5 | 1173.66 | 6840.1 | 24.76 | 0 | 0 | 170.5 | 1718.56 | 10015.8 | 33.73 | 0 | 0 | 170.5 | 1819 | 10601.2 | 35.41
[6 empty] | 0 | 0 | 211.5 | 1168.43 | 6809.6 | 24.64 | 0 | 0 | 171.5 | 1707.49 | 9951.3 | 33.51 | 0 | 0 | 171.5 | 1806.62 | 10529 | 35.18
[6 empty] | 0 | 0 | 212.5 | 1163.21 | 6779.2 | 24.52 | 0 | 0 | 172.5 | 1696.42 | 9886.8 | 33.3 | 0 | 0 | 172.5 | 1794.25 | 10456.9 | 34.94
[6 empty] | 0 | 0 | 213.5 | 1157.98 | 6748.7 | 24.4 | 0 | 0 | 173.5 | 1685.35 | 9822.3 | 33.08 | 0 | 0 | 173.5 | 1781.87 | 10384.8 | 34.71
[6 empty] | 0 | 0 | 214.5 | 1152.75 | 6718.3 | 24.29 | 0 | 0 | 174.5 | 1674.28 | 9757.8 | 32.86 | 0 | 0 | 174.5 | 1769.49 | 10312.6 | 34.47
[6 empty] | 0 | 0 | 215.5 | 1147.66 | 6688.6 | 24.18 | 0 | 0 | 175.5 | 1663.32 | 9693.9 | 32.65 | 0 | 0 | 175.5 | 1758 | 10245.7 | 34.25
[6 empty] | 0 | 0 | 216.5 | 1142.72 | 6659.8 | 24.08 | 0 | 0 | 176.5 | 1652.49 | 9630.7 | 32.43 | 0 | 0 | 176.5 | 1747.71 | 10185.7 | 34.05
[6 empty] | 0 | 0 | 217.5 | 1137.78 | 6631 | 23.98 | 0 | 0 | 177.5 | 1641.65 | 9567.6 | 32.21 | 0 | 0 | 177.5 | 1737.41 | 10125.7 | 33.84
[6 empty] | 0 | 0 | 218.5 | 1132.84 | 6602.2 | 23.88 | 0 | 0 | 178.5 | 1630.82 | 9504.5 | 31.99 | 0 | 0 | 178.5 | 1727.11 | 10065.7 | 33.64
[6 empty] | 0 | 0 | 219.5 | 1127.9 | 6573.4 | 23.78 | 0 | 0 | 179.5 | 1619.99 | 9441.3 | 31.77 | 0 | 0 | 179.5 | 1716.82 | 10005.6 | 33.43
[6 empty] | 0 | 0 | 220.5 | 1122.16 | 6540 | 23.66 | 0 | 0 | 180.5 | 1609.59 | 9380.7 | 31.57 | 0 | 0 | 180.5 | 1705.69 | 9940.8 | 33.21
[6 empty] | 0 | 0 | 221.5 | 1115.41 | 6500.6 | 23.52 | 0 | 0 | 181.5 | 1599.71 | 9323.2 | 31.38 | 0 | 0 | 181.5 | 1693.64 | 9870.6 | 32.98
[6 empty] | 0 | 0 | 222.5 | 1108.67 | 6461.3 | 23.38 | 0 | 0 | 182.5 | 1589.84 | 9265.6 | 31.19 | 0 | 0 | 182.5 | 1681.6 | 9800.4 | 32.75
[6 empty] | 0 | 0 | 223.5 | 1101.92 | 6422 | 23.24 | 0 | 0 | 183.5 | 1579.97 | 9208.1 | 31 | 0 | 0 | 183.5 | 1669.55 | 9730.2 | 32.51
[6 empty] | 0 | 0 | 224.5 | 1095.17 | 6382.7 | 23.1 | 0 | 0 | 184.5 | 1570.1 | 9150.6 | 30.81 | 0 | 0 | 184.5 | 1657.5 | 9660 | 32.28
[6 empty] | 0 | 0 | 225.5 | 1087.96 | 6340.7 | 22.95 | 0 | 0 | 185.5 | 1559.26 | 9087.4 | 30.6 | 0 | 0 | 185.5 | 1645.8 | 9591.7 | 32.05
[6 empty] | 0 | 0 | 226.5 | 1080.17 | 6295.3 | 22.79 | 0 | 0 | 186.5 | 1547.16 | 9016.9 | 30.37 | 0 | 0 | 186.5 | 1634.49 | 9525.8 | 31.83
[6 empty] | 0 | 0 | 227.5 | 1072.38 | 6249.9 | 22.63 | 0 | 0 | 187.5 | 1535.07 | 8946.4 | 30.13 | 0 | 0 | 187.5 | 1623.18 | 9459.9 | 31.61
[6 empty] | 0 | 0 | 228.5 | 1064.59 | 6204.5 | 22.47 | 0 | 0 | 188.5 | 1522.98 | 8876 | 29.89 | 0 | 0 | 188.5 | 1611.86 | 9394 | 31.39
[6 empty] | 0 | 0 | 229.5 | 1056.8 | 6159.1 | 22.31 | 0 | 0 | 189.5 | 1510.89 | 8805.5 | 29.66 | 0 | 0 | 189.5 | 1600.55 | 9328 | 31.17
[6 empty] | 0 | 0 | 230.5 | 1048.68 | 6111.8 | 22.14 | 0 | 0 | 190.5 | 1499.55 | 8739.4 | 29.43 | 0 | 0 | 190.5 | 1590.16 | 9267.5 | 30.97
[6 empty] | 0 | 0 | 231.5 | 1040.18 | 6062.2 | 21.96 | 0 | 0 | 191.5 | 1489.11 | 8678.6 | 29.22 | 0 | 0 | 191.5 | 1580.9 | 9213.5 | 30.79
[6 empty] | 0 | 0 | 232.5 | 1031.68 | 6012.6 | 21.79 | 0 | 0 | 192.5 | 1478.66 | 8617.7 | 29.01 | 0 | 0 | 192.5 | 1571.64 | 9159.6 | 30.61
[6 empty] | 0 | 0 | 233.5 | 1023.17 | 5963.1 | 21.61 | 0 | 0 | 193.5 | 1468.22 | 8556.8 | 28.8 | 0 | 0 | 193.5 | 1562.39 | 9105.6 | 30.43
[6 empty] | 0 | 0 | 234.5 | 1014.67 | 5913.5 | 21.43 | 0 | 0 | 194.5 | 1457.78 | 8495.9 | 28.59 | 0 | 0 | 194.5 | 1553.13 | 9051.7 | 30.25
[6 empty] | 0 | 0 | 235.5 | 1008.07 | 5875 | 21.29 | 0 | 0 | 195.5 | 1447.86 | 8438.2 | 28.39 | 0 | 0 | 195.5 | 1543.77 | 8997.1 | 30.06
[6 empty] | 0 | 0 | 236.5 | 1003.93 | 5850.9 | 21.19 | 0 | 0 | 196.5 | 1438.63 | 8384.4 | 28.21 | 0 | 0 | 196.5 | 1534.28 | 8941.8 | 29.88
[6 empty] | 0 | 0 | 237.5 | 999.8 | 5826.8 | 21.09 | 0 | 0 | 197.5 | 1429.41 | 8330.6 | 28.04 | 0 | 0 | 197.5 | 1524.79 | 8886.5 | 29.69
[6 empty] | 0 | 0 | 238.5 | 995.66 | 5802.7 | 21 | 0 | 0 | 198.5 | 1420.18 | 8276.8 | 27.86 | 0 | 0 | 198.5 | 1515.3 | 8831.2 | 29.5
[6 empty] | 0 | 0 | 239.5 | 991.53 | 5778.6 | 20.9 | 0 | 0 | 199.5 | 1410.95 | 8223.1 | 27.68 | 0 | 0 | 199.5 | 1505.8 | 8775.9 | 29.31
[6 empty] | 0 | 0 | 240.5 | 985.89 | 5745.8 | 20.78 | 0 | 0 | 200.5 | 1402.05 | 8171.2 | 27.5 | 0 | 0 | 200.5 | 1495.84 | 8717.8 | 29.12
[6 empty] | 0 | 0 | 241.5 | 978.36 | 5701.9 | 20.62 | 0 | 0 | 201.5 | 1393.54 | 8121.6 | 27.33 | 0 | 0 | 201.5 | 1485.25 | 8656.1 | 28.91
[6 empty] | 0 | 0 | 242.5 | 970.83 | 5658 | 20.47 | 0 | 0 | 202.5 | 1385.03 | 8072 | 27.16 | 0 | 0 | 202.5 | 1474.67 | 8594.4 | 28.71
[6 empty] | 0 | 0 | 243.5 | 963.3 | 5614.1 | 20.32 | 0 | 0 | 203.5 | 1376.53 | 8022.4 | 26.99 | 0 | 0 | 203.5 | 1464.09 | 8532.7 | 28.5
[6 empty] | 0 | 0 | 244.5 | 955.77 | 5570.3 | 20.17 | 0 | 0 | 204.5 | 1368.02 | 7972.9 | 26.82 | 0 | 0 | 204.5 | 1453.51 | 8471.1 | 28.29
[6 empty] | 0 | 0 | 245.5 | 949.28 | 5532.4 | 20.04 | 0 | 0 | 205.5 | 1360.25 | 7927.6 | 26.67 | 0 | 0 | 205.5 | 1442.87 | 8409.1 | 28.09
[6 empty] | 0 | 0 | 246.5 | 944.02 | 5501.8 | 19.92 | 0 | 0 | 206.5 | 1353.38 | 7887.5 | 26.54 | 0 | 0 | 206.5 | 1432.17 | 8346.7 | 27.88
[6 empty] | 0 | 0 | 247.5 | 938.76 | 5471.1 | 19.81 | 0 | 0 | 207.5 | 1346.5 | 7847.4 | 26.42 | 0 | 0 | 207.5 | 1421.47 | 8284.4 | 27.68
[6 empty] | 0 | 0 | 248.5 | 933.51 | 5440.5 | 19.7 | 0 | 0 | 208.5 | 1339.62 | 7807.3 | 26.29 | 0 | 0 | 208.5 | 1410.77 | 8222 | 27.47
[6 empty] | 0 | 0 | 249.5 | 928.25 | 5409.9 | 19.59 | 0 | 0 | 209.5 | 1332.75 | 7767.3 | 26.16 | 0 | 0 | 209.5 | 1400.07 | 8159.6 | 27.26
[12 empty] | 0 | 0 | 210.5 | 1325.29 | 7723.8 | 26.02 | 0 | 0 | 210.5 | 1390.85 | 8105.9 | 27.09
[12 empty] | 0 | 0 | 211.5 | 1317.1 | 7676.1 | 25.85 | 0 | 0 | 211.5 | 1383.52 | 8063.2 | 26.94
[12 empty] | 0 | 0 | 212.5 | 1308.9 | 7628.3 | 25.69 | 0 | 0 | 212.5 | 1376.18 | 8020.4 | 26.8
[12 empty] | 0 | 0 | 213.5 | 1300.71 | 7580.6 | 25.53 | 0 | 0 | 213.5 | 1368.84 | 7977.6 | 26.65
[12 empty] | 0 | 0 | 214.5 | 1292.52 | 7532.8 | 25.37 | 0 | 0 | 214.5 | 1361.5 | 7934.9 | 26.51
[12 empty] | 0 | 0 | 215.5 | 1284.23 | 7484.5 | 25.2 | 0 | 0 | 215.5 | 1353.61 | 7888.9 | 26.35
[12 empty] | 0 | 0 | 216.5 | 1275.84 | 7435.6 | 25.04 | 0 | 0 | 216.5 | 1345.05 | 7839 | 26.19
[12 empty] | 0 | 0 | 217.5 | 1267.44 | 7386.7 | 24.88 | 0 | 0 | 217.5 | 1336.5 | 7789.1 | 26.02
[12 empty] | 0 | 0 | 218.5 | 1259.04 | 7337.7 | 24.71 | 0 | 0 | 218.5 | 1327.94 | 7739.3 | 25.85
[12 empty] | 0 | 0 | 219.5 | 1250.64 | 7288.8 | 24.55 | 0 | 0 | 219.5 | 1319.39 | 7689.4 | 25.68
[12 empty] | 0 | 0 | 220.5 | 1242.49 | 7241.3 | 24.39 | 0 | 0 | 220.5 | 1311.14 | 7641.4 | 25.52
[12 empty] | 0 | 0 | 221.5 | 1234.66 | 7195.6 | 24.23 | 0 | 0 | 221.5 | 1303.3 | 7595.7 | 25.37
[12 empty] | 0 | 0 | 222.5 | 1226.83 | 7150 | 24.08 | 0 | 0 | 222.5 | 1295.46 | 7550 | 25.21
[12 empty] | 0 | 0 | 223.5 | 1219 | 7104.4 | 23.92 | 0 | 0 | 223.5 | 1287.63 | 7504.3 | 25.06
[12 empty] | 0 | 0 | 224.5 | 1211.17 | 7058.7 | 23.77 | 0 | 0 | 224.5 | 1279.79 | 7458.6 | 24.9
[12 empty] | 0 | 0 | 225.5 | 1202.61 | 7008.9 | 23.6 | 0 | 0 | 225.5 | 1272.29 | 7414.9 | 24.76
[12 empty] | 0 | 0 | 226.5 | 1193.12 | 6953.5 | 23.41 | 0 | 0 | 226.5 | 1265.21 | 7373.7 | 24.62
[12 empty] | 0 | 0 | 227.5 | 1183.63 | 6898.2 | 23.22 | 0 | 0 | 227.5 | 1258.14 | 7332.4 | 24.49
[12 empty] | 0 | 0 | 228.5 | 1174.13 | 6842.9 | 23.03 | 0 | 0 | 228.5 | 1251.06 | 7291.2 | 24.35
[12 empty] | 0 | 0 | 229.5 | 1164.64 | 6787.5 | 22.84 | 0 | 0 | 229.5 | 1243.98 | 7249.9 | 24.22
[12 empty] | 0 | 0 | 230.5 | 1156.4 | 6739.5 | 22.68 | 0 | 0 | 230.5 | 1237.3 | 7211 | 24.09
[12 empty] | 0 | 0 | 231.5 | 1149.72 | 6700.6 | 22.55 | 0 | 0 | 231.5 | 1231.09 | 7174.8 | 23.96
[12 empty] | 0 | 0 | 232.5 | 1143.04 | 6661.6 | 22.42 | 0 | 0 | 232.5 | 1224.88 | 7138.6 | 23.84
[12 empty] | 0 | 0 | 233.5 | 1136.35 | 6622.7 | 22.3 | 0 | 0 | 233.5 | 1218.67 | 7102.4 | 23.72
[12 empty] | 0 | 0 | 234.5 | 1129.67 | 6583.7 | 22.17 | 0 | 0 | 234.5 | 1212.46 | 7066.2 | 23.59
[12 empty] | 0 | 0 | 235.5 | 1123.52 | 6547.9 | 22.05 | 0 | 0 | 235.5 | 1205.62 | 7026.4 | 23.46
[12 empty] | 0 | 0 | 236.5 | 1118.03 | 6515.9 | 21.94 | 0 | 0 | 236.5 | 1197.96 | 6981.7 | 23.32
[12 empty] | 0 | 0 | 237.5 | 1112.54 | 6483.9 | 21.83 | 0 | 0 | 237.5 | 1190.29 | 6937 | 23.17
[12 empty] | 0 | 0 | 238.5 | 1107.05 | 6451.9 | 21.73 | 0 | 0 | 238.5 | 1182.63 | 6892.4 | 23.03
[12 empty] | 0 | 0 | 239.5 | 1101.56 | 6419.9 | 21.62 | 0 | 0 | 239.5 | 1174.96 | 6847.7 | 22.88
[12 empty] | 0 | 0 | 240.5 | 1095.15 | 6382.6 | 21.49 | 0 | 0 | 240.5 | 1166.87 | 6800.5 | 22.73
[12 empty] | 0 | 0 | 241.5 | 1087.64 | 6338.8 | 21.34 | 0 | 0 | 241.5 | 1158.27 | 6750.4 | 22.56
[12 empty] | 0 | 0 | 242.5 | 1080.13 | 6295 | 21.2 | 0 | 0 | 242.5 | 1149.67 | 6700.3 | 22.39
[12 empty] | 0 | 0 | 243.5 | 1072.61 | 6251.2 | 21.05 | 0 | 0 | 243.5 | 1141.07 | 6650.2 | 22.22
[12 empty] | 0 | 0 | 244.5 | 1065.1 | 6207.4 | 20.9 | 0 | 0 | 244.5 | 1132.47 | 6600 | 22.05
[12 empty] | 0 | 0 | 245.5 | 1058.61 | 6169.6 | 20.77 | 0 | 0 | 245.5 | 1125.56 | 6559.8 | 21.91
[12 empty] | 0 | 0 | 246.5 | 1053.45 | 6139.5 | 20.67 | 0 | 0 | 246.5 | 1120.74 | 6531.7 | 21.82
[12 empty] | 0 | 0 | 247.5 | 1048.29 | 6109.5 | 20.57 | 0 | 0 | 247.5 | 1115.93 | 6503.7 | 21.72
[12 empty] | 0 | 0 | 248.5 | 1043.14 | 6079.4 | 20.47 | 0 | 0 | 248.5 | 1111.11 | 6475.6 | 21.63
[12 empty] | 0 | 0 | 249.5 | 1037.98 | 6049.4 | 20.37 | 0 | 0 | 249.5 | 1106.29 | 6447.5 | 21.53

## Valeurs relevées
[9 empty] | INLINE |  |  |  |  | CROSSLINE
Détecteur | Collimation | Taille de champ (cm x cm) | Energie (MV) | Pas de mesure (cm) | R100 (cm) | R50 (cm) | D10 (%) | D20 (%) | Symétrie (%) | Taille de champ (cm) | Pénombre G (cm) | Pénombre D (cm) | Centre (cm) | Symétrie (%)2 | Taille de champ (cm)3 | Pénombre G (cm)4 | Pénombre D (cm)5 | Centre (cm)6
µD | MLC | 0.5 | 6 FFF | 0.05 | 1 | 10.03 | 50.12 | 24.63 | 107.67 | 0.48 | 0.2 | 0.19 | -0.01 | 107.19 | 0.58 | 0.22 | 0.22 | 0.02
NanoRazor | MLC | 0.5 | 6 FFF | 0.05 | 1 | 10.13 | 50.39 | 24.86 | 107.84 | 0.5 | 0.24 | 0.22 | 0.01 | 102.29 | 0.59 | 0.25 | 0.23 | 0
SFD | MLC | 0.5 | 6 FFF | 0.05 | 0.91 | 9.59 | 48.41 | 22.99 | 105.69 | 0.46 | 0.17 | 0.15 | -0.01 | 103.43 | 0.58 | 0.19 | 0.18 | 0
µD | MLC | 1 | 6 FFF | 0.05 | 1 | 10.64 | 52.29 | 26.05 | 105.39 | 0.95 | 0.22 | 0.22 | -0.02 | 100.65 | 1.06 | 0.24 | 0.24 | 0.02
SFD | MLC | 1 | 6 FFF | 0.05 | 1.01 | 10.29 | 51.06 | 24.8 | 105.54 | 0.95 | 0.19 | 0.19 | -0.02 | 101.97 | 1.06 | 0.21 | 0.22 | 0
SFD | MLC | 1 | 6 | 0.05 | 1.2 | 11.34 | 55.08 | 27.99 | 0 | 0.95 | 0.2 | 0.2 | -0.04 | 0 | 1.07 | 0.23 | 0.22 | 0.01
SFD | CONE | 1 | 6 FFF | 0.05 | 1.01 | 10.2 | 50.75 | 24.7 | 102.19 | 0.97 | 0.15 | 0.14 | -0.02 | 100.53 | 0.97 | 0.15 | 0.14 | 0
NanoRazor | MLC | 1 | 6 FFF | 0.05 | 1.1 | 10.73 | 52.74 | 26.72 | 100.62 | 0.97 | 0.25 | 0.26 | 0.01 | 102.37 | 1.07 | 0.27 | 0.28 | 0
SFD | MLC | 1 | 6 FFF | 0.1 | X | X | X | X | 106.26 | 0.95 | 0.19 | 0.22 | -0.02 | 102.86 | 1.07 | 0.23 | 0.23 | 0
µD | MLC | 2 | 6 FFF | 0.05 | 1.3 | 11.13 | 54.2 | 27.05 | 101 | 1.93 | 0.26 | 0.26 | -0.03 | 100.5 | 2.07 | 0.28 | 0.28 | 0.02
NanoRazor | MLC | 2 | 6 FFF | 0.05 | 1.3 | 11.24 | 54.52 | 27.58 | 101.99 | 1.95 | 0.3 | 0.3 | 0 | 100.43 | 2.07 | 0.32 | 0.32 | 0
SFD | MLC | 2 | 6 FFF | 0.05 | 1.31 | 10.82 | 53.2 | 25.96 | 101.77 | 1.94 | 0.22 | 0.24 | -0.04 | 100.6 | 2.06 | 0.25 | 0.26 | -0.01
µD | MLC | 3 | 6 FFF | 0.05 | 1.41 | 11.5 | 55.52 | 27.87 | 100.72 | 2.94 | 0.28 | 0.28 | -0.03 | 100.45 | 3.06 | 0.3 | 0.3 | 0.02
NanoRazor | MLC | 3 | 6 FFF | 0.05 | 1.5 | 11.81 | 56.63 | 29.21 | 101 | 2.94 | 0.33 | 0.32 | -0.01 | 100.47 | 3.08 | 0.38 | 0.39 | -0.01
SFD | MLC | 3 | 6 FFF | 0.05 | 1.31 | 11.25 | 54.71 | 26.94 | 101.18 | 2.94 | 0.25 | 0.26 | -0.05 | 100.41 | 3.07 | 0.28 | 0.28 | -0.01