## SFD
0,5 cm² | [5 empty] | 1 cm² | [5 empty] | 2cm x 2cm
Measurement time: | 12/04/2023 18:06:08 |  |  |  |  | Measurement time: | 12/04/2023 18:28:42 |  |  |  |  | Measurement time: | 12/04/2023 18:56:46 |  |  |  |  | Measurement time: | 12/04/2023 19:15:44
Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele
Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 5 x 5 mm |  |  |  |  | Field size: | 10 x 10 mm |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
5 | 0 | 100 | 97.71 | 34144.8 | 3.97 | -10 | 0 | 100 | 2.24 | 23220.6 | 2.24 | 20 | 0 | 100 | 2.13 | 22053.6 | 1.94 | 30 | 0 | 100 | 2.53 | 29960.2 | 2.54
4.5 | 0 | 100 | 140.14 | 48973.5 | 5.7 | -9.5 | 0 | 100 | 2.81 | 29134.9 | 2.81 | 18.9 | 0 | 100 | 2.36 | 24456.3 | 2.15 | 29.1 | 0 | 100 | 2.69 | 31861.5 | 2.7
4 | 0 | 100 | 194.29 | 67896 | 7.9 | -8.9 | 0 | 100 | 3.55 | 36812.3 | 3.55 | 18 | 0 | 100 | 2.69 | 27836 | 2.45 | 28 | 0 | 100 | 2.86 | 33912.9 | 2.87
3.5 | 0 | 100 | 308.49 | 107801.2 | 12.55 | -8.5 | 0 | 100 | 4.44 | 46032.2 | 4.44 | 17 | 0 | 100 | 3.09 | 32030.1 | 2.82 | 27 | 0 | 100 | 3.06 | 36297 | 3.07
3 | 0 | 100 | 499.06 | 174398.1 | 20.26 | -8 | 0 | 100 | 5.67 | 58807.9 | 5.68 | 16 | 0 | 100 | 3.53 | 36619.3 | 3.22 | 26.1 | 0 | 100 | 3.3 | 39177.6 | 3.32
2.5 | 0 | 100 | 866.43 | 302775.5 | 35.21 | -7.5 | 0 | 100 | 7.28 | 75471.7 | 7.28 | 15 | 0 | 100 | 4.39 | 45541.6 | 4.01 | 25 | 0 | 100 | 3.55 | 42139.1 | 3.57
2 | 0 | 100 | 1452.08 | 507428.6 | 59.05 | -7 | 0 | 100 | 9.78 | 101327.9 | 9.78 | 14 | 0 | 100 | 5.39 | 55826.9 | 4.91 | 24.1 | 0 | 100 | 3.84 | 45524.4 | 3.85
1.5 | 0 | 100 | 2003.44 | 700104.6 | 81.38 | -6.5 | 0 | 100 | 13.43 | 139196.1 | 13.44 | 13 | 0 | 100 | 7.04 | 72998.1 | 6.42 | 23 | 0 | 100 | 4.2 | 49819.6 | 4.22
1 | 0 | 100 | 2275.87 | 795304.4 | 92.42 | -5.9 | 0 | 100 | 20.42 | 211640.3 | 20.43 | 12 | 0 | 100 | 9.88 | 102436.1 | 9.01 | 22.1 | 0 | 100 | 4.66 | 55270.6 | 4.68
0.6 | 0 | 100 | 2407.5 | 841303.4 | 97.85 | -5.5 | 0 | 100 | 30.94 | 320728.2 | 30.96 | 11.1 | 0 | 100 | 14.71 | 152454.4 | 13.41 | 21.1 | 0 | 100 | 5.13 | 60868.4 | 5.15
-0.1 | 0 | 100 | 2464.92 | 861366.9 | 100.27 | -5 | 0 | 100 | 50.52 | 523631.7 | 50.55 | 10 | 0 | 100 | 28.34 | 293760 | 25.84 | 20 | 0 | 100 | 5.88 | 69760.5 | 5.91
-0.5 | 0 | 100 | 2439.73 | 852564.1 | 99.25 | -4.5 | 0 | 100 | 70.45 | 730144.2 | 70.48 | 9.1 | 0 | 100 | 63.93 | 662650.2 | 58.3 | 19 | 0 | 100 | 6.9 | 81851.7 | 6.93
-1 | 0 | 100 | 2354.74 | 822866.6 | 95.73 | -4 | 0 | 100 | 84.37 | 874466.9 | 84.37 | 8.1 | 0 | 100 | 89.92 | 931979.2 | 81.99 | 18 | 0 | 100 | 8.61 | 102123 | 8.65
-1.5 | 0 | 100 | 2133.14 | 745428.1 | 86.74 | -3.5 | 0 | 100 | 90.47 | 937651 | 90.56 | 7 | 0 | 100 | 100.07 | 1037173 | 91.23 | 17 | 0 | 100 | 10.76 | 127616.7 | 10.8
-2 | 0 | 100 | 1694.82 | 592255.7 | 68.88 | -3 | 0 | 100 | 94.39 | 978284.7 | 94.47 | 6.1 | 0 | 100 | 104.37 | 1081755.1 | 95.11 | 16 | 0 | 100 | 15.36 | 182192.3 | 15.42
-2.5 | 0 | 100 | 1142.05 | 399091.4 | 46.39 | -2.5 | 0 | 100 | 96.56 | 1000806.1 | 96.62 | 5.1 | 0 | 100 | 106.37 | 1102512.7 | 96.97 | 15 | 0 | 100 | 27.33 | 324069.3 | 27.43
-3 | 0 | 100 | 662.79 | 231612.8 | 26.97 | -2 | 0 | 100 | 98.26 | 1018395.3 | 98.3 | 4.1 | 0 | 100 | 107.9 | 1118337.4 | 98.36 | 14.1 | 0 | 100 | 55.96 | 663559.5 | 56.17
-3.5 | 0 | 100 | 408.1 | 142610.5 | 16.59 | -1.5 | 0 | 100 | 99.17 | 1027821.6 | 99.22 | 3.1 | 0 | 100 | 108.61 | 1125681.3 | 99.07 | 13 | 0 | 100 | 80.19 | 950948.9 | 80.5
-4 | 0 | 100 | 266.3 | 93057.5 | 10.84 | -1 | 0 | 100 | 99.69 | 1033220 | 99.76 | 2.1 | 0 | 100 | 109.19 | 1131732.7 | 99.58 | 12 | 0 | 100 | 89.12 | 1056862.4 | 89.47
-4.5 | 0 | 100 | 175.73 | 61408.8 | 7.14 | -0.4 | 0 | 100 | 100.08 | 1037245.6 | 100.12 | 1 | 0 | 100 | 109.62 | 1136126.6 | 100 | 11.1 | 0 | 100 | 93.03 | 1103210.7 | 93.39
-5 | 0 | 100 | 122.57 | 42830.6 | 4.99 | 0 | 0 | 100 | 99.94 | 1035803.1 | 100 | 0 | 0 | 100 | 109.63 | 1136307.4 | 100 | 10 | 0 | 100 | 95.33 | 1130418 | 95.7
[6 empty] | 0.5 | 0 | 100 | 99.41 | 1030325.9 | 99.45 | -1 | 0 | 100 | 109.53 | 1135231.2 | 99.91 | 9 | 0 | 100 | 96.41 | 1143310.3 | 96.79
[6 empty] | 1 | 0 | 100 | 99.05 | 1026603.6 | 99.11 | -2.1 | 0 | 100 | 109.62 | 1136140.4 | 99.97 | 8 | 0 | 100 | 97.56 | 1156875.2 | 97.94
[6 empty] | 1.5 | 0 | 100 | 97.91 | 1014791.9 | 98.06 | -3 | 0 | 100 | 109.13 | 1131054.8 | 99.54 | 7.1 | 0 | 100 | 98.27 | 1165369.5 | 98.66
Measurement time: | 12/04/2023 18:06:57 |  |  |  |  | 2 | 0 | 100 | 96.17 | 996789.1 | 96.25 | -4 | 0 | 100 | 108.83 | 1128016 | 99.25 | 6.1 | 0 | 100 | 98.78 | 1171322.5 | 99.16
Radiation device: | Novalis TB Accelerator |  |  |  |  | 2.5 | 0 | 100 | 93.43 | 968307.9 | 93.45 | -5 | 0 | 100 | 107.78 | 1117130 | 98.28 | 5 | 0 | 100 | 99.02 | 1174265.7 | 99.41
Energy: | 6 MV (FFF) |  |  |  |  | 3 | 0 | 100 | 89.17 | 924157.1 | 89.23 | -6 | 0 | 100 | 106.26 | 1101322.2 | 96.94 | 4 | 0 | 100 | 99.32 | 1177726.5 | 99.7
Controller: | CCU |  |  |  |  | 3.5 | 0 | 100 | 82.73 | 857417.4 | 82.78 | -7.1 | 0 | 100 | 103.64 | 1074132.1 | 94.52 | 3 | 0 | 100 | 99.91 | 1184731.6 | 100.3
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 4 | 0 | 100 | 70.87 | 734548.8 | 70.92 | -8.1 | 0 | 100 | 99.41 | 1030372.7 | 90.68 | 2.1 | 0 | 100 | 99.71 | 1182338.9 | 100.09
Field detector: | SFD parallele |  |  |  |  | 4.5 | 0 | 100 | 52.02 | 539187.9 | 52.07 | -9 | 0 | 100 | 89.75 | 930210.3 | 81.86 | 1 | 0 | 100 | 99.57 | 1180734.2 | 99.96
Reference detector: | Stealth CI transmission |  |  |  |  | 5 | 0 | 100 | 32.11 | 332774.7 | 32.14 | -9.9 | 0 | 100 | 58.72 | 608609.6 | 53.55 | 0.1 | 0 | 100 | 99.59 | 1180936.3 | 99.97
SAD: | 1000 mm |  |  |  |  | 5.5 | 0 | 100 | 20.66 | 214162.9 | 20.68 | -10.9 | 0 | 100 | 26.49 | 274576.5 | 24.16 | -1 | 0 | 100 | 100.06 | 1186498.6 | 100.44
SSD: | 900 mm |  |  |  |  | 6 | 0 | 100 | 13.29 | 137747.3 | 13.3 | -11.9 | 0 | 100 | 14.65 | 151850.9 | 13.37 | -2 | 0 | 100 | 99.71 | 1182394.6 | 100.1
Field size: | 5 x 5 mm |  |  |  |  | 6.5 | 0 | 100 | 9.52 | 98680.8 | 9.53 | -12.9 | 0 | 100 | 9.56 | 99098.5 | 8.72 | -3 | 0 | 100 | 99.89 | 1184528.5 | 100.28
Gantry angle: | 0 |  |  |  |  | 7 | 0 | 100 | 7.04 | 72998.1 | 7.05 | -14 | 0 | 100 | 6.85 | 71017.3 | 6.25 | -4 | 0 | 100 | 99.64 | 1181611.1 | 100.03
Measurement medium: | Water |  |  |  |  | 7.4 | 0 | 100 | 5.49 | 56952.5 | 5.5 | -15 | 0 | 100 | 5.3 | 54925.3 | 4.84 | -5 | 0 | 100 | 99.32 | 1177819.4 | 99.71
Scan type: | Crossline |  |  |  |  | 8 | 0 | 100 | 4.29 | 44483.7 | 4.3 | -16 | 0 | 100 | 4.27 | 44292.6 | 3.9 | -6 | 0 | 100 | 99.16 | 1175814.3 | 99.54
Scan mode: | StepByStep |  |  |  |  | 8.5 | 0 | 100 | 3.44 | 35604.4 | 3.44 | -17.1 | 0 | 100 | 3.45 | 35796 | 3.15 | -7.1 | 0 | 100 | 98.67 | 1170072.5 | 99.05
Wedge: | none |  |  |  |  | 9 | 0 | 100 | 2.71 | 28130.3 | 2.72 | -18 | 0 | 100 | 3 | 31086.6 | 2.74 | -8 | 0 | 100 | 98.37 | 1166449 | 98.75
[6 empty] | 9.5 | 0 | 100 | 2.18 | 22588.5 | 2.18 | -19 | 0 | 100 | 2.67 | 27696.4 | 2.44 | -9 | 0 | 100 | 97.62 | 1157574.6 | 98
Points [mm]: | [5 empty] | 10 | 0 | 100 | 1.76 | 18254.1 | 1.76 | -20 | 0 | 100 | 2.36 | 24477.1 | 2.16 | -10 | 0 | 100 | 96.74 | 1147138.9 | 97.11
Inline | Crossline | Depth | Normalized field | Current field | Ratio | [12 empty] | -11 | 0 | 100 | 95.22 | 1129166.6 | 95.59
0 | 5 | 100 | 155.02 | 54172.8 | 6.3 | [12 empty] | -12 | 0 | 100 | 93.23 | 1105556 | 93.59
0 | 4.5 | 100 | 257.82 | 90094.1 | 10.49 | [12 empty] | -12.9 | 0 | 100 | 89.73 | 1064106.5 | 90.08
0 | 4 | 100 | 420.45 | 146926.6 | 17.11 | Measurement time: | 12/04/2023 18:30:13 |  |  |  |  | Measurement time: | 12/04/2023 18:59:35 |  |  |  |  | -14 | 0 | 100 | 81.17 | 962486.5 | 81.48
0 | 3.5 | 100 | 693.21 | 242243.6 | 28.19 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | -15 | 0 | 100 | 55.4 | 656919.7 | 55.61
0 | 3 | 100 | 1110.45 | 388048.1 | 45.16 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | -16 | 0 | 100 | 26.19 | 310594.1 | 26.29
0 | 2.4 | 100 | 1664.66 | 581714.5 | 67.72 | Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | -17.1 | 0 | 100 | 14.83 | 175820.8 | 14.88
0 | 1.9 | 100 | 2023.11 | 706976.4 | 82.3 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | -18 | 0 | 100 | 10.73 | 127216.3 | 10.77
0 | 1.4 | 100 | 2250.34 | 786383.1 | 91.46 | Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | -18.9 | 0 | 100 | 8.54 | 101243.2 | 8.57
0 | 1 | 100 | 2363.39 | 825886.7 | 95.95 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | -20 | 0 | 100 | 6.99 | 82912.8 | 7.02
0 | 0.5 | 100 | 2428.48 | 848633.3 | 98.7 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | -21 | 0 | 100 | 5.92 | 70167.5 | 5.94
0 | 0 | 100 | 2461.33 | 860112.3 | 100.06 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | -22 | 0 | 100 | 5.23 | 62076.2 | 5.26
0 | -0.5 | 100 | 2458.56 | 859144.7 | 99.84 | Field size: | 10 x 10 mm |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | -23 | 0 | 100 | 4.68 | 55464.5 | 4.7
0 | -1 | 100 | 2416.3 | 844378.3 | 98.19 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | -23.9 | 0 | 100 | 4.26 | 50477 | 4.27
0 | -1.5 | 100 | 2301.1 | 804120.1 | 93.6 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | -25 | 0 | 100 | 3.85 | 45614.5 | 3.86
0 | -2 | 100 | 2074.6 | 724970.8 | 84.37 | Scan type: | Crossline |  |  |  |  | Scan type: | Crossline |  |  |  |  | -26 | 0 | 100 | 3.58 | 42415.8 | 3.59
0 | -2.5 | 100 | 1669.76 | 583497.5 | 67.9 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | -27 | 0 | 100 | 3.28 | 38913.9 | 3.29
0 | -3 | 100 | 1167.72 | 408061.8 | 47.51 | Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | -28 | 0 | 100 | 3.06 | 36258.1 | 3.07
0 | -3.5 | 100 | 775.27 | 270919.4 | 31.54 | [12 empty] | -29 | 0 | 100 | 2.84 | 33680.2 | 2.85
0 | -4 | 100 | 502.35 | 175547.5 | 20.43 | Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | -30 | 0 | 100 | 2.67 | 31606.2 | 2.68
0 | -4.5 | 100 | 322.33 | 112637.9 | 13.11 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -5 | 100 | 210.22 | 73462 | 8.55 | 0 | -10 | 100 | 2.55 | 26458.6 | 2.56 | 0 | 20 | 100 | 8.04 | 24252.4 | 2.14
[6 empty] | 0 | -9.5 | 100 | 3.22 | 33400.9 | 3.23 | 0 | 19 | 100 | 9.2 | 27731.8 | 2.44
[6 empty] | 0 | -9 | 100 | 3.97 | 41189.2 | 3.98 | 0 | 18 | 100 | 10.6 | 31967 | 2.81 | Measurement time: | 12/04/2023 19:18:35
[6 empty] | 0 | -8.5 | 100 | 5.28 | 54759.5 | 5.29 | 0 | 17 | 100 | 12.29 | 37065.7 | 3.26 | Radiation device: | Novalis TB Accelerator
Measurement time: | 12/04/2023 18:07:44 |  |  |  |  | 0 | -8 | 100 | 7.17 | 74319 | 7.19 | 0 | 16 | 100 | 14.97 | 45157.1 | 3.97 | Energy: | 6 MV (FFF)
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -7.5 | 100 | 9.63 | 99804 | 9.65 | 0 | 15 | 100 | 18.35 | 55327.8 | 4.87 | Controller: | CCU
Energy: | 6 MV (FFF) |  |  |  |  | 0 | -7 | 100 | 13.33 | 138150.2 | 13.36 | 0 | 14 | 100 | 24.16 | 72867.7 | 6.42 | Measurement device: | Blue Phantom 2 (48*48*41)
Controller: | CCU |  |  |  |  | 0 | -6.4 | 100 | 20.27 | 210139.7 | 20.32 | 0 | 13 | 100 | 33.86 | 102116.3 | 8.99 | Field detector: | SFD parallele
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -6 | 100 | 29.43 | 305029.8 | 29.52 | 0 | 12 | 100 | 52.6 | 158625.6 | 13.96 | Reference detector: | Stealth CI transmission
Field detector: | SFD parallele |  |  |  |  | 0 | -5.5 | 100 | 43.21 | 447869.2 | 43.31 | 0 | 11.1 | 100 | 99.58 | 300326.6 | 26.44 | SAD: | 1000 mm
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -5 | 100 | 61.58 | 638248.4 | 61.76 | 0 | 10 | 100 | 212.24 | 640106.2 | 56.35 | SSD: | 900 mm
SAD: | 1000 mm |  |  |  |  | 0 | -4.5 | 100 | 75.85 | 786119.9 | 75.98 | 0 | 9 | 100 | 301.03 | 907859 | 79.95 | Field size: | 30 x 30 mm
SSD: | 900 mm |  |  |  |  | 0 | -4.1 | 100 | 86.35 | 894976.4 | 86.52 | 0 | 8 | 100 | 338.09 | 1019655.2 | 89.76 | Gantry angle: | 0
Field size: | 5 x 5 mm |  |  |  |  | 0 | -3.6 | 100 | 92.01 | 953651.3 | 92.26 | 0 | 7.1 | 100 | 355.45 | 1072010.1 | 94.35 | Measurement medium: | Water
Gantry angle: | 0 |  |  |  |  | 0 | -3 | 100 | 95.17 | 986387.3 | 95.43 | 0 | 6 | 100 | 364.06 | 1097953.2 | 96.68 | Scan type: | Crossline
Measurement medium: | Water |  |  |  |  | 0 | -2.6 | 100 | 96.71 | 1002404.2 | 96.96 | 0 | 5 | 100 | 369.31 | 1113809.5 | 98 | Scan mode: | StepByStep
Scan type: | Beam |  |  |  |  | 0 | -2 | 100 | 98.04 | 1016119.6 | 98.33 | 0 | 4 | 100 | 372.97 | 1124844.8 | 99.01 | Wedge: | none
Scan mode: | StepByStep |  |  |  |  | 0 | -1.4 | 100 | 99.1 | 1027173 | 99.37 | 0 | 3.1 | 100 | 374.28 | 1128776.3 | 99.42
Wedge: | none |  |  |  |  | 0 | -1.1 | 100 | 99.27 | 1028878.4 | 99.52 | 0 | 2.1 | 100 | 375.94 | 1133784.6 | 99.81 | Points [mm]:
[6 empty] | 0 | -0.6 | 100 | 99.66 | 1032900.6 | 99.94 | 0 | 1.1 | 100 | 376.1 | 1134281.7 | 99.92 | Inline | Crossline | Depth | Normalized field | Current field | Ratio
Points [mm]: | [5 empty] | 0 | 0 | 100 | 99.73 | 1033632.5 | 100.01 | 0 | 0.1 | 100 | 376.61 | 1135797.8 | 99.99 | 0 | 30 | 100 | 2.66 | 31505.5 | 2.67
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 0.5 | 100 | 99.66 | 1032965.2 | 99.91 | 0 | -0.9 | 100 | 376.81 | 1136408.2 | 100.11 | 0 | 29 | 100 | 2.83 | 33549.7 | 2.84
0 | 0 | 250 | 819.69 | 286442 | 16.13 | 0 | 0.9 | 100 | 99.41 | 1030332 | 99.67 | 0 | -1.9 | 100 | 375.69 | 1133035.5 | 99.79 | 0 | 28.1 | 100 | 3.03 | 35917.4 | 3.04
0 | 0 | 245 | 845.62 | 295502.1 | 16.67 | 0 | 1.4 | 100 | 98.7 | 1023011.3 | 99 | 0 | -2.9 | 100 | 374.78 | 1130283.3 | 99.58 | 0 | 27.1 | 100 | 3.25 | 38549.8 | 3.26
0 | 0 | 240 | 878.37 | 306946 | 17.3 | 0 | 2 | 100 | 97.66 | 1012214.8 | 97.9 | 0 | -3.9 | 100 | 373.48 | 1126380.5 | 99.16 | 0 | 26.1 | 100 | 3.48 | 41259.3 | 3.49
0 | 0 | 235 | 907.5 | 317125.8 | 17.85 | 0 | 2.4 | 100 | 95.99 | 994904.5 | 96.26 | 0 | -4.9 | 100 | 370.06 | 1116059.5 | 98.28 | 0 | 25 | 100 | 3.81 | 45149.9 | 3.82
0 | 0 | 230 | 940.27 | 328578.4 | 18.52 | 0 | 3 | 100 | 93.47 | 968799.1 | 93.75 | 0 | -5.9 | 100 | 366.94 | 1106654.3 | 97.43 | 0 | 24 | 100 | 4.16 | 49320.9 | 4.18
0 | 0 | 225 | 977 | 341413.7 | 19.21 | 0 | 3.5 | 100 | 89.49 | 927488.6 | 89.7 | 0 | -6.9 | 100 | 359.21 | 1083335.9 | 95.37 | 0 | 23.1 | 100 | 4.58 | 54284.5 | 4.6
0 | 0 | 220 | 1011.19 | 353359.9 | 19.92 | 0 | 4 | 100 | 83.71 | 867614.6 | 83.99 | 0 | -7.9 | 100 | 349.27 | 1053343.3 | 92.76 | 0 | 21.9 | 100 | 5.18 | 61470.4 | 5.21
0 | 0 | 215.1 | 1045.77 | 365445.3 | 20.62 | 0 | 4.4 | 100 | 75.95 | 787173.4 | 76.11 | 0 | -9 | 100 | 323.21 | 974760.8 | 85.87 | 0 | 21 | 100 | 5.84 | 69230 | 5.86
0 | 0 | 210 | 1081.97 | 378093.9 | 21.32 | 0 | 5 | 100 | 62.6 | 648801.1 | 62.79 | 0 | -9.9 | 100 | 241.9 | 729549.8 | 64.23 | 0 | 20.1 | 100 | 6.91 | 81939.7 | 6.94
0 | 0 | 205.1 | 1124.98 | 393125.8 | 22.16 | 0 | 5.4 | 100 | 45.69 | 473582.2 | 45.84 | 0 | -10.9 | 100 | 129.73 | 391255.8 | 34.47 | 0 | 19 | 100 | 8.46 | 100350.2 | 8.5
0 | 0 | 200.1 | 1166.3 | 407562.5 | 22.98 | 0 | 6 | 100 | 28.23 | 292540.2 | 28.31 | 0 | -11.9 | 100 | 65.9 | 198741.7 | 17.5 | 0 | 18 | 100 | 10.93 | 129650 | 10.98
0 | 0 | 195 | 1202.94 | 420368.3 | 23.7 | 0 | 6.5 | 100 | 17.79 | 184337.6 | 17.83 | 0 | -13 | 100 | 39.43 | 118903 | 10.47 | 0 | 17 | 100 | 15.79 | 187273.3 | 15.86
0 | 0 | 190.1 | 1247.41 | 435909.5 | 24.56 | 0 | 7 | 100 | 12.2 | 126413.8 | 12.23 | 0 | -14 | 100 | 27.21 | 82055.8 | 7.23 | 0 | 16 | 100 | 27.96 | 331596.9 | 28.08
0 | 0 | 185.1 | 1294.22 | 452265.8 | 25.53 | 0 | 7.5 | 100 | 8.76 | 90768.4 | 8.79 | 0 | -15 | 100 | 20.36 | 61393.2 | 5.41 | 0 | 15 | 100 | 57.05 | 676469.2 | 57.28
0 | 0 | 180.1 | 1341.53 | 468796.8 | 26.46 | 0 | 7.9 | 100 | 6.52 | 67526.4 | 6.54 | 0 | -15.9 | 100 | 16.02 | 48329.4 | 4.25 | 0 | 14 | 100 | 78.25 | 927943.5 | 78.58
0 | 0 | 175.1 | 1398.11 | 488569 | 27.54 | 0 | 8.4 | 100 | 5.02 | 52025.8 | 5.04 | 0 | -17 | 100 | 13.01 | 39229 | 3.45 | 0 | 13.1 | 100 | 87.64 | 1039301.4 | 88.01
0 | 0 | 170.1 | 1448.57 | 506204.9 | 28.52 | 0 | 8.9 | 100 | 3.85 | 39936.5 | 3.86 | 0 | -17.9 | 100 | 10.99 | 33133 | 2.92 | 0 | 12 | 100 | 92.58 | 1097801.9 | 92.96
0 | 0 | 165.1 | 1505.46 | 526081.9 | 29.64 | 0 | 9.4 | 100 | 3.06 | 31711 | 3.07 | 0 | -18.9 | 100 | 9.59 | 28919.8 | 2.55 | 0 | 11 | 100 | 94.9 | 1125399.4 | 95.3
0 | 0 | 160.1 | 1557.3 | 544198.7 | 30.69 | 0 | 10 | 100 | 2.43 | 25216.2 | 2.44 | 0 | -20 | 100 | 8.37 | 25242.5 | 2.22 | 0 | 10 | 100 | 96.49 | 1144199.1 | 96.89
0 | 0 | 155.1 | 1617.01 | 565064.5 | 31.84 | [12 empty] | 0 | 9 | 100 | 97.25 | 1153223.2 | 97.65
0 | 0 | 150 | 1680.16 | 587133.5 | 33.06 | [12 empty] | 0 | 8 | 100 | 98 | 1162096.7 | 98.4
0 | 0 | 145.1 | 1742.55 | 608935.8 | 34.3 | [12 empty] | 0 | 7.1 | 100 | 98.54 | 1168534.2 | 98.95
0 | 0 | 140 | 1809.47 | 632321 | 35.72 | Measurement time: | 12/04/2023 18:31:48 |  |  |  |  | Measurement time: | 12/04/2023 19:01:28 |  |  |  |  | 0 | 6 | 100 | 99.08 | 1174895.4 | 99.49
0 | 0 | 135.1 | 1880.68 | 657203.3 | 37.07 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | 5.1 | 100 | 99.57 | 1180703.8 | 99.98
0 | 0 | 130.1 | 1954.86 | 683127 | 38.52 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | 0 | 4 | 100 | 99.5 | 1179902.2 | 99.91
0 | 0 | 125 | 2028.16 | 708741.4 | 39.96 | Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | 0 | 3 | 100 | 99.49 | 1179831.8 | 99.91
0 | 0 | 120.1 | 2103.13 | 734939.7 | 41.43 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | 2 | 100 | 99.49 | 1179775.4 | 99.9
0 | 0 | 115.1 | 2187.41 | 764390.5 | 43.04 | Field detector: | SFD parallele |  |  |  |  | Field detector: | SFD parallele |  |  |  |  | 0 | 1 | 100 | 99.77 | 1183129.3 | 100.19
0 | 0 | 110.1 | 2275.84 | 795293.3 | 44.77 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | 0 | 100 | 99.58 | 1180883.2 | 100
0 | 0 | 105.1 | 2362.57 | 825602.3 | 46.56 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | 0 | -0.9 | 100 | 99.67 | 1181917.5 | 100.08
0 | 0 | 100.1 | 2457.19 | 858664.6 | 48.38 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | 0 | -2 | 100 | 99.69 | 1182166.6 | 100.1
0 | 0 | 95.1 | 2556.97 | 893533.5 | 50.32 | Field size: | 10 x 10 mm |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | 0 | -3 | 100 | 99.28 | 1177299.2 | 99.69
0 | 0 | 90.1 | 2653.69 | 927333.8 | 52.32 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | 0 | -4 | 100 | 99.39 | 1178615.9 | 99.8
0 | 0 | 85 | 2770.43 | 968129.1 | 54.6 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | 0 | -4.9 | 100 | 99.22 | 1176569.1 | 99.63
0 | 0 | 80 | 2882.02 | 1007124.3 | 56.77 | Scan type: | Beam |  |  |  |  | Scan type: | Beam |  |  |  |  | 0 | -6 | 100 | 98.84 | 1172042 | 99.25
0 | 0 | 75 | 2999.47 | 1048165.9 | 59.18 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | 0 | -7 | 100 | 98.46 | 1167582.8 | 98.87
0 | 0 | 70 | 3121.72 | 1090885.1 | 61.48 | Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | 0 | -7.9 | 100 | 98.09 | 1163181.7 | 98.5
0 | 0 | 65 | 3247.79 | 1134940.1 | 63.93 | [12 empty] | 0 | -8.9 | 100 | 97.31 | 1153980 | 97.72
0 | 0 | 60 | 3376.25 | 1179831.3 | 66.49 | Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | 0 | -9.9 | 100 | 96.41 | 1143312.3 | 96.81
0 | 0 | 55 | 3524.92 | 1231784.6 | 69.47 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | -11 | 100 | 95.49 | 1132338.9 | 95.88
0 | 0 | 50 | 3667.95 | 1281768.2 | 72.26 | 0 | 0 | 250 | 34.3 | 355518.6 | 17.57 | 0 | 0 | 250 | 130.81 | 394497.1 | 18.48 | 0 | -12 | 100 | 93.65 | 1110544.8 | 94.04
0 | 0 | 44.9 | 3827.2 | 1337415.1 | 75.31 | 0 | 0 | 245 | 35.42 | 367060.7 | 18.16 | 0 | 0 | 245 | 135.93 | 409950.4 | 19.19 | 0 | -12.9 | 100 | 90.81 | 1076835.2 | 91.18
0 | 0 | 39.9 | 3989.64 | 1394179.8 | 78.58 | 0 | 0 | 240 | 36.74 | 380763.2 | 18.84 | 0 | 0 | 240 | 139.35 | 420271.5 | 19.68 | 0 | -13.9 | 100 | 84.51 | 1002184.3 | 84.86
0 | 0 | 34.9 | 4162.14 | 1454463.3 | 82.01 | 0 | 0 | 235 | 37.9 | 392811.9 | 19.42 | 0 | 0 | 235 | 145.04 | 437426.3 | 20.49 | 0 | -15 | 100 | 64.66 | 766706.2 | 64.92
0 | 0 | 29.9 | 4334.49 | 1514688 | 85.36 | 0 | 0 | 230 | 39.31 | 407400.6 | 20.14 | 0 | 0 | 230 | 149.65 | 451336.2 | 21.13 | 0 | -16 | 100 | 36.68 | 434920.1 | 36.83
0 | 0 | 24.9 | 4513.08 | 1577098.8 | 88.93 | 0 | 0 | 225 | 40.66 | 421446.3 | 20.84 | 0 | 0 | 225 | 155.04 | 467598 | 21.9 | 0 | -17 | 100 | 19.59 | 232279.2 | 19.67
0 | 0 | 20.1 | 4697.01 | 1641371.9 | 92.47 | 0 | 0 | 220 | 42.13 | 436631.7 | 21.59 | 0 | 0 | 220.1 | 159.99 | 482500.8 | 22.58 | 0 | -17.9 | 100 | 12.65 | 149992.7 | 12.7
0 | 0 | 19 | 4737.67 | 1655582.1 | 93.39 | 0 | 0 | 215.1 | 43.64 | 452288 | 22.37 | 0 | 0 | 215.1 | 166.05 | 500779.3 | 23.45 | 0 | -19 | 100 | 9.26 | 109835.7 | 9.3
0 | 0 | 18.1 | 4781.98 | 1671065.1 | 94.21 | 0 | 0 | 210 | 44.95 | 465895.6 | 23.04 | 0 | 0 | 210.1 | 171.79 | 518108.3 | 24.26 | 0 | -20 | 100 | 7.39 | 87598.7 | 7.42
0 | 0 | 17.1 | 4809.76 | 1680771 | 94.79 | 0 | 0 | 205.1 | 46.63 | 483313.8 | 23.9 | 0 | 0 | 205.1 | 177.19 | 534371.5 | 25.03 | 0 | -20.9 | 100 | 6.17 | 73188.5 | 6.2
0 | 0 | 16 | 4864.04 | 1699742.5 | 95.76 | 0 | 0 | 200 | 48.38 | 501413 | 24.79 | 0 | 0 | 200.1 | 183.73 | 554094.7 | 25.94 | 0 | -21.9 | 100 | 5.37 | 63632 | 5.39
0 | 0 | 15 | 4897.24 | 1711342.6 | 96.32 | 0 | 0 | 195 | 49.92 | 517418.5 | 25.59 | 0 | 0 | 195 | 189.78 | 572349.4 | 26.8 | 0 | -23 | 100 | 4.77 | 56510.7 | 4.79
0 | 0 | 14.1 | 4930.09 | 1722822.5 | 97.18 | 0 | 0 | 190 | 51.87 | 537600.6 | 26.6 | 0 | 0 | 190 | 197.07 | 594351.2 | 27.84 | 0 | -23.9 | 100 | 4.25 | 50442.4 | 4.27
0 | 0 | 13 | 4961.71 | 1733872.8 | 97.8 | 0 | 0 | 185.1 | 53.43 | 553819.9 | 27.39 | 0 | 0 | 185.1 | 203.73 | 614439.3 | 28.77 | 0 | -24.9 | 100 | 3.89 | 46096.7 | 3.9
0 | 0 | 12.1 | 4990.73 | 1744013.3 | 98.15 | 0 | 0 | 180 | 55.51 | 575366.4 | 28.47 | 0 | 0 | 180.1 | 211.23 | 637054.7 | 29.84 | 0 | -26 | 100 | 3.58 | 42415.2 | 3.59
0 | 0 | 11.1 | 5037.59 | 1760389.1 | 99.19 | 0 | 0 | 175.1 | 57.63 | 597271.8 | 29.54 | 0 | 0 | 175.1 | 218.74 | 659683.4 | 30.9 | 0 | -27 | 100 | 3.29 | 39060.3 | 3.31
0 | 0 | 10.1 | 5047.11 | 1763714.1 | 99.33 | 0 | 0 | 170.1 | 59.62 | 617937.3 | 30.56 | 0 | 0 | 170.1 | 226.6 | 683398.6 | 32.02 | 0 | -28 | 100 | 3.1 | 36714.8 | 3.11
0 | 0 | 9.1 | 5073.04 | 1772777.6 | 100 | 0 | 0 | 165 | 61.72 | 639679.3 | 31.63 | 0 | 0 | 165.1 | 234.26 | 706502.6 | 33.09 | 0 | -29 | 100 | 2.85 | 33811 | 2.86
0 | 0 | 8 | 5077.09 | 1774190.6 | 99.91 | 0 | 0 | 160.1 | 64.14 | 664760.7 | 32.86 | 0 | 0 | 160.1 | 242.66 | 731842.8 | 34.29 | 0 | -30 | 100 | 2.65 | 31448 | 2.66
0 | 0 | 7 | 5064.11 | 1769655 | 99.79 | 0 | 0 | 155.1 | 66.28 | 686954.3 | 33.97 | 0 | 0 | 155 | 251.3 | 757901.2 | 35.5
0 | 0 | 6.1 | 5012.86 | 1751744.4 | 98.86 | 0 | 0 | 150 | 68.63 | 711291.2 | 35.2 | 0 | 0 | 150.1 | 260.24 | 784865.1 | 36.77
0 | 0 | 5 | 4938.8 | 1725866.6 | 97.3 | 0 | 0 | 145.1 | 71.31 | 739144.9 | 36.57 | 0 | 0 | 145.1 | 269.89 | 813952.8 | 38.14
0 | 0 | 4 | 4777.05 | 1669341 | 94.12 | 0 | 0 | 140 | 73.9 | 765970.9 | 37.89 | 0 | 0 | 140 | 279.78 | 843774 | 39.54 | Measurement time: | 12/04/2023 19:22:22
0 | 0 | 3 | 4507.73 | 1575228.3 | 88.84 | 0 | 0 | 135 | 76.67 | 794652.7 | 39.31 | 0 | 0 | 135.1 | 291.05 | 877785.4 | 41.11 | Radiation device: | Novalis TB Accelerator
0 | 0 | 2 | 4040.66 | 1412009 | 79.62 | 0 | 0 | 130.1 | 79.49 | 823867 | 40.78 | 0 | 0 | 130.1 | 301.14 | 908190.9 | 42.56 | Energy: | 6 MV (FFF)
0 | 0 | 1 | 3147.73 | 1099976.8 | 62.09 | 0 | 0 | 125.1 | 82.55 | 855544.6 | 42.33 | 0 | 0 | 125.1 | 312.46 | 942343.4 | 44.15 | Controller: | CCU
0 | 0 | 0 | 2775.49 | 969896.4 | 54.67 | 0 | 0 | 120.1 | 85.7 | 888188.7 | 43.94 | 0 | 0 | 120.1 | 324.76 | 979452.1 | 45.9 | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | -0.5 | 2773.72 | 969277.8 | 54.57 | 0 | 0 | 115.1 | 89.01 | 922513.7 | 45.65 | 0 | 0 | 115 | 336.26 | 1014113.4 | 47.52 | Field detector: | SFD parallele
[6 empty] | 0 | 0 | 110.1 | 92.48 | 958543.4 | 47.44 | 0 | 0 | 110.1 | 349.19 | 1053109.3 | 49.34 | Reference detector: | Stealth CI transmission
[6 empty] | 0 | 0 | 105 | 95.92 | 994135.3 | 49.2 | 0 | 0 | 105.1 | 361.44 | 1090072.1 | 51.08 | SAD: | 1000 mm
[6 empty] | 0 | 0 | 100 | 99.51 | 1031374.2 | 51.04 | 0 | 0 | 100.1 | 376.21 | 1134618.6 | 53.17 | SSD: | 900 mm
[6 empty] | 0 | 0 | 95 | 103.36 | 1071300.1 | 53.02 | 0 | 0 | 95.1 | 390.44 | 1177525.9 | 55.18 | Field size: | 30 x 30 mm
[6 empty] | 0 | 0 | 90.1 | 107.29 | 1111985.8 | 55.02 | 0 | 0 | 90.1 | 405.06 | 1221607.9 | 57.25 | Gantry angle: | 0
[6 empty] | 0 | 0 | 85 | 111.47 | 1155313.9 | 57.18 | 0 | 0 | 85 | 421.24 | 1270418.4 | 59.53 | Measurement medium: | Water
[6 empty] | 0 | 0 | 80 | 115.61 | 1198238.5 | 59.29 | 0 | 0 | 80 | 436.63 | 1316841.1 | 61.73 | Scan type: | Beam
[6 empty] | 0 | 0 | 75 | 120.6 | 1249946.2 | 61.86 | 0 | 0 | 75.1 | 453.46 | 1367581.3 | 64.09 | Scan mode: | StepByStep
[6 empty] | 0 | 0 | 70 | 125.17 | 1297294.8 | 64.18 | 0 | 0 | 70 | 470.72 | 1419628.5 | 66.55 | Wedge: | none
[6 empty] | 0 | 0 | 65 | 130.29 | 1350440.8 | 66.83 | 0 | 0 | 65 | 489.48 | 1476222 | 69.2
[6 empty] | 0 | 0 | 60 | 135.18 | 1401101.9 | 69.33 | 0 | 0 | 60 | 508.68 | 1534123.9 | 71.94 | Points [mm]:
[6 empty] | 0 | 0 | 55 | 140.71 | 1458385.5 | 72.16 | 0 | 0 | 55 | 527.3 | 1590268.3 | 74.56 | Inline | Crossline | Depth | Normalized field | Current field | Ratio
[6 empty] | 0 | 0 | 49.9 | 146.68 | 1520316 | 75.22 | 0 | 0 | 50 | 548.69 | 1654774.4 | 77.56 | 0 | 0 | 250 | 34.88 | 413631.7 | 19.19
[6 empty] | 0 | 0 | 44.9 | 152.47 | 1580291 | 78.2 | 0 | 0 | 44.9 | 569.99 | 1719032 | 80.58 | 0 | 0 | 245 | 36.07 | 427733 | 19.84
[6 empty] | 0 | 0 | 39.9 | 158.38 | 1641555.2 | 81.26 | 0 | 0 | 39.9 | 591.42 | 1783651.4 | 83.62 | 0 | 0 | 240 | 37.22 | 441414.7 | 20.48
[6 empty] | 0 | 0 | 34.9 | 165 | 1710095.6 | 84.62 | 0 | 0 | 34.9 | 614.39 | 1852928 | 86.86 | 0 | 0 | 235 | 38.53 | 456928.4 | 21.2
[6 empty] | 0 | 0 | 29.9 | 171.57 | 1778208.3 | 88.03 | 0 | 0 | 29.9 | 637.85 | 1923695.3 | 90.17 | 0 | 0 | 230 | 39.79 | 471830.7 | 21.89
[6 empty] | 0 | 0 | 24.9 | 178.68 | 1851947 | 91.64 | 0 | 0 | 24.9 | 661.73 | 1995715.4 | 93.58 | 0 | 0 | 225 | 41.29 | 489607.4 | 22.71
[6 empty] | 0 | 0 | 20.1 | 185.35 | 1921047.4 | 95.07 | 0 | 0 | 20 | 684.93 | 2065685.6 | 96.87 | 0 | 0 | 220.1 | 42.46 | 503545.1 | 23.36
[6 empty] | 0 | 0 | 19.1 | 186.57 | 1933744.1 | 95.72 | 0 | 0 | 19.1 | 689.6 | 2079761.8 | 97.47 | 0 | 0 | 215.1 | 44.08 | 522768.6 | 24.25
[6 empty] | 0 | 0 | 18.1 | 188.24 | 1950973.1 | 96.59 | 0 | 0 | 18.1 | 693.19 | 2090569.5 | 98.01 | 0 | 0 | 210 | 45.72 | 542159.8 | 25.15
[6 empty] | 0 | 0 | 17.1 | 189.28 | 1961833.8 | 97.1 | 0 | 0 | 17.1 | 696.62 | 2100918.5 | 98.56 | 0 | 0 | 205.1 | 47.32 | 561079.4 | 26.03
[6 empty] | 0 | 0 | 16.1 | 190.6 | 1975478 | 97.76 | 0 | 0 | 16.1 | 700.68 | 2113176.1 | 99.05 | 0 | 0 | 200.1 | 48.95 | 580410.9 | 26.92
[6 empty] | 0 | 0 | 15.1 | 191.83 | 1988182.3 | 98.41 | 0 | 0 | 15.1 | 703.6 | 2121986.4 | 99.47 | 0 | 0 | 195 | 50.7 | 601241.2 | 27.89
[6 empty] | 0 | 0 | 14.1 | 192.83 | 1998557.5 | 98.96 | 0 | 0 | 14.1 | 705.71 | 2128336.5 | 99.78 | 0 | 0 | 190.1 | 52.37 | 621028.8 | 28.81
[6 empty] | 0 | 0 | 13.1 | 193.81 | 2008732.4 | 99.47 | 0 | 0 | 13.1 | 707.03 | 2132334.9 | 100 | 0 | 0 | 185.1 | 54.08 | 641354.9 | 29.75
[6 empty] | 0 | 0 | 12.1 | 194.49 | 2015753 | 99.78 | 0 | 0 | 12.1 | 707.09 | 2132496.7 | 99.96 | 0 | 0 | 180.1 | 56.14 | 665748 | 30.88
[6 empty] | 0 | 0 | 11.1 | 194.69 | 2017865 | 99.89 | 0 | 0 | 11.1 | 706.23 | 2129905.9 | 99.85 | 0 | 0 | 175.1 | 57.93 | 686905.9 | 31.86
[6 empty] | 0 | 0 | 10.1 | 194.86 | 2019680.6 | 100 | 0 | 0 | 10.1 | 704.77 | 2125506.1 | 99.61 | 0 | 0 | 170.1 | 60.26 | 714583.4 | 33.15
[6 empty] | 0 | 0 | 9 | 194.27 | 2013542.3 | 99.67 | 0 | 0 | 9.1 | 698.29 | 2105965.2 | 98.73 | 0 | 0 | 165.1 | 62.47 | 740830.1 | 34.37
[6 empty] | 0 | 0 | 8 | 192.67 | 1996900.2 | 98.85 | 0 | 0 | 8 | 689.43 | 2079233 | 97.51 | 0 | 0 | 160.1 | 64.64 | 766518.6 | 35.56
[6 empty] | 0 | 0 | 7 | 190.66 | 1976127.2 | 97.85 | 0 | 0 | 7 | 676.34 | 2039749.6 | 95.68 | 0 | 0 | 155.1 | 66.78 | 791860.9 | 36.73
[6 empty] | 0 | 0 | 6 | 186.78 | 1935894.6 | 95.91 | 0 | 0 | 6 | 661.52 | 1995062.8 | 93.52 | 0 | 0 | 150.1 | 69.3 | 821750.4 | 38.12
[6 empty] | 0 | 0 | 5.1 | 181.55 | 1881681.3 | 93.18 | 0 | 0 | 5 | 637.08 | 1921357.8 | 90.12 | 0 | 0 | 145.1 | 71.81 | 851554.8 | 39.5
[6 empty] | 0 | 0 | 4 | 172.41 | 1786952.2 | 88.46 | 0 | 0 | 4 | 605.88 | 1827272.6 | 85.7 | 0 | 0 | 140.1 | 74.3 | 881068.5 | 40.87
[6 empty] | 0 | 0 | 3 | 159.76 | 1655787.1 | 81.95 | 0 | 0 | 3 | 556.97 | 1679759.1 | 78.8 | 0 | 0 | 135.1 | 77.21 | 915550.3 | 42.47
[6 empty] | 0 | 0 | 2 | 141.01 | 1461473.2 | 72.35 | 0 | 0 | 2 | 486.47 | 1467145.4 | 68.82 | 0 | 0 | 130.1 | 79.95 | 948018.6 | 43.98
[6 empty] | 0 | 0 | 1 | 108.86 | 1128290.5 | 55.86 | 0 | 0 | 1 | 370.24 | 1116591.2 | 52.37 | 0 | 0 | 125 | 82.91 | 983151.5 | 45.61
[6 empty] | 0 | 0 | 0 | 97.47 | 1010231.6 | 50.03 | 0 | 0 | 0 | 343.63 | 1036362.3 | 48.61 | 0 | 0 | 120.1 | 85.97 | 1019469.4 | 47.29
[6 empty] | 0 | 0 | -0.5 | 97.67 | 1012313 | 50.13 | 0 | 0 | -0.5 | 343.52 | 1036029.9 | 48.6 | 0 | 0 | 115.1 | 89.26 | 1058515.6 | 49.1
[18 empty] | 0 | 0 | 110.1 | 92.39 | 1095555.7 | 50.82
[18 empty] | 0 | 0 | 105.1 | 95.97 | 1137995.6 | 52.79
[18 empty] | 0 | 0 | 100 | 99.43 | 1179027 | 54.69
[18 empty] | 0 | 0 | 95.1 | 103.24 | 1224206.8 | 56.79
[18 empty] | 0 | 0 | 90.1 | 106.98 | 1268568.6 | 58.85
[18 empty] | 0 | 0 | 85 | 111.13 | 1317761.6 | 61.13
[18 empty] | 0 | 0 | 80 | 115.17 | 1365774 | 63.36
[18 empty] | 0 | 0 | 75 | 119.59 | 1418152.3 | 65.79
[18 empty] | 0 | 0 | 70 | 124.2 | 1472774.7 | 68.32
[18 empty] | 0 | 0 | 65 | 128.82 | 1527614.8 | 70.86
[18 empty] | 0 | 0 | 60 | 133.57 | 1583929.5 | 73.48
[18 empty] | 0 | 0 | 55 | 138.73 | 1645072.8 | 76.31
[18 empty] | 0 | 0 | 49.9 | 143.87 | 1706103.5 | 79.14
[18 empty] | 0 | 0 | 44.9 | 149.29 | 1770329.5 | 82.12
[18 empty] | 0 | 0 | 39.9 | 154.76 | 1835207.4 | 85.13
[18 empty] | 0 | 0 | 34.9 | 160.5 | 1903217.7 | 88.29
[18 empty] | 0 | 0 | 29.9 | 165.87 | 1966919.3 | 91.24
[18 empty] | 0 | 0 | 24.9 | 171.79 | 2037127.7 | 94.5
[18 empty] | 0 | 0 | 20.1 | 177.35 | 2103076.4 | 97.56
[18 empty] | 0 | 0 | 19.1 | 178.09 | 2111876.1 | 97.97
[18 empty] | 0 | 0 | 18.1 | 179.15 | 2124390.7 | 98.55
[18 empty] | 0 | 0 | 17.1 | 179.8 | 2132146.7 | 98.91
[18 empty] | 0 | 0 | 16.1 | 180.41 | 2139363.8 | 99.24
[18 empty] | 0 | 0 | 15.1 | 181.34 | 2150432.3 | 99.76
[18 empty] | 0 | 0 | 14.1 | 181.41 | 2151191.4 | 99.79
[18 empty] | 0 | 0 | 13.1 | 181.79 | 2155701.9 | 100
[18 empty] | 0 | 0 | 12 | 181.58 | 2153216.2 | 99.88
[18 empty] | 0 | 0 | 11.1 | 181.27 | 2149580.5 | 99.72
[18 empty] | 0 | 0 | 10.1 | 180.66 | 2142340.2 | 99.38
[18 empty] | 0 | 0 | 9.1 | 178.91 | 2121608.9 | 98.42
[18 empty] | 0 | 0 | 8 | 177.1 | 2100061.1 | 97.42
[18 empty] | 0 | 0 | 7 | 173.69 | 2059723.3 | 95.55
[18 empty] | 0 | 0 | 6 | 168.78 | 2001443.2 | 92.84
[18 empty] | 0 | 0 | 5 | 162.83 | 1930915.1 | 89.57
[18 empty] | 0 | 0 | 4 | 154.59 | 1833233.6 | 85.04
[18 empty] | 0 | 0 | 3 | 142.7 | 1692205 | 78.5
[18 empty] | 0 | 0 | 2 | 125.15 | 1484111.3 | 68.85
[18 empty] | 0 | 0 | 1 | 94.97 | 1126221.8 | 52.24
[18 empty] | 0 | 0 | 0 | 90.04 | 1067711.6 | 49.53
[18 empty] | 0 | 0 | -0.5 | 89.98 | 1066990.7 | 49.5

## µD
0,5 x 0,5
Measurement time: | 27/05/2023 10:11:08 |  |  |  |  | Measurement time: | 27/05/2023 10:34:11 |  |  |  |  | Measurement time: | 27/05/2023 10:55:28 |  |  |  |  | Measurement time: | 27/05/2023 11:21:09
Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant
Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 5 x 5 mm |  |  |  |  | Field size: | 10 x 10 mm |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 54.94 | 713 | 0.95 | 20 | 0 | 100 | 63.26 | 534.4 | 0.58 | -30 | 0 | 100 | 130.51 | 1102.5 | 1.08 | -30 | 0 | 100 | 324.34 | 2739.7 | 2.56
-9.5 | 0 | 100 | 58.36 | 757.5 | 1.01 | 19.5 | 0 | 100 | 64.68 | 546.3 | 0.59 | -29 | 0 | 100 | 154.35 | 1303.8 | 1.28 | -29.1 | 0 | 100 | 351.13 | 2966 | 2.78
-9 | 0 | 100 | 67.74 | 879.2 | 1.17 | 19 | 0 | 100 | 71.45 | 603.5 | 0.65 | -28 | 0 | 100 | 155.29 | 1311.7 | 1.28 | -28.1 | 0 | 100 | 382.72 | 3232.9 | 3.03
-8.4 | 0 | 100 | 83.09 | 1078.4 | 1.44 | 18.5 | 0 | 100 | 67.96 | 574.1 | 0.62 | -27 | 0 | 100 | 168.83 | 1426.2 | 1.4 | -27.1 | 0 | 100 | 404.92 | 3420.4 | 3.2
-8 | 0 | 100 | 96.87 | 1257.2 | 1.68 | 18 | 0 | 100 | 78.11 | 659.8 | 0.71 | -26 | 0 | 100 | 182.91 | 1545 | 1.51 | -26.1 | 0 | 100 | 431.65 | 3646.2 | 3.41
-7.5 | 0 | 100 | 116.72 | 1514.9 | 2.03 | 17.5 | 0 | 100 | 80.81 | 682.6 | 0.74 | -25 | 0 | 100 | 196.83 | 1662.6 | 1.63 | -25.1 | 0 | 100 | 463.89 | 3918.5 | 3.67
-7 | 0 | 100 | 134.31 | 1743.2 | 2.33 | 17 | 0 | 100 | 82.8 | 699.4 | 0.76 | -24 | 0 | 100 | 208.07 | 1757.6 | 1.72 | -24 | 0 | 100 | 509.22 | 4301.4 | 4.03
-6.6 | 0 | 100 | 158.31 | 2054.7 | 2.75 | 16.5 | 0 | 100 | 91.61 | 773.8 | 0.84 | -23 | 0 | 100 | 217.29 | 1835.5 | 1.8 | -23.1 | 0 | 100 | 555.09 | 4688.9 | 4.39
-6 | 0 | 100 | 206.94 | 2685.9 | 3.59 | 16 | 0 | 100 | 88.95 | 751.4 | 0.81 | -22 | 0 | 100 | 237.47 | 2006 | 1.96 | -22 | 0 | 100 | 620.71 | 5243.2 | 4.91
-5.5 | 0 | 100 | 260.45 | 3380.3 | 4.52 | 15.5 | 0 | 100 | 97.15 | 820.6 | 0.89 | -21 | 0 | 100 | 255.7 | 2159.9 | 2.11 | -21 | 0 | 100 | 698.78 | 5902.7 | 5.53
-5 | 0 | 100 | 362.46 | 4704.3 | 6.3 | 15 | 0 | 100 | 98.78 | 834.4 | 0.9 | -20 | 0 | 100 | 287.85 | 2431.5 | 2.38 | -20.1 | 0 | 100 | 808.12 | 6826.3 | 6.39
-4.5 | 0 | 100 | 499.22 | 6479.3 | 8.65 | 14.5 | 0 | 100 | 105.25 | 889.1 | 0.96 | -19 | 0 | 100 | 308.09 | 2602.5 | 2.55 | -19.1 | 0 | 100 | 978.23 | 8263.2 | 7.74
-4.1 | 0 | 100 | 733.22 | 9516.4 | 12.69 | 14 | 0 | 100 | 106.81 | 902.2 | 0.97 | -18 | 0 | 100 | 349.48 | 2952.1 | 2.89 | -18.1 | 0 | 100 | 1250.76 | 10565.3 | 9.9
-3.4 | 0 | 100 | 1329.75 | 17258.7 | 23.1 | 13.5 | 0 | 100 | 122.44 | 1034.3 | 1.12 | -17 | 0 | 100 | 395.66 | 3342.1 | 3.27 | -17 | 0 | 100 | 1785.1 | 15078.9 | 14.13
-3 | 0 | 100 | 1917.02 | 24880.8 | 33.25 | 13 | 0 | 100 | 124.8 | 1054.2 | 1.14 | -16 | 0 | 100 | 471.45 | 3982.3 | 3.9 | -16 | 0 | 100 | 3143.7 | 26555.1 | 24.89
-2.5 | 0 | 100 | 2850.84 | 37000.9 | 49.32 | 12.5 | 0 | 100 | 137.17 | 1158.6 | 1.25 | -15 | 0 | 100 | 565.45 | 4776.4 | 4.68 | -15.1 | 0 | 100 | 6234.12 | 52660.2 | 49.32
-2 | 0 | 100 | 3856.07 | 50047.6 | 66.9 | 12 | 0 | 100 | 161.03 | 1360.3 | 1.47 | -14 | 0 | 100 | 716.66 | 6053.7 | 5.93 | -14 | 0 | 100 | 9561.94 | 80770.6 | 75.65
-1.5 | 0 | 100 | 4777.49 | 62006.6 | 82.48 | 11.5 | 0 | 100 | 177.43 | 1498.7 | 1.62 | -13 | 0 | 100 | 970.4 | 8197.1 | 8.03 | -13.1 | 0 | 100 | 11156.14 | 94236.9 | 88.23
-1 | 0 | 100 | 5432.54 | 70508.5 | 93.71 | 11 | 0 | 100 | 192.47 | 1625.8 | 1.76 | -12 | 0 | 100 | 1456.6 | 12304 | 12.05 | -12.1 | 0 | 100 | 11728.36 | 99070.5 | 92.85
-0.5 | 0 | 100 | 5742.58 | 74532.4 | 99.12 | 10.5 | 0 | 100 | 230.15 | 1944.1 | 2.1 | -11 | 0 | 100 | 2793.83 | 23599.7 | 23.1 | -11 | 0 | 100 | 12060.13 | 101873 | 95.46
0 | 0 | 100 | 5802.08 | 75304.7 | 100.04 | 10 | 0 | 100 | 254.87 | 2152.9 | 2.33 | -10.1 | 0 | 100 | 5656 | 47776.7 | 46.78 | -10.1 | 0 | 100 | 12243.9 | 103425.3 | 96.92
0.6 | 0 | 100 | 5644.8 | 73263.4 | 97.56 | 9.5 | 0 | 100 | 299.51 | 2530 | 2.74 | -9.1 | 0 | 100 | 9033.92 | 76310.3 | 74.76 | -9.1 | 0 | 100 | 12340.32 | 104239.8 | 97.67
1 | 0 | 100 | 5287.35 | 68624.1 | 91.58 | 9 | 0 | 100 | 354.01 | 2990.3 | 3.23 | -8 | 0 | 100 | 10792.29 | 91163.5 | 89.33 | -8 | 0 | 100 | 12445.44 | 105127.8 | 98.5
1.5 | 0 | 100 | 4471.65 | 58037.2 | 77.52 | 8.5 | 0 | 100 | 429.12 | 3624.9 | 3.92 | -7.1 | 0 | 100 | 11380.42 | 96131.4 | 94.21 | -7 | 0 | 100 | 12508.24 | 105658.2 | 99.03
2 | 0 | 100 | 3472.06 | 45063.6 | 59.96 | 8 | 0 | 100 | 524.09 | 4427 | 4.79 | -5.9 | 0 | 100 | 11690.25 | 98748.6 | 96.76 | -6 | 0 | 100 | 12565.22 | 106139.6 | 99.41
2.5 | 0 | 100 | 2359.08 | 30618.3 | 40.7 | 7.5 | 0 | 100 | 660.51 | 5579.4 | 6.03 | -4.9 | 0 | 100 | 11856.62 | 100154 | 98.15 | -5.1 | 0 | 100 | 12580.58 | 106269.3 | 99.53
3 | 0 | 100 | 1521.5 | 19747.4 | 26.29 | 7 | 0 | 100 | 874.68 | 7388.5 | 7.99 | -4 | 0 | 100 | 11987.23 | 101257.2 | 99.17 | -4 | 0 | 100 | 12608.71 | 106506.9 | 99.77
3.5 | 0 | 100 | 890.33 | 11555.5 | 15.36 | 6.5 | 0 | 100 | 1165.61 | 9846 | 10.66 | -3.1 | 0 | 100 | 12033.42 | 101647.4 | 99.57 | -3.1 | 0 | 100 | 12623.63 | 106632.9 | 99.92
4 | 0 | 100 | 564.82 | 7330.7 | 9.78 | 6 | 0 | 100 | 1597.97 | 13498.2 | 14.61 | -2 | 0 | 100 | 12065.59 | 101919.1 | 99.88 | -2 | 0 | 100 | 12641.55 | 106784.3 | 100.05
4.5 | 0 | 100 | 401.14 | 5206.4 | 6.93 | 5.5 | 0 | 100 | 2427.12 | 20502.1 | 22.17 | -1.1 | 0 | 100 | 12104.83 | 102250.6 | 100.18 | -1.1 | 0 | 100 | 12636.34 | 106740.3 | 100
5 | 0 | 100 | 291.24 | 3780 | 5.03 | 5 | 0 | 100 | 3772.34 | 31865.3 | 34.46 | 0 | 0 | 100 | 12087.29 | 102102.4 | 99.99 | 0 | 0 | 100 | 12635.64 | 106734.4 | 100
5.4 | 0 | 100 | 221.85 | 2879.4 | 3.83 | 4.6 | 0 | 100 | 5388.47 | 45516.9 | 49.25 | 1 | 0 | 100 | 12063.37 | 101900.4 | 99.83 | 1.1 | 0 | 100 | 12618.01 | 106585.5 | 99.93
6 | 0 | 100 | 164.01 | 2128.7 | 2.84 | 4.1 | 0 | 100 | 7030.9 | 59390.6 | 64.27 | 2 | 0 | 100 | 12062.91 | 101896.5 | 99.79 | 2 | 0 | 100 | 12606.32 | 106486.7 | 99.77
6.5 | 0 | 100 | 131.5 | 1706.7 | 2.28 | 3.5 | 0 | 100 | 8644.18 | 73018.2 | 78.97 | 3 | 0 | 100 | 11977.76 | 101177.2 | 99.16 | 3 | 0 | 100 | 12586.76 | 106321.5 | 99.67
7 | 0 | 100 | 103.39 | 1341.9 | 1.79 | 3 | 0 | 100 | 9572.27 | 80857.8 | 87.47 | 4 | 0 | 100 | 11919.25 | 100683 | 98.69 | 4.1 | 0 | 100 | 12558.92 | 106086.3 | 99.44
7.5 | 0 | 100 | 84.03 | 1090.6 | 1.45 | 2.6 | 0 | 100 | 10165.77 | 85871.1 | 92.86 | 5 | 0 | 100 | 11801.6 | 99689.2 | 97.64 | 5 | 0 | 100 | 12528.97 | 105833.3 | 99.21
8 | 0 | 100 | 71.77 | 931.5 | 1.24 | 2 | 0 | 100 | 10535.72 | 88996.2 | 96.36 | 6 | 0 | 100 | 11556.19 | 97616.2 | 95.61 | 6 | 0 | 100 | 12502.21 | 105607.3 | 98.97
8.5 | 0 | 100 | 63.14 | 819.5 | 1.09 | 1.5 | 0 | 100 | 10759.03 | 90882.5 | 98.43 | 7 | 0 | 100 | 11104.11 | 93797.5 | 91.9 | 7 | 0 | 100 | 12435.05 | 105040 | 98.52
9 | 0 | 100 | 54.7 | 710 | 0.94 | 1 | 0 | 100 | 10868.21 | 91804.7 | 99.24 | 8 | 0 | 100 | 9977.81 | 84283.5 | 82.55 | 8 | 0 | 100 | 12375.73 | 104538.9 | 98.01
9.5 | 0 | 100 | 51.78 | 672 | 0.89 | 0.6 | 0 | 100 | 10894.97 | 92030.8 | 99.55 | 9 | 0 | 100 | 7523.5 | 63551.7 | 62.27 | 9 | 0 | 100 | 12252.23 | 103495.6 | 97.03
10 | 0 | 100 | 41.03 | 532.6 | 0.71 | -0.1 | 0 | 100 | 10956.36 | 92549.4 | 100.07 | 10 | 0 | 100 | 4009.49 | 33868.5 | 33.18 | 10 | 0 | 100 | 12094.59 | 102164.1 | 95.72
[6 empty] | -0.5 | 0 | 100 | 10957.86 | 92562 | 100.01 | 11 | 0 | 100 | 1893.7 | 15996.2 | 15.67 | 11 | 0 | 100 | 11840.54 | 100018.1 | 93.77
[6 empty] | -1 | 0 | 100 | 10903.44 | 92102.3 | 99.58 | 12 | 0 | 100 | 1152.24 | 9733.1 | 9.53 | 12 | 0 | 100 | 11376.23 | 96096.1 | 90.1
[6 empty] | -1.5 | 0 | 100 | 10821.78 | 91412.6 | 98.9 | 13.1 | 0 | 100 | 803.36 | 6786.1 | 6.65 | 13 | 0 | 100 | 10275.13 | 86794.9 | 81.37
Measurement time: | 27/05/2023 10:12:41 |  |  |  |  | -2 | 0 | 100 | 10709.65 | 90465.4 | 97.82 | 14 | 0 | 100 | 606.85 | 5126.2 | 5.02 | 14 | 0 | 100 | 7744.08 | 65414.9 | 61.3
Radiation device: | Novalis TB Accelerator |  |  |  |  | -2.5 | 0 | 100 | 10512.66 | 88801.4 | 96.14 | 15 | 0 | 100 | 493.33 | 4167.2 | 4.08 | 15 | 0 | 100 | 4026.55 | 34012.6 | 31.89
Energy: | 6 MV (FFF) |  |  |  |  | -2.9 | 0 | 100 | 10213.84 | 86277.2 | 93.29 | 16 | 0 | 100 | 415.37 | 3508.7 | 3.44 | 15.9 | 0 | 100 | 2213.7 | 18699.3 | 17.54
Controller: | CCU |  |  |  |  | -3.5 | 0 | 100 | 9658.49 | 81586.1 | 88.19 | 17 | 0 | 100 | 352.88 | 2980.8 | 2.92 | 17 | 0 | 100 | 1420.86 | 12002.1 | 11.25
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | -4 | 0 | 100 | 8519.2 | 71962.4 | 77.85 | 18 | 0 | 100 | 306.94 | 2592.8 | 2.54 | 17.9 | 0 | 100 | 1062.5 | 8975 | 8.41
Field detector: | PTW60019 Diamant |  |  |  |  | -4.5 | 0 | 100 | 7065.97 | 59686.9 | 64.54 | 19 | 0 | 100 | 285.56 | 2412.2 | 2.36 | 18.9 | 0 | 100 | 851.35 | 7191.5 | 6.74
Reference detector: | Stealth CI transmission |  |  |  |  | -5 | 0 | 100 | 5316.07 | 44905.3 | 48.51 | 20 | 0 | 100 | 260.04 | 2196.6 | 2.15 | 19.9 | 0 | 100 | 727.83 | 6148 | 5.76
SAD: | 1000 mm |  |  |  |  | -5.5 | 0 | 100 | 3759.7 | 31758.5 | 34.33 | 21.1 | 0 | 100 | 229.25 | 1936.5 | 1.9 | 21 | 0 | 100 | 625.96 | 5287.6 | 4.96
SSD: | 900 mm |  |  |  |  | -6 | 0 | 100 | 2394.97 | 20230.5 | 21.87 | 21.9 | 0 | 100 | 220.06 | 1858.9 | 1.82 | 21.9 | 0 | 100 | 560.11 | 4731.3 | 4.44
Field size: | 5 x 5 mm |  |  |  |  | -6.4 | 0 | 100 | 1625.12 | 13727.5 | 14.85 | 23 | 0 | 100 | 209.83 | 1772.4 | 1.74 | 23 | 0 | 100 | 514.2 | 4343.5 | 4.07
Gantry angle: | 0 |  |  |  |  | -7 | 0 | 100 | 1106.42 | 9346 | 10.11 | 23.9 | 0 | 100 | 195.49 | 1651.3 | 1.62 | 24 | 0 | 100 | 469.64 | 3967.1 | 3.72
Measurement medium: | Water |  |  |  |  | -7.5 | 0 | 100 | 824.88 | 6967.8 | 7.53 | 25 | 0 | 100 | 182.35 | 1540.3 | 1.51 | 25 | 0 | 100 | 425.67 | 3595.7 | 3.37
Scan type: | Crossline |  |  |  |  | -8 | 0 | 100 | 638.47 | 5393.2 | 5.84 | 26 | 0 | 100 | 173.65 | 1466.8 | 1.44 | 26 | 0 | 100 | 399.44 | 3374.1 | 3.16
Scan mode: | StepByStep |  |  |  |  | -8.5 | 0 | 100 | 503.4 | 4252.3 | 4.6 | 27.1 | 0 | 100 | 165 | 1393.8 | 1.37 | 27 | 0 | 100 | 382.76 | 3233.2 | 3.03
Wedge: | none |  |  |  |  | -8.9 | 0 | 100 | 418.02 | 3531.1 | 3.82 | 28 | 0 | 100 | 162.47 | 1372.4 | 1.34 | 27.9 | 0 | 100 | 363.91 | 3074 | 2.88
[6 empty] | -9.5 | 0 | 100 | 336.97 | 2846.4 | 3.08 | 28.9 | 0 | 100 | 166.75 | 1408.5 | 1.38 | 29 | 0 | 100 | 332.24 | 2806.5 | 2.63
Points [mm]: | [5 empty] | -10 | 0 | 100 | 287.8 | 2431.1 | 2.63 | 30 | 0 | 100 | 150.42 | 1270.6 | 1.25 | 30 | 0 | 100 | 318.4 | 2689.6 | 2.52
Inline | Crossline | Depth | Normalized field | Current field | Ratio | -10.4 | 0 | 100 | 250.41 | 2115.2 | 2.29
0 | -10 | 100 | 47.36 | 614.6 | 0.81 | -11 | 0 | 100 | 210.03 | 1774.1 | 1.92
0 | -9.5 | 100 | 56.05 | 727.5 | 0.96 | -11.5 | 0 | 100 | 181.01 | 1529 | 1.65 | Measurement time: | 27/05/2023 10:58:20 |  |  |  |  | Measurement time: | 27/05/2023 11:23:54
0 | -9 | 100 | 63.16 | 819.8 | 1.08 | -12 | 0 | 100 | 164.6 | 1390.4 | 1.51 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
0 | -8.5 | 100 | 70.65 | 917 | 1.21 | -12.5 | 0 | 100 | 153.91 | 1300.1 | 1.41 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
0 | -8 | 100 | 85.68 | 1112.1 | 1.47 | -12.9 | 0 | 100 | 140.96 | 1190.7 | 1.28 | Controller: | CCU |  |  |  |  | Controller: | CCU
0 | -7.4 | 100 | 112.71 | 1462.9 | 1.93 | -13.5 | 0 | 100 | 130.4 | 1101.5 | 1.19 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
0 | -7 | 100 | 139.68 | 1812.9 | 2.39 | -14 | 0 | 100 | 115.96 | 979.6 | 1.06 | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant
0 | -6.5 | 100 | 181.19 | 2351.7 | 3.11 | -14.4 | 0 | 100 | 113.16 | 955.8 | 1.03 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
0 | -6 | 100 | 229.39 | 2977.2 | 3.93 | -15 | 0 | 100 | 109.32 | 923.4 | 1 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
0 | -5.5 | 100 | 326.53 | 4238 | 5.6 | -15.5 | 0 | 100 | 102.12 | 862.7 | 0.93 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
0 | -5 | 100 | 462.81 | 6006.8 | 7.92 | -16 | 0 | 100 | 92.74 | 783.4 | 0.85 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
0 | -4.5 | 100 | 680.4 | 8830.9 | 11.67 | -16.4 | 0 | 100 | 94.81 | 800.8 | 0.87 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
0 | -4 | 100 | 1035.67 | 13441.9 | 17.81 | -17 | 0 | 100 | 86.83 | 733.5 | 0.79 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
0 | -3.5 | 100 | 1525.9 | 19804.6 | 26.15 | -17.5 | 0 | 100 | 75.67 | 639.2 | 0.69 | Scan type: | Crossline |  |  |  |  | Scan type: | Crossline
0 | -3 | 100 | 2265.59 | 29405 | 38.92 | -18 | 0 | 100 | 82.82 | 699.6 | 0.76 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
0 | -2.6 | 100 | 3155.86 | 40959.6 | 54.15 | -18.5 | 0 | 100 | 78.71 | 664.9 | 0.72 | Wedge: | none |  |  |  |  | Wedge: | none
0 | -2 | 100 | 4157.88 | 53964.7 | 71.33 | -19 | 0 | 100 | 71.71 | 605.8 | 0.66
0 | -1.5 | 100 | 4896.35 | 63549.3 | 84.29 | -19.5 | 0 | 100 | 74.84 | 632.2 | 0.68 | Points [mm]: | [5 empty] | Points [mm]:
0 | -1 | 100 | 5431.68 | 70497.4 | 93.34 | -20 | 0 | 100 | 75.1 | 634.4 | 0.69 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -0.5 | 100 | 5701.08 | 73993.9 | 97.77 | [6 empty] | 0 | 30 | 100 | 151.88 | 1283 | 1.26 | 0 | 30 | 100 | 338.42 | 2858.7 | 2.68
0 | 0 | 100 | 5800.98 | 75290.4 | 100.07 | [6 empty] | 0 | 29 | 100 | 159.62 | 1348.3 | 1.32 | 0 | 29.1 | 100 | 358.4 | 3027.4 | 2.84
0 | 0.4 | 100 | 5781.65 | 75039.5 | 99.4 | Measurement time: | 27/05/2023 10:37:20 |  |  |  |  | 0 | 28 | 100 | 165.77 | 1400.3 | 1.37 | 0 | 28 | 100 | 376.09 | 3176.9 | 2.98
0 | 0.9 | 100 | 5632.02 | 73097.5 | 96.47 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | 27 | 100 | 179.85 | 1519.2 | 1.49 | 0 | 27.1 | 100 | 406.96 | 3437.6 | 3.22
0 | 1.4 | 100 | 5375.74 | 69771.3 | 92.33 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | 26 | 100 | 190.31 | 1607.5 | 1.57 | 0 | 26 | 100 | 435.67 | 3680.1 | 3.45
0 | 2 | 100 | 4877.15 | 63300.1 | 83.81 | Controller: | CCU |  |  |  |  | 0 | 25 | 100 | 213.39 | 1802.6 | 1.77 | 0 | 25 | 100 | 474.96 | 4012.1 | 3.76
0 | 2.5 | 100 | 4090.82 | 53094.4 | 70.05 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | 24.1 | 100 | 210.96 | 1782 | 1.75 | 0 | 24 | 100 | 521.18 | 4402.4 | 4.12
0 | 3 | 100 | 3310.46 | 42966.1 | 56.91 | Field detector: | PTW60019 Diamant |  |  |  |  | 0 | 22.9 | 100 | 230.68 | 1948.6 | 1.91 | 0 | 22.9 | 100 | 586.47 | 4954 | 4.64
0 | 3.5 | 100 | 2261.32 | 29349.5 | 38.91 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | 21.9 | 100 | 247.39 | 2089.7 | 2.05 | 0 | 22 | 100 | 634.25 | 5357.5 | 5.02
0 | 4 | 100 | 1530.51 | 19864.3 | 26.21 | SAD: | 1000 mm |  |  |  |  | 0 | 21 | 100 | 274.78 | 2321.1 | 2.27 | 0 | 21 | 100 | 736.26 | 6219.3 | 5.83
0 | 4.6 | 100 | 869.64 | 11287 | 14.92 | SSD: | 900 mm |  |  |  |  | 0 | 19.9 | 100 | 299 | 2525.7 | 2.47 | 0 | 19.9 | 100 | 893.73 | 7549.4 | 7.07
0 | 5.1 | 100 | 526.86 | 6838 | 9.03 | Field size: | 10 x 10 mm |  |  |  |  | 0 | 19 | 100 | 328.77 | 2777.1 | 2.72 | 0 | 19 | 100 | 1091.19 | 9217.4 | 8.64
0 | 5.5 | 100 | 377.07 | 4893.9 | 6.49 | Gantry angle: | 0 |  |  |  |  | 0 | 18 | 100 | 373.56 | 3155.5 | 3.09 | 0 | 18 | 100 | 1492.37 | 12606.2 | 11.81
0 | 5.9 | 100 | 277.16 | 3597.2 | 4.77 | Measurement medium: | Water |  |  |  |  | 0 | 17 | 100 | 432.28 | 3651.5 | 3.58 | 0 | 17 | 100 | 2596.19 | 21930.3 | 20.55
0 | 6.5 | 100 | 197.43 | 2562.4 | 3.39 | Scan type: | Crossline |  |  |  |  | 0 | 16 | 100 | 504.58 | 4262.2 | 4.18 | 0 | 16 | 100 | 4967.12 | 41957.7 | 39.33
0 | 6.9 | 100 | 146.27 | 1898.4 | 2.51 | Scan mode: | StepByStep |  |  |  |  | 0 | 15 | 100 | 637.36 | 5383.8 | 5.28 | 0 | 15 | 100 | 8226.1 | 69486.6 | 65.12
0 | 7.4 | 100 | 121.74 | 1580 | 2.09 | Wedge: | none |  |  |  |  | 0 | 14.1 | 100 | 834.91 | 7052.6 | 6.91 | 0 | 14 | 100 | 10353.5 | 87456.9 | 81.98
0 | 7.9 | 100 | 96.84 | 1256.9 | 1.66 | [6 empty] | 0 | 13.1 | 100 | 1219.63 | 10302.3 | 10.1 | 0 | 13.1 | 100 | 11342.08 | 95807.5 | 89.8
0 | 8.5 | 100 | 80.78 | 1048.4 | 1.39 | Points [mm]: | [5 empty] | 0 | 12 | 100 | 2220.32 | 18755.3 | 18.38 | 0 | 12 | 100 | 11861.22 | 100192.8 | 93.94
0 | 9 | 100 | 66.2 | 859.2 | 1.13 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 11.1 | 100 | 4460.1 | 37674.8 | 36.9 | 0 | 11 | 100 | 12116.81 | 102351.8 | 95.93
0 | 9.5 | 100 | 60.41 | 784.1 | 1.03 | 0 | 20 | 100 | 72.57 | 613 | 0.66 | 0 | 10.1 | 100 | 7643.31 | 64563.7 | 63.28 | 0 | 10 | 100 | 12294.49 | 103852.6 | 97.34
0 | 10 | 100 | 54.14 | 702.7 | 0.93 | 0 | 19.6 | 100 | 71.96 | 607.9 | 0.66 | 0 | 9 | 100 | 9969.48 | 84213.1 | 82.53 | 0 | 9 | 100 | 12390.16 | 104660.8 | 98.11
[6 empty] | 0 | 19 | 100 | 82.01 | 692.7 | 0.75 | 0 | 8.1 | 100 | 11016.91 | 93060.8 | 91.19 | 0 | 8 | 100 | 12459.73 | 105248.5 | 98.66
[6 empty] | 0 | 18.4 | 100 | 81.59 | 689.2 | 0.75 | 0 | 7 | 100 | 11519.36 | 97305 | 95.4 | 0 | 7 | 100 | 12516.93 | 105731.7 | 99.08
[6 empty] | 0 | 17.9 | 100 | 84.62 | 714.8 | 0.77 | 0 | 6.1 | 100 | 11754.95 | 99295.1 | 97.35 | 0 | 6.1 | 100 | 12554.7 | 106050.7 | 99.41
Measurement time: | 27/05/2023 10:14:20 |  |  |  |  | 0 | 17.4 | 100 | 86.7 | 732.3 | 0.79 | 0 | 5 | 100 | 11895.97 | 100486.4 | 98.48 | 0 | 5 | 100 | 12571.41 | 106191.8 | 99.56
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | 17 | 100 | 94.4 | 797.4 | 0.86 | 0 | 4 | 100 | 11993.51 | 101310.3 | 99.27 | 0 | 4 | 100 | 12606.17 | 106485.5 | 99.79
Energy: | 6 MV (FFF) |  |  |  |  | 0 | 16.4 | 100 | 103.76 | 876.4 | 0.95 | 0 | 3 | 100 | 12047.72 | 101768.2 | 99.75 | 0 | 3.1 | 100 | 12612.21 | 106536.5 | 99.89
Controller: | CCU |  |  |  |  | 0 | 15.9 | 100 | 106.34 | 898.3 | 0.97 | 0 | 2 | 100 | 12072.71 | 101979.3 | 99.94 | 0 | 2 | 100 | 12646.7 | 106827.8 | 100.14
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | 15.4 | 100 | 106.65 | 900.9 | 0.97 | 0 | 1 | 100 | 12081.42 | 102052.9 | 99.99 | 0 | 1 | 100 | 12633.67 | 106717.7 | 100.01
Field detector: | PTW60019 Diamant |  |  |  |  | 0 | 15 | 100 | 114.11 | 963.9 | 1.04 | 0 | 0 | 100 | 12075.73 | 102004.8 | 100 | 0 | 0 | 100 | 12629.5 | 106682.5 | 100
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | 14.5 | 100 | 125.42 | 1059.4 | 1.14 | 0 | -0.9 | 100 | 12081.22 | 102051.1 | 99.95 | 0 | -1 | 100 | 12624.95 | 106644.1 | 99.99
SAD: | 1000 mm |  |  |  |  | 0 | 13.9 | 100 | 137.57 | 1162.1 | 1.25 | 0 | -1.9 | 100 | 12085.04 | 102083.4 | 99.96 | 0 | -2 | 100 | 12628.4 | 106673.2 | 99.97
SSD: | 900 mm |  |  |  |  | 0 | 13.5 | 100 | 144.8 | 1223.1 | 1.32 | 0 | -3 | 100 | 12025.75 | 101582.6 | 99.5 | 0 | -2.9 | 100 | 12622.67 | 106624.9 | 99.95
Field size: | 5 x 5 mm |  |  |  |  | 0 | 13 | 100 | 153.63 | 1297.7 | 1.4 | 0 | -4 | 100 | 11949.78 | 100940.9 | 98.89 | 0 | -3.9 | 100 | 12567.55 | 106159.3 | 99.56
Gantry angle: | 0 |  |  |  |  | 0 | 12.4 | 100 | 171.96 | 1452.6 | 1.57 | 0 | -4.9 | 100 | 11882.61 | 100373.4 | 98.41 | 0 | -4.9 | 100 | 12556.29 | 106064.1 | 99.43
Measurement medium: | Water |  |  |  |  | 0 | 11.9 | 100 | 191.17 | 1614.8 | 1.74 | 0 | -6 | 100 | 11711.42 | 98927.4 | 96.97 | 0 | -6 | 100 | 12527.29 | 105819.1 | 99.19
Scan type: | Beam |  |  |  |  | 0 | 11.4 | 100 | 213.55 | 1803.9 | 1.95 | 0 | -6.9 | 100 | 11472.81 | 96911.8 | 94.94 | 0 | -6.9 | 100 | 12462.81 | 105274.5 | 98.74
Scan mode: | StepByStep |  |  |  |  | 0 | 11 | 100 | 246.11 | 2078.9 | 2.25 | 0 | -8 | 100 | 10924.65 | 92281.5 | 90.4 | 0 | -8 | 100 | 12408.76 | 104817.9 | 98.3
Wedge: | none |  |  |  |  | 0 | 10.5 | 100 | 288.1 | 2433.6 | 2.63 | 0 | -9 | 100 | 9393.8 | 79350.3 | 77.76 | 0 | -9 | 100 | 12325.93 | 104118.3 | 97.7
[6 empty] | 0 | 10 | 100 | 350.66 | 2962 | 3.2 | 0 | -10 | 100 | 6456.28 | 54536.8 | 53.41 | 0 | -10 | 100 | 12213.94 | 103172.2 | 96.81
Points [mm]: | [5 empty] | 0 | 9.4 | 100 | 423.92 | 3580.9 | 3.87 | 0 | -11 | 100 | 3489.12 | 29472.9 | 28.89 | 0 | -11 | 100 | 12053.01 | 101812.9 | 95.44
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 9 | 100 | 511.73 | 4322.7 | 4.66 | 0 | -11.9 | 100 | 1858.89 | 15702.2 | 15.39 | 0 | -11.9 | 100 | 11786.93 | 99565.2 | 93.41
0 | 0 | 250 | 2041.77 | 26500 | 17.7 | 0 | 8.5 | 100 | 651.87 | 5506.4 | 5.94 | 0 | -13 | 100 | 1113.09 | 9402.4 | 9.21 | 0 | -12.9 | 100 | 11230.69 | 94866.6 | 89
0 | 0 | 245 | 2102.98 | 27294.4 | 18.28 | 0 | 8 | 100 | 852.27 | 7199.2 | 7.77 | 0 | -13.9 | 100 | 770.08 | 6504.9 | 6.37 | 0 | -14 | 100 | 9614.77 | 81216.8 | 76.15
0 | 0 | 240 | 2171.95 | 28189.5 | 18.97 | 0 | 7.5 | 100 | 1190.31 | 10054.7 | 10.85 | 0 | -15 | 100 | 584.81 | 4939.9 | 4.84 | 0 | -14.9 | 100 | 6808.77 | 57514.3 | 53.93
0 | 0 | 235.1 | 2245.66 | 29146.2 | 19.61 | 0 | 7 | 100 | 1736.82 | 14671.1 | 15.83 | 0 | -15.9 | 100 | 477.78 | 4035.9 | 3.95 | 0 | -15.9 | 100 | 3724.47 | 31460.9 | 29.52
0 | 0 | 230 | 2318.35 | 30089.7 | 20.24 | 0 | 6.5 | 100 | 2618.92 | 22122.3 | 23.86 | 0 | -17 | 100 | 396.25 | 3347.2 | 3.28 | 0 | -17 | 100 | 2086.08 | 17621.3 | 16.54
0 | 0 | 225 | 2401.59 | 31170 | 21 | 0 | 6 | 100 | 4010.56 | 33877.5 | 36.57 | 0 | -17.9 | 100 | 350.53 | 2961 | 2.9 | 0 | -17.9 | 100 | 1379.3 | 11651.1 | 10.93
0 | 0 | 220.1 | 2481.16 | 32202.8 | 21.66 | 0 | 5.5 | 100 | 5591.35 | 47230.7 | 51 | 0 | -18.9 | 100 | 316.87 | 2676.6 | 2.62 | 0 | -18.9 | 100 | 1019.57 | 8612.4 | 8.08
0 | 0 | 215.1 | 2566.02 | 33304.2 | 22.39 | 0 | 5 | 100 | 7271.25 | 61420.9 | 66.29 | 0 | -20 | 100 | 282.44 | 2385.8 | 2.34 | 0 | -20 | 100 | 815.68 | 6890.1 | 6.46
0 | 0 | 210.1 | 2650.7 | 34403.2 | 23.12 | 0 | 4.5 | 100 | 8413.67 | 71071 | 76.87 | 0 | -21 | 100 | 251.97 | 2128.4 | 2.09 | 0 | -21 | 100 | 685.5 | 5790.5 | 5.43
0 | 0 | 205.1 | 2747.65 | 35661.5 | 23.95 | 0 | 4 | 100 | 9348.24 | 78965.4 | 85.24 | 0 | -22 | 100 | 236.2 | 1995.2 | 1.96 | 0 | -21.9 | 100 | 606.1 | 5119.7 | 4.8
0 | 0 | 200 | 2833.62 | 36777.3 | 24.62 | 0 | 3.5 | 100 | 9903.18 | 83653.1 | 90.32 | 0 | -22.9 | 100 | 216.76 | 1831 | 1.8 | 0 | -23 | 100 | 537.04 | 4536.4 | 4.25
0 | 0 | 195 | 2938.18 | 38134.4 | 25.59 | 0 | 3 | 100 | 10266.42 | 86721.4 | 93.72 | 0 | -23.9 | 100 | 205 | 1731.6 | 1.7 | 0 | -24 | 100 | 489.28 | 4133 | 3.88
0 | 0 | 190.1 | 3039.07 | 39443.9 | 26.53 | 0 | 2.5 | 100 | 10584.09 | 89404.8 | 96.56 | 0 | -25 | 100 | 196.8 | 1662.4 | 1.63 | 0 | -24.9 | 100 | 453.59 | 3831.5 | 3.59
0 | 0 | 185 | 3130.61 | 40631.9 | 27.24 | 0 | 2 | 100 | 10745.41 | 90767.4 | 98.06 | 0 | -25.9 | 100 | 181.5 | 1533.1 | 1.5 | 0 | -26 | 100 | 413.88 | 3496.1 | 3.28
0 | 0 | 180.1 | 3243.25 | 42093.9 | 28.28 | 0 | 1.6 | 100 | 10834.82 | 91522.7 | 98.98 | 0 | -27 | 100 | 173.62 | 1466.6 | 1.44 | 0 | -26.9 | 100 | 395.23 | 3338.6 | 3.13
0 | 0 | 175 | 3367.49 | 43706.4 | 29.38 | 0 | 1 | 100 | 10899.04 | 92065.2 | 99.43 | 0 | -28 | 100 | 162.78 | 1375 | 1.35 | 0 | -27.9 | 100 | 369.53 | 3121.5 | 2.93
0 | 0 | 170.1 | 3481.02 | 45179.9 | 30.39 | 0 | 0.5 | 100 | 10947.68 | 92476.1 | 99.9 | 0 | -29 | 100 | 147.75 | 1248.1 | 1.22 | 0 | -29 | 100 | 342.38 | 2892.1 | 2.71
0 | 0 | 165.1 | 3610.49 | 46860.3 | 31.45 | 0 | 0 | 100 | 10969.48 | 92660.1 | 100 | 0 | -30 | 100 | 141.02 | 1191.2 | 1.17 | 0 | -30 | 100 | 321.71 | 2717.5 | 2.55
0 | 0 | 160.1 | 3735.08 | 48477.3 | 32.42 | 0 | -0.5 | 100 | 10936.56 | 92382.1 | 99.85
0 | 0 | 155.1 | 3874.53 | 50287.2 | 33.59 | 0 | -1 | 100 | 10891.17 | 91998.7 | 99.31
0 | 0 | 150 | 4011.27 | 52062 | 34.87 | 0 | -1.5 | 100 | 10802.05 | 91245.9 | 98.46 | Measurement time: | 27/05/2023 11:00:52 |  |  |  |  | Measurement time: | 27/05/2023 11:26:24
0 | 0 | 145 | 4164.81 | 54054.7 | 36.06 | 0 | -2 | 100 | 10726.6 | 90608.5 | 97.84 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
0 | 0 | 140.1 | 4307.2 | 55902.8 | 37.34 | 0 | -2.5 | 100 | 10568.52 | 89273.2 | 96.26 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
0 | 0 | 135.1 | 4472.71 | 58050.9 | 38.94 | 0 | -3 | 100 | 10238.92 | 86489.1 | 93.3 | Controller: | CCU |  |  |  |  | Controller: | CCU
0 | 0 | 130 | 4628.57 | 60073.8 | 40.23 | 0 | -3.5 | 100 | 9600.94 | 81100 | 87.53 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | 125.1 | 4808.12 | 62404.3 | 41.82 | 0 | -4 | 100 | 8793.4 | 74278.7 | 80.17 | Field detector: | PTW60019 Diamant |  |  |  |  | Field detector: | PTW60019 Diamant
0 | 0 | 120.1 | 4987.49 | 64732.2 | 43.33 | 0 | -4.5 | 100 | 7413.78 | 62624.8 | 67.55 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
0 | 0 | 115.1 | 5179.67 | 67226.5 | 45.08 | 0 | -5 | 100 | 5770.67 | 48745.4 | 52.58 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
0 | 0 | 110.1 | 5361.71 | 69589.2 | 46.6 | 0 | -5.5 | 100 | 4232.6 | 35753.2 | 38.64 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
0 | 0 | 105.1 | 5566.02 | 72240.9 | 48.28 | 0 | -6 | 100 | 2968.08 | 25071.7 | 27.06 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
0 | 0 | 100 | 5767.54 | 74856.5 | 50.1 | 0 | -6.5 | 100 | 2063.8 | 17433.1 | 18.82 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
0 | 0 | 95.1 | 5996.28 | 77825.2 | 52.22 | 0 | -6.9 | 100 | 1418.62 | 11983.2 | 12.92 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
0 | 0 | 90 | 6220.37 | 80733.7 | 54.24 | 0 | -7.5 | 100 | 981.39 | 8289.9 | 8.95 | Scan type: | Beam |  |  |  |  | Scan type: | Beam
0 | 0 | 85.1 | 6469.11 | 83962.1 | 56.26 | 0 | -8.1 | 100 | 729.32 | 6160.6 | 6.65 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
0 | 0 | 80 | 6710.1 | 87089.9 | 58.11 | 0 | -8.5 | 100 | 571.03 | 4823.5 | 5.21 | Wedge: | none |  |  |  |  | Wedge: | none
0 | 0 | 75 | 6975.03 | 90528.3 | 60.61 | 0 | -8.9 | 100 | 455.51 | 3847.8 | 4.15
0 | 0 | 70 | 7251.27 | 94113.6 | 62.79 | 0 | -9.4 | 100 | 377.52 | 3188.9 | 3.44 | Points [mm]: | [5 empty] | Points [mm]:
0 | 0 | 65 | 7531.6 | 97752 | 65.29 | 0 | -10 | 100 | 312.61 | 2640.7 | 2.85 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 60 | 7843.2 | 101796.3 | 68.39 | 0 | -10.4 | 100 | 255.83 | 2161 | 2.34 | 0 | 0 | 250 | 4352.31 | 36764.4 | 19.51 | 0 | 0 | 250 | 4567.21 | 38579.7 | 20.09
0 | 0 | 55 | 8127.34 | 105484 | 70.47 | 0 | -11 | 100 | 222.57 | 1880.1 | 2.03 | 0 | 0 | 245 | 4497.78 | 37993.2 | 20.17 | 0 | 0 | 245.1 | 4715.48 | 39832.1 | 20.73
0 | 0 | 49.9 | 8460.37 | 109806.4 | 73.31 | 0 | -11.5 | 100 | 188.51 | 1592.4 | 1.72 | 0 | 0 | 240 | 4645.93 | 39244.6 | 20.83 | 0 | 0 | 240 | 4871.62 | 41151 | 21.42
0 | 0 | 44.9 | 8789.6 | 114079.4 | 76.28 | 0 | -11.9 | 100 | 160.77 | 1358 | 1.47 | 0 | 0 | 235 | 4799.86 | 40544.9 | 21.53 | 0 | 0 | 235.1 | 5022.58 | 42426.2 | 22.09
0 | 0 | 39.9 | 9141.47 | 118646.4 | 79.44 | 0 | -12.5 | 100 | 151 | 1275.5 | 1.38 | 0 | 0 | 230.1 | 4949.08 | 41805.4 | 22.19 | 0 | 0 | 230 | 5205.61 | 43972.3 | 22.9
0 | 0 | 34.9 | 9524.35 | 123615.8 | 83.26 | 0 | -12.9 | 100 | 141.81 | 1197.9 | 1.29 | 0 | 0 | 225 | 5104.27 | 43116.2 | 22.89 | 0 | 0 | 225.1 | 5369.54 | 45357 | 23.63
0 | 0 | 29.9 | 9897.09 | 128453.5 | 86.13 | 0 | -13.5 | 100 | 125.57 | 1060.7 | 1.15 | 0 | 0 | 220.1 | 5297.96 | 44752.3 | 23.75 | 0 | 0 | 220.1 | 5552.18 | 46899.8 | 24.41
0 | 0 | 25.1 | 10301.59 | 133703.5 | 90.06 | 0 | -14 | 100 | 121.94 | 1030.1 | 1.11 | 0 | 0 | 215.1 | 5465.56 | 46168.1 | 24.52 | 0 | 0 | 215 | 5742.7 | 48509.1 | 25.26
0 | 0 | 24.1 | 10358.98 | 134448.3 | 90.51 | 0 | -14.5 | 100 | 115.85 | 978.6 | 1.06 | 0 | 0 | 210.1 | 5654.48 | 47763.9 | 25.35 | 0 | 0 | 210 | 5917.15 | 49982.7 | 26.03
0 | 0 | 23.1 | 10465.61 | 135832.2 | 91.26 | 0 | -15 | 100 | 98.49 | 832 | 0.9 | 0 | 0 | 205.1 | 5838.19 | 49315.7 | 26.18 | 0 | 0 | 205.1 | 6128.3 | 51766.3 | 26.97
0 | 0 | 22.1 | 10522.85 | 136575.2 | 91.43 | 0 | -15.5 | 100 | 97.28 | 821.7 | 0.89 | 0 | 0 | 200.1 | 6025.44 | 50897.4 | 27.03 | 0 | 0 | 200 | 6333.85 | 53502.6 | 27.86
0 | 0 | 21 | 10613.55 | 137752.3 | 92.5 | 0 | -15.9 | 100 | 95.28 | 804.8 | 0.87 | 0 | 0 | 195 | 6255.02 | 52836.7 | 28.06 | 0 | 0 | 195 | 6558.37 | 55399.1 | 28.85
0 | 0 | 20.1 | 10697.67 | 138844.1 | 93.19 | 0 | -16.5 | 100 | 97.04 | 819.7 | 0.89 | 0 | 0 | 190.1 | 6461 | 54576.6 | 28.97 | 0 | 0 | 190.1 | 6772.9 | 57211.3 | 29.79
0 | 0 | 19 | 10789.65 | 140037.9 | 94.01 | 0 | -17 | 100 | 82.85 | 699.8 | 0.76 | 0 | 0 | 185 | 6683.61 | 56457.1 | 29.98 | 0 | 0 | 185.1 | 6988.05 | 59028.7 | 30.74
0 | 0 | 18 | 10875.78 | 141155.8 | 95.12 | 0 | -17.5 | 100 | 79.47 | 671.3 | 0.73 | 0 | 0 | 180 | 6912.66 | 58391.9 | 31.02 | 0 | 0 | 180.1 | 7244.52 | 61195.1 | 31.88
0 | 0 | 17 | 10955.38 | 142189 | 95.19 | 0 | -18 | 100 | 72.48 | 612.2 | 0.66 | 0 | 0 | 175 | 7140.61 | 60317.4 | 32.04 | 0 | 0 | 175.1 | 7486.14 | 63236.1 | 32.95
0 | 0 | 16.1 | 11047.24 | 143381.2 | 96.55 | 0 | -18.5 | 100 | 75.42 | 637.1 | 0.69 | 0 | 0 | 170.1 | 7393.65 | 62454.8 | 33.16 | 0 | 0 | 170.1 | 7755.56 | 65511.9 | 34.12
0 | 0 | 15 | 11116.86 | 144284.9 | 96.83 | 0 | -19 | 100 | 77.73 | 656.6 | 0.71 | 0 | 0 | 165.1 | 7656.36 | 64674 | 34.35 | 0 | 0 | 165.1 | 8025.91 | 67795.6 | 35.32
0 | 0 | 14 | 11185.79 | 145179.4 | 97.4 | 0 | -19.4 | 100 | 66.59 | 562.5 | 0.61 | 0 | 0 | 160.1 | 7928.62 | 66973.8 | 35.57 | 0 | 0 | 160.1 | 8286.23 | 69994.5 | 36.45
0 | 0 | 13.1 | 11286.23 | 146483.1 | 98.51 | 0 | -20 | 100 | 65.87 | 556.4 | 0.6 | 0 | 0 | 155.1 | 8206.1 | 69317.7 | 36.82 | 0 | 0 | 155.1 | 8593.54 | 72590.4 | 37.81
0 | 0 | 12 | 11346.92 | 147270.8 | 98.64 | [6 empty] | 0 | 0 | 150.1 | 8491.5 | 71728.4 | 38.09 | 0 | 0 | 150.1 | 8882.5 | 75031.3 | 39.09
0 | 0 | 11 | 11404.63 | 148019.8 | 99.37 | [6 empty] | 0 | 0 | 145.1 | 8773.12 | 74107.4 | 39.36 | 0 | 0 | 145.1 | 9204.44 | 77750.7 | 40.48
0 | 0 | 10 | 11456.19 | 148688.9 | 100 | Measurement time: | 27/05/2023 10:40:15 |  |  |  |  | 0 | 0 | 140.1 | 9107.87 | 76935 | 40.86 | 0 | 0 | 140 | 9535.39 | 80546.3 | 41.94
0 | 0 | 9 | 11468.93 | 148854.3 | 99.37 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | 0 | 135 | 9428.28 | 79641.5 | 42.3 | 0 | 0 | 135.1 | 9853.46 | 83233.1 | 43.36
0 | 0 | 8 | 11447.09 | 148570.9 | 99.73 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | 0 | 130.1 | 9763.45 | 82472.7 | 43.82 | 0 | 0 | 130.1 | 10211.16 | 86254.6 | 44.94
0 | 0 | 7 | 11399.99 | 147959.5 | 98.9 | Controller: | CCU |  |  |  |  | 0 | 0 | 125.1 | 10102.28 | 85334.9 | 45.34 | 0 | 0 | 125.1 | 10586.27 | 89423.1 | 46.6
0 | 0 | 6 | 11288.43 | 146511.7 | 98.05 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | 0 | 120.1 | 10485.5 | 88571.9 | 47.04 | 0 | 0 | 120.1 | 10949.35 | 92490.1 | 48.17
0 | 0 | 5 | 11065.76 | 143621.5 | 96.26 | Field detector: | PTW60019 Diamant |  |  |  |  | 0 | 0 | 115 | 10854.07 | 91685.3 | 48.68 | 0 | 0 | 115 | 11356.22 | 95927 | 50
0 | 0 | 4 | 10651.02 | 138238.7 | 92.54 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | 0 | 110 | 11233.44 | 94889.8 | 50.43 | 0 | 0 | 110.1 | 11741.9 | 99184.9 | 51.69
0 | 0 | 3 | 9971.7 | 129421.9 | 86.86 | SAD: | 1000 mm |  |  |  |  | 0 | 0 | 105.1 | 11691.03 | 98755.2 | 52.45 | 0 | 0 | 105.1 | 12192.92 | 102994.7 | 53.64
0 | 0 | 2 | 8740.51 | 113442.3 | 76.18 | SSD: | 900 mm |  |  |  |  | 0 | 0 | 100.1 | 12068.58 | 101944.4 | 54.17 | 0 | 0 | 100.1 | 12612.01 | 106534.8 | 55.49
0 | 0 | 1 | 7309.28 | 94866.6 | 63.39 | Field size: | 10 x 10 mm |  |  |  |  | 0 | 0 | 95.1 | 12521.74 | 105772.3 | 56.16 | 0 | 0 | 95.1 | 13091.22 | 110582.7 | 57.6
0 | 0 | 0 | 7280.43 | 94492.1 | 63.33 | Gantry angle: | 0 |  |  |  |  | 0 | 0 | 90.1 | 12969.42 | 109553.9 | 58.19 | 0 | 0 | 90.1 | 13565.62 | 114590 | 59.67
0 | 0 | -0.5 | 7277.92 | 94459.5 | 63.34 | Measurement medium: | Water |  |  |  |  | 0 | 0 | 85 | 13482.85 | 113890.8 | 60.5 | 0 | 0 | 85 | 14047.46 | 118660.1 | 61.85
[6 empty] | Scan type: | Beam |  |  |  |  | 0 | 0 | 80 | 13985.7 | 118138.5 | 62.8 | 0 | 0 | 80.1 | 14550.88 | 122912.6 | 64.08
[6 empty] | Scan mode: | StepByStep |  |  |  |  | 0 | 0 | 75 | 14515.67 | 122615.1 | 65.12 | 0 | 0 | 75 | 15091.33 | 127477.8 | 66.48
[6 empty] | Wedge: | none |  |  |  |  | 0 | 0 | 70 | 15046.98 | 127103.2 | 67.56 | 0 | 0 | 70 | 15661.39 | 132293.2 | 68.96
[12 empty] | 0 | 0 | 65 | 15608.38 | 131845.4 | 70.08 | 0 | 0 | 65 | 16237.96 | 137163.5 | 71.44
[6 empty] | Points [mm]: | [5 empty] | 0 | 0 | 60 | 16192.39 | 136778.6 | 72.67 | 0 | 0 | 60 | 16803.42 | 141940 | 74
[6 empty] | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 0 | 55 | 16781.97 | 141758.8 | 75.33 | 0 | 0 | 55 | 17442.05 | 147334.6 | 76.79
[6 empty] | 0 | 0 | 250 | 3940.16 | 33282.9 | 18.84 | 0 | 0 | 50 | 17456.78 | 147459 | 78.35 | 0 | 0 | 49.9 | 18075.36 | 152684.2 | 79.63
[6 empty] | 0 | 0 | 245 | 4058.79 | 34285 | 19.4 | 0 | 0 | 44.9 | 18105.95 | 152942.6 | 81.26 | 0 | 0 | 45 | 18757.13 | 158443.2 | 82.6
[6 empty] | 0 | 0 | 240 | 4186.39 | 35362.8 | 20.01 | 0 | 0 | 39.9 | 18803.16 | 158832 | 84.38 | 0 | 0 | 39.9 | 19416.32 | 164011.4 | 85.48
[6 empty] | 0 | 0 | 235 | 4334.03 | 36610 | 20.72 | 0 | 0 | 34.9 | 19492.48 | 164654.7 | 87.46 | 0 | 0 | 34.9 | 20103.26 | 169814 | 88.54
[6 empty] | 0 | 0 | 230 | 4456.26 | 37642.4 | 21.33 | 0 | 0 | 29.9 | 20198.31 | 170617 | 90.66 | 0 | 0 | 29.9 | 20816.95 | 175842.7 | 91.64
[6 empty] | 0 | 0 | 225 | 4627.96 | 39092.8 | 22.13 | 0 | 0 | 25.1 | 20907.6 | 176608.3 | 93.87 | 0 | 0 | 25.1 | 21448.91 | 181180.9 | 94.48
[6 empty] | 0 | 0 | 220 | 4759.24 | 40201.8 | 22.77 | 0 | 0 | 24.1 | 21064.68 | 177935.2 | 94.47 | 0 | 0 | 24.1 | 21643.68 | 182826.1 | 95.34
[6 empty] | 0 | 0 | 215 | 4928.91 | 41635 | 23.56 | 0 | 0 | 23.1 | 21210.77 | 179169.3 | 95.16 | 0 | 0 | 23.1 | 21732.07 | 183572.7 | 95.72
[6 empty] | 0 | 0 | 210.1 | 5104.65 | 43119.4 | 24.4 | 0 | 0 | 22 | 21324.14 | 180126.9 | 95.79 | 0 | 0 | 22.1 | 21865 | 184695.6 | 96.36
[6 empty] | 0 | 0 | 205 | 5272.45 | 44536.8 | 25.22 | 0 | 0 | 21.1 | 21485.02 | 181485.9 | 96.43 | 0 | 0 | 21.1 | 22047.36 | 186236 | 97.09
[6 empty] | 0 | 0 | 200 | 5452.95 | 46061.6 | 26.04 | 0 | 0 | 20.1 | 21639.65 | 182792.1 | 97.18 | 0 | 0 | 20.1 | 22156.32 | 187156.4 | 97.55
[6 empty] | 0 | 0 | 195 | 5635.35 | 47602.3 | 26.94 | 0 | 0 | 19 | 21770.89 | 183900.7 | 97.73 | 0 | 0 | 19 | 22303.18 | 188396.9 | 98.23
[6 empty] | 0 | 0 | 190.1 | 5821.42 | 49174.1 | 27.84 | 0 | 0 | 18 | 21886.54 | 184877.5 | 98.24 | 0 | 0 | 18 | 22390.39 | 189133.7 | 98.66
[6 empty] | 0 | 0 | 185 | 6018.02 | 50834.8 | 28.82 | 0 | 0 | 17.1 | 21963.72 | 185529.5 | 98.6 | 0 | 0 | 17.1 | 22490.58 | 189979.9 | 99
[6 empty] | 0 | 0 | 180 | 6231.62 | 52639 | 29.84 | 0 | 0 | 16 | 22065.62 | 186390.3 | 99.13 | 0 | 0 | 16 | 22538.07 | 190381.1 | 99.28
[6 empty] | 0 | 0 | 175.1 | 6449.2 | 54476.9 | 30.86 | 0 | 0 | 15 | 22152.32 | 187122.6 | 99.47 | 0 | 0 | 15 | 22649.64 | 191323.5 | 99.81
[6 empty] | 0 | 0 | 170.1 | 6680.98 | 56434.8 | 31.95 | 0 | 0 | 14.1 | 22215.74 | 187658.3 | 99.77 | 0 | 0 | 14.1 | 22713.77 | 191865.3 | 100
[6 empty] | 0 | 0 | 165.1 | 6899.43 | 58280.1 | 33.01 | 0 | 0 | 13 | 22263.61 | 188062.7 | 100 | 0 | 0 | 13.1 | 22691.8 | 191679.7 | 99.98
[6 empty] | 0 | 0 | 160.1 | 7133.02 | 60253.3 | 34.05 | 0 | 0 | 12 | 22274.29 | 188152.9 | 99.96 | 0 | 0 | 12 | 22634.23 | 191193.3 | 99.74
[6 empty] | 0 | 0 | 155.1 | 7392.67 | 62446.5 | 35.37 | 0 | 0 | 11 | 22210.39 | 187613.2 | 99.72 | 0 | 0 | 11 | 22557.9 | 190548.6 | 99.4
[6 empty] | 0 | 0 | 150.1 | 7652.34 | 64640 | 36.62 | 0 | 0 | 10 | 22040.62 | 186179.1 | 98.97 | 0 | 0 | 10 | 22409.04 | 189291.2 | 98.67
[6 empty] | 0 | 0 | 145.1 | 7917.75 | 66881.9 | 37.86 | 0 | 0 | 9 | 21882.97 | 184847.4 | 98.2 | 0 | 0 | 9 | 22171.32 | 187283.1 | 97.68
[6 empty] | 0 | 0 | 140 | 8236.07 | 69570.8 | 39.35 | 0 | 0 | 8 | 21572.79 | 182227.2 | 96.78 | 0 | 0 | 8 | 21896.91 | 184965.2 | 96.44
[6 empty] | 0 | 0 | 135.1 | 8500.29 | 71802.7 | 40.62 | 0 | 0 | 7 | 21125.13 | 178445.9 | 94.79 | 0 | 0 | 7 | 21402.12 | 180785.7 | 94.3
[6 empty] | 0 | 0 | 130 | 8817.3 | 74480.5 | 42.19 | 0 | 0 | 6 | 20485.48 | 173042.7 | 92 | 0 | 0 | 6 | 20783.24 | 175557.9 | 91.58
[6 empty] | 0 | 0 | 125.1 | 9156.32 | 77344.2 | 43.76 | 0 | 0 | 5 | 19598.06 | 165546.6 | 87.97 | 0 | 0 | 5 | 19881.38 | 167939.8 | 87.56
[6 empty] | 0 | 0 | 120.1 | 9461.91 | 79925.6 | 45.22 | 0 | 0 | 4 | 18447.34 | 155826.3 | 82.88 | 0 | 0 | 4 | 18719.77 | 158127.5 | 82.46
[6 empty] | 0 | 0 | 115.1 | 9822.65 | 82972.8 | 46.92 | 0 | 0 | 3 | 16738.45 | 141391.2 | 75.17 | 0 | 0 | 3 | 17014.9 | 143726.4 | 74.95
[6 empty] | 0 | 0 | 110 | 10197.76 | 86141.4 | 48.73 | 0 | 0 | 2 | 14495.92 | 122448.3 | 65.09 | 0 | 0 | 2 | 14549.96 | 122904.8 | 64.13
[6 empty] | 0 | 0 | 105.1 | 10567.56 | 89265.1 | 50.47 | 0 | 0 | 1 | 12391.12 | 104668.9 | 55.6 | 0 | 0 | 1 | 12711.14 | 107372.1 | 56.01
[6 empty] | 0 | 0 | 100.1 | 10936.53 | 92381.8 | 52.26 | 0 | 0 | 0 | 12385.15 | 104618.5 | 55.61 | 0 | 0 | 0 | 12766.32 | 107838.3 | 56.28
[6 empty] | 0 | 0 | 95.1 | 11337.41 | 95768.1 | 54.2 | 0 | 0 | -0.5 | 12354.2 | 104357 | 55.46 | 0 | 0 | -0.5 | 12697.36 | 107255.7 | 55.93
[6 empty] | 0 | 0 | 90 | 11776.11 | 99473.9 | 56.31
[6 empty] | 0 | 0 | 85 | 12221.93 | 103239.7 | 58.38
[6 empty] | 0 | 0 | 80 | 12678.03 | 107092.4 | 60.59
[6 empty] | 0 | 0 | 75 | 13163.49 | 111193.1 | 62.84
[6 empty] | 0 | 0 | 70 | 13671.4 | 115483.6 | 65.26
[6 empty] | 0 | 0 | 65 | 14160.25 | 119612.9 | 67.63
[6 empty] | 0 | 0 | 60 | 14736.89 | 124483.8 | 70.41
[6 empty] | 0 | 0 | 55 | 15297.12 | 129216.1 | 73.03
[6 empty] | 0 | 0 | 49.9 | 15881.24 | 134150.2 | 75.88
[6 empty] | 0 | 0 | 44.9 | 16504.11 | 139411.7 | 78.88
[6 empty] | 0 | 0 | 39.9 | 17162.27 | 144971.3 | 82.05
[6 empty] | 0 | 0 | 34.9 | 17831.74 | 150626.3 | 85.25
[6 empty] | 0 | 0 | 29.8 | 18510.95 | 156363.6 | 88.53
[6 empty] | 0 | 0 | 25.1 | 19250.56 | 162611.2 | 92.09
[6 empty] | 0 | 0 | 24.1 | 19380.25 | 163706.7 | 92.69
[6 empty] | 0 | 0 | 23 | 19527.54 | 164950.9 | 93.34
[6 empty] | 0 | 0 | 22.1 | 19666.99 | 166128.8 | 94.06
[6 empty] | 0 | 0 | 21 | 19813.48 | 167366.3 | 94.67
[6 empty] | 0 | 0 | 20.1 | 19993.08 | 168883.4 | 95.54
[6 empty] | 0 | 0 | 19 | 20099.32 | 169780.8 | 96.12
[6 empty] | 0 | 0 | 18.1 | 20250.07 | 171054.1 | 96.78
[6 empty] | 0 | 0 | 17 | 20361.72 | 171997.3 | 97.48
[6 empty] | 0 | 0 | 16 | 20488.14 | 173065.1 | 98.09
[6 empty] | 0 | 0 | 15 | 20631.36 | 174275 | 98.71
[6 empty] | 0 | 0 | 14 | 20733.26 | 175135.7 | 99.27
[6 empty] | 0 | 0 | 13 | 20833.26 | 175980.4 | 99.76
[6 empty] | 0 | 0 | 12 | 20884.28 | 176411.4 | 99.96
[6 empty] | 0 | 0 | 11 | 20901.53 | 176557.1 | 100
[6 empty] | 0 | 0 | 10 | 20883.12 | 176401.6 | 100
[6 empty] | 0 | 0 | 9 | 20775.85 | 175495.4 | 99.53
[6 empty] | 0 | 0 | 8 | 20576.48 | 173811.4 | 98.46
[6 empty] | 0 | 0 | 7 | 20307.11 | 171536 | 97.26
[6 empty] | 0 | 0 | 6 | 19825.52 | 167467.9 | 94.93
[6 empty] | 0 | 0 | 5 | 19065.08 | 161044.5 | 91.27
[6 empty] | 0 | 0 | 4 | 17962.42 | 151730.1 | 86.01
[6 empty] | 0 | 0 | 3 | 16440.86 | 138877.4 | 78.71
[6 empty] | 0 | 0 | 2 | 14113.38 | 119217 | 67.44
[6 empty] | 0 | 0 | 1 | 11993.23 | 101307.9 | 57.34
[6 empty] | 0 | 0 | 0 | 11947.06 | 100917.9 | 57.2
[6 empty] | 0 | 0 | -0.5 | 11919.65 | 100686.4 | 57.02

## NanoRazor
Measurement time: | 27/05/2023 12:25:45 |  |  |  |  | Measurement time: | 27/05/2023 12:42:01 |  |  |  |  | Measurement time: | 27/05/2023 13:03:36 |  |  |  |  | Measurement time: | 27/05/2023 13:34:18
Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
Field size: | 5 x 5 mm |  |  |  |  | Field size: | 10 x 10 mm |  |  |  |  | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 23.3 | 135.8 | 1.18 | -20 | 0 | 100 | 68.08 | 396.8 | 2.71 | -30 | 0 | 100 | 97.06 | 565.7 | 3.47 | -30 | 0 | 100 | 134.13 | 781.7 | 4.62
-9.4 | 0 | 100 | 27.27 | 158.9 | 1.39 | -19.5 | 0 | 100 | 68.68 | 400.3 | 2.73 | -29.1 | 0 | 100 | 98.07 | 571.6 | 3.51 | -29 | 0 | 100 | 136.11 | 793.2 | 4.68
-8.9 | 0 | 100 | 39.6 | 230.8 | 2.02 | -18.9 | 0 | 100 | 86.72 | 505.4 | 3.45 | -28.1 | 0 | 100 | 93.04 | 542.2 | 3.32 | -28.1 | 0 | 100 | 154.86 | 902.5 | 5.33
-8.4 | 0 | 100 | 9.34 | 54.4 | 0.48 | -18.5 | 0 | 100 | 100.72 | 587 | 4.01 | -27 | 0 | 100 | 97.4 | 567.7 | 3.48 | -27 | 0 | 100 | 171.72 | 1000.8 | 5.91
-8 | 0 | 100 | 36.96 | 215.4 | 1.88 | -18 | 0 | 100 | 97.29 | 567 | 3.87 | -26 | 0 | 100 | 118.35 | 689.7 | 4.23 | -26 | 0 | 100 | 163.64 | 953.7 | 5.63
-7.5 | 0 | 100 | 67.52 | 393.5 | 3.45 | -17.5 | 0 | 100 | 89.89 | 523.9 | 3.58 | -25.1 | 0 | 100 | 116.51 | 679 | 4.16 | -25.1 | 0 | 100 | 153.83 | 896.5 | 5.3
-7 | 0 | 100 | 77.93 | 454.2 | 3.98 | -17 | 0 | 100 | 92.52 | 539.2 | 3.68 | -24.1 | 0 | 100 | 127.31 | 742 | 4.55 | -24 | 0 | 100 | 150.35 | 876.2 | 5.17
-6.5 | 0 | 100 | 110.63 | 644.7 | 5.63 | -16.5 | 0 | 100 | 99.41 | 579.4 | 3.96 | -23 | 0 | 100 | 136.11 | 793.3 | 4.86 | -23 | 0 | 100 | 163.92 | 955.3 | 5.64
-6 | 0 | 100 | 133.58 | 778.5 | 6.8 | -16 | 0 | 100 | 99.82 | 581.7 | 3.97 | -22.1 | 0 | 100 | 145.83 | 849.9 | 5.21 | -22.1 | 0 | 100 | 175.17 | 1020.9 | 6.03
-5.5 | 0 | 100 | 168.55 | 982.3 | 8.59 | -15.6 | 0 | 100 | 87.61 | 510.6 | 3.48 | -21 | 0 | 100 | 138.1 | 804.8 | 4.94 | -21 | 0 | 100 | 189.76 | 1105.9 | 6.53
-5 | 0 | 100 | 202.38 | 1179.5 | 10.36 | -14.9 | 0 | 100 | 68.46 | 399 | 2.72 | -20 | 0 | 100 | 127.25 | 741.6 | 4.55 | -20 | 0 | 100 | 228.14 | 1329.6 | 7.85
-4.5 | 0 | 100 | 257.98 | 1503.5 | 13.18 | -14.4 | 0 | 100 | 68.27 | 397.9 | 2.72 | -19.1 | 0 | 100 | 132.87 | 774.4 | 4.75 | -19.1 | 0 | 100 | 276.5 | 1611.4 | 9.51
-4 | 0 | 100 | 341.85 | 1992.3 | 17.49 | -14.1 | 0 | 100 | 65.13 | 379.6 | 2.59 | -18 | 0 | 100 | 135.24 | 788.2 | 4.84 | -18 | 0 | 100 | 345.26 | 2012.2 | 11.88
-3.5 | 0 | 100 | 462.27 | 2694.1 | 23.67 | -13.4 | 0 | 100 | 72.68 | 423.6 | 2.89 | -17 | 0 | 100 | 134.41 | 783.4 | 4.8 | -17 | 0 | 100 | 475.41 | 2770.7 | 16.36
-3 | 0 | 100 | 646.48 | 3767.7 | 33.03 | -13 | 0 | 100 | 75.06 | 437.4 | 2.99 | -16 | 0 | 100 | 144.86 | 844.2 | 5.18 | -16.1 | 0 | 100 | 731.18 | 4261.3 | 25.15
-2.4 | 0 | 100 | 972.84 | 5669.7 | 49.49 | -12.6 | 0 | 100 | 71.7 | 417.9 | 2.85 | -15 | 0 | 100 | 173.25 | 1009.7 | 6.2 | -15.1 | 0 | 100 | 1284.84 | 7488 | 44.21
-2 | 0 | 100 | 1226.65 | 7148.9 | 62.34 | -11.9 | 0 | 100 | 71.13 | 414.6 | 2.83 | -14 | 0 | 100 | 234.86 | 1368.8 | 8.4 | -14 | 0 | 100 | 2014.86 | 11742.6 | 69.34
-1.5 | 0 | 100 | 1542.3 | 8988.6 | 78.31 | -11.5 | 0 | 100 | 67.82 | 395.2 | 2.7 | -13 | 0 | 100 | 294.84 | 1718.3 | 10.54 | -13.1 | 0 | 100 | 2462.53 | 14351.7 | 84.74
-1 | 0 | 100 | 1772.41 | 10329.6 | 90.31 | -11 | 0 | 100 | 69.82 | 406.9 | 2.78 | -12 | 0 | 100 | 415.45 | 2421.2 | 14.87 | -12 | 0 | 100 | 2654.21 | 15468.8 | 91.33
-0.6 | 0 | 100 | 1922.48 | 11204.2 | 98.06 | -10.5 | 0 | 100 | 64.5 | 375.9 | 2.57 | -11 | 0 | 100 | 659.5 | 3843.6 | 23.58 | -11.1 | 0 | 100 | 2751.37 | 16035 | 94.71
-0.1 | 0 | 100 | 1966.97 | 11463.6 | 100.09 | -10 | 0 | 100 | 85.06 | 495.7 | 3.39 | -10 | 0 | 100 | 1204.72 | 7021.1 | 43.06 | -10 | 0 | 100 | 2797.99 | 16306.8 | 96.32
0.5 | 0 | 100 | 1943.51 | 11326.8 | 99.18 | -9.5 | 0 | 100 | 101.12 | 589.3 | 4.03 | -9 | 0 | 100 | 1917.39 | 11174.6 | 68.56 | -9 | 0 | 100 | 2831.29 | 16500.8 | 97.46
1 | 0 | 100 | 1864.16 | 10864.3 | 94.52 | -9 | 0 | 100 | 115.62 | 673.9 | 4.6 | -8 | 0 | 100 | 2414.26 | 14070.4 | 86.31 | -8 | 0 | 100 | 2849.25 | 16605.5 | 98.07
1.5 | 0 | 100 | 1669.65 | 9730.7 | 85.05 | -8.5 | 0 | 100 | 138.34 | 806.2 | 5.51 | -7.1 | 0 | 100 | 2578.94 | 15030.1 | 92.24 | -7.1 | 0 | 100 | 2865.28 | 16698.9 | 98.54
2 | 0 | 100 | 1369.06 | 7978.9 | 69.73 | -8 | 0 | 100 | 176.72 | 1029.9 | 7.04 | -6 | 0 | 100 | 2669.22 | 15556.3 | 95.45 | -6.1 | 0 | 100 | 2884.5 | 16810.9 | 99.24
2.5 | 0 | 100 | 1044.55 | 6087.7 | 53.28 | -7.5 | 0 | 100 | 217.48 | 1267.5 | 8.65 | -5 | 0 | 100 | 2716.6 | 15832.4 | 97.11 | -5 | 0 | 100 | 2905.48 | 16933.2 | 99.99
2.9 | 0 | 100 | 735.46 | 4286.3 | 37.52 | -7.1 | 0 | 100 | 279.75 | 1630.4 | 11.15 | -4.1 | 0 | 100 | 2732.65 | 15925.9 | 97.69 | -4.1 | 0 | 100 | 2910.47 | 16962.3 | 100.12
3.4 | 0 | 100 | 504.2 | 2938.5 | 25.69 | -6.5 | 0 | 100 | 381.89 | 2225.7 | 15.2 | -3.1 | 0 | 100 | 2737.86 | 15956.3 | 97.93 | -3 | 0 | 100 | 2909.59 | 16957.1 | 100.07
4.1 | 0 | 100 | 316.78 | 1846.2 | 16.16 | -6 | 0 | 100 | 537.68 | 3133.6 | 21.41 | -2 | 0 | 100 | 2753.36 | 16046.6 | 98.43 | -2 | 0 | 100 | 2905.79 | 16935 | 99.99
4.5 | 0 | 100 | 253.98 | 1480.2 | 12.93 | -5.4 | 0 | 100 | 829.84 | 4836.3 | 33.08 | -1.1 | 0 | 100 | 2778.32 | 16192.1 | 99.34 | -1 | 0 | 100 | 2895.79 | 16876.7 | 99.67
5 | 0 | 100 | 199.65 | 1163.6 | 10.18 | -5 | 0 | 100 | 1064.27 | 6202.6 | 42.34 | -0.1 | 0 | 100 | 2794.04 | 16283.7 | 99.96 | 0 | 0 | 100 | 2904.68 | 16928.5 | 100
5.6 | 0 | 100 | 165.69 | 965.6 | 8.45 | -4.5 | 0 | 100 | 1410.57 | 8220.8 | 56.2 | 0.9 | 0 | 100 | 2811.34 | 16384.5 | 100.54 | 1 | 0 | 100 | 2882.65 | 16800.2 | 99.21
6 | 0 | 100 | 146.94 | 856.4 | 7.49 | -3.9 | 0 | 100 | 1844.77 | 10751.4 | 73.45 | 2 | 0 | 100 | 2800.76 | 16322.9 | 100.13 | 2 | 0 | 100 | 2880.43 | 16787.2 | 99.18
6.5 | 0 | 100 | 122.02 | 711.1 | 6.23 | -3.5 | 0 | 100 | 2088.23 | 12170.3 | 83.19 | 2.9 | 0 | 100 | 2793.92 | 16283 | 99.93 | 2.9 | 0 | 100 | 2888.82 | 16836.1 | 99.42
7 | 0 | 100 | 108.46 | 632.1 | 5.54 | -3 | 0 | 100 | 2249.88 | 13112.3 | 89.56 | 3.9 | 0 | 100 | 2779.82 | 16200.8 | 99.44 | 4 | 0 | 100 | 2891.23 | 16850.2 | 99.49
7.5 | 0 | 100 | 106.72 | 622 | 5.46 | -2.5 | 0 | 100 | 2356.71 | 13734.9 | 93.88 | 5 | 0 | 100 | 2740.64 | 15972.5 | 97.99 | 5 | 0 | 100 | 2899.47 | 16898.2 | 99.81
8.1 | 0 | 100 | 87.62 | 510.6 | 4.49 | -2.1 | 0 | 100 | 2422.55 | 14118.7 | 96.47 | 5.9 | 0 | 100 | 2690.09 | 15677.9 | 96.24 | 6 | 0 | 100 | 2886.78 | 16824.2 | 99.35
8.5 | 0 | 100 | 100.73 | 587 | 5.14 | -1.5 | 0 | 100 | 2473.1 | 14413.3 | 98.51 | 7 | 0 | 100 | 2594 | 15117.9 | 92.75 | 7 | 0 | 100 | 2892.7 | 16858.7 | 99.6
9 | 0 | 100 | 91.84 | 535.2 | 4.69 | -1.1 | 0 | 100 | 2484.57 | 14480.1 | 98.96 | 8 | 0 | 100 | 2392.21 | 13941.9 | 85.54 | 8 | 0 | 100 | 2873.49 | 16746.7 | 98.91
9.5 | 0 | 100 | 93.77 | 546.5 | 4.78 | -0.5 | 0 | 100 | 2491.96 | 14523.2 | 99.34 | 9 | 0 | 100 | 1869.37 | 10894.7 | 66.85 | 9 | 0 | 100 | 2842.44 | 16565.8 | 97.83
10 | 0 | 100 | 78.07 | 455 | 3.97 | -0.1 | 0 | 100 | 2511.6 | 14637.6 | 100.03 | 10 | 0 | 100 | 1170.8 | 6823.4 | 41.89 | 9.9 | 0 | 100 | 2804.86 | 16346.8 | 96.55
[6 empty] | 0.5 | 0 | 100 | 2503.23 | 14588.9 | 99.74 | 11.1 | 0 | 100 | 600.28 | 3498.4 | 21.47 | 11 | 0 | 100 | 2726.56 | 15890.4 | 93.88
[6 empty] | 1 | 0 | 100 | 2502.61 | 14585.3 | 99.59 | 12 | 0 | 100 | 397.38 | 2316 | 14.21 | 11.9 | 0 | 100 | 2646.35 | 15423 | 91.04
Measurement time: | 27/05/2023 12:27:28 |  |  |  |  | 1.4 | 0 | 100 | 2481.04 | 14459.6 | 98.74 | 12.9 | 0 | 100 | 273.43 | 1593.6 | 9.77 | 13 | 0 | 100 | 2409.73 | 14044 | 82.93
Radiation device: | Novalis TB Accelerator |  |  |  |  | 2 | 0 | 100 | 2431.28 | 14169.5 | 96.92 | 14 | 0 | 100 | 223.7 | 1303.7 | 8 | 14 | 0 | 100 | 1922.01 | 11201.5 | 66.15
Energy: | 6 MV (FFF) |  |  |  |  | 2.4 | 0 | 100 | 2373.33 | 13831.8 | 94.53 | 15.1 | 0 | 100 | 176.31 | 1027.5 | 6.3 | 14.9 | 0 | 100 | 1223.67 | 7131.6 | 42.12
Controller: | CCU |  |  |  |  | 2.9 | 0 | 100 | 2277.98 | 13276.1 | 90.74 | 16 | 0 | 100 | 158.86 | 925.8 | 5.68 | 16 | 0 | 100 | 678.84 | 3956.3 | 23.37
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 3.5 | 0 | 100 | 2092.99 | 12198 | 83.35 | 17.1 | 0 | 100 | 138.78 | 808.8 | 4.96 | 17 | 0 | 100 | 437.9 | 2552.1 | 15.07
Field detector: | Razor Nano |  |  |  |  | 4 | 0 | 100 | 1867.59 | 10884.3 | 74.41 | 18 | 0 | 100 | 124.98 | 728.4 | 4.47 | 18 | 0 | 100 | 333.44 | 1943.3 | 11.48
Reference detector: | Stealth CI transmission |  |  |  |  | 4.5 | 0 | 100 | 1561.06 | 9097.9 | 62.16 | 19 | 0 | 100 | 125.84 | 733.4 | 4.5 | 18.9 | 0 | 100 | 276.59 | 1612 | 9.52
SAD: | 1000 mm |  |  |  |  | 5 | 0 | 100 | 1204.05 | 7017.2 | 47.95 | 19.9 | 0 | 100 | 119.29 | 695.2 | 4.27 | 20.1 | 0 | 100 | 214.44 | 1249.8 | 7.39
SSD: | 900 mm |  |  |  |  | 5.5 | 0 | 100 | 851.93 | 4965.1 | 33.94 | 21 | 0 | 100 | 113.27 | 660.1 | 4.05 | 20.9 | 0 | 100 | 206.46 | 1203.2 | 7.11
Field size: | 5 x 5 mm |  |  |  |  | 6 | 0 | 100 | 597.48 | 3482.1 | 23.82 | 22 | 0 | 100 | 111.16 | 647.8 | 3.97 | 22 | 0 | 100 | 189.82 | 1106.3 | 6.53
Gantry angle: | 0 |  |  |  |  | 6.4 | 0 | 100 | 451.95 | 2633.9 | 17.99 | 23 | 0 | 100 | 94.93 | 553.3 | 3.4 | 23 | 0 | 100 | 166.13 | 968.2 | 5.72
Measurement medium: | Water |  |  |  |  | 7 | 0 | 100 | 331.31 | 1930.9 | 13.2 | 24 | 0 | 100 | 91.01 | 530.4 | 3.26 | 24 | 0 | 100 | 154.42 | 899.9 | 5.32
Scan type: | Crossline |  |  |  |  | 7.4 | 0 | 100 | 278.72 | 1624.4 | 11.09 | 25 | 0 | 100 | 92.88 | 541.3 | 3.32 | 25 | 0 | 100 | 161.14 | 939.1 | 5.55
Scan mode: | StepByStep |  |  |  |  | 7.9 | 0 | 100 | 227.72 | 1327.1 | 9.08 | 26 | 0 | 100 | 92.81 | 540.9 | 3.32 | 26 | 0 | 100 | 140.86 | 820.9 | 4.85
Wedge: | none |  |  |  |  | 8.5 | 0 | 100 | 200.23 | 1166.9 | 7.98 | 26.9 | 0 | 100 | 101.31 | 590.4 | 3.62 | 27 | 0 | 100 | 142.46 | 830.3 | 4.91
[6 empty] | 9 | 0 | 100 | 174.38 | 1016.3 | 6.95 | 27.9 | 0 | 100 | 93.03 | 542.2 | 3.33 | 28 | 0 | 100 | 132.35 | 771.4 | 4.56
Points [mm]: | [5 empty] | 9.5 | 0 | 100 | 168.3 | 980.8 | 6.71 | 29 | 0 | 100 | 93.62 | 545.6 | 3.35 | 29 | 0 | 100 | 125.99 | 734.3 | 4.34
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 10 | 0 | 100 | 128.87 | 751.1 | 5.13 | 30 | 0 | 100 | 89.9 | 524 | 3.22 | 30 | 0 | 100 | 120.82 | 704.2 | 4.16
0 | 10 | 100 | 75.67 | 441 | 3.84 | 10.5 | 0 | 100 | 141.43 | 824.3 | 5.63
0 | 9.5 | 100 | 77.9 | 454 | 3.96 | 11 | 0 | 100 | 117.41 | 684.3 | 4.68
0 | 9 | 100 | 78.12 | 455.3 | 3.99 | 11.6 | 0 | 100 | 106.33 | 619.7 | 4.24 | Measurement time: | 27/05/2023 13:06:18 |  |  |  |  | Measurement time: | 27/05/2023 13:37:05
0 | 8.5 | 100 | 74.62 | 434.9 | 3.8 | 12 | 0 | 100 | 104.41 | 608.5 | 4.17 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
0 | 7.9 | 100 | 100.19 | 583.9 | 5.09 | 12.5 | 0 | 100 | 101.18 | 589.7 | 4.03 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
0 | 7.5 | 100 | 97.9 | 570.6 | 4.98 | 13 | 0 | 100 | 105.82 | 616.7 | 4.21 | Controller: | CCU |  |  |  |  | Controller: | CCU
0 | 7 | 100 | 112.6 | 656.2 | 5.73 | 13.4 | 0 | 100 | 121.43 | 707.7 | 4.84 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 6.5 | 100 | 119.43 | 696.1 | 6.08 | 14 | 0 | 100 | 101.43 | 591.2 | 4.04 | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano
0 | 6 | 100 | 146.75 | 855.2 | 7.45 | 14.4 | 0 | 100 | 97.33 | 567.2 | 3.88 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
0 | 5.6 | 100 | 189.09 | 1102 | 9.66 | 15 | 0 | 100 | 87.88 | 512.2 | 3.5 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
0 | 5 | 100 | 246.84 | 1438.6 | 12.58 | 15.5 | 0 | 100 | 94.38 | 550.1 | 3.76 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
0 | 4.5 | 100 | 325.98 | 1899.8 | 16.6 | 15.9 | 0 | 100 | 97.39 | 567.6 | 3.88 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
0 | 4 | 100 | 469.56 | 2736.6 | 23.87 | 16.4 | 0 | 100 | 98.76 | 575.6 | 3.93 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
0 | 3.6 | 100 | 692.01 | 4033 | 35.37 | 17 | 0 | 100 | 82.29 | 479.6 | 3.27 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
0 | 3 | 100 | 967.02 | 5635.8 | 49.32 | 17.5 | 0 | 100 | 83.51 | 486.7 | 3.33 | Scan type: | Crossline |  |  |  |  | Scan type: | Crossline
0 | 2.5 | 100 | 1310.16 | 7635.6 | 66.68 | 18 | 0 | 100 | 84.12 | 490.3 | 3.35 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
0 | 2 | 100 | 1556.34 | 9070.4 | 79.11 | 18.4 | 0 | 100 | 72 | 419.6 | 2.86 | Wedge: | none |  |  |  |  | Wedge: | none
0 | 1.4 | 100 | 1750.19 | 10200.1 | 88.92 | 18.9 | 0 | 100 | 81.19 | 473.2 | 3.24
0 | 1 | 100 | 1887.94 | 11002.9 | 95.96 | 19.5 | 0 | 100 | 80.67 | 470.2 | 3.22 | Points [mm]: | [5 empty] | Points [mm]:
0 | 0.5 | 100 | 1953.79 | 11386.7 | 99.12 | 20 | 0 | 100 | 90.01 | 524.6 | 3.59 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -0.1 | 100 | 1970.57 | 11484.5 | 100.11 | [6 empty] | 0 | -30 | 100 | 94.09 | 548.4 | 3.39 | 0 | 30 | 100 | 128.32 | 747.8 | 4.4
0 | -0.5 | 100 | 1960.1 | 11423.5 | 99.79 | [6 empty] | 0 | -29 | 100 | 96.76 | 563.9 | 3.49 | 0 | 29 | 100 | 138.89 | 809.5 | 4.76
0 | -1 | 100 | 1910.58 | 11134.9 | 97.13 | Measurement time: | 27/05/2023 12:45:21 |  |  |  |  | 0 | -28 | 100 | 97.91 | 570.6 | 3.53 | 0 | 28.1 | 100 | 145.25 | 846.5 | 4.98
0 | -1.5 | 100 | 1781.32 | 10381.6 | 90.54 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -27 | 100 | 101.13 | 589.4 | 3.65 | 0 | 27.1 | 100 | 145.64 | 848.8 | 5
0 | -2 | 100 | 1575.68 | 9183.1 | 79.66 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | -26.1 | 100 | 99.4 | 579.3 | 3.59 | 0 | 26 | 100 | 160.55 | 935.7 | 5.51
0 | -2.5 | 100 | 1280.38 | 7462.1 | 64.94 | Controller: | CCU |  |  |  |  | 0 | -25.1 | 100 | 100.53 | 585.9 | 3.63 | 0 | 25 | 100 | 169.46 | 987.6 | 5.81
0 | -3 | 100 | 972.69 | 5668.8 | 49.33 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -24.1 | 100 | 116.19 | 677.2 | 4.19 | 0 | 24 | 100 | 182.55 | 1063.9 | 6.26
0 | -3.4 | 100 | 729.57 | 4252 | 37.14 | Field detector: | Razor Nano |  |  |  |  | 0 | -23 | 100 | 91.92 | 535.7 | 3.32 | 0 | 23.1 | 100 | 190.61 | 1110.9 | 6.54
0 | -3.9 | 100 | 527.41 | 3073.8 | 26.85 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -22 | 100 | 111.24 | 648.3 | 4.01 | 0 | 22 | 100 | 207.04 | 1206.7 | 7.1
0 | -4.5 | 100 | 357.37 | 2082.7 | 18.21 | SAD: | 1000 mm |  |  |  |  | 0 | -21.1 | 100 | 115.64 | 674 | 4.17 | 0 | 21 | 100 | 227.32 | 1324.8 | 7.8
0 | -4.9 | 100 | 274.4 | 1599.2 | 13.99 | SSD: | 900 mm |  |  |  |  | 0 | -20.1 | 100 | 123.24 | 718.2 | 4.44 | 0 | 19.9 | 100 | 271.82 | 1584.2 | 9.32
0 | -5.5 | 100 | 199.57 | 1163.1 | 10.23 | Field size: | 10 x 10 mm |  |  |  |  | 0 | -19 | 100 | 130.68 | 761.6 | 4.72 | 0 | 19 | 100 | 342.23 | 1994.5 | 11.74
0 | -6 | 100 | 164.18 | 956.8 | 8.37 | Gantry angle: | 0 |  |  |  |  | 0 | -18 | 100 | 144.16 | 840.2 | 5.2 | 0 | 18 | 100 | 447.96 | 2610.7 | 15.37
0 | -6.4 | 100 | 136.73 | 796.9 | 6.97 | Measurement medium: | Water |  |  |  |  | 0 | -17 | 100 | 148.3 | 864.3 | 5.35 | 0 | 17 | 100 | 667.82 | 3892 | 22.92
0 | -7 | 100 | 118.98 | 693.4 | 6.06 | Scan type: | Crossline |  |  |  |  | 0 | -16 | 100 | 166.97 | 973.1 | 6.02 | 0 | 16.1 | 100 | 1045.89 | 6095.5 | 35.89
0 | -7.5 | 100 | 115.03 | 670.4 | 5.85 | Scan mode: | StepByStep |  |  |  |  | 0 | -15 | 100 | 193.31 | 1126.6 | 6.98 | 0 | 15 | 100 | 1633.82 | 9521.9 | 56.06
0 | -8 | 100 | 105.07 | 612.3 | 5.34 | Wedge: | none |  |  |  |  | 0 | -14 | 100 | 255.63 | 1489.8 | 9.22 | 0 | 14 | 100 | 2186.84 | 12744.9 | 75.01
0 | -8.5 | 100 | 103.98 | 606 | 5.3 | [6 empty] | 0 | -13 | 100 | 368.37 | 2146.9 | 13.29 | 0 | 13 | 100 | 2499.78 | 14568.8 | 85.79
0 | -9 | 100 | 85.52 | 498.4 | 4.35 | Points [mm]: | [5 empty] | 0 | -12 | 100 | 568.13 | 3311.1 | 20.48 | 0 | 12 | 100 | 2654.83 | 15472.4 | 91.08
0 | -9.5 | 100 | 88.69 | 516.9 | 4.52 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | -11 | 100 | 967.15 | 5636.6 | 34.87 | 0 | 11 | 100 | 2744.72 | 15996.3 | 94.18
0 | -10 | 100 | 79.93 | 465.8 | 4.07 | 0 | -20 | 100 | 89.11 | 519.3 | 3.56 | 0 | -10 | 100 | 1608.83 | 9376.3 | 58.07 | 0 | 10 | 100 | 2809.04 | 16371.2 | 96.36
[6 empty] | 0 | -19.5 | 100 | 89.85 | 523.7 | 3.59 | 0 | -9 | 100 | 2177.96 | 12693.2 | 78.56 | 0 | 9 | 100 | 2836.18 | 16529.3 | 97.31
[6 empty] | 0 | -19 | 100 | 81.15 | 473 | 3.24 | 0 | -8 | 100 | 2471.72 | 14405.2 | 89.24 | 0 | 8.1 | 100 | 2861.67 | 16677.9 | 98.22
Measurement time: | 27/05/2023 12:28:57 |  |  |  |  | 0 | -18.5 | 100 | 88.33 | 514.8 | 3.53 | 0 | -7 | 100 | 2609.3 | 15207.1 | 94.16 | 0 | 7.1 | 100 | 2869.35 | 16722.7 | 98.46
Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | -18 | 100 | 85.07 | 495.8 | 3.4 | 0 | -6.1 | 100 | 2683.84 | 15641.5 | 96.76 | 0 | 6 | 100 | 2895.02 | 16872.3 | 99.25
Energy: | 6 MV (FFF) |  |  |  |  | 0 | -17.5 | 100 | 83.92 | 489.1 | 3.36 | 0 | -5 | 100 | 2722.98 | 15869.6 | 98.17 | 0 | 5.1 | 100 | 2892.8 | 16859.3 | 99.23
Controller: | CCU |  |  |  |  | 0 | -17 | 100 | 89.87 | 523.8 | 3.59 | 0 | -4.1 | 100 | 2739.34 | 15964.9 | 98.77 | 0 | 4 | 100 | 2897.87 | 16888.8 | 99.45
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | -16.5 | 100 | 92.05 | 536.5 | 3.67 | 0 | -3 | 100 | 2760.89 | 16090.5 | 99.57 | 0 | 3.1 | 100 | 2904.14 | 16925.4 | 99.67
Field detector: | Razor Nano |  |  |  |  | 0 | -16 | 100 | 84.08 | 490 | 3.36 | 0 | -2 | 100 | 2764.97 | 16114.3 | 99.78 | 0 | 2 | 100 | 2920.71 | 17021.9 | 100.23
Reference detector: | Stealth CI transmission |  |  |  |  | 0 | -15.5 | 100 | 89.78 | 523.2 | 3.58 | 0 | -1 | 100 | 2787.84 | 16247.6 | 100.58 | 0 | 1 | 100 | 2914.67 | 16986.7 | 100.03
SAD: | 1000 mm |  |  |  |  | 0 | -14.9 | 100 | 90.27 | 526.1 | 3.6 | 0 | -0.1 | 100 | 2773.7 | 16165.2 | 99.99 | 0 | 0.1 | 100 | 2914.75 | 16987.2 | 99.99
SSD: | 900 mm |  |  |  |  | 0 | -14.5 | 100 | 92.36 | 538.3 | 3.69 | 0 | 0.9 | 100 | 2776.95 | 16184.1 | 100.15 | 0 | -0.9 | 100 | 2915.39 | 16990.9 | 100.08
Field size: | 5 x 5 mm |  |  |  |  | 0 | -14 | 100 | 88.69 | 516.9 | 3.54 | 0 | 1.9 | 100 | 2777.49 | 16187.3 | 100.25 | 0 | -2 | 100 | 2915.4 | 16991 | 99.98
Gantry angle: | 0 |  |  |  |  | 0 | -13.5 | 100 | 93.05 | 542.3 | 3.72 | 0 | 3 | 100 | 2764.27 | 16110.2 | 99.75 | 0 | -2.9 | 100 | 2920.93 | 17023.2 | 100.18
Measurement medium: | Water |  |  |  |  | 0 | -13 | 100 | 106.07 | 618.2 | 4.23 | 0 | 4.1 | 100 | 2743.57 | 15989.6 | 98.96 | 0 | -4 | 100 | 2909.16 | 16954.7 | 99.81
Scan type: | Beam |  |  |  |  | 0 | -12.5 | 100 | 102.87 | 599.5 | 4.11 | 0 | 4.9 | 100 | 2712.01 | 15805.7 | 97.8 | 0 | -5 | 100 | 2896.03 | 16878.1 | 99.33
Scan mode: | StepByStep |  |  |  |  | 0 | -12 | 100 | 113.01 | 658.6 | 4.51 | 0 | 5.9 | 100 | 2681.45 | 15627.5 | 96.71 | 0 | -5.9 | 100 | 2900.74 | 16905.6 | 99.54
Wedge: | none |  |  |  |  | 0 | -11.5 | 100 | 120.44 | 701.9 | 4.81 | 0 | 6.9 | 100 | 2600.13 | 15153.6 | 93.76 | 0 | -6.9 | 100 | 2879.79 | 16783.4 | 98.83
[6 empty] | 0 | -11 | 100 | 124.93 | 728.1 | 4.99 | 0 | 7.9 | 100 | 2452.92 | 14295.7 | 88.51 | 0 | -8 | 100 | 2862.08 | 16680.2 | 98.23
Points [mm]: | [5 empty] | 0 | -10.5 | 100 | 138.69 | 808.3 | 5.54 | 0 | 9 | 100 | 2179.13 | 12700 | 78.59 | 0 | -9 | 100 | 2835.35 | 16524.4 | 97.22
Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | -10 | 100 | 154.4 | 899.8 | 6.17 | 0 | 9.9 | 100 | 1642.94 | 9575.1 | 59.28 | 0 | -9.9 | 100 | 2800.08 | 16318.9 | 96.11
0 | 0 | 250 | 730.96 | 4260 | 18.52 | 0 | -9.5 | 100 | 166.22 | 968.7 | 6.64 | 0 | 10.9 | 100 | 1005.01 | 5857.2 | 36.26 | 0 | -11 | 100 | 2759.79 | 16084.1 | 94.71
0 | 0 | 245.1 | 749.53 | 4368.3 | 19 | 0 | -9 | 100 | 192.1 | 1119.6 | 7.68 | 0 | 12 | 100 | 560.17 | 3264.7 | 20.21 | 0 | -12 | 100 | 2680.76 | 15623.5 | 92
0 | 0 | 240 | 788.2 | 4593.6 | 20.03 | 0 | -8.5 | 100 | 222.62 | 1297.5 | 8.89 | 0 | 12.9 | 100 | 372.14 | 2168.9 | 13.43 | 0 | -12.9 | 100 | 2566.47 | 14957.4 | 88.06
0 | 0 | 235.1 | 805.69 | 4695.6 | 20.48 | 0 | -8 | 100 | 255.26 | 1487.7 | 10.2 | 0 | 13.9 | 100 | 268.7 | 1566 | 9.69 | 0 | -13.9 | 100 | 2275.92 | 13264.1 | 78.08
0 | 0 | 230 | 824.87 | 4807.4 | 21.08 | 0 | -7.5 | 100 | 348.43 | 2030.6 | 13.91 | 0 | 15 | 100 | 207.86 | 1211.4 | 7.5 | 0 | -15 | 100 | 1735.94 | 10117.1 | 59.62
0 | 0 | 225 | 857.01 | 4994.7 | 21.79 | 0 | -7 | 100 | 450.47 | 2625.3 | 17.98 | 0 | 15.9 | 100 | 167.3 | 975.1 | 6.04 | 0 | -16 | 100 | 1156.94 | 6742.7 | 39.71
0 | 0 | 220.1 | 879.09 | 5123.4 | 22.32 | 0 | -6.5 | 100 | 625.27 | 3644.1 | 24.97 | 0 | 17 | 100 | 152.11 | 886.5 | 5.49 | 0 | -17 | 100 | 721.15 | 4202.9 | 24.75
0 | 0 | 215.1 | 900.15 | 5246.1 | 22.97 | 0 | -6 | 100 | 867.36 | 5055 | 34.68 | 0 | 18 | 100 | 133.04 | 775.4 | 4.8 | 0 | -17.9 | 100 | 483.74 | 2819.2 | 16.6
0 | 0 | 210.1 | 923.5 | 5382.2 | 23.52 | 0 | -5.5 | 100 | 1160.91 | 6765.8 | 46.36 | 0 | 18.9 | 100 | 121.09 | 705.7 | 4.37 | 0 | -19 | 100 | 356.73 | 2079 | 12.24
0 | 0 | 205.1 | 960.18 | 5596 | 24.36 | 0 | -5 | 100 | 1515.76 | 8833.9 | 60.66 | 0 | 20 | 100 | 116.69 | 680.1 | 4.21 | 0 | -19.9 | 100 | 266.89 | 1555.4 | 9.17
0 | 0 | 200.1 | 976.85 | 5693.1 | 24.84 | 0 | -4.5 | 100 | 1848.04 | 10770.4 | 73.84 | 0 | 21 | 100 | 113.77 | 663.1 | 4.11 | 0 | -21 | 100 | 237.07 | 1381.7 | 8.14
0 | 0 | 195.1 | 1035.82 | 6036.8 | 26.39 | 0 | -4 | 100 | 2075.08 | 12093.6 | 82.92 | 0 | 21.9 | 100 | 96.66 | 563.3 | 3.49 | 0 | -22 | 100 | 223.34 | 1301.6 | 7.67
0 | 0 | 190.1 | 1064.65 | 6204.8 | 27.07 | 0 | -3.6 | 100 | 2273.46 | 13249.7 | 90.8 | 0 | 23 | 100 | 95.89 | 558.8 | 3.46 | 0 | -23 | 100 | 203.57 | 1186.4 | 6.99
0 | 0 | 185.1 | 1100.51 | 6413.8 | 28.03 | 0 | -3 | 100 | 2365.07 | 13783.7 | 94.45 | 0 | 23.9 | 100 | 91.89 | 535.6 | 3.32 | 0 | -24 | 100 | 188.37 | 1097.8 | 6.46
0 | 0 | 180.1 | 1129.19 | 6580.9 | 28.79 | 0 | -2.6 | 100 | 2409.69 | 14043.7 | 96.28 | 0 | 24.9 | 100 | 90.72 | 528.7 | 3.27 | 0 | -24.9 | 100 | 164.45 | 958.4 | 5.64
0 | 0 | 175 | 1180.72 | 6881.2 | 30.09 | 0 | -2 | 100 | 2451.99 | 14290.2 | 97.91 | 0 | 25.9 | 100 | 83.99 | 489.5 | 3.03 | 0 | -25.9 | 100 | 169.9 | 990.2 | 5.83
0 | 0 | 170 | 1225.88 | 7144.4 | 31.23 | 0 | -1.5 | 100 | 2485.58 | 14486 | 99.26 | 0 | 27 | 100 | 97.86 | 570.3 | 3.53 | 0 | -26.9 | 100 | 165.4 | 964 | 5.68
0 | 0 | 165.1 | 1265.48 | 7375.3 | 32.12 | 0 | -1 | 100 | 2497.45 | 14555.2 | 99.76 | 0 | 28 | 100 | 91.7 | 534.4 | 3.31 | 0 | -28 | 100 | 159.8 | 931.3 | 5.49
0 | 0 | 160 | 1290.3 | 7519.9 | 32.82 | 0 | -0.6 | 100 | 2500.06 | 14570.4 | 99.89 | 0 | 29 | 100 | 74.19 | 432.4 | 2.68 | 0 | -28.9 | 100 | 155.33 | 905.3 | 5.33
0 | 0 | 155.1 | 1352.79 | 7884.1 | 34.44 | 0 | -0.1 | 100 | 2506.37 | 14607.2 | 100.04 | 0 | 30 | 100 | 66.15 | 385.5 | 2.39 | 0 | -30 | 100 | 143.31 | 835.2 | 4.92
0 | 0 | 150.1 | 1390.75 | 8105.3 | 35.54 | 0 | 0.5 | 100 | 2502.2 | 14582.9 | 99.83
0 | 0 | 145.1 | 1443.13 | 8410.6 | 36.71 | 0 | 1 | 100 | 2489.33 | 14507.8 | 99.37
0 | 0 | 140.1 | 1497.74 | 8728.9 | 38.22 | 0 | 1.6 | 100 | 2465.95 | 14371.6 | 98.46 | Measurement time: | 27/05/2023 13:09:05 |  |  |  |  | Measurement time: | 27/05/2023 13:39:45
0 | 0 | 135.1 | 1547.21 | 9017.2 | 39.42 | 0 | 2 | 100 | 2427.05 | 14144.9 | 96.96 | Radiation device: | Novalis TB Accelerator |  |  |  |  | Radiation device: | Novalis TB Accelerator
0 | 0 | 130.1 | 1604.01 | 9348.2 | 40.89 | 0 | 2.5 | 100 | 2382.18 | 13883.4 | 95.26 | Energy: | 6 MV (FFF) |  |  |  |  | Energy: | 6 MV (FFF)
0 | 0 | 125.1 | 1667.96 | 9720.9 | 42.61 | 0 | 3 | 100 | 2303.84 | 13426.8 | 91.99 | Controller: | CCU |  |  |  |  | Controller: | CCU
0 | 0 | 120.1 | 1731.99 | 10094.1 | 44.16 | 0 | 3.5 | 100 | 2200.7 | 12825.7 | 87.91 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 0 | 115 | 1796.56 | 10470.4 | 45.68 | 0 | 4 | 100 | 2035.92 | 11865.4 | 81.43 | Field detector: | Razor Nano |  |  |  |  | Field detector: | Razor Nano
0 | 0 | 110.1 | 1859.84 | 10839.2 | 47.37 | 0 | 4.5 | 100 | 1797.66 | 10476.8 | 71.86 | Reference detector: | Stealth CI transmission |  |  |  |  | Reference detector: | Stealth CI transmission
0 | 0 | 105.1 | 1925.95 | 11224.5 | 48.94 | 0 | 5 | 100 | 1498.37 | 8732.5 | 59.85 | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
0 | 0 | 100.1 | 1981.59 | 11548.7 | 50.36 | 0 | 5.5 | 100 | 1175.96 | 6853.5 | 46.96 | SSD: | 900 mm |  |  |  |  | SSD: | 900 mm
0 | 0 | 95.1 | 2062.1 | 12018 | 52.45 | 0 | 5.9 | 100 | 865.44 | 5043.8 | 34.58 | Field size: | 20 x 20 mm |  |  |  |  | Field size: | 30 x 30 mm
0 | 0 | 90.1 | 2142.19 | 12484.7 | 54.5 | 0 | 6.5 | 100 | 607.52 | 3540.6 | 24.25 | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
0 | 0 | 85 | 2231.64 | 13006.1 | 57.01 | 0 | 7.1 | 100 | 439.16 | 2559.4 | 17.55 | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
0 | 0 | 80.1 | 2319.74 | 13519.5 | 58.97 | 0 | 7.5 | 100 | 335.14 | 1953.2 | 13.39 | Scan type: | Beam |  |  |  |  | Scan type: | Beam
0 | 0 | 75 | 2396.41 | 13966.3 | 61.17 | 0 | 8 | 100 | 268.64 | 1565.6 | 10.73 | Scan mode: | StepByStep |  |  |  |  | Scan mode: | StepByStep
0 | 0 | 70 | 2484.59 | 14480.2 | 63.37 | 0 | 8.5 | 100 | 235.84 | 1374.5 | 9.42 | Wedge: | none |  |  |  |  | Wedge: | none
0 | 0 | 65 | 2577.38 | 15021 | 65.68 | 0 | 9 | 100 | 190.78 | 1111.9 | 7.62
0 | 0 | 60 | 2672.51 | 15575.5 | 67.78 | 0 | 9.4 | 100 | 158.66 | 924.7 | 6.34 | Points [mm]: | [5 empty] | Points [mm]:
0 | 0 | 55 | 2787.61 | 16246.3 | 70.85 | 0 | 10 | 100 | 142.92 | 833 | 5.7 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 49.9 | 2892.03 | 16854.8 | 73.35 | 0 | 10.5 | 100 | 130.91 | 763 | 5.23 | 0 | 0 | 250 | 1035.4 | 6034.3 | 20.31 | 0 | 0 | 250 | 1103.89 | 6433.5 | 21.48
0 | 0 | 45 | 3016.92 | 17582.7 | 77.02 | 0 | 11 | 100 | 121.31 | 707 | 4.84 | 0 | 0 | 245.1 | 1060.83 | 6182.5 | 20.81 | 0 | 0 | 245.1 | 1127.72 | 6572.4 | 21.95
0 | 0 | 39.9 | 3142.82 | 18316.4 | 79.94 | 0 | 11.4 | 100 | 122.25 | 712.5 | 4.89 | 0 | 0 | 240 | 1098.57 | 6402.5 | 21.55 | 0 | 0 | 240 | 1170.78 | 6823.3 | 22.8
0 | 0 | 34.9 | 3258.98 | 18993.4 | 83.24 | 0 | 12 | 100 | 106.28 | 619.4 | 4.24 | 0 | 0 | 235.1 | 1125.96 | 6562.1 | 22.09 | 0 | 0 | 235.1 | 1208.94 | 7045.7 | 23.52
0 | 0 | 29.9 | 3385.07 | 19728.3 | 86.31 | 0 | 12.5 | 100 | 98.57 | 574.5 | 3.94 | 0 | 0 | 230.1 | 1159.4 | 6757 | 22.73 | 0 | 0 | 230 | 1240.15 | 7227.6 | 24.14
0 | 0 | 25 | 3517.49 | 20500 | 89.56 | 0 | 13 | 100 | 90.38 | 526.7 | 3.61 | 0 | 0 | 225.1 | 1206.75 | 7033 | 23.67 | 0 | 0 | 225.1 | 1275.46 | 7433.4 | 24.81
0 | 0 | 24.1 | 3545.17 | 20661.3 | 89.76 | 0 | 13.4 | 100 | 84.33 | 491.5 | 3.37 | 0 | 0 | 220.1 | 1245.89 | 7261.1 | 24.45 | 0 | 0 | 220.1 | 1314.54 | 7661.2 | 25.58
0 | 0 | 23.1 | 3567.34 | 20790.5 | 90.85 | 0 | 13.9 | 100 | 81.69 | 476.1 | 3.27 | 0 | 0 | 215 | 1288.13 | 7507.2 | 25.27 | 0 | 0 | 215 | 1357.51 | 7911.6 | 26.42
0 | 0 | 22 | 3605.9 | 21015.3 | 91.81 | 0 | 14.5 | 100 | 85.7 | 499.5 | 3.42 | 0 | 0 | 210.1 | 1328.91 | 7744.9 | 26.08 | 0 | 0 | 210.1 | 1394.1 | 8124.8 | 27.14
0 | 0 | 21 | 3625.3 | 21128.3 | 92.4 | 0 | 15 | 100 | 82.21 | 479.1 | 3.29 | 0 | 0 | 205 | 1363.35 | 7945.6 | 26.72 | 0 | 0 | 205.1 | 1447.49 | 8436 | 28.17
0 | 0 | 20 | 3661.28 | 21338 | 93.47 | 0 | 15.4 | 100 | 76.99 | 448.7 | 3.08 | 0 | 0 | 200.1 | 1405.82 | 8193.1 | 27.57 | 0 | 0 | 200.1 | 1500.46 | 8744.7 | 29.2
0 | 0 | 19 | 3696.32 | 21542.2 | 93.71 | 0 | 16 | 100 | 81.83 | 476.9 | 3.27 | 0 | 0 | 195.1 | 1451.88 | 8461.6 | 28.46 | 0 | 0 | 195.1 | 1547.95 | 9021.5 | 30.14
0 | 0 | 18 | 3727.72 | 21725.2 | 95.08 | 0 | 16.5 | 100 | 71.44 | 416.3 | 2.86 | 0 | 0 | 190 | 1504.36 | 8767.4 | 29.52 | 0 | 0 | 190.1 | 1594.29 | 9291.5 | 31.05
0 | 0 | 17 | 3766.16 | 21949.3 | 95.76 | 0 | 17 | 100 | 66.61 | 388.2 | 2.67 | 0 | 0 | 185.1 | 1564.54 | 9118.2 | 30.7 | 0 | 0 | 185 | 1651.08 | 9622.5 | 32.14
0 | 0 | 16.1 | 3775.15 | 22001.6 | 95.87 | 0 | 17.5 | 100 | 66.72 | 388.9 | 2.67 | 0 | 0 | 180 | 1614.03 | 9406.6 | 31.64 | 0 | 0 | 180 | 1711.4 | 9974.1 | 33.32
0 | 0 | 15 | 3801.24 | 22153.7 | 96.18 | 0 | 17.9 | 100 | 72.25 | 421.1 | 2.89 | 0 | 0 | 175.1 | 1668.05 | 9721.4 | 32.73 | 0 | 0 | 175.1 | 1762.4 | 10271.3 | 34.33
0 | 0 | 14 | 3839.33 | 22375.7 | 97.37 | 0 | 18.5 | 100 | 60.72 | 353.9 | 2.43 | 0 | 0 | 170 | 1723.56 | 10044.9 | 33.81 | 0 | 0 | 170.1 | 1824.34 | 10632.3 | 35.51
0 | 0 | 13 | 3862.17 | 22508.8 | 98.16 | 0 | 19 | 100 | 88.23 | 514.2 | 3.53 | 0 | 0 | 165.1 | 1785.2 | 10404.2 | 35.03 | 0 | 0 | 165.1 | 1882.32 | 10970.2 | 36.65
0 | 0 | 12 | 3895.21 | 22701.4 | 99.07 | 0 | 19.4 | 100 | 72.11 | 420.2 | 2.88 | 0 | 0 | 160.1 | 1829.57 | 10662.7 | 35.91 | 0 | 0 | 160.1 | 1936.29 | 11284.7 | 37.71
0 | 0 | 11 | 3894.48 | 22697.1 | 99.17 | 0 | 20 | 100 | 71.69 | 417.8 | 2.87 | 0 | 0 | 155.1 | 1890.73 | 11019.2 | 37.08 | 0 | 0 | 155.1 | 2015.07 | 11743.9 | 39.26
0 | 0 | 10 | 3922.58 | 22860.8 | 100 | [6 empty] | 0 | 0 | 150.1 | 1952.77 | 11380.8 | 38.3 | 0 | 0 | 150.1 | 2071.26 | 12071.4 | 40.34
0 | 0 | 9 | 3921.19 | 22852.8 | 99.65 | [6 empty] | 0 | 0 | 145.1 | 2027.22 | 11814.7 | 39.78 | 0 | 0 | 145.1 | 2141.54 | 12480.9 | 41.71
0 | 0 | 8 | 3919.32 | 22841.9 | 99.75 | Measurement time: | 27/05/2023 12:48:22 |  |  |  |  | 0 | 0 | 140.1 | 2105.85 | 12272.9 | 41.34 | 0 | 0 | 140.1 | 2213.24 | 12898.8 | 43.1
0 | 0 | 7 | 3909.97 | 22787.4 | 99.49 | Radiation device: | Novalis TB Accelerator |  |  |  |  | 0 | 0 | 135.1 | 2182.39 | 12719 | 42.81 | 0 | 0 | 135.1 | 2288.79 | 13339.1 | 44.59
0 | 0 | 6 | 3853 | 22455.3 | 97.89 | Energy: | 6 MV (FFF) |  |  |  |  | 0 | 0 | 130 | 2246.43 | 13092.3 | 44.06 | 0 | 0 | 130.1 | 2363.23 | 13773 | 46.01
0 | 0 | 5 | 3775.56 | 22004 | 96.19 | Controller: | CCU |  |  |  |  | 0 | 0 | 125.1 | 2330.61 | 13582.8 | 45.71 | 0 | 0 | 125.1 | 2448.79 | 14271.6 | 47.69
0 | 0 | 4 | 3603.5 | 21001.3 | 92.08 | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | 0 | 0 | 120 | 2408.64 | 14037.6 | 47.24 | 0 | 0 | 120.1 | 2530.84 | 14749.8 | 49.3
0 | 0 | 3 | 3339.42 | 19462.2 | 84.85 | Field detector: | Razor Nano |  |  |  |  | 0 | 0 | 115 | 2490.8 | 14516.4 | 48.9 | 0 | 0 | 115.1 | 2623.35 | 15288.9 | 51.11
0 | 0 | 2 | 2867.21 | 16710.2 | 72.94 | Reference detector: | Stealth CI transmission |  |  |  |  | 0 | 0 | 110 | 2596.48 | 15132.3 | 50.96 | 0 | 0 | 110.1 | 2719.52 | 15849.4 | 53
0 | 0 | 1 | 2443.37 | 14240 | 61.96 | SAD: | 1000 mm |  |  |  |  | 0 | 0 | 105.1 | 2689.44 | 15674.1 | 52.79 | 0 | 0 | 105.1 | 2817.27 | 16419.1 | 54.87
0 | 0 | 0 | 2379.89 | 13870.1 | 60.28 | SSD: | 900 mm |  |  |  |  | 0 | 0 | 100.1 | 2778.28 | 16191.9 | 54.5 | 0 | 0 | 100.1 | 2906.58 | 16939.6 | 56.6
0 | 0 | -0.5 | 2373.54 | 13833 | 60.43 | Field size: | 10 x 10 mm |  |  |  |  | 0 | 0 | 95.1 | 2889.48 | 16839.9 | 56.69 | 0 | 0 | 95.1 | 3007.59 | 17528.3 | 58.58
[6 empty] | Gantry angle: | 0 |  |  |  |  | 0 | 0 | 90.1 | 2980.24 | 17368.9 | 58.47 | 0 | 0 | 90.1 | 3122.75 | 18199.5 | 60.81
[6 empty] | Measurement medium: | Water |  |  |  |  | 0 | 0 | 85 | 3092.84 | 18025.1 | 60.68 | 0 | 0 | 85 | 3235.83 | 18858.5 | 63.04
[6 empty] | Scan type: | Beam |  |  |  |  | 0 | 0 | 80 | 3213.61 | 18729 | 63.05 | 0 | 0 | 80 | 3340.91 | 19470.9 | 65.08
[6 empty] | Scan mode: | StepByStep |  |  |  |  | 0 | 0 | 75 | 3336.45 | 19444.9 | 65.43 | 0 | 0 | 75 | 3459.8 | 20163.8 | 67.4
[6 empty] | Wedge: | none |  |  |  |  | 0 | 0 | 70 | 3450.27 | 20108.2 | 67.71 | 0 | 0 | 70 | 3584.7 | 20891.7 | 69.86
[12 empty] | 0 | 0 | 65 | 3581.51 | 20873.1 | 70.27 | 0 | 0 | 65 | 3723.92 | 21703.1 | 72.55
[6 empty] | Points [mm]: | [5 empty] | 0 | 0 | 60 | 3715.4 | 21653.4 | 72.87 | 0 | 0 | 60 | 3855.43 | 22469.5 | 75.13
[6 empty] | Inline | Crossline | Depth | Normalized field | Current field | Ratio | 0 | 0 | 55 | 3859.69 | 22494.3 | 75.73 | 0 | 0 | 55 | 3977.08 | 23178.5 | 77.49
[6 empty] | 0 | 0 | 250 | 925.63 | 5394.6 | 19.51 | 0 | 0 | 49.9 | 3997.03 | 23294.8 | 78.45 | 0 | 0 | 49.9 | 4125.48 | 24043.4 | 80.37
[6 empty] | 0 | 0 | 245 | 951.66 | 5546.3 | 20.06 | 0 | 0 | 44.9 | 4158.19 | 24234 | 81.58 | 0 | 0 | 44.9 | 4285.42 | 24975.5 | 83.51
[6 empty] | 0 | 0 | 240.1 | 989.23 | 5765.2 | 20.82 | 0 | 0 | 39.9 | 4303.6 | 25081.5 | 84.44 | 0 | 0 | 39.9 | 4436.94 | 25858.5 | 86.44
[6 empty] | 0 | 0 | 235.1 | 1009.86 | 5885.5 | 21.3 | 0 | 0 | 34.9 | 4459.93 | 25992.6 | 87.54 | 0 | 0 | 34.9 | 4581.94 | 26703.6 | 89.29
[6 empty] | 0 | 0 | 230 | 1052.58 | 6134.4 | 22.19 | 0 | 0 | 29.9 | 4629.01 | 26977.9 | 90.87 | 0 | 0 | 29.9 | 4740.54 | 27628 | 92.39
[6 empty] | 0 | 0 | 225.1 | 1091.46 | 6361 | 22.99 | 0 | 0 | 25 | 4782.62 | 27873.2 | 93.83 | 0 | 0 | 25.1 | 4882.94 | 28457.9 | 95.17
[6 empty] | 0 | 0 | 220.1 | 1125.15 | 6557.4 | 23.69 | 0 | 0 | 24 | 4810.47 | 28035.5 | 94.4 | 0 | 0 | 24.1 | 4913.75 | 28637.4 | 95.79
[6 empty] | 0 | 0 | 215.1 | 1149.85 | 6701.3 | 24.19 | 0 | 0 | 23.1 | 4842.47 | 28222 | 95.1 | 0 | 0 | 23 | 4962.23 | 28919.9 | 96.71
[6 empty] | 0 | 0 | 210.1 | 1175.91 | 6853.3 | 24.78 | 0 | 0 | 22 | 4884.73 | 28468.3 | 95.85 | 0 | 0 | 22 | 4986.01 | 29058.6 | 97.17
[6 empty] | 0 | 0 | 205 | 1222.82 | 7126.6 | 25.73 | 0 | 0 | 21.1 | 4926.69 | 28712.8 | 96.69 | 0 | 0 | 21.1 | 5014.36 | 29223.8 | 97.72
[6 empty] | 0 | 0 | 200.1 | 1267.49 | 7386.9 | 26.7 | 0 | 0 | 20.1 | 4962.25 | 28920.1 | 97.36 | 0 | 0 | 20.1 | 5028.61 | 29306.9 | 97.99
[6 empty] | 0 | 0 | 195.1 | 1318.03 | 7681.5 | 27.74 | 0 | 0 | 19.1 | 4992.31 | 29095.3 | 98.01 | 0 | 0 | 19 | 5047.84 | 29418.9 | 98.38
[6 empty] | 0 | 0 | 190.1 | 1351.29 | 7875.4 | 28.43 | 0 | 0 | 18 | 5013.13 | 29216.6 | 98.42 | 0 | 0 | 18 | 5074.19 | 29572.5 | 98.85
[6 empty] | 0 | 0 | 185.1 | 1399.38 | 8155.6 | 29.45 | 0 | 0 | 17.1 | 5031.56 | 29324 | 98.79 | 0 | 0 | 17.1 | 5091.75 | 29674.8 | 99.23
[6 empty] | 0 | 0 | 180.1 | 1440.77 | 8396.8 | 30.34 | 0 | 0 | 16.1 | 5073.24 | 29567 | 99.57 | 0 | 0 | 16.1 | 5108.64 | 29773.2 | 99.59
[6 empty] | 0 | 0 | 175.1 | 1501.09 | 8748.4 | 31.6 | 0 | 0 | 15 | 5091.62 | 29674.1 | 99.92 | 0 | 0 | 15 | 5129.36 | 29894 | 100
[6 empty] | 0 | 0 | 170.1 | 1536.75 | 8956.2 | 32.39 | 0 | 0 | 14 | 5092.9 | 29681.5 | 99.93 | 0 | 0 | 14 | 5127.91 | 29885.6 | 99.96
[6 empty] | 0 | 0 | 165 | 1612.24 | 9396.2 | 33.95 | 0 | 0 | 13 | 5092.99 | 29682.1 | 100 | 0 | 0 | 13 | 5126.17 | 29875.4 | 99.9
[6 empty] | 0 | 0 | 160.1 | 1665.87 | 9708.7 | 35.11 | 0 | 0 | 12 | 5085.68 | 29639.5 | 99.86 | 0 | 0 | 12 | 5110.91 | 29786.5 | 99.6
[6 empty] | 0 | 0 | 155.1 | 1701.37 | 9915.6 | 35.81 | 0 | 0 | 11 | 5068.65 | 29540.2 | 99.47 | 0 | 0 | 11 | 5068.06 | 29536.7 | 98.79
[6 empty] | 0 | 0 | 150.1 | 1763.12 | 10275.5 | 37.16 | 0 | 0 | 10 | 5024 | 29280 | 98.63 | 0 | 0 | 10 | 5038 | 29361.6 | 98.23
[6 empty] | 0 | 0 | 145.1 | 1817.66 | 10593.4 | 38.25 | 0 | 0 | 9 | 4962.8 | 28923.3 | 97.47 | 0 | 0 | 9 | 4962.48 | 28921.4 | 96.69
[6 empty] | 0 | 0 | 140.1 | 1904.79 | 11101.2 | 40.09 | 0 | 0 | 8 | 4887.58 | 28484.9 | 95.93 | 0 | 0 | 8 | 4879.19 | 28436 | 95.11
[6 empty] | 0 | 0 | 135 | 1955.74 | 11398.1 | 41.2 | 0 | 0 | 7 | 4768.25 | 27789.4 | 93.58 | 0 | 0 | 7 | 4742.64 | 27640.2 | 92.45
[6 empty] | 0 | 0 | 130 | 2017.38 | 11757.3 | 42.53 | 0 | 0 | 6 | 4585.88 | 26726.6 | 90.09 | 0 | 0 | 6 | 4577.36 | 26676.9 | 89.19
[6 empty] | 0 | 0 | 125.1 | 2101.15 | 12245.5 | 44.29 | 0 | 0 | 5 | 4374.46 | 25494.4 | 85.87 | 0 | 0 | 5 | 4322.54 | 25191.9 | 84.24
[6 empty] | 0 | 0 | 120.1 | 2171.34 | 12654.6 | 45.69 | 0 | 0 | 4 | 4082.86 | 23795 | 80.22 | 0 | 0 | 4 | 3990.29 | 23255.5 | 77.81
[6 empty] | 0 | 0 | 115 | 2254.11 | 13137 | 47.51 | 0 | 0 | 3 | 3652.52 | 21286.9 | 71.71 | 0 | 0 | 3 | 3558.93 | 20741.5 | 69.37
[6 empty] | 0 | 0 | 110.1 | 2322.28 | 13534.3 | 48.97 | 0 | 0 | 2 | 3037.09 | 17700.2 | 59.63 | 0 | 0 | 2 | 2958.95 | 17244.8 | 57.68
[6 empty] | 0 | 0 | 105.1 | 2409.98 | 14045.4 | 50.81 | 0 | 0 | 1 | 2558.47 | 14910.8 | 50.27 | 0 | 0 | 1 | 2495.77 | 14545.4 | 48.66
[6 empty] | 0 | 0 | 100.1 | 2502.42 | 14584.1 | 52.72 | 0 | 0 | 0 | 2472.71 | 14411 | 48.52 | 0 | 0 | 0 | 2404.99 | 14016.3 | 46.9
[6 empty] | 0 | 0 | 95.1 | 2601.17 | 15159.7 | 54.84 | 0 | 0 | -0.5 | 2444.68 | 14247.6 | 48 | 0 | 0 | -0.5 | 2388.81 | 13922 | 46.58
[6 empty] | 0 | 0 | 90.1 | 2694.03 | 15700.8 | 56.76
[6 empty] | 0 | 0 | 85.1 | 2789.24 | 16255.8 | 58.77
[6 empty] | 0 | 0 | 80 | 2895.52 | 16875.2 | 61.02
[6 empty] | 0 | 0 | 75 | 3005.46 | 17515.9 | 63.24
[6 empty] | 0 | 0 | 70 | 3115.34 | 18156.3 | 65.66
[6 empty] | 0 | 0 | 65 | 3236.51 | 18862.4 | 68.11
[6 empty] | 0 | 0 | 60 | 3353.87 | 19546.4 | 70.59
[6 empty] | 0 | 0 | 55 | 3481.52 | 20290.4 | 73.37
[6 empty] | 0 | 0 | 49.9 | 3608.72 | 21031.7 | 76.01
[6 empty] | 0 | 0 | 44.9 | 3750 | 21855.1 | 79.09
[6 empty] | 0 | 0 | 39.9 | 3909.63 | 22785.4 | 82.36
[6 empty] | 0 | 0 | 34.9 | 4043.07 | 23563.1 | 85.15
[6 empty] | 0 | 0 | 29.9 | 4199.47 | 24474.6 | 88.58
[6 empty] | 0 | 0 | 25.1 | 4372.27 | 25481.7 | 92.14
[6 empty] | 0 | 0 | 24.1 | 4410.32 | 25703.4 | 92.99
[6 empty] | 0 | 0 | 23.1 | 4446.63 | 25915.1 | 93.56
[6 empty] | 0 | 0 | 22 | 4480.58 | 26112.9 | 94.39
[6 empty] | 0 | 0 | 21.1 | 4517.08 | 26325.7 | 95.18
[6 empty] | 0 | 0 | 20.1 | 4537.78 | 26446.3 | 95.54
[6 empty] | 0 | 0 | 19 | 4579.23 | 26687.8 | 96.41
[6 empty] | 0 | 0 | 18 | 4621.18 | 26932.3 | 97.2
[6 empty] | 0 | 0 | 17 | 4641.07 | 27048.3 | 97.76
[6 empty] | 0 | 0 | 16.1 | 4660.48 | 27161.4 | 98.17
[6 empty] | 0 | 0 | 15 | 4686.98 | 27315.8 | 98.66
[6 empty] | 0 | 0 | 14 | 4721.13 | 27514.9 | 99.32
[6 empty] | 0 | 0 | 13 | 4724.49 | 27534.4 | 99.48
[6 empty] | 0 | 0 | 12 | 4732.19 | 27579.3 | 99.59
[6 empty] | 0 | 0 | 11 | 4747.63 | 27669.3 | 100
[6 empty] | 0 | 0 | 10 | 4736.18 | 27602.5 | 99.75
[6 empty] | 0 | 0 | 9 | 4725.03 | 27537.6 | 99.54
[6 empty] | 0 | 0 | 8 | 4669.53 | 27214.1 | 98.33
[6 empty] | 0 | 0 | 7 | 4590.63 | 26754.3 | 96.72
[6 empty] | 0 | 0 | 6 | 4478.5 | 26100.8 | 94.42
[6 empty] | 0 | 0 | 5 | 4304.1 | 25084.4 | 90.71
[6 empty] | 0 | 0 | 4 | 4039.03 | 23539.5 | 85.21
[6 empty] | 0 | 0 | 3 | 3647.13 | 21255.5 | 76.88
[6 empty] | 0 | 0 | 2 | 3093.52 | 18029.1 | 65.22
[6 empty] | 0 | 0 | 1 | 2566.23 | 14956 | 54.12
[6 empty] | 0 | 0 | 0 | 2465.49 | 14368.9 | 52.04
[6 empty] | 0 | 0 | -0.5 | 2467.99 | 14383.5 | 52.06

## Cône
Measurement time: | 12/04/2023 19:58:03
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
10 | 0 | 100 | 2.08 | 21021.6 | 0
9.5 | 0 | 100 | 2.41 | 24354.7 | 0
9 | 0 | 100 | 2.83 | 28669.4 | 0
8.5 | 0 | 100 | 3.35 | 33941.4 | 0
8 | 0 | 100 | 4.05 | 40970.2 | 0
7.6 | 0 | 100 | 4.91 | 49654.4 | 0
7 | 0 | 100 | 6.42 | 65004.7 | 0
6.5 | 0 | 100 | 8.27 | 83677 | 0
6.1 | 0 | 100 | 11.06 | 111900.2 | 0
5.5 | 0 | 100 | 17.15 | 173530.5 | 0
5.1 | 0 | 100 | 29.47 | 298161.4 | 0
4.5 | 0 | 100 | 61.54 | 622620 | 0
4 | 0 | 100 | 80.5 | 814487.6 | 0
3.5 | 0 | 100 | 87.7 | 887289.9 | 0
3 | 0 | 100 | 92.19 | 932793.8 | 0
2.5 | 0 | 100 | 94.94 | 960575.2 | 0
2.1 | 0 | 100 | 96.7 | 978414.8 | 0
1.5 | 0 | 100 | 98.3 | 994599.5 | 0
1 | 0 | 100 | 99.02 | 1001844.1 | 0
0.5 | 0 | 100 | 99.66 | 1008313.5 | 0
0 | 0 | 100 | 99.98 | 1011554.7 | 0
-0.5 | 0 | 100 | 100.35 | 1015283.7 | 0
-0.9 | 0 | 100 | 99.6 | 1007688.9 | 0
-1.5 | 0 | 100 | 99.2 | 1003647.2 | 0
-2.1 | 0 | 100 | 97.82 | 989755.4 | 0
-2.5 | 0 | 100 | 96.47 | 976078.7 | 0
-3 | 0 | 100 | 94.39 | 955021.1 | 0
-3.4 | 0 | 100 | 91.64 | 927204 | 0
-4 | 0 | 100 | 86.23 | 872471.4 | 0
-4.5 | 0 | 100 | 77.03 | 779381.5 | 0
-5 | 0 | 100 | 51.61 | 522224.6 | 0
-5.4 | 0 | 100 | 28.35 | 286861.8 | 0
-5.9 | 0 | 100 | 16.44 | 166380.5 | 0
-6.5 | 0 | 100 | 11.19 | 113182.6 | 0
-7 | 0 | 100 | 8.09 | 81868.3 | 0
-7.6 | 0 | 100 | 6.03 | 60997.7 | 0
-8 | 0 | 100 | 4.89 | 49456.3 | 0
-8.5 | 0 | 100 | 4.01 | 40607.1 | 0
-9 | 0 | 100 | 3.32 | 33610.5 | 0
-9.5 | 0 | 100 | 2.84 | 28699.8 | 0
-10 | 0 | 100 | 2.45 | 24823.8 | 0
Measurement time: | 12/04/2023 19:59:40
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -10 | 100 | 2.24 | 22637.2 | 0
0 | -9.5 | 100 | 2.67 | 26990 | 0
0 | -9 | 100 | 3.1 | 31379.2 | 0
0 | -8.5 | 100 | 3.7 | 37450.4 | 0
0 | -7.9 | 100 | 4.75 | 48051.9 | 0
0 | -7.5 | 100 | 5.59 | 56554.4 | 0
0 | -7 | 100 | 7.14 | 72245.7 | 0
0 | -6.4 | 100 | 10.04 | 101547.8 | 0
0 | -5.9 | 100 | 14.01 | 141680 | 0
0 | -5.5 | 100 | 22.09 | 223492.8 | 0
0 | -5 | 100 | 38.72 | 391739.9 | 0
0 | -4.5 | 100 | 63.13 | 638587.5 | 0
0 | -4 | 100 | 82.12 | 830757.5 | 0
0 | -3.5 | 100 | 89.22 | 902581.1 | 0
0 | -3 | 100 | 92.62 | 936941.4 | 0
0 | -2.5 | 100 | 95.26 | 963698.3 | 0
0 | -2 | 100 | 97.18 | 983065.7 | 0
0 | -1.5 | 100 | 98.58 | 997206.5 | 0
0 | -1 | 100 | 99.08 | 1002351 | 0
0 | -0.5 | 100 | 99.87 | 1010311.7 | 0
0 | 0 | 100 | 100.01 | 1011696.6 | 0
0 | 0.5 | 100 | 99.86 | 1010195.3 | 0
0 | 1 | 100 | 99.46 | 1006122.9 | 0
0 | 1.4 | 100 | 98.95 | 1001030.1 | 0
0 | 2 | 100 | 97.76 | 988933.7 | 0
0 | 2.6 | 100 | 95.33 | 964374.2 | 0
0 | 2.9 | 100 | 93.57 | 946587.6 | 0
0 | 3.4 | 100 | 90.8 | 918552.3 | 0
0 | 4 | 100 | 84.54 | 855241.6 | 0
0 | 4.5 | 100 | 70.62 | 714373.3 | 0
0 | 5 | 100 | 42.56 | 430561.6 | 0
0 | 5.5 | 100 | 21.77 | 220232.7 | 0
0 | 6 | 100 | 13.5 | 136551.9 | 0
0 | 6.4 | 100 | 9.63 | 97407.6 | 0
0 | 7 | 100 | 7.15 | 72288.5 | 0
0 | 7.5 | 100 | 5.63 | 56969.8 | 0
0 | 7.9 | 100 | 4.59 | 46407 | 0
0 | 8.4 | 100 | 3.75 | 37979.7 | 0
0 | 9 | 100 | 3.05 | 30805.9 | 0
0 | 9.5 | 100 | 2.56 | 25915.6 | 0
0 | 10 | 100 | 2.19 | 22175.6 | 0
Measurement time: | 12/04/2023 20:01:27
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | <none>
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 250 | 17.64 | 351239.6 | 0
0 | 0 | 245.1 | 18.15 | 361275.4 | 0
0 | 0 | 240 | 18.84 | 375142.6 | 0
0 | 0 | 235 | 19.47 | 387650.3 | 0
0 | 0 | 230 | 20.16 | 401417.3 | 0
0 | 0 | 225 | 20.86 | 415292.4 | 0
0 | 0 | 220.1 | 21.59 | 429884.7 | 0
0 | 0 | 215.1 | 22.32 | 444419.9 | 0
0 | 0 | 210.1 | 23.07 | 459230 | 0
0 | 0 | 205.1 | 23.82 | 474244.8 | 0
0 | 0 | 200.1 | 24.68 | 491378.6 | 0
0 | 0 | 195.1 | 25.56 | 508790.5 | 0
0 | 0 | 190.1 | 26.44 | 526435.7 | 0
0 | 0 | 185.1 | 27.42 | 545894 | 0
0 | 0 | 180 | 28.36 | 564511.4 | 0
0 | 0 | 175.1 | 29.4 | 585358.9 | 0
0 | 0 | 170.1 | 30.41 | 605441.2 | 0
0 | 0 | 165.1 | 31.65 | 629998.4 | 0
0 | 0 | 160.1 | 32.67 | 650324.5 | 0
0 | 0 | 155.1 | 33.86 | 674099.8 | 0
0 | 0 | 150.1 | 35.09 | 698567 | 0
0 | 0 | 145.1 | 36.34 | 723535.6 | 0
0 | 0 | 140 | 37.85 | 753512.5 | 0
0 | 0 | 135.1 | 39.12 | 778777.1 | 0
0 | 0 | 130.1 | 40.48 | 805778.5 | 0
0 | 0 | 125 | 42.13 | 838817.2 | 0
0 | 0 | 120.1 | 43.74 | 870753.8 | 0
0 | 0 | 115.1 | 45.35 | 902822.2 | 0
0 | 0 | 110.1 | 47.01 | 935793.5 | 0
0 | 0 | 105.1 | 48.84 | 972280.3 | 0
0 | 0 | 100.1 | 50.72 | 1009654.3 | 0
0 | 0 | 95.1 | 52.64 | 1047872 | 0
0 | 0 | 90.1 | 54.68 | 1088499.8 | 0
0 | 0 | 85 | 56.83 | 1131429.6 | 0
0 | 0 | 80 | 59.08 | 1176089.6 | 0
0 | 0 | 75 | 61.32 | 1220818 | 0
0 | 0 | 70 | 63.76 | 1269418 | 0
0 | 0 | 65 | 66.34 | 1320731.1 | 0
0 | 0 | 60 | 69.1 | 1375596.9 | 0
0 | 0 | 54.9 | 71.66 | 1426562.8 | 0
0 | 0 | 50 | 74.69 | 1486986.2 | 0
0 | 0 | 44.9 | 77.86 | 1550065.9 | 0
0 | 0 | 39.9 | 80.88 | 1610136.8 | 0
0 | 0 | 34.9 | 84.08 | 1673811.3 | 0
0 | 0 | 29.9 | 87.63 | 1744620.9 | 0
0 | 0 | 24.9 | 91.27 | 1817027.2 | 0
0 | 0 | 20.1 | 94.74 | 1885999.7 | 0
0 | 0 | 19.1 | 95.31 | 1897468.8 | 0
0 | 0 | 18.1 | 96.08 | 1912857.3 | 0
0 | 0 | 17.1 | 96.8 | 1927058.1 | 0
0 | 0 | 16.1 | 97.45 | 1939968.1 | 0
0 | 0 | 15.1 | 98.26 | 1956085.7 | 0
0 | 0 | 14.1 | 98.81 | 1967050.1 | 0
0 | 0 | 13.1 | 99.28 | 1976384.9 | 0
0 | 0 | 12.1 | 99.62 | 1983253.8 | 0
0 | 0 | 11.1 | 99.85 | 1987881.1 | 0
0 | 0 | 10.1 | 100 | 1990801.1 | 0
0 | 0 | 9 | 99.64 | 1983605.7 | 0
0 | 0 | 8 | 99.19 | 1974644.6 | 0
0 | 0 | 7 | 97.86 | 1948193.3 | 0
0 | 0 | 6 | 96.09 | 1913028.7 | 0
0 | 0 | 5 | 93.24 | 1856301.7 | 0
0 | 0 | 4 | 88.87 | 1769324 | 0
0 | 0 | 3 | 82.07 | 1633882.6 | 0
0 | 0 | 2 | 71.71 | 1427586.1 | 0
0 | 0 | 1 | 54.14 | 1077751.1 | 0
0 | 0 | 0 | 51.49 | 1024999.5 | 0
0 | 0 | -0.5 | 51.55 | 1026298.4 | 0

## X6
Measurement time: | 12/04/2023 18:39:08
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | Stealth CI transmission
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-10 | 0 | 100 | 0.34 | 3551.6 | 1.24
-9.5 | 0 | 100 | 0.55 | 5677.7 | 1.98
-9 | 0 | 100 | 0.8 | 8282.9 | 2.89
-8.5 | 0 | 100 | 1.15 | 11867.7 | 4.13
-8 | 0 | 100 | 1.57 | 16311.2 | 5.69
-7.5 | 0 | 100 | 2.19 | 22670.9 | 7.91
-7 | 0 | 100 | 2.98 | 30891.1 | 10.77
-6.6 | 0 | 100 | 4.25 | 44065 | 15.35
-6 | 0 | 100 | 6.4 | 66341.8 | 23.16
-5.5 | 0 | 100 | 9.65 | 100055.6 | 34.87
-5 | 0 | 100 | 15.62 | 161850.1 | 56.38
-4.5 | 0 | 100 | 20.78 | 215368.5 | 75.19
-4 | 0 | 100 | 23.58 | 244412 | 85.3
-3.5 | 0 | 100 | 25.26 | 261765.8 | 91.27
-3 | 0 | 100 | 26.29 | 272445.3 | 95.02
-2.5 | 0 | 100 | 26.92 | 278971.3 | 97.27
-2 | 0 | 100 | 27.36 | 283605.8 | 98.87
-1.5 | 0 | 100 | 27.52 | 285203.6 | 99.46
-1 | 0 | 100 | 27.75 | 287628.6 | 100.28
-0.6 | 0 | 100 | 27.77 | 287860.7 | 100.46
-0.1 | 0 | 100 | 27.66 | 286654.5 | 100.04
0.5 | 0 | 100 | 27.58 | 285876.2 | 99.7
1 | 0 | 100 | 27.4 | 284036.8 | 99.07
1.5 | 0 | 100 | 27.15 | 281404.3 | 98.07
1.9 | 0 | 100 | 26.44 | 274023 | 95.6
2.6 | 0 | 100 | 25.31 | 262284.8 | 91.4
3 | 0 | 100 | 23.78 | 246517.9 | 86.01
3.5 | 0 | 100 | 21.58 | 223673.3 | 77.92
4 | 0 | 100 | 17.37 | 180072.1 | 62.8
4.6 | 0 | 100 | 11.12 | 115299 | 40.24
5 | 0 | 100 | 7.58 | 78542.2 | 27.37
5.5 | 0 | 100 | 4.95 | 51330.2 | 17.89
6 | 0 | 100 | 3.36 | 34785.7 | 12.13
6.5 | 0 | 100 | 2.31 | 23896.9 | 8.33
6.9 | 0 | 100 | 1.71 | 17711.5 | 6.18
7.5 | 0 | 100 | 1.2 | 12434.1 | 4.33
8 | 0 | 100 | 0.81 | 8351.5 | 2.91
8.5 | 0 | 100 | 0.57 | 5906.6 | 2.06
9 | 0 | 100 | 0.37 | 3851.9 | 1.34
9.5 | 0 | 100 | 0.22 | 2231.3 | 0.78
10 | 0 | 100 | 0.08 | 815.2 | 0.28
Measurement time: | 12/04/2023 18:40:43
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | Stealth CI transmission
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | -10 | 100 | 0.37 | 3843.1 | 1.34
0 | -9.6 | 100 | 0.59 | 6131.6 | 2.14
0 | -9 | 100 | 0.86 | 8874.8 | 3.1
0 | -8.5 | 100 | 1.3 | 13454.4 | 4.7
0 | -8 | 100 | 1.83 | 19002.7 | 6.64
0 | -7.5 | 100 | 2.63 | 27299 | 9.52
0 | -7.1 | 100 | 3.54 | 36731.7 | 12.83
0 | -6.5 | 100 | 5.56 | 57589.2 | 20.07
0 | -6 | 100 | 8.38 | 86815.9 | 30.28
0 | -5.5 | 100 | 12.22 | 126641.7 | 44.16
0 | -4.9 | 100 | 17.08 | 177044.2 | 61.84
0 | -4.5 | 100 | 20.8 | 215533.6 | 75.22
0 | -3.9 | 100 | 23.72 | 245828.6 | 85.77
0 | -3.5 | 100 | 25.27 | 261888.7 | 91.45
0 | -3.1 | 100 | 26.26 | 272131.3 | 94.9
0 | -2.5 | 100 | 26.74 | 277143.4 | 96.7
0 | -2 | 100 | 27.12 | 281074.7 | 98.09
0 | -1.5 | 100 | 27.46 | 284609.4 | 99.33
0 | -1 | 100 | 27.56 | 285643.3 | 99.62
0 | -0.5 | 100 | 27.78 | 287889 | 100.6
0 | 0 | 100 | 27.65 | 286548.4 | 100
0 | 0.5 | 100 | 27.67 | 286759 | 100.12
0 | 1 | 100 | 27.64 | 286462.8 | 100
0 | 1.5 | 100 | 27.36 | 283581.3 | 98.97
0 | 1.9 | 100 | 27.01 | 279930.7 | 97.68
0 | 2.5 | 100 | 26.5 | 274634.5 | 95.75
0 | 3 | 100 | 25.68 | 266158 | 92.86
0 | 3.5 | 100 | 24.67 | 255705.9 | 89.25
0 | 4 | 100 | 23.14 | 239822.6 | 83.7
0 | 4.5 | 100 | 20.82 | 215776.9 | 75.28
0 | 5 | 100 | 17.12 | 177485.4 | 61.97
0 | 5.5 | 100 | 12.95 | 134226.8 | 46.88
0 | 6 | 100 | 8.33 | 86306.5 | 30.13
0 | 6.4 | 100 | 5.54 | 57468.5 | 20.07
0 | 6.9 | 100 | 3.68 | 38145.3 | 13.32
0 | 7.4 | 100 | 2.51 | 26037.1 | 9.08
0 | 8 | 100 | 1.72 | 17820.4 | 6.22
0 | 8.4 | 100 | 1.24 | 12803.6 | 4.47
0 | 9 | 100 | 0.86 | 8929.2 | 3.12
0 | 9.4 | 100 | 0.59 | 6096.5 | 2.13
0 | 10 | 100 | 0.37 | 3826.7 | 1.34
Measurement time: | 12/04/2023 18:43:02
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | Stealth CI transmission
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 250 | 10.11 | 104818.5 | 20.08
0 | 0 | 245.1 | 10.53 | 109140.1 | 20.91
0 | 0 | 240 | 10.8 | 111906.1 | 21.42
0 | 0 | 235 | 11.19 | 115931.5 | 22.2
0 | 0 | 230.1 | 11.58 | 119975.7 | 22.98
0 | 0 | 225 | 11.93 | 123600.8 | 23.66
0 | 0 | 220.1 | 12.33 | 127788.4 | 24.46
0 | 0 | 215.1 | 12.72 | 131808.3 | 25.25
0 | 0 | 210.1 | 13.22 | 137016.7 | 26.28
0 | 0 | 205.1 | 13.6 | 140986 | 27.03
0 | 0 | 200.1 | 14.09 | 146079.6 | 27.98
0 | 0 | 195 | 14.54 | 150721.1 | 28.86
0 | 0 | 190.1 | 14.99 | 155395.1 | 29.75
0 | 0 | 185.1 | 15.59 | 161604.4 | 30.89
0 | 0 | 180 | 16.08 | 166692.6 | 31.98
0 | 0 | 175.1 | 16.66 | 172664.2 | 33.05
0 | 0 | 170 | 17.19 | 178174 | 34.14
0 | 0 | 165.1 | 17.81 | 184591.3 | 35.37
0 | 0 | 160.1 | 18.37 | 190369.5 | 36.45
0 | 0 | 155.1 | 19 | 196917 | 37.7
0 | 0 | 150.1 | 19.61 | 203299.7 | 38.94
0 | 0 | 145.1 | 20.29 | 210264.4 | 40.26
0 | 0 | 140 | 20.96 | 217194.1 | 41.64
0 | 0 | 135.1 | 21.8 | 225917.7 | 43.2
0 | 0 | 130 | 22.44 | 232618.6 | 44.62
0 | 0 | 125.1 | 23.19 | 240374.4 | 46.04
0 | 0 | 120.1 | 24.16 | 250385.2 | 47.99
0 | 0 | 115.1 | 24.88 | 257856.6 | 49.37
0 | 0 | 110.1 | 25.81 | 267513.7 | 51.27
0 | 0 | 105.1 | 26.73 | 277058.7 | 53.04
0 | 0 | 100.1 | 27.72 | 287253.8 | 55.06
0 | 0 | 95.1 | 28.58 | 296213.5 | 56.72
0 | 0 | 90.1 | 29.54 | 306211.2 | 58.55
0 | 0 | 85 | 30.75 | 318702.1 | 60.99
0 | 0 | 80 | 31.77 | 329310.1 | 63.03
0 | 0 | 75 | 33.02 | 342277.1 | 65.51
0 | 0 | 70 | 34.18 | 354295.5 | 67.84
0 | 0 | 65 | 35.3 | 365862.5 | 70.01
0 | 0 | 60 | 36.54 | 378756.3 | 72.45
0 | 0 | 54.9 | 37.99 | 393733.2 | 75.46
0 | 0 | 50 | 39.45 | 408902.4 | 78.27
0 | 0 | 44.9 | 40.79 | 422743.1 | 80.95
0 | 0 | 39.9 | 42.26 | 438003.8 | 83.82
0 | 0 | 34.9 | 43.91 | 455147.9 | 87.21
0 | 0 | 29.9 | 45.38 | 470325.1 | 90.06
0 | 0 | 24.9 | 47.02 | 487355.1 | 93.36
0 | 0 | 20 | 48.58 | 503460.1 | 96.42
0 | 0 | 19.1 | 48.82 | 505984.3 | 96.79
0 | 0 | 18.1 | 49.27 | 510619.7 | 97.73
0 | 0 | 17.1 | 49.56 | 513693.9 | 98.29
0 | 0 | 16.1 | 49.68 | 514889.3 | 98.65
0 | 0 | 15.1 | 50.02 | 518472.6 | 99.31
0 | 0 | 14.1 | 50.28 | 521137.4 | 99.62
0 | 0 | 13 | 50.35 | 521874.6 | 99.93
0 | 0 | 12 | 50.47 | 523086.9 | 100
0 | 0 | 11 | 50.3 | 521299.6 | 99.84
0 | 0 | 10 | 50.11 | 519394.1 | 99.36
0 | 0 | 9.1 | 49.74 | 515496.3 | 98.76
0 | 0 | 8 | 49.14 | 509313 | 97.55
0 | 0 | 7 | 48.03 | 497779.2 | 95.28
0 | 0 | 6.1 | 46.66 | 483644.4 | 92.62
0 | 0 | 5 | 44.43 | 460491.7 | 88.01
0 | 0 | 4 | 41.12 | 426235.5 | 81.62
0 | 0 | 3 | 36.83 | 381747.8 | 73.01
0 | 0 | 2 | 30.79 | 319163.8 | 61.12
0 | 0 | 1 | 21.42 | 221969.8 | 42.49
0 | 0 | 0 | 18.5 | 191745.4 | 36.71
0 | 0 | -0.5 | 18.51 | 191797 | 36.7

## Pas de mesure
Measurement time: | 12/04/2023 18:35:42
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | Stealth CI transmission
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
10 | 0 | 100 | 1.77 | 18383.9 | 1.79
9 | 0 | 100 | 2.71 | 28073.6 | 2.73
8 | 0 | 100 | 4.2 | 43558.2 | 4.24
7 | 0 | 100 | 6.88 | 71303.6 | 6.94
6 | 0 | 100 | 12.86 | 133278.1 | 12.98
5 | 0 | 100 | 30.3 | 313993.6 | 30.57
4 | 0 | 100 | 65.91 | 683095.6 | 66.48
3 | 0 | 100 | 88 | 912065.8 | 88.76
2 | 0 | 100 | 95.15 | 986163.3 | 96.02
1 | 0 | 100 | 98.08 | 1016582.5 | 98.96
0 | 0 | 100 | 99.11 | 1027179.5 | 100
-1 | 0 | 100 | 98.93 | 1025349.3 | 99.87
-2 | 0 | 100 | 97.44 | 1009891.6 | 98.31
-3 | 0 | 100 | 93.95 | 973793.2 | 94.81
-4 | 0 | 100 | 83.59 | 866420.1 | 84.39
-5 | 0 | 100 | 49.93 | 517509.1 | 50.38
-5.9 | 0 | 100 | 20.8 | 215631.1 | 20.99
-7 | 0 | 100 | 9.99 | 103510.4 | 10.08
-8 | 0 | 100 | 5.77 | 59838.5 | 5.83
-9 | 0 | 100 | 3.58 | 37092.4 | 3.61
-10 | 0 | 100 | 2.38 | 24667.3 | 2.4
Measurement time: | 12/04/2023 18:36:37
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele
Reference detector: | Stealth CI transmission
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 10 x 10 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | StepByStep
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 10 | 100 | 2.4 | 24885.2 | 2.43
0 | 9 | 100 | 3.8 | 39376.5 | 3.84
0 | 8.1 | 100 | 6.17 | 63977.9 | 6.24
0 | 7.1 | 100 | 11.49 | 119091.2 | 11.63
0 | 6 | 100 | 26.6 | 275692.9 | 26.91
0 | 5 | 100 | 59.85 | 620296.5 | 60.52
0 | 4 | 100 | 83.11 | 861434.2 | 84.08
0 | 3 | 100 | 92.29 | 956592.1 | 93.34
0 | 2.1 | 100 | 96.87 | 1003967 | 98.05
0 | 1.1 | 100 | 98.55 | 1021391.3 | 99.7
0 | 0.1 | 100 | 98.89 | 1024948.3 | 100
0 | -0.9 | 100 | 98.75 | 1023489.6 | 99.95
0 | -2 | 100 | 97.56 | 1011147.2 | 98.68
0 | -3 | 100 | 95.08 | 985431.7 | 96.19
0 | -4 | 100 | 86.11 | 892522.9 | 87.11
0 | -5 | 100 | 60.5 | 627095.6 | 61.22
0 | -5.9 | 100 | 30.99 | 321157.7 | 31.35
0 | -7 | 100 | 13.93 | 144372.8 | 14.1
0 | -8 | 100 | 7.22 | 74859.9 | 7.31
0 | -8.9 | 100 | 4.24 | 43955.8 | 4.29
0 | -10 | 100 | 2.69 | 27846.1 | 2.72

## Graphes
Taille de champ