## 10_Profils position chambre ref
Measurement time: | 12/08/2022 19:20:27
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
-100 | 0 | 100 | 2.58 | 6387.8 | 3.81 | 3.8565
-99.3 | 0 | 100 | 2.65 | 6554.7 | 3.95 | 3.9611
-97.9 | 0 | 100 | 2.79 | 6909.7 | 4.17 | 4.1704
-96.9 | 0 | 100 | 2.88 | 7114.7 | 4.29 | 4.3049
-95.7 | 0 | 100 | 2.97 | 7333.3 | 4.42 | 4.4395
-94.4 | 0 | 100 | 3.05 | 7531.9 | 4.54 | 4.559
-93.2 | 0 | 100 | 3.17 | 7833.9 | 4.73 | 4.7384
-92 | 0 | 100 | 3.23 | 7986.9 | 4.82 | 4.8281
-90.8 | 0 | 100 | 3.38 | 8356.4 | 5.05 | 5.0523
-89.6 | 0 | 100 | 3.5 | 8651.5 | 5.22 | 5.2317
-88.4 | 0 | 100 | 3.59 | 8868 | 5.34 | 5.3662
-87.2 | 0 | 100 | 3.71 | 9179.2 | 5.54 | 5.5456
-86 | 0 | 100 | 3.87 | 9566.5 | 5.77 | 5.7848
-84.8 | 0 | 100 | 4.01 | 9901.9 | 5.97 | 5.994
-83.6 | 0 | 100 | 4.1 | 10126.5 | 6.11 | 6.1286
-82.4 | 0 | 100 | 4.24 | 10488.8 | 6.33 | 6.3378
-81.2 | 0 | 100 | 4.42 | 10929.9 | 6.59 | 6.6069
-80 | 0 | 100 | 4.59 | 11337.9 | 6.84 | 6.861
-78.8 | 0 | 100 | 4.76 | 11758.5 | 7.11 | 7.1151
-77.6 | 0 | 100 | 4.93 | 12186 | 7.36 | 7.3692
-76.3 | 0 | 100 | 5.16 | 12759.8 | 7.71 | 7.713
-75.2 | 0 | 100 | 5.35 | 13216.4 | 7.97 | 7.997
-73.9 | 0 | 100 | 5.55 | 13719 | 8.29 | 8.296
-72.8 | 0 | 100 | 5.81 | 14375.5 | 8.67 | 8.6846
-71.6 | 0 | 100 | 5.99 | 14800.6 | 8.93 | 8.9537
-70.3 | 0 | 100 | 6.26 | 15486.5 | 9.35 | 9.3572
-69.1 | 0 | 100 | 6.51 | 16105.2 | 9.73 | 9.7309
-67.9 | 0 | 100 | 6.83 | 16878.9 | 10.19 | 10.2093
-66.7 | 0 | 100 | 7.19 | 17773.9 | 10.75 | 10.7474
-65.5 | 0 | 100 | 7.53 | 18608 | 11.25 | 11.2556
-64.3 | 0 | 100 | 8.05 | 19895.9 | 12.02 | 12.0329
-63.1 | 0 | 100 | 8.73 | 21574.5 | 13.04 | 13.0493
-61.9 | 0 | 100 | 9.76 | 24125.2 | 14.58 | 14.5889
-60.7 | 0 | 100 | 11.57 | 28615.6 | 17.27 | 17.2945
-59.5 | 0 | 100 | 14.5 | 35856.5 | 21.66 | 21.6741
-58.3 | 0 | 100 | 19.01 | 46987.8 | 28.39 | 28.4155
-57.1 | 0 | 100 | 24.96 | 61706.7 | 37.21 | 37.3094
-55.9 | 0 | 100 | 32.41 | 80126.8 | 48.45 | 48.4454
-54.7 | 0 | 100 | 40.85 | 100991.3 | 61.05 | 61.0613
-53.5 | 0 | 100 | 48.28 | 119363.6 | 72.24 | 72.1674
-52.3 | 0 | 100 | 54.88 | 135681.9 | 81.99 | 82.0329
-51 | 0 | 100 | 59.11 | 146148.3 | 88.33 | 88.3558
-49.9 | 0 | 100 | 61.44 | 151895.3 | 91.87 | 91.8386
-48.6 | 0 | 100 | 62.7 | 155013 | 93.74 | 93.722
-47.5 | 0 | 100 | 63.36 | 156649.1 | 94.68 | 94.7085
-46.2 | 0 | 100 | 63.78 | 157681.8 | 95.4 | 95.3363
-45 | 0 | 100 | 64.26 | 158885 | 96.08 | 96.0538
-43.8 | 0 | 100 | 64.48 | 159416 | 96.52 | 96.3827
-42.6 | 0 | 100 | 64.72 | 160018.6 | 96.62 | 96.7414
-41.4 | 0 | 100 | 64.79 | 160190.4 | 96.87 | 96.846
-40.2 | 0 | 100 | 64.95 | 160572.3 | 97.19 | 97.0852
-39 | 0 | 100 | 65.21 | 161219 | 97.52 | 97.4738
-37.8 | 0 | 100 | 65.45 | 161806.1 | 97.69 | 97.8326
-36.6 | 0 | 100 | 65.49 | 161910.3 | 97.97 | 97.8924
-35.4 | 0 | 100 | 65.68 | 162387.3 | 98.22 | 98.1764
-34.2 | 0 | 100 | 65.77 | 162602.6 | 98.51 | 98.3109
-33 | 0 | 100 | 65.78 | 162633 | 98.47 | 98.3259
-31.8 | 0 | 100 | 66.09 | 163393.7 | 98.71 | 98.7892
-30.6 | 0 | 100 | 66.09 | 163395.2 | 98.79 | 98.7892
-29.4 | 0 | 100 | 66.2 | 163666 | 98.98 | 98.9537
-28.2 | 0 | 100 | 66.35 | 164041 | 99.43 | 99.1779
-27 | 0 | 100 | 66.33 | 164002.1 | 99.08 | 99.148
-25.8 | 0 | 100 | 66.34 | 164026.6 | 99.18 | 99.1629
-24.6 | 0 | 100 | 66.49 | 164394.1 | 99.37 | 99.3871
-23.4 | 0 | 100 | 66.55 | 164536.7 | 99.48 | 99.4768
-22.2 | 0 | 100 | 66.63 | 164729.5 | 99.57 | 99.5964
-21 | 0 | 100 | 66.76 | 165051.5 | 99.7 | 99.7907
-19.8 | 0 | 100 | 66.72 | 164945.4 | 99.78 | 99.7309
-18.6 | 0 | 100 | 66.9 | 165409.2 | 100 | 100
-17.4 | 0 | 100 | 66.79 | 165136.3 | 99.77 | 99.8356
-16.2 | 0 | 100 | 66.73 | 164973.7 | 99.85 | 99.7459
-15 | 0 | 100 | 66.72 | 164945.8 | 99.62 | 99.7309
-13.7 | 0 | 100 | 66.71 | 164919.9 | 99.66 | 99.716
-12.6 | 0 | 100 | 66.77 | 165083.5 | 99.91 | 99.8057
-11.3 | 0 | 100 | 66.72 | 164955.9 | 99.7 | 99.7309
-10.1 | 0 | 100 | 66.57 | 164593.8 | 99.71 | 99.5067
-8.9 | 0 | 100 | 66.56 | 164570.5 | 99.5 | 99.4918
-7.7 | 0 | 100 | 66.55 | 164536.5 | 99.67 | 99.4768
-6.5 | 0 | 100 | 66.52 | 164460.5 | 99.42 | 99.432
-5.3 | 0 | 100 | 66.43 | 164233.8 | 99.36 | 99.2975
-4.1 | 0 | 100 | 66.33 | 163986.3 | 99.15 | 99.148
-2.9 | 0 | 100 | 66.4 | 164162.3 | 99.27 | 99.2526
-1.7 | 0 | 100 | 66.41 | 164196.4 | 99.41 | 99.2676
-0.5 | 0 | 100 | 66.42 | 164217.2 | 99.38 | 99.2825
0.7 | 0 | 100 | 66.43 | 164237 | 99.36 | 99.2975
1.9 | 0 | 100 | 66.27 | 163838.2 | 99.24 | 99.0583
3.1 | 0 | 100 | 66.24 | 163780.8 | 99.22 | 99.0135
4.3 | 0 | 100 | 66.33 | 163996.4 | 99.25 | 99.148
5.5 | 0 | 100 | 66.24 | 163774.7 | 99.18 | 99.0135
6.7 | 0 | 100 | 66.31 | 163951.3 | 99.27 | 99.1181
7.9 | 0 | 100 | 66.25 | 163782.4 | 99.2 | 99.0284
9.1 | 0 | 100 | 66.31 | 163943 | 99.13 | 99.1181
10.3 | 0 | 100 | 66.22 | 163708.7 | 99.07 | 98.9836
11.5 | 0 | 100 | 66.24 | 163773.5 | 99.15 | 99.0135
12.7 | 0 | 100 | 66.16 | 163564.4 | 99.09 | 98.8939
13.9 | 0 | 100 | 66.16 | 163561.2 | 99.08 | 98.8939
15.2 | 0 | 100 | 66.16 | 163566.1 | 98.84 | 98.8939
16.3 | 0 | 100 | 66.16 | 163571.2 | 98.89 | 98.8939
17.6 | 0 | 100 | 66.23 | 163756.1 | 98.99 | 98.9985
18.7 | 0 | 100 | 66.12 | 163482.8 | 99.06 | 98.8341
20 | 0 | 100 | 66.12 | 163482.4 | 98.77 | 98.8341
21.2 | 0 | 100 | 66.02 | 163229.3 | 98.79 | 98.6846
22.4 | 0 | 100 | 65.94 | 163038.3 | 98.67 | 98.565
23.6 | 0 | 100 | 65.99 | 163160.1 | 98.64 | 98.6398
24.8 | 0 | 100 | 65.95 | 163059.8 | 98.62 | 98.58
26 | 0 | 100 | 65.77 | 162603.5 | 98.34 | 98.3109
27.2 | 0 | 100 | 65.53 | 162025.3 | 98.19 | 97.9522
28.4 | 0 | 100 | 65.35 | 161563.4 | 97.89 | 97.6831
29.6 | 0 | 100 | 65.37 | 161620.9 | 97.78 | 97.713
30.8 | 0 | 100 | 65.16 | 161093.9 | 97.6 | 97.3991
32 | 0 | 100 | 65.15 | 161072.1 | 97.65 | 97.3842
33.2 | 0 | 100 | 64.96 | 160612.2 | 97.23 | 97.1001
34.4 | 0 | 100 | 64.84 | 160296.2 | 97.09 | 96.9208
35.6 | 0 | 100 | 64.68 | 159908.7 | 96.8 | 96.6816
36.8 | 0 | 100 | 64.5 | 159457.9 | 96.4 | 96.4126
38 | 0 | 100 | 64.41 | 159240 | 96.35 | 96.278
39.2 | 0 | 100 | 64.11 | 158492 | 95.88 | 95.8296
40.4 | 0 | 100 | 63.95 | 158097.1 | 95.68 | 95.5904
41.6 | 0 | 100 | 63.76 | 157629.8 | 95.48 | 95.3064
42.8 | 0 | 100 | 63.53 | 157068.5 | 95.13 | 94.9626
44.1 | 0 | 100 | 63.25 | 156370 | 94.74 | 94.5441
45.2 | 0 | 100 | 62.83 | 155336.1 | 94.08 | 93.9163
46.5 | 0 | 100 | 62.39 | 154243.3 | 93.3 | 93.2586
47.7 | 0 | 100 | 61.6 | 152302.7 | 92.28 | 92.0777
48.9 | 0 | 100 | 60.43 | 149412.2 | 90.43 | 90.3288
50.1 | 0 | 100 | 58.2 | 143896.7 | 87.11 | 86.9955
51.3 | 0 | 100 | 54.8 | 135473.6 | 82.02 | 81.9133
52.5 | 0 | 100 | 49.64 | 122730.1 | 74.33 | 74.2003
53.7 | 0 | 100 | 43.17 | 106719.9 | 64.63 | 64.5291
54.9 | 0 | 100 | 35.47 | 87695.9 | 53.15 | 53.0194
56.1 | 0 | 100 | 27.46 | 67899.5 | 41.14 | 41.0463
57.3 | 0 | 100 | 20.45 | 50548.5 | 30.6 | 30.568
58.5 | 0 | 100 | 14.92 | 36880.5 | 22.33 | 22.3019
59.7 | 0 | 100 | 11.66 | 28838.6 | 17.46 | 17.429
60.9 | 0 | 100 | 9.84 | 24324.8 | 14.73 | 14.7085
62.1 | 0 | 100 | 8.73 | 21595.8 | 13.07 | 13.0493
63.3 | 0 | 100 | 8.06 | 19929 | 12.05 | 12.0478
64.5 | 0 | 100 | 7.49 | 18521.6 | 11.2 | 11.1958
65.7 | 0 | 100 | 7.15 | 17680 | 10.69 | 10.6876
66.9 | 0 | 100 | 6.77 | 16745.8 | 10.12 | 10.1196
68.1 | 0 | 100 | 6.44 | 15931.4 | 9.62 | 9.6263
69.3 | 0 | 100 | 6.21 | 15355 | 9.3 | 9.2825
70.5 | 0 | 100 | 5.93 | 14661.7 | 8.87 | 8.864
71.7 | 0 | 100 | 5.67 | 14016.6 | 8.49 | 8.4753
72.9 | 0 | 100 | 5.43 | 13434.6 | 8.14 | 8.1166
74.1 | 0 | 100 | 5.3 | 13102.3 | 7.92 | 7.9223
75.3 | 0 | 100 | 5.07 | 12539.5 | 7.59 | 7.5785
76.5 | 0 | 100 | 4.87 | 12042.4 | 7.29 | 7.2795
77.7 | 0 | 100 | 4.71 | 11641 | 7.06 | 7.0404
79 | 0 | 100 | 4.55 | 11248.6 | 6.81 | 6.8012
80.2 | 0 | 100 | 4.4 | 10872.4 | 6.6 | 6.577
81.4 | 0 | 100 | 4.24 | 10484 | 6.35 | 6.3378
82.6 | 0 | 100 | 4.06 | 10031.7 | 6.08 | 6.0688
83.8 | 0 | 100 | 3.94 | 9737.7 | 5.9 | 5.8894
85 | 0 | 100 | 3.79 | 9372.5 | 5.69 | 5.6652
86.2 | 0 | 100 | 3.71 | 9171.7 | 5.56 | 5.5456
87.4 | 0 | 100 | 3.54 | 8761.1 | 5.32 | 5.2915
88.6 | 0 | 100 | 3.45 | 8525.8 | 5.17 | 5.157
89.8 | 0 | 100 | 3.32 | 8220.4 | 4.98 | 4.9626
91 | 0 | 100 | 3.23 | 7997.7 | 4.85 | 4.8281
92.2 | 0 | 100 | 3.12 | 7706.9 | 4.66 | 4.6637
93.4 | 0 | 100 | 3.05 | 7538.5 | 4.56 | 4.559
94.6 | 0 | 100 | 2.9 | 7172.7 | 4.34 | 4.3348
95.8 | 0 | 100 | 2.77 | 6855.5 | 4.15 | 4.1405
97 | 0 | 100 | 2.76 | 6829.9 | 4.13 | 4.1256
98.2 | 0 | 100 | 2.65 | 6563.1 | 3.97 | 3.9611
100 | 0 | 100 | 2.55 | 6315.7 | 3.82 | 3.8117
Measurement time: | 12/08/2022 19:23:34
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
100 | 0 | 100 | 2.7 | 6669 | 3.82 | 4.0425
99.4 | 0 | 100 | 2.6 | 6417 | 3.87 | 3.8928
98.1 | 0 | 100 | 2.7 | 6672.2 | 4.03 | 4.0425
97 | 0 | 100 | 2.8 | 6917.8 | 4.17 | 4.1922
95.9 | 0 | 100 | 2.87 | 7087.7 | 4.28 | 4.2971
94.7 | 0 | 100 | 2.95 | 7286.8 | 4.4 | 4.4168
93.4 | 0 | 100 | 3.07 | 7599.2 | 4.59 | 4.5965
92.3 | 0 | 100 | 3.14 | 7775.4 | 4.69 | 4.7013
91 | 0 | 100 | 3.26 | 8064.2 | 4.86 | 4.881
89.8 | 0 | 100 | 3.39 | 8389.3 | 5.06 | 5.0756
88.6 | 0 | 100 | 3.5 | 8661.8 | 5.22 | 5.2403
87.4 | 0 | 100 | 3.62 | 8945.4 | 5.4 | 5.42
86.2 | 0 | 100 | 3.77 | 9321.1 | 5.63 | 5.6446
85 | 0 | 100 | 3.87 | 9567.9 | 5.78 | 5.7943
83.8 | 0 | 100 | 3.98 | 9838.7 | 5.93 | 5.959
82.6 | 0 | 100 | 4.11 | 10150 | 6.13 | 6.1536
81.4 | 0 | 100 | 4.3 | 10622.2 | 6.41 | 6.4381
80.2 | 0 | 100 | 4.45 | 11009.3 | 6.65 | 6.6627
79 | 0 | 100 | 4.6 | 11367.9 | 6.87 | 6.8873
77.8 | 0 | 100 | 4.78 | 11819.5 | 7.13 | 7.1568
76.6 | 0 | 100 | 4.93 | 12197.4 | 7.38 | 7.3813
75.4 | 0 | 100 | 5.11 | 12640.6 | 7.64 | 7.6508
74.2 | 0 | 100 | 5.3 | 13115.3 | 7.93 | 7.9353
73 | 0 | 100 | 5.49 | 13567.8 | 8.21 | 8.2198
71.8 | 0 | 100 | 5.77 | 14276.3 | 8.64 | 8.639
70.6 | 0 | 100 | 5.93 | 14671.6 | 8.86 | 8.8786
69.4 | 0 | 100 | 6.19 | 15315.6 | 9.26 | 9.2679
68.2 | 0 | 100 | 6.51 | 16095.7 | 9.73 | 9.747
67 | 0 | 100 | 6.82 | 16850.5 | 10.18 | 10.2111
65.8 | 0 | 100 | 7.17 | 17727.2 | 10.71 | 10.7351
64.6 | 0 | 100 | 7.57 | 18722.9 | 11.32 | 11.334
63.4 | 0 | 100 | 8.1 | 20022.4 | 12.09 | 12.1276
62.2 | 0 | 100 | 8.78 | 21702.4 | 13.11 | 13.1457
60.9 | 0 | 100 | 9.91 | 24499.1 | 14.83 | 14.8376
59.8 | 0 | 100 | 11.69 | 28897.6 | 17.5 | 17.5026
58.5 | 0 | 100 | 15.05 | 37201.4 | 22.48 | 22.5333
57.3 | 0 | 100 | 20.72 | 51227.6 | 31.02 | 31.0226
56.2 | 0 | 100 | 27.69 | 68448.8 | 41.33 | 41.4583
54.9 | 0 | 100 | 35.9 | 88754.5 | 53.62 | 53.7506
53.8 | 0 | 100 | 43.61 | 107831.1 | 65.1 | 65.2942
52.5 | 0 | 100 | 49.8 | 123110.9 | 74.4 | 74.5621
51.3 | 0 | 100 | 55.03 | 136041.4 | 82.2 | 82.3926
50.2 | 0 | 100 | 58.33 | 144206.9 | 87.25 | 87.3334
48.9 | 0 | 100 | 60.52 | 149634.6 | 90.55 | 90.6124
47.7 | 0 | 100 | 61.73 | 152616 | 92.2 | 92.424
46.5 | 0 | 100 | 62.38 | 154225.5 | 93.49 | 93.3972
45.3 | 0 | 100 | 62.87 | 155426 | 94.2 | 94.1309
44.1 | 0 | 100 | 63.27 | 156426 | 94.69 | 94.7297
42.9 | 0 | 100 | 63.55 | 157112.8 | 94.97 | 95.149
41.7 | 0 | 100 | 63.88 | 157931.2 | 95.59 | 95.6431
40.5 | 0 | 100 | 63.96 | 158129.7 | 95.4 | 95.7628
39.3 | 0 | 100 | 64.01 | 158264.6 | 95.77 | 95.8377
38.1 | 0 | 100 | 64.15 | 158589.8 | 95.93 | 96.0473
36.8 | 0 | 100 | 64.45 | 159353.6 | 96.44 | 96.4965
35.7 | 0 | 100 | 64.58 | 159657.2 | 96.74 | 96.6911
34.5 | 0 | 100 | 64.84 | 160314.3 | 96.99 | 97.0804
33.3 | 0 | 100 | 64.94 | 160565.4 | 97.23 | 97.2301
32.1 | 0 | 100 | 65.24 | 161288.9 | 97.35 | 97.6793
30.8 | 0 | 100 | 65.24 | 161286.6 | 97.59 | 97.6793
29.7 | 0 | 100 | 65.46 | 161848.9 | 97.92 | 98.0087
28.4 | 0 | 100 | 65.6 | 162177.8 | 98.05 | 98.2183
27.2 | 0 | 100 | 65.65 | 162312.1 | 98.26 | 98.2932
26.1 | 0 | 100 | 65.75 | 162554.9 | 98.42 | 98.4429
24.8 | 0 | 100 | 65.86 | 162833 | 98.66 | 98.6076
23.6 | 0 | 100 | 65.98 | 163118.3 | 98.64 | 98.7872
22.4 | 0 | 100 | 65.9 | 162932.3 | 98.59 | 98.6675
21.2 | 0 | 100 | 66.05 | 163291.9 | 98.91 | 98.892
20 | 0 | 100 | 66.11 | 163435.3 | 99.01 | 98.9819
18.8 | 0 | 100 | 66.3 | 163914.4 | 99.17 | 99.2664
17.6 | 0 | 100 | 66.08 | 163379.3 | 98.75 | 98.937
16.4 | 0 | 100 | 66.26 | 163808 | 99.13 | 99.2065
15.2 | 0 | 100 | 66.24 | 163779.7 | 99.15 | 99.1765
14 | 0 | 100 | 66.27 | 163845 | 99.16 | 99.2214
12.8 | 0 | 100 | 66.3 | 163920.5 | 99.25 | 99.2664
11.6 | 0 | 100 | 66.31 | 163939.8 | 99.25 | 99.2813
10.4 | 0 | 100 | 66.33 | 163984.2 | 99.26 | 99.3113
9.1 | 0 | 100 | 66.26 | 163809.7 | 99.13 | 99.2065
8 | 0 | 100 | 66.3 | 163920.2 | 99.38 | 99.2664
6.7 | 0 | 100 | 66.38 | 164111.3 | 99.27 | 99.3861
5.6 | 0 | 100 | 66.27 | 163854.2 | 99.13 | 99.2214
4.4 | 0 | 100 | 66.22 | 163727.4 | 99.2 | 99.1466
3.1 | 0 | 100 | 66.31 | 163946.7 | 99.2 | 99.2813
2 | 0 | 100 | 66.18 | 163615 | 99.02 | 99.0867
0.7 | 0 | 100 | 66.35 | 164050.7 | 99.16 | 99.3412
-0.4 | 0 | 100 | 66.36 | 164074.3 | 99.25 | 99.3562
-1.7 | 0 | 100 | 66.29 | 163897.6 | 99.29 | 99.2514
-2.9 | 0 | 100 | 66.37 | 164093.1 | 99.33 | 99.3712
-4 | 0 | 100 | 66.31 | 163953.4 | 99.34 | 99.2813
-5.3 | 0 | 100 | 66.37 | 164099.4 | 99.56 | 99.3712
-6.4 | 0 | 100 | 66.46 | 164301.7 | 99.62 | 99.5059
-7.7 | 0 | 100 | 66.5 | 164411.9 | 99.5 | 99.5658
-8.9 | 0 | 100 | 66.5 | 164400.1 | 99.53 | 99.5658
-10.1 | 0 | 100 | 66.48 | 164351.3 | 99.5 | 99.5359
-11.3 | 0 | 100 | 66.61 | 164693.5 | 99.54 | 99.7305
-12.5 | 0 | 100 | 66.67 | 164827.3 | 99.77 | 99.8203
-13.7 | 0 | 100 | 66.64 | 164759.7 | 99.97 | 99.7754
-14.9 | 0 | 100 | 66.66 | 164796.5 | 99.82 | 99.8054
-16.1 | 0 | 100 | 66.67 | 164839 | 99.75 | 99.8203
-17.3 | 0 | 100 | 66.79 | 165127.2 | 99.8 | 100
-18.5 | 0 | 100 | 66.79 | 165127 | 100 | 100
-19.7 | 0 | 100 | 66.53 | 164487 | 99.76 | 99.6107
-20.9 | 0 | 100 | 66.61 | 164685.3 | 99.74 | 99.7305
-22.1 | 0 | 100 | 66.63 | 164727.3 | 99.81 | 99.7604
-23.3 | 0 | 100 | 66.57 | 164591.8 | 99.67 | 99.6706
-24.5 | 0 | 100 | 66.45 | 164285.7 | 99.45 | 99.4909
-25.7 | 0 | 100 | 66.24 | 163761.4 | 99.38 | 99.1765
-27 | 0 | 100 | 66.36 | 164060.7 | 99.34 | 99.3562
-28.1 | 0 | 100 | 66.13 | 163501.1 | 99.06 | 99.0118
-29.4 | 0 | 100 | 66.14 | 163511.2 | 99.03 | 99.0268
-30.5 | 0 | 100 | 66.16 | 163571.9 | 99.1 | 99.0567
-31.7 | 0 | 100 | 65.81 | 162706 | 98.67 | 98.5327
-32.9 | 0 | 100 | 65.75 | 162548.1 | 98.43 | 98.4429
-34.1 | 0 | 100 | 65.71 | 162461.3 | 98.44 | 98.383
-35.4 | 0 | 100 | 65.67 | 162360.6 | 98.36 | 98.3231
-36.5 | 0 | 100 | 65.52 | 161976.6 | 98.02 | 98.0985
-37.8 | 0 | 100 | 65.2 | 161200.3 | 97.57 | 97.6194
-39 | 0 | 100 | 65.11 | 160986.8 | 97.47 | 97.4847
-40.2 | 0 | 100 | 64.97 | 160620.5 | 97.27 | 97.275
-41.4 | 0 | 100 | 64.85 | 160338.5 | 97.19 | 97.0954
-42.5 | 0 | 100 | 64.69 | 159941.4 | 96.94 | 96.8558
-43.8 | 0 | 100 | 64.47 | 159399.3 | 96.59 | 96.5264
-45 | 0 | 100 | 64.1 | 158472.6 | 96.14 | 95.9725
-46.2 | 0 | 100 | 63.72 | 157531.7 | 95.45 | 95.4035
-47.4 | 0 | 100 | 63.3 | 156511 | 94.81 | 94.7747
-48.6 | 0 | 100 | 62.53 | 154603.4 | 93.69 | 93.6218
-49.8 | 0 | 100 | 61.38 | 151763.6 | 91.98 | 91.9
-51 | 0 | 100 | 58.9 | 145620.6 | 88.11 | 88.1869
-52.2 | 0 | 100 | 54.77 | 135418.6 | 81.99 | 82.0033
-53.4 | 0 | 100 | 48.2 | 119163.5 | 72.13 | 72.1665
-54.6 | 0 | 100 | 40.52 | 100179.3 | 60.6 | 60.6678
-55.9 | 0 | 100 | 32.25 | 79727.1 | 48.28 | 48.2857
-57.1 | 0 | 100 | 24.66 | 60964.1 | 36.93 | 36.9217
-58.2 | 0 | 100 | 18.9 | 46735 | 28.32 | 28.2976
-59.5 | 0 | 100 | 14.34 | 35463.9 | 21.48 | 21.4703
-60.6 | 0 | 100 | 11.47 | 28345.9 | 17.15 | 17.1732
-61.9 | 0 | 100 | 9.75 | 24112 | 14.64 | 14.598
-63.1 | 0 | 100 | 8.75 | 21635.2 | 13.12 | 13.1008
-64.3 | 0 | 100 | 8.06 | 19936.5 | 12.07 | 12.0677
-65.5 | 0 | 100 | 7.54 | 18642 | 11.29 | 11.2891
-66.7 | 0 | 100 | 7.14 | 17660.2 | 10.71 | 10.6902
-67.9 | 0 | 100 | 6.84 | 16910.7 | 10.24 | 10.2411
-69.1 | 0 | 100 | 6.54 | 16180.7 | 9.82 | 9.7919
-70.3 | 0 | 100 | 6.29 | 15555.3 | 9.42 | 9.4176
-71.5 | 0 | 100 | 6 | 14825.9 | 8.99 | 8.9834
-72.7 | 0 | 100 | 5.78 | 14294.1 | 8.66 | 8.654
-73.9 | 0 | 100 | 5.55 | 13711.2 | 8.32 | 8.3096
-75.1 | 0 | 100 | 5.34 | 13204.5 | 7.99 | 7.9952
-76.3 | 0 | 100 | 5.14 | 12699.1 | 7.69 | 7.6958
-77.5 | 0 | 100 | 4.96 | 12273.8 | 7.43 | 7.4263
-78.7 | 0 | 100 | 4.82 | 11921.7 | 7.22 | 7.2166
-80 | 0 | 100 | 4.59 | 11340.7 | 6.86 | 6.8723
-81.1 | 0 | 100 | 4.45 | 11003.3 | 6.66 | 6.6627
-82.3 | 0 | 100 | 4.32 | 10687.7 | 6.48 | 6.468
-83.6 | 0 | 100 | 4.12 | 10193.9 | 6.17 | 6.1686
-84.7 | 0 | 100 | 4 | 9896.3 | 5.99 | 5.9889
-86 | 0 | 100 | 3.86 | 9533.2 | 5.78 | 5.7793
-87.2 | 0 | 100 | 3.75 | 9269.3 | 5.61 | 5.6146
-88.4 | 0 | 100 | 3.62 | 8942.4 | 5.42 | 5.42
-89.6 | 0 | 100 | 3.49 | 8621 | 5.22 | 5.2253
-90.8 | 0 | 100 | 3.35 | 8284.7 | 5.02 | 5.0157
-92 | 0 | 100 | 3.24 | 8019.5 | 4.86 | 4.851
-93.2 | 0 | 100 | 3.17 | 7836.6 | 4.75 | 4.7462
-94.4 | 0 | 100 | 3.04 | 7525.3 | 4.56 | 4.5516
-95.6 | 0 | 100 | 2.98 | 7367.3 | 4.47 | 4.4617
-96.8 | 0 | 100 | 2.83 | 6998.3 | 4.24 | 4.2372
-98 | 0 | 100 | 2.78 | 6866.6 | 4.16 | 4.1623
-100 | 0 | 100 | 2.63 | 6499.4 | 3.94 | 3.9377
Measurement time: | 12/08/2022 19:28:12
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-98.7 | 0 | 100 | 2.69 | 6649.1 | 4.02 | 4.0306
-96.6 | 0 | 100 | 2.88 | 7130.8 | 4.32 | 4.3153
-95.4 | 0 | 100 | 3.01 | 7430.4 | 4.51 | 4.51
-94.2 | 0 | 100 | 3.06 | 7560.4 | 4.57 | 4.585
-93 | 0 | 100 | 3.24 | 8022.2 | 4.86 | 4.8547
-91.8 | 0 | 100 | 3.29 | 8134.3 | 4.92 | 4.9296
-90.5 | 0 | 100 | 3.4 | 8413.5 | 5.1 | 5.0944
-89.3 | 0 | 100 | 3.53 | 8717.3 | 5.28 | 5.2892
-88.1 | 0 | 100 | 3.62 | 8950.1 | 5.4 | 5.424
-86.9 | 0 | 100 | 3.76 | 9305.6 | 5.63 | 5.6338
-85.7 | 0 | 100 | 3.92 | 9703.3 | 5.87 | 5.8735
-84.5 | 0 | 100 | 4.06 | 10038.9 | 6.09 | 6.0833
-83.3 | 0 | 100 | 4.19 | 10367.3 | 6.26 | 6.2781
-82.1 | 0 | 100 | 4.28 | 10592.5 | 6.41 | 6.4129
-80.9 | 0 | 100 | 4.53 | 11187.7 | 6.77 | 6.7875
-79.7 | 0 | 100 | 4.68 | 11571.9 | 6.99 | 7.0123
-78.5 | 0 | 100 | 4.9 | 12107.9 | 7.33 | 7.3419
-77.3 | 0 | 100 | 5.07 | 12524.7 | 7.59 | 7.5966
-76.1 | 0 | 100 | 5.21 | 12882.9 | 7.79 | 7.8064
-74.9 | 0 | 100 | 5.39 | 13330.9 | 8.06 | 8.0761
-73.7 | 0 | 100 | 5.58 | 13806.8 | 8.36 | 8.3608
-72.5 | 0 | 100 | 5.82 | 14389.7 | 8.7 | 8.7204
-71.3 | 0 | 100 | 6.06 | 14985.8 | 9.06 | 9.08
-70.1 | 0 | 100 | 6.32 | 15626.4 | 9.43 | 9.4696
-68.9 | 0 | 100 | 6.53 | 16138.8 | 9.75 | 9.7842
-67.7 | 0 | 100 | 6.86 | 16965.2 | 10.27 | 10.2787
-66.5 | 0 | 100 | 7.23 | 17881.5 | 10.82 | 10.8331
-65.3 | 0 | 100 | 7.68 | 18978.9 | 11.47 | 11.5073
-64.1 | 0 | 100 | 8.22 | 20315.6 | 12.27 | 12.3165
-62.9 | 0 | 100 | 8.97 | 22171 | 13.41 | 13.4402
-61.6 | 0 | 100 | 10.13 | 25044.7 | 15.13 | 15.1783
-60.5 | 0 | 100 | 12.09 | 29879.3 | 18.05 | 18.1151
-59.2 | 0 | 100 | 15.34 | 37936.3 | 22.96 | 22.9847
-58 | 0 | 100 | 20.3 | 50191 | 30.36 | 30.4165
-56.8 | 0 | 100 | 26.5 | 65509.3 | 39.72 | 39.7063
-55.6 | 0 | 100 | 34.1 | 84308.9 | 50.9 | 51.0938
-54.4 | 0 | 100 | 42.39 | 104793.9 | 63.79 | 63.5151
-53.2 | 0 | 100 | 49.91 | 123397.3 | 74.82 | 74.7827
-52 | 0 | 100 | 55.95 | 138327.1 | 83.53 | 83.8328
-50.8 | 0 | 100 | 59.7 | 147589.5 | 89.64 | 89.4516
-49.6 | 0 | 100 | 61.74 | 152651.7 | 92.3 | 92.5082
-48.4 | 0 | 100 | 62.77 | 155177.9 | 93.58 | 94.0515
-47.2 | 0 | 100 | 63.43 | 156817.4 | 94.66 | 95.0405
-46 | 0 | 100 | 64.05 | 158360 | 95.44 | 95.9694
-44.8 | 0 | 100 | 64.27 | 158903.6 | 96.02 | 96.2991
-43.6 | 0 | 100 | 64.54 | 159564.7 | 96.55 | 96.7036
-42.4 | 0 | 100 | 64.72 | 160020.3 | 96.9 | 96.9733
-41.2 | 0 | 100 | 64.9 | 160457.5 | 97.08 | 97.243
-40 | 0 | 100 | 65.07 | 160882.8 | 97.56 | 97.4978
-38.8 | 0 | 100 | 65.2 | 161195.4 | 97.75 | 97.6925
-37.6 | 0 | 100 | 65.49 | 161912.5 | 97.41 | 98.1271
-36.4 | 0 | 100 | 65.52 | 161979.5 | 97.67 | 98.172
-35.2 | 0 | 100 | 65.64 | 162288.6 | 98.18 | 98.3518
-34 | 0 | 100 | 65.77 | 162606.3 | 98.23 | 98.5466
-32.8 | 0 | 100 | 65.8 | 162683.1 | 98.78 | 98.5915
-31.6 | 0 | 100 | 66 | 163165.5 | 98.2 | 98.8912
-30.4 | 0 | 100 | 66.14 | 163524.9 | 98.75 | 99.101
-29.1 | 0 | 100 | 66.23 | 163744.9 | 98.93 | 99.2358
-28 | 0 | 100 | 66.3 | 163916.9 | 99.12 | 99.3407
-26.8 | 0 | 100 | 66.32 | 163954.1 | 99.16 | 99.3707
-25.5 | 0 | 100 | 66.59 | 164632.8 | 99.72 | 99.7752
-24.3 | 0 | 100 | 66.71 | 164922.6 | 99.73 | 99.955
-23.1 | 0 | 100 | 66.47 | 164347 | 99.53 | 99.5954
-21.9 | 0 | 100 | 66.57 | 164580 | 99.48 | 99.7453
-20.7 | 0 | 100 | 66.64 | 164763.2 | 99.82 | 99.8502
-19.5 | 0 | 100 | 66.69 | 164891.3 | 99.88 | 99.9251
-18.3 | 0 | 100 | 66.67 | 164823.6 | 99.43 | 99.8951
-17.1 | 0 | 100 | 66.72 | 164963.4 | 99.8 | 99.97
-15.9 | 0 | 100 | 66.74 | 164992.5 | 99.69 | 100
-14.7 | 0 | 100 | 66.6 | 164650 | 99.59 | 99.7902
-13.5 | 0 | 100 | 66.69 | 164888.4 | 99.97 | 99.9251
-12.3 | 0 | 100 | 66.68 | 164852.9 | 99.82 | 99.9101
-11.1 | 0 | 100 | 66.61 | 164694.4 | 100 | 99.8052
-9.9 | 0 | 100 | 66.53 | 164483.9 | 99.44 | 99.6853
-8.7 | 0 | 100 | 66.56 | 164554.1 | 99.73 | 99.7303
-7.5 | 0 | 100 | 66.51 | 164444.8 | 99.32 | 99.6554
-6.3 | 0 | 100 | 66.39 | 164148.1 | 99.73 | 99.4756
-5.1 | 0 | 100 | 66.61 | 164693.8 | 99.36 | 99.8052
-3.9 | 0 | 100 | 66.35 | 164045.5 | 99.12 | 99.4156
-2.7 | 0 | 100 | 66.3 | 163922.7 | 99.24 | 99.3407
-1.5 | 0 | 100 | 66.25 | 163801.7 | 99.14 | 99.2658
-0.3 | 0 | 100 | 66.26 | 163820.4 | 99.43 | 99.2808
0.9 | 0 | 100 | 66.3 | 163926.7 | 98.96 | 99.3407
2.1 | 0 | 100 | 66.25 | 163781.8 | 99.01 | 99.2658
3.4 | 0 | 100 | 66.35 | 164028.4 | 99.31 | 99.4156
4.6 | 0 | 100 | 66.31 | 163941.9 | 99.15 | 99.3557
5.8 | 0 | 100 | 66.35 | 164032.3 | 99.39 | 99.4156
7 | 0 | 100 | 66.31 | 163952.8 | 99.36 | 99.3557
8.2 | 0 | 100 | 66.31 | 163950.3 | 99.4 | 99.3557
9.4 | 0 | 100 | 66.37 | 164093.6 | 98.86 | 99.4456
10.6 | 0 | 100 | 66.26 | 163821.2 | 99.35 | 99.2808
11.8 | 0 | 100 | 66.3 | 163917.5 | 99.11 | 99.3407
13 | 0 | 100 | 66.32 | 163962.7 | 99.29 | 99.3707
14.2 | 0 | 100 | 66.16 | 163569 | 99.07 | 99.131
15.4 | 0 | 100 | 66.17 | 163604.9 | 98.98 | 99.1459
16.6 | 0 | 100 | 66.27 | 163839.7 | 98.99 | 99.2958
17.8 | 0 | 100 | 66.12 | 163483.2 | 98.94 | 99.071
19 | 0 | 100 | 66.11 | 163447.3 | 98.51 | 99.056
20.2 | 0 | 100 | 66.02 | 163222.8 | 99 | 98.9212
21.4 | 0 | 100 | 65.99 | 163155.2 | 98.62 | 98.8762
22.6 | 0 | 100 | 65.97 | 163095.4 | 98.83 | 98.8463
23.8 | 0 | 100 | 65.98 | 163133.8 | 98.7 | 98.8613
25 | 0 | 100 | 65.68 | 162393.5 | 98.19 | 98.4117
26.2 | 0 | 100 | 65.74 | 162525.4 | 97.97 | 98.5016
27.4 | 0 | 100 | 65.5 | 161938.6 | 97.88 | 98.142
28.7 | 0 | 100 | 65.5 | 161942.8 | 97.9 | 98.142
29.8 | 0 | 100 | 65.23 | 161280.6 | 97.69 | 97.7375
31.1 | 0 | 100 | 65.21 | 161223.3 | 97.79 | 97.7075
32.3 | 0 | 100 | 65.01 | 160736.9 | 97.8 | 97.4079
33.5 | 0 | 100 | 64.94 | 160558.8 | 97.31 | 97.303
34.7 | 0 | 100 | 64.77 | 160131 | 96.7 | 97.0482
35.9 | 0 | 100 | 64.51 | 159480.6 | 96.74 | 96.6587
37.1 | 0 | 100 | 64.41 | 159239.3 | 96.73 | 96.5088
38.3 | 0 | 100 | 64.16 | 158624.7 | 96.25 | 96.1343
39.5 | 0 | 100 | 64 | 158221.1 | 95.9 | 95.8945
40.7 | 0 | 100 | 63.76 | 157631 | 95.74 | 95.5349
41.9 | 0 | 100 | 63.65 | 157360 | 95.55 | 95.3701
43.1 | 0 | 100 | 63.46 | 156885.5 | 95.06 | 95.0854
44.3 | 0 | 100 | 63.09 | 155984.6 | 94.33 | 94.531
45.5 | 0 | 100 | 62.75 | 155147.4 | 93.72 | 94.0216
46.7 | 0 | 100 | 62.29 | 153991 | 93.33 | 93.3323
47.9 | 0 | 100 | 61.33 | 151618.8 | 91.94 | 91.8939
49.1 | 0 | 100 | 60.04 | 148442.1 | 89.86 | 89.961
50.3 | 0 | 100 | 57.65 | 142534.4 | 86.43 | 86.38
51.5 | 0 | 100 | 53.74 | 132872.9 | 80.46 | 80.5214
52.7 | 0 | 100 | 48.33 | 119481.6 | 72.14 | 72.4153
53.9 | 0 | 100 | 41.48 | 102563.9 | 61.94 | 62.1516
55.1 | 0 | 100 | 33.75 | 83454.1 | 50.57 | 50.5694
56.4 | 0 | 100 | 25.8 | 63780.3 | 38.73 | 38.6575
57.6 | 0 | 100 | 18.86 | 46616.4 | 28.34 | 28.2589
58.8 | 0 | 100 | 13.93 | 34447.3 | 20.92 | 20.872
60 | 0 | 100 | 11.07 | 27376.7 | 16.6 | 16.5868
61.2 | 0 | 100 | 9.53 | 23565.7 | 14.27 | 14.2793
62.4 | 0 | 100 | 8.55 | 21148 | 12.8 | 12.8109
63.6 | 0 | 100 | 7.92 | 19571.5 | 11.82 | 11.8669
64.8 | 0 | 100 | 7.42 | 18335.6 | 11.12 | 11.1178
66 | 0 | 100 | 7.02 | 17358.3 | 10.5 | 10.5184
67.2 | 0 | 100 | 6.71 | 16590.3 | 10.07 | 10.0539
68.4 | 0 | 100 | 6.39 | 15789.1 | 9.55 | 9.5745
69.6 | 0 | 100 | 6.13 | 15147.1 | 9.16 | 9.1849
70.8 | 0 | 100 | 5.89 | 14563.6 | 8.83 | 8.8253
72 | 0 | 100 | 5.59 | 13815.3 | 8.38 | 8.3758
73.2 | 0 | 100 | 5.43 | 13434.1 | 8.14 | 8.1361
74.4 | 0 | 100 | 5.2 | 12860.6 | 7.77 | 7.7914
75.6 | 0 | 100 | 5.03 | 12428.4 | 7.53 | 7.5367
76.8 | 0 | 100 | 4.87 | 12029 | 7.27 | 7.297
78 | 0 | 100 | 4.69 | 11606.2 | 7.03 | 7.0273
79.2 | 0 | 100 | 4.57 | 11306.1 | 6.82 | 6.8475
80.5 | 0 | 100 | 4.36 | 10772.3 | 6.51 | 6.5328
81.6 | 0 | 100 | 4.22 | 10426.6 | 6.31 | 6.323
82.9 | 0 | 100 | 4.08 | 10079.6 | 6.11 | 6.1133
84 | 0 | 100 | 3.95 | 9775.1 | 5.92 | 5.9185
85.2 | 0 | 100 | 3.76 | 9286.2 | 5.62 | 5.6338
86.5 | 0 | 100 | 3.64 | 8990.5 | 5.42 | 5.454
87.6 | 0 | 100 | 3.58 | 8841.3 | 5.36 | 5.3641
88.9 | 0 | 100 | 3.45 | 8519.8 | 5.16 | 5.1693
90.1 | 0 | 100 | 3.3 | 8156.8 | 4.94 | 4.9446
91.3 | 0 | 100 | 3.2 | 7920.6 | 4.8 | 4.7947
92.5 | 0 | 100 | 3.13 | 7731.8 | 4.7 | 4.6898
93.7 | 0 | 100 | 3.04 | 7507.9 | 4.56 | 4.555
94.9 | 0 | 100 | 2.95 | 7300.6 | 4.42 | 4.4201
96.1 | 0 | 100 | 2.85 | 7040.3 | 4.26 | 4.2703
97.3 | 0 | 100 | 2.74 | 6780.5 | 4.1 | 4.1055
98.5 | 0 | 100 | 2.64 | 6538.4 | 3.95 | 3.9556
100 | 0 | 100 | 2.53 | 6244.3 | 3.79 | 3.7908