## 11_ Profil détecteur
Measurement time: | 12/08/2022 20:26:32 | [5 empty] | Measurement time: | 12/08/2022 18:12:50
Radiation device: | Clinac2 Accelerator | [5 empty] | Radiation device: | Clinac2 Accelerator
Energy: | 6 MV | [5 empty] | Energy: | 6 MV
Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | FC 65-P | [5 empty] | Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie | [5 empty] | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm
SSD: | 1000 mm | [5 empty] | SSD: | 1000 mm
Field size: | 100 x 100 mm | [5 empty] | Field size: | 100 x 100 mm
Gantry angle: | 0 | [5 empty] | Gantry angle: | 0
Measurement medium: | Water | [5 empty] | Measurement medium: | Water
Scan type: | Inline | [5 empty] | Scan type: | Inline
Scan mode: | Continuous | [5 empty] | Scan mode: | Continuous
Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
155 | 0 | 100 | 2.17 | 13220.5 | 1.46 | 1.4371 | -100 | 0 | 100 | 2.43 | 6017.8 | 3.78 | 3.6215
154.1 | 0 | 100 | 2.29 | 13965.7 | 1.52 | 1.5166 | -99.1 | 0 | 100 | 2.98 | 7363.3 | 4.43 | 4.4411
152.6 | 0 | 100 | 2.38 | 14502.1 | 1.58 | 1.5762 | -97.7 | 0 | 100 | 3.03 | 7502 | 4.51 | 4.5156
151.5 | 0 | 100 | 2.39 | 14532.6 | 1.58 | 1.5828 | -96.6 | 0 | 100 | 3.16 | 7801.4 | 4.69 | 4.7094
150.3 | 0 | 100 | 2.44 | 14854.5 | 1.61 | 1.6159 | -95.4 | 0 | 100 | 3.24 | 8009.7 | 4.81 | 4.8286
149.1 | 0 | 100 | 2.5 | 15202.7 | 1.65 | 1.6556 | -94.1 | 0 | 100 | 3.32 | 8198.1 | 4.92 | 4.9478
147.9 | 0 | 100 | 2.53 | 15405.8 | 1.67 | 1.6755 | -93 | 0 | 100 | 3.43 | 8467.9 | 5.08 | 5.1118
146.7 | 0 | 100 | 2.6 | 15813.2 | 1.72 | 1.7219 | -91.8 | 0 | 100 | 3.57 | 8821.7 | 5.3 | 5.3204
145.5 | 0 | 100 | 2.65 | 16122.7 | 1.75 | 1.755 | -90.6 | 0 | 100 | 3.68 | 9097.9 | 5.47 | 5.4844
144.3 | 0 | 100 | 2.68 | 16336 | 1.77 | 1.7748 | -89.4 | 0 | 100 | 3.76 | 9301.1 | 5.59 | 5.6036
143.1 | 0 | 100 | 2.76 | 16829.3 | 1.83 | 1.8278 | -88.1 | 0 | 100 | 3.92 | 9693.8 | 5.82 | 5.842
141.9 | 0 | 100 | 2.82 | 17194.5 | 1.87 | 1.8675 | -86.9 | 0 | 100 | 4.02 | 9930.5 | 5.97 | 5.9911
140.6 | 0 | 100 | 2.88 | 17530.9 | 1.9 | 1.9073 | -85.7 | 0 | 100 | 4.18 | 10341.6 | 6.21 | 6.2295
139.5 | 0 | 100 | 2.93 | 17833.5 | 1.94 | 1.9404 | -84.5 | 0 | 100 | 4.32 | 10689.8 | 6.44 | 6.4382
138.2 | 0 | 100 | 3.01 | 18298.6 | 1.99 | 1.9934 | -83.3 | 0 | 100 | 4.43 | 10940.6 | 6.6 | 6.6021
137 | 0 | 100 | 3.06 | 18625.3 | 2.02 | 2.0265 | -82.1 | 0 | 100 | 4.62 | 11423.2 | 6.87 | 6.8852
135.9 | 0 | 100 | 3.13 | 19042.3 | 2.07 | 2.0728 | -80.9 | 0 | 100 | 4.72 | 11681.8 | 7.02 | 7.0343
134.6 | 0 | 100 | 3.2 | 19460 | 2.12 | 2.1192 | -79.7 | 0 | 100 | 4.92 | 12167.6 | 7.31 | 7.3323
133.4 | 0 | 100 | 3.26 | 19867.1 | 2.16 | 2.1589 | -78.5 | 0 | 100 | 5.09 | 12589.6 | 7.56 | 7.5857
132.2 | 0 | 100 | 3.33 | 20280.8 | 2.2 | 2.2053 | -77.3 | 0 | 100 | 5.25 | 12972.6 | 7.79 | 7.8241
131 | 0 | 100 | 3.41 | 20780.7 | 2.26 | 2.2583 | -76.1 | 0 | 100 | 5.45 | 13463.5 | 8.1 | 8.1222
129.8 | 0 | 100 | 3.51 | 21375.6 | 2.32 | 2.3245 | -74.9 | 0 | 100 | 5.65 | 13957.7 | 8.39 | 8.4203
128.6 | 0 | 100 | 3.56 | 21672.2 | 2.35 | 2.3576 | -73.7 | 0 | 100 | 5.8 | 14339.6 | 8.63 | 8.6438
127.4 | 0 | 100 | 3.64 | 22181.5 | 2.41 | 2.4106 | -72.5 | 0 | 100 | 6.06 | 14991.6 | 9.01 | 9.0313
126.2 | 0 | 100 | 3.76 | 22905.7 | 2.49 | 2.4901 | -71.3 | 0 | 100 | 6.33 | 15644.6 | 9.4 | 9.4337
125 | 0 | 100 | 3.85 | 23449.4 | 2.55 | 2.5497 | -70.1 | 0 | 100 | 6.54 | 16178.4 | 9.72 | 9.7466
123.8 | 0 | 100 | 3.93 | 23923.3 | 2.6 | 2.6026 | -68.9 | 0 | 100 | 6.88 | 17008.1 | 10.22 | 10.2534
122.6 | 0 | 100 | 4.02 | 24471.9 | 2.66 | 2.6623 | -67.7 | 0 | 100 | 7.14 | 17659.6 | 10.6 | 10.6408
121.4 | 0 | 100 | 4.12 | 25073.1 | 2.73 | 2.7285 | -66.5 | 0 | 100 | 7.49 | 18510.3 | 11.14 | 11.1624
120.2 | 0 | 100 | 4.24 | 25800.6 | 2.8 | 2.8079 | -65.3 | 0 | 100 | 7.92 | 19573.7 | 11.78 | 11.8033
119 | 0 | 100 | 4.34 | 26385.8 | 2.87 | 2.8742 | -64.1 | 0 | 100 | 8.44 | 20854.5 | 12.54 | 12.5782
117.8 | 0 | 100 | 4.45 | 27075.7 | 2.94 | 2.947 | -62.9 | 0 | 100 | 9.15 | 22615.7 | 13.61 | 13.6364
116.6 | 0 | 100 | 4.56 | 27771.5 | 3.02 | 3.0199 | -61.7 | 0 | 100 | 10.29 | 25446.7 | 15.29 | 15.3353
115.4 | 0 | 100 | 4.67 | 28431.7 | 3.09 | 3.0927 | -60.5 | 0 | 100 | 12.2 | 30162.1 | 18.14 | 18.1818
114.2 | 0 | 100 | 4.78 | 29120.2 | 3.17 | 3.1656 | -59.3 | 0 | 100 | 15.3 | 37836.1 | 22.75 | 22.8018
113 | 0 | 100 | 4.92 | 29940.3 | 3.25 | 3.2583 | -58 | 0 | 100 | 20.1 | 49689.6 | 29.9 | 29.9553
111.8 | 0 | 100 | 5.04 | 30688.5 | 3.33 | 3.3377 | -56.8 | 0 | 100 | 26.33 | 65108.2 | 39.15 | 39.2399
110.6 | 0 | 100 | 5.19 | 31570.8 | 3.43 | 3.4371 | -55.6 | 0 | 100 | 33.92 | 83864.5 | 50.51 | 50.5514
109.4 | 0 | 100 | 5.31 | 32320.5 | 3.52 | 3.5166 | -54.4 | 0 | 100 | 42.21 | 104351.1 | 62.76 | 62.9061
108.2 | 0 | 100 | 5.46 | 33248.5 | 3.62 | 3.6159 | -53.2 | 0 | 100 | 49.84 | 123212 | 74.1 | 74.2772
107 | 0 | 100 | 5.62 | 34193.1 | 3.72 | 3.7219 | -52 | 0 | 100 | 55.97 | 138366.6 | 83.2 | 83.4128
105.8 | 0 | 100 | 5.76 | 35074.4 | 3.81 | 3.8146 | -50.8 | 0 | 100 | 59.72 | 147654.1 | 88.93 | 89.0015
104.5 | 0 | 100 | 5.94 | 36180.4 | 3.94 | 3.9338 | -49.6 | 0 | 100 | 61.87 | 152956.1 | 92.23 | 92.2057
103.4 | 0 | 100 | 6.07 | 36938.1 | 4.02 | 4.0199 | -48.4 | 0 | 100 | 63.12 | 156064.7 | 93.91 | 94.0686
102.1 | 0 | 100 | 6.28 | 38254.6 | 4.16 | 4.1589 | -47.2 | 0 | 100 | 63.78 | 157694.1 | 94.91 | 95.0522
100.9 | 0 | 100 | 6.47 | 39382.4 | 4.28 | 4.2848 | -46 | 0 | 100 | 64.21 | 158744.5 | 95.5 | 95.693
99.7 | 0 | 100 | 6.61 | 40233.5 | 4.37 | 4.3775 | -44.8 | 0 | 100 | 64.44 | 159316.2 | 96.1 | 96.0358
98.5 | 0 | 100 | 6.82 | 41497.2 | 4.52 | 4.5166 | -43.6 | 0 | 100 | 64.63 | 159781.4 | 96.33 | 96.3189
97.3 | 0 | 100 | 7.04 | 42835.8 | 4.66 | 4.6623 | -42.4 | 0 | 100 | 64.97 | 160634.4 | 96.73 | 96.8256
96.1 | 0 | 100 | 7.22 | 43951.6 | 4.78 | 4.7815 | -41.2 | 0 | 100 | 65.12 | 161010.6 | 96.94 | 97.0492
94.9 | 0 | 100 | 7.43 | 45206.9 | 4.92 | 4.9205 | -40 | 0 | 100 | 65.34 | 161543.3 | 97.23 | 97.377
93.7 | 0 | 100 | 7.63 | 46414 | 5.05 | 5.053 | -38.8 | 0 | 100 | 65.32 | 161498.9 | 97.26 | 97.3472
92.5 | 0 | 100 | 7.9 | 48113.6 | 5.23 | 5.2318 | -37.6 | 0 | 100 | 65.44 | 161799.3 | 97.57 | 97.5261
91.3 | 0 | 100 | 8.13 | 49501.7 | 5.39 | 5.3841 | -36.4 | 0 | 100 | 65.73 | 162515.8 | 97.83 | 97.9583
90.1 | 0 | 100 | 8.41 | 51195.6 | 5.56 | 5.5695 | -35.2 | 0 | 100 | 65.95 | 163057.7 | 97.99 | 98.2861
88.9 | 0 | 100 | 8.65 | 52622.2 | 5.72 | 5.7285 | -34 | 0 | 100 | 66.13 | 163495.7 | 98.34 | 98.5544
87.7 | 0 | 100 | 8.97 | 54591.5 | 5.94 | 5.9404 | -32.8 | 0 | 100 | 66.15 | 163542.7 | 98.48 | 98.5842
86.5 | 0 | 100 | 9.25 | 56331.3 | 6.13 | 6.1258 | -31.6 | 0 | 100 | 66.36 | 164057.9 | 98.66 | 98.8972
85.3 | 0 | 100 | 9.57 | 58222.3 | 6.33 | 6.3377 | -30.4 | 0 | 100 | 66.53 | 164473.8 | 98.97 | 99.1505
84.1 | 0 | 100 | 9.89 | 60225.8 | 6.55 | 6.5497 | -29.2 | 0 | 100 | 66.57 | 164588.7 | 99.17 | 99.2101
82.9 | 0 | 100 | 10.21 | 62129.1 | 6.76 | 6.7616 | -28 | 0 | 100 | 66.74 | 165015.9 | 99.27 | 99.4635
81.7 | 0 | 100 | 10.57 | 64315.4 | 7 | 7 | -26.8 | 0 | 100 | 66.82 | 165191.2 | 99.49 | 99.5827
80.5 | 0 | 100 | 11 | 66931.2 | 7.28 | 7.2848 | -25.5 | 0 | 100 | 66.71 | 164924.1 | 99.31 | 99.4188
79.3 | 0 | 100 | 11.35 | 69071 | 7.52 | 7.5166 | -24.4 | 0 | 100 | 66.96 | 165550.5 | 99.67 | 99.7914
78 | 0 | 100 | 11.79 | 71778.8 | 7.8 | 7.8079 | -23.1 | 0 | 100 | 66.93 | 165476.1 | 99.79 | 99.7466
76.9 | 0 | 100 | 12.27 | 74670.9 | 8.12 | 8.1258 | -21.9 | 0 | 100 | 66.96 | 165553 | 99.61 | 99.7914
75.7 | 0 | 100 | 12.76 | 77638.8 | 8.44 | 8.4503 | -20.7 | 0 | 100 | 66.98 | 165608.6 | 99.68 | 99.8212
74.4 | 0 | 100 | 13.32 | 81073.6 | 8.8 | 8.8212 | -19.5 | 0 | 100 | 67 | 165652 | 99.66 | 99.851
73.3 | 0 | 100 | 13.92 | 84708 | 9.21 | 9.2185 | -18.3 | 0 | 100 | 67.08 | 165847.8 | 99.88 | 99.9702
72 | 0 | 100 | 14.71 | 89563.1 | 9.75 | 9.7417 | -17.1 | 0 | 100 | 67.04 | 165748.2 | 99.74 | 99.9106
70.8 | 0 | 100 | 15.7 | 95588.1 | 10.4 | 10.3974 | -15.9 | 0 | 100 | 66.99 | 165617.2 | 99.87 | 99.8361
69.6 | 0 | 100 | 17.06 | 103846.8 | 11.3 | 11.298 | -14.7 | 0 | 100 | 67.1 | 165890.4 | 99.93 | 100
68.4 | 0 | 100 | 19.38 | 117970.7 | 12.83 | 12.8344 | -13.5 | 0 | 100 | 66.98 | 165599.5 | 99.8 | 99.8212
67.3 | 0 | 100 | 22.77 | 138612.3 | 15.08 | 15.0795 | -12.3 | 0 | 100 | 66.93 | 165483 | 99.86 | 99.7466
66 | 0 | 100 | 27.38 | 166679.1 | 18.13 | 18.1325 | -11.1 | 0 | 100 | 66.99 | 165627.2 | 100 | 99.8361
64.8 | 0 | 100 | 32.63 | 198600.8 | 21.59 | 21.6093 | -9.9 | 0 | 100 | 66.89 | 165373 | 99.72 | 99.687
63.6 | 0 | 100 | 38.06 | 231656.3 | 25.23 | 25.2053 | -8.7 | 0 | 100 | 66.99 | 165628.2 | 99.66 | 99.8361
62.4 | 0 | 100 | 44.33 | 269826.7 | 29.41 | 29.3576 | -7.5 | 0 | 100 | 66.89 | 165363.8 | 99.56 | 99.687
61.2 | 0 | 100 | 50.39 | 306732.9 | 33.4 | 33.3709 | -6.3 | 0 | 100 | 66.8 | 165157.8 | 99.51 | 99.5529
60 | 0 | 100 | 56.87 | 346143.5 | 37.66 | 37.6623 | -5.1 | 0 | 100 | 66.98 | 165588.2 | 99.75 | 99.8212
58.8 | 0 | 100 | 63.37 | 385740.8 | 41.95 | 41.9669 | -3.8 | 0 | 100 | 66.83 | 165232.3 | 99.55 | 99.5976
57.6 | 0 | 100 | 69.85 | 425127.2 | 46.29 | 46.2583 | -2.7 | 0 | 100 | 66.71 | 164931.2 | 99.43 | 99.4188
56.4 | 0 | 100 | 76.63 | 466448.5 | 50.73 | 50.7483 | -1.4 | 0 | 100 | 66.65 | 164781.7 | 99.11 | 99.3294
55.2 | 0 | 100 | 83.09 | 505715.3 | 54.97 | 55.0265 | -0.3 | 0 | 100 | 66.62 | 164718.3 | 99.28 | 99.2846
54 | 0 | 100 | 89.73 | 546149.6 | 59.43 | 59.4238 | 1 | 0 | 100 | 66.68 | 164850.7 | 99.24 | 99.3741
52.8 | 0 | 100 | 96.25 | 585832.3 | 63.82 | 63.7417 | 2.2 | 0 | 100 | 66.68 | 164859.5 | 99.29 | 99.3741
51.6 | 0 | 100 | 102.49 | 623816.5 | 67.85 | 67.8742 | 3.4 | 0 | 100 | 66.53 | 164476.8 | 99.08 | 99.1505
50.3 | 0 | 100 | 109.31 | 665340 | 72.35 | 72.3907 | 4.6 | 0 | 100 | 66.48 | 164354.3 | 99.1 | 99.076
49.2 | 0 | 100 | 115.46 | 702759.5 | 76.35 | 76.4636 | 5.8 | 0 | 100 | 66.62 | 164704.4 | 99.24 | 99.2846
47.9 | 0 | 100 | 121.45 | 739203.5 | 80.49 | 80.4305 | 7 | 0 | 100 | 66.46 | 164323.9 | 99.09 | 99.0462
46.7 | 0 | 100 | 127.3 | 774854.8 | 84.37 | 84.3046 | 8.2 | 0 | 100 | 66.64 | 164766.1 | 99.27 | 99.3145
45.6 | 0 | 100 | 132.21 | 804728.3 | 87.7 | 87.5563 | 9.4 | 0 | 100 | 66.65 | 164776.5 | 99.38 | 99.3294
44.3 | 0 | 100 | 136.76 | 832392.4 | 90.79 | 90.5695 | 10.6 | 0 | 100 | 66.52 | 164451.7 | 99.15 | 99.1356
43.2 | 0 | 100 | 140.23 | 853559 | 92.9 | 92.8675 | 11.8 | 0 | 100 | 66.58 | 164614.9 | 99.24 | 99.225
41.9 | 0 | 100 | 142.73 | 868732.8 | 94.48 | 94.5232 | 13 | 0 | 100 | 66.58 | 164600.8 | 99.29 | 99.225
40.7 | 0 | 100 | 144.39 | 878876.8 | 95.64 | 95.6225 | 14.2 | 0 | 100 | 66.55 | 164545.8 | 99.3 | 99.1803
39.5 | 0 | 100 | 145.3 | 884389.1 | 96.04 | 96.2252 | 15.4 | 0 | 100 | 66.52 | 164468 | 99.29 | 99.1356
38.3 | 0 | 100 | 145.91 | 888103.4 | 96.6 | 96.6291 | 16.6 | 0 | 100 | 66.5 | 164414.9 | 99.05 | 99.1058
37.1 | 0 | 100 | 146.36 | 890870.7 | 96.85 | 96.9272 | 17.8 | 0 | 100 | 66.44 | 164259.1 | 98.88 | 99.0164
35.9 | 0 | 100 | 146.66 | 892655.7 | 97.12 | 97.1258 | 19 | 0 | 100 | 66.51 | 164446 | 98.84 | 99.1207
34.7 | 0 | 100 | 147.14 | 895602.4 | 97.43 | 97.4437 | 20.2 | 0 | 100 | 66.36 | 164075.3 | 98.83 | 98.8972
33.5 | 0 | 100 | 147.5 | 897774.9 | 97.58 | 97.6821 | 21.4 | 0 | 100 | 66.28 | 163856.4 | 98.77 | 98.7779
32.3 | 0 | 100 | 147.56 | 898121.6 | 97.72 | 97.7219 | 22.6 | 0 | 100 | 66.1 | 163433.9 | 98.45 | 98.5097
31.1 | 0 | 100 | 147.95 | 900505.5 | 97.86 | 97.9801 | 23.8 | 0 | 100 | 66.01 | 163189.2 | 98.38 | 98.3756
29.9 | 0 | 100 | 148.35 | 902953.1 | 98.15 | 98.245 | 25 | 0 | 100 | 65.99 | 163152.9 | 98.27 | 98.3458
28.7 | 0 | 100 | 148.34 | 902877.8 | 98.23 | 98.2384 | 26.2 | 0 | 100 | 66.05 | 163310.6 | 98.48 | 98.4352
27.5 | 0 | 100 | 148.43 | 903464 | 98.31 | 98.298 | 27.4 | 0 | 100 | 65.85 | 162796.9 | 98.16 | 98.1371
26.3 | 0 | 100 | 148.75 | 905411.5 | 98.5 | 98.5099 | 28.7 | 0 | 100 | 65.75 | 162550.3 | 97.94 | 97.9881
25.1 | 0 | 100 | 149.08 | 907400.6 | 98.36 | 98.7285 | 29.8 | 0 | 100 | 65.63 | 162264.2 | 97.81 | 97.8092
23.8 | 0 | 100 | 149.31 | 908830.5 | 98.7 | 98.8808 | 31 | 0 | 100 | 65.44 | 161779.4 | 97.41 | 97.5261
22.7 | 0 | 100 | 149.43 | 909504.6 | 98.97 | 98.9603 | 32.3 | 0 | 100 | 65.28 | 161400.1 | 97.26 | 97.2876
21.5 | 0 | 100 | 149.58 | 910434.4 | 99.05 | 99.0596 | 33.4 | 0 | 100 | 65.22 | 161236.3 | 97.24 | 97.1982
20.2 | 0 | 100 | 149.65 | 910874.4 | 99.17 | 99.106 | 34.7 | 0 | 100 | 65.01 | 160737.9 | 97.07 | 96.8852
19.1 | 0 | 100 | 149.8 | 911806 | 99.1 | 99.2053 | 35.9 | 0 | 100 | 64.78 | 160150.6 | 96.5 | 96.5425
17.8 | 0 | 100 | 150.19 | 914146.6 | 99.32 | 99.4636 | 37.1 | 0 | 100 | 64.71 | 159980.3 | 96.4 | 96.4382
16.7 | 0 | 100 | 149.87 | 912204.7 | 99.42 | 99.2517 | 38.3 | 0 | 100 | 64.61 | 159749.8 | 96.24 | 96.2891
15.4 | 0 | 100 | 150.1 | 913633.2 | 99.42 | 99.404 | 39.5 | 0 | 100 | 64.35 | 159099 | 96 | 95.9016
14.2 | 0 | 100 | 150.16 | 913964.4 | 99.54 | 99.4437 | 40.7 | 0 | 100 | 64.18 | 158674.8 | 95.71 | 95.6483
13 | 0 | 100 | 150.05 | 913320 | 99.5 | 99.3709 | 41.9 | 0 | 100 | 63.85 | 157854.8 | 95.26 | 95.1565
11.8 | 0 | 100 | 150.32 | 914948.5 | 99.46 | 99.5497 | 43.1 | 0 | 100 | 63.7 | 157493.9 | 94.9 | 94.9329
10.6 | 0 | 100 | 150.21 | 914305.1 | 99.6 | 99.4768 | 44.3 | 0 | 100 | 63.36 | 156657.5 | 94.59 | 94.4262
9.4 | 0 | 100 | 150.45 | 915755.9 | 99.53 | 99.6358 | 45.5 | 0 | 100 | 62.85 | 155375 | 93.58 | 93.6662
8.2 | 0 | 100 | 150.34 | 915082.5 | 99.67 | 99.5629 | 46.7 | 0 | 100 | 62.32 | 154069.3 | 92.74 | 92.8763
7 | 0 | 100 | 150.58 | 916529.3 | 99.56 | 99.7219 | 47.9 | 0 | 100 | 61.3 | 151555.5 | 91.37 | 91.3562
5.7 | 0 | 100 | 150.45 | 915745.9 | 99.4 | 99.6358 | 49.1 | 0 | 100 | 59.78 | 147802.8 | 89.09 | 89.0909
4.6 | 0 | 100 | 150.22 | 914344.9 | 99.53 | 99.4834 | 50.3 | 0 | 100 | 56.9 | 140672.5 | 84.87 | 84.7988
3.4 | 0 | 100 | 150.08 | 913464.6 | 99.34 | 99.3907 | 51.5 | 0 | 100 | 52.65 | 130164.1 | 78.45 | 78.465
2.2 | 0 | 100 | 150.28 | 914723.5 | 99.43 | 99.5232 | 52.7 | 0 | 100 | 47.05 | 116319.4 | 70.12 | 70.1192
1 | 0 | 100 | 150.45 | 915736.2 | 99.61 | 99.6358 | 53.9 | 0 | 100 | 39.68 | 98110 | 59.23 | 59.1356
-0.3 | 0 | 100 | 150.52 | 916171 | 99.62 | 99.6821 | 55.1 | 0 | 100 | 31.62 | 78172.9 | 47.16 | 47.1237
-1.4 | 0 | 100 | 150.36 | 915199.6 | 99.57 | 99.5762 | 56.3 | 0 | 100 | 23.99 | 59312.8 | 35.79 | 35.7526
-2.7 | 0 | 100 | 150.35 | 915112.1 | 99.56 | 99.5695 | 57.6 | 0 | 100 | 17.64 | 43602.6 | 26.32 | 26.2891
-3.9 | 0 | 100 | 150.64 | 916868.2 | 99.77 | 99.7616 | 58.8 | 0 | 100 | 13.28 | 32838.7 | 19.77 | 19.7914
-5 | 0 | 100 | 150.67 | 917063.9 | 99.81 | 99.7815 | 60 | 0 | 100 | 10.87 | 26864.8 | 16.21 | 16.1997
-6.3 | 0 | 100 | 150.71 | 917345.4 | 99.77 | 99.8079 | 61.2 | 0 | 100 | 9.49 | 23454.1 | 14.14 | 14.1431
-7.5 | 0 | 100 | 150.68 | 917152.1 | 99.7 | 99.7881 | 62.4 | 0 | 100 | 8.66 | 21415.8 | 12.9 | 12.9061
-8.7 | 0 | 100 | 150.82 | 917987.6 | 99.81 | 99.8808 | 63.6 | 0 | 100 | 8.03 | 19856.9 | 11.96 | 11.9672
-9.9 | 0 | 100 | 150.7 | 917256.4 | 99.79 | 99.8013 | 64.8 | 0 | 100 | 7.56 | 18702.2 | 11.26 | 11.2668
-11.1 | 0 | 100 | 150.94 | 918738.7 | 99.81 | 99.9603 | 66 | 0 | 100 | 7.19 | 17781 | 10.73 | 10.7154
-12.3 | 0 | 100 | 150.89 | 918429.4 | 99.94 | 99.9272 | 67.2 | 0 | 100 | 6.86 | 16957.3 | 10.22 | 10.2235
-13.5 | 0 | 100 | 150.92 | 918576.3 | 99.81 | 99.947 | 68.4 | 0 | 100 | 6.54 | 16175 | 9.75 | 9.7466
-14.7 | 0 | 100 | 151 | 919100 | 100 | 100 | 69.6 | 0 | 100 | 6.37 | 15741.7 | 9.49 | 9.4933
-15.9 | 0 | 100 | 150.62 | 916792 | 99.91 | 99.7483 | 70.8 | 0 | 100 | 6.06 | 14971 | 9.02 | 9.0313
-17.1 | 0 | 100 | 150.77 | 917685.9 | 99.73 | 99.8477 | 72 | 0 | 100 | 5.86 | 14493.3 | 8.74 | 8.7332
-18.3 | 0 | 100 | 150.82 | 917976.5 | 99.76 | 99.8808 | 73.2 | 0 | 100 | 5.65 | 13979.3 | 8.44 | 8.4203
-19.5 | 0 | 100 | 150.57 | 916488.5 | 99.76 | 99.7152 | 74.4 | 0 | 100 | 5.48 | 13539.2 | 8.17 | 8.1669
-20.7 | 0 | 100 | 150.69 | 917223.8 | 99.75 | 99.7947 | 75.6 | 0 | 100 | 5.28 | 13061.8 | 7.87 | 7.8689
-21.9 | 0 | 100 | 150.25 | 914528.2 | 99.63 | 99.5033 | 76.8 | 0 | 100 | 5.08 | 12551.8 | 7.55 | 7.5708
-23.1 | 0 | 100 | 150.32 | 914939.7 | 99.45 | 99.5497 | 78.1 | 0 | 100 | 4.88 | 12069.3 | 7.28 | 7.2727
-24.4 | 0 | 100 | 150.3 | 914802.7 | 99.35 | 99.5364 | 79.3 | 0 | 100 | 4.79 | 11833 | 7.14 | 7.1386
-25.5 | 0 | 100 | 150.12 | 913707.7 | 99.27 | 99.4172 | 80.5 | 0 | 100 | 4.57 | 11288.3 | 6.81 | 6.8107
-26.8 | 0 | 100 | 149.77 | 911612.7 | 99.1 | 99.1854 | 81.7 | 0 | 100 | 4.41 | 10906.6 | 6.58 | 6.5723
-28 | 0 | 100 | 149.85 | 912059.2 | 99.29 | 99.2384 | 82.9 | 0 | 100 | 4.26 | 10542.5 | 6.37 | 6.3487
-29.1 | 0 | 100 | 149.68 | 911041.7 | 99.18 | 99.1258 | 84.1 | 0 | 100 | 4.18 | 10343.8 | 6.24 | 6.2295
-30.4 | 0 | 100 | 149.36 | 909117.2 | 98.92 | 98.9139 | 85.3 | 0 | 100 | 4.04 | 9992.8 | 6.03 | 6.0209
-31.5 | 0 | 100 | 149.12 | 907639.4 | 98.8 | 98.755 | 86.5 | 0 | 100 | 3.92 | 9685.3 | 5.83 | 5.842
-32.8 | 0 | 100 | 148.77 | 905492.7 | 98.63 | 98.5232 | 87.7 | 0 | 100 | 3.84 | 9498.3 | 5.73 | 5.7228
-34 | 0 | 100 | 148.83 | 905886.1 | 98.37 | 98.5629 | 88.9 | 0 | 100 | 3.67 | 9064.5 | 5.47 | 5.4694
-35.1 | 0 | 100 | 148.43 | 903453.5 | 98.2 | 98.298 | 90.1 | 0 | 100 | 3.51 | 8684.6 | 5.24 | 5.231
-36.4 | 0 | 100 | 147.86 | 899997.8 | 97.87 | 97.9205 | 91.3 | 0 | 100 | 3.43 | 8491.5 | 5.11 | 5.1118
-37.5 | 0 | 100 | 147.34 | 896787.4 | 97.42 | 97.5762 | 92.5 | 0 | 100 | 3.34 | 8265.1 | 4.99 | 4.9776
-38.8 | 0 | 100 | 146.78 | 893391.6 | 97.18 | 97.2053 | 93.7 | 0 | 100 | 3.25 | 8032.2 | 4.85 | 4.8435
-40 | 0 | 100 | 145.75 | 887162.6 | 96.55 | 96.5232 | 94.9 | 0 | 100 | 3.18 | 7874.3 | 4.75 | 4.7392
-41.2 | 0 | 100 | 144.33 | 878470.9 | 95.67 | 95.5828 | 96.1 | 0 | 100 | 3.04 | 7527.2 | 4.55 | 4.5306
-42.4 | 0 | 100 | 141.97 | 864132.7 | 93.92 | 94.0199 | 97.3 | 0 | 100 | 2.93 | 7240.6 | 4.37 | 4.3666
-43.6 | 0 | 100 | 138.02 | 840092.7 | 91.38 | 91.404 | 98.5 | 0 | 100 | 2.93 | 7255.9 | 4.37 | 4.3666
-44.8 | 0 | 100 | 133.13 | 810296.2 | 88.13 | 88.1656 | 100 | 0 | 100 | 2.81 | 6937.7 | 4.18 | 4.1878
-46 | 0 | 100 | 127.6 | 776652.8 | 84.65 | 84.5033
-47.2 | 0 | 100 | 122.08 | 743043.1 | 80.85 | 80.8477
-48.4 | 0 | 100 | 115.71 | 704301.9 | 76.62 | 76.6291
-49.6 | 0 | 100 | 109.41 | 665972.1 | 72.37 | 72.457
-50.8 | 0 | 100 | 102.79 | 625680.3 | 68 | 68.0728
-52 | 0 | 100 | 95.9 | 583700 | 63.46 | 63.5099
-53.2 | 0 | 100 | 89.45 | 544448.3 | 59.27 | 59.2384
-54.4 | 0 | 100 | 82.54 | 502401.2 | 54.6 | 54.6623
-55.6 | 0 | 100 | 76.11 | 463266.2 | 50.37 | 50.404
-56.8 | 0 | 100 | 69.44 | 422653.1 | 45.98 | 45.9868
-58 | 0 | 100 | 62.52 | 380539.6 | 41.49 | 41.404
-59.2 | 0 | 100 | 56.2 | 342063.3 | 37.14 | 37.2185
-60.4 | 0 | 100 | 49.44 | 300940.3 | 32.72 | 32.7417
-61.6 | 0 | 100 | 43.34 | 263796.5 | 28.69 | 28.702
-62.8 | 0 | 100 | 37.3 | 227007.1 | 24.7 | 24.702
-64 | 0 | 100 | 31.63 | 192528 | 20.93 | 20.947
-65.2 | 0 | 100 | 26.76 | 162899.9 | 17.71 | 17.7219
-66.5 | 0 | 100 | 22.43 | 136545.6 | 14.85 | 14.8543
-67.6 | 0 | 100 | 19.3 | 117475.6 | 12.76 | 12.7815
-68.9 | 0 | 100 | 17.09 | 103993.2 | 11.29 | 11.3179
-70.1 | 0 | 100 | 15.65 | 95275.5 | 10.36 | 10.3642
-71.3 | 0 | 100 | 14.71 | 89561.3 | 9.73 | 9.7417
-72.5 | 0 | 100 | 13.9 | 84578.5 | 9.21 | 9.2053
-73.7 | 0 | 100 | 13.33 | 81143.8 | 8.83 | 8.8278
-74.9 | 0 | 100 | 12.78 | 77785.7 | 8.46 | 8.4636
-76.1 | 0 | 100 | 12.31 | 74951.7 | 8.16 | 8.1523
-77.3 | 0 | 100 | 11.9 | 72437.7 | 7.87 | 7.8808
-78.5 | 0 | 100 | 11.49 | 69943.4 | 7.61 | 7.6093
-79.7 | 0 | 100 | 11.09 | 67497.3 | 7.34 | 7.3444
-80.9 | 0 | 100 | 10.69 | 65054.9 | 7.07 | 7.0795
-82.1 | 0 | 100 | 10.34 | 62961.6 | 6.84 | 6.8477
-83.3 | 0 | 100 | 10.06 | 61212 | 6.65 | 6.6623
-84.5 | 0 | 100 | 9.7 | 59053.6 | 6.43 | 6.4238
-85.7 | 0 | 100 | 9.4 | 57184.9 | 6.22 | 6.2252
-86.9 | 0 | 100 | 9.11 | 55475.5 | 6.03 | 6.0331
-88.1 | 0 | 100 | 8.8 | 53548.2 | 5.83 | 5.8278
-89.3 | 0 | 100 | 8.54 | 51951.4 | 5.64 | 5.6556
-90.6 | 0 | 100 | 8.26 | 50269.4 | 5.47 | 5.4702
-91.7 | 0 | 100 | 8.01 | 48784.2 | 5.3 | 5.3046
-92.9 | 0 | 100 | 7.8 | 47467.7 | 5.15 | 5.1656
-94.1 | 0 | 100 | 7.53 | 45810.1 | 4.98 | 4.9868
-95.3 | 0 | 100 | 7.31 | 44481.4 | 4.84 | 4.8411
-96.6 | 0 | 100 | 7.13 | 43399.7 | 4.72 | 4.7219
-97.7 | 0 | 100 | 6.92 | 42107.1 | 4.58 | 4.5828
-99 | 0 | 100 | 6.72 | 40919.8 | 4.45 | 4.4503
-100.2 | 0 | 100 | 6.56 | 39922 | 4.34 | 4.3444
-101.3 | 0 | 100 | 6.34 | 38567.2 | 4.19 | 4.1987
-102.6 | 0 | 100 | 6.14 | 37386.5 | 4.06 | 4.0662
-103.8 | 0 | 100 | 6.02 | 36612.3 | 3.98 | 3.9868
-105 | 0 | 100 | 5.86 | 35644.7 | 3.88 | 3.8808
-106.2 | 0 | 100 | 5.72 | 34843.3 | 3.79 | 3.7881
-107.4 | 0 | 100 | 5.55 | 33796.5 | 3.67 | 3.6755
-108.6 | 0 | 100 | 5.37 | 32674.8 | 3.55 | 3.5563
-109.8 | 0 | 100 | 5.25 | 31933 | 3.47 | 3.4768
-111 | 0 | 100 | 5.09 | 30975.5 | 3.36 | 3.3709
-112.2 | 0 | 100 | 4.97 | 30261.1 | 3.29 | 3.2914
-113.4 | 0 | 100 | 4.87 | 29625.1 | 3.22 | 3.2252
-114.6 | 0 | 100 | 4.74 | 28844.9 | 3.13 | 3.1391
-115.8 | 0 | 100 | 4.6 | 28013.1 | 3.04 | 3.0464
-117 | 0 | 100 | 4.51 | 27426.3 | 2.98 | 2.9868
-118.2 | 0 | 100 | 4.38 | 26662.3 | 2.9 | 2.9007
-119.4 | 0 | 100 | 4.27 | 25988 | 2.83 | 2.8278
-120.7 | 0 | 100 | 4.17 | 25352.1 | 2.76 | 2.7616
-121.8 | 0 | 100 | 4.06 | 24712.1 | 2.69 | 2.6887
-123.1 | 0 | 100 | 3.98 | 24207.3 | 2.63 | 2.6358
-124.3 | 0 | 100 | 3.88 | 23640 | 2.57 | 2.5695
-125.4 | 0 | 100 | 3.79 | 23052.9 | 2.5 | 2.5099
-126.7 | 0 | 100 | 3.69 | 22456 | 2.44 | 2.4437
-127.8 | 0 | 100 | 3.62 | 22024.6 | 2.39 | 2.3974
-129.1 | 0 | 100 | 3.54 | 21522.1 | 2.34 | 2.3444
-130.3 | 0 | 100 | 3.43 | 20852.5 | 2.27 | 2.2715
-131.5 | 0 | 100 | 3.35 | 20419.1 | 2.22 | 2.2185
-132.7 | 0 | 100 | 3.29 | 20036.6 | 2.18 | 2.1788
-133.9 | 0 | 100 | 3.22 | 19583.4 | 2.13 | 2.1325
-135.1 | 0 | 100 | 3.14 | 19093.8 | 2.08 | 2.0795
-136.3 | 0 | 100 | 3.08 | 18752.3 | 2.04 | 2.0397
-137.5 | 0 | 100 | 3.01 | 18314.4 | 1.99 | 1.9934
-138.7 | 0 | 100 | 2.93 | 17821.9 | 1.94 | 1.9404
-139.9 | 0 | 100 | 2.87 | 17449.6 | 1.9 | 1.9007
-141.1 | 0 | 100 | 2.82 | 17149 | 1.86 | 1.8675
-142.3 | 0 | 100 | 2.75 | 16763 | 1.82 | 1.8212
-143.5 | 0 | 100 | 2.69 | 16395 | 1.78 | 1.7815
-144.7 | 0 | 100 | 2.62 | 15918.2 | 1.73 | 1.7351
-145.9 | 0 | 100 | 2.6 | 15798.7 | 1.72 | 1.7219
-147.1 | 0 | 100 | 2.52 | 15365.9 | 1.67 | 1.6689
-148.3 | 0 | 100 | 2.47 | 15009.9 | 1.63 | 1.6358
-149.5 | 0 | 100 | 2.41 | 14697.2 | 1.6 | 1.596
-150.7 | 0 | 100 | 2.39 | 14566.4 | 1.58 | 1.5828
-151.9 | 0 | 100 | 2.33 | 14173 | 1.54 | 1.543
-153.1 | 0 | 100 | 2.27 | 13830.6 | 1.5 | 1.5033
-155 | 0 | 100 | 2.23 | 13557.5 | 1.47 | 1.4768
Measurement time: | 12/08/2022 20:27:53
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | FC 65-P
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 43.65 | 265700.6 | 20.53
0 | 0 | 309.1 | 46.61 | 283690.7 | 20.64
0 | 0 | 307.6 | 47.15 | 287015.7 | 20.85
0 | 0 | 306.5 | 47.41 | 288596.1 | 20.97
0 | 0 | 305.3 | 47.77 | 290762.1 | 21.11
0 | 0 | 304.1 | 48.05 | 292479.6 | 21.27
0 | 0 | 302.9 | 48.39 | 294521.6 | 21.43
0 | 0 | 301.7 | 48.69 | 296376.8 | 21.56
0 | 0 | 300.5 | 49.05 | 298552.2 | 21.72
0 | 0 | 299.3 | 49.28 | 299976.2 | 21.83
0 | 0 | 298.1 | 49.62 | 302006.1 | 21.99
0 | 0 | 296.9 | 49.9 | 303750.5 | 22.09
0 | 0 | 295.7 | 50.3 | 306185.6 | 22.26
0 | 0 | 294.5 | 50.7 | 308577 | 22.45
0 | 0 | 293.3 | 51.07 | 310865.9 | 22.58
0 | 0 | 292 | 51.35 | 312524.1 | 22.72
0 | 0 | 290.9 | 51.56 | 313804.6 | 22.83
0 | 0 | 289.7 | 52.04 | 316739.2 | 23.02
0 | 0 | 288.5 | 52.49 | 319478.5 | 23.22
0 | 0 | 287.3 | 52.74 | 320982.1 | 23.34
0 | 0 | 286 | 53.14 | 323434.7 | 23.5
0 | 0 | 284.9 | 53.46 | 325418.8 | 23.65
0 | 0 | 283.7 | 53.93 | 328267.4 | 23.88
0 | 0 | 282.4 | 54.24 | 330160.8 | 24.02
0 | 0 | 281.3 | 54.53 | 331911.8 | 24.15
0 | 0 | 280 | 55.02 | 334898.7 | 24.36
0 | 0 | 278.8 | 55.25 | 336311.2 | 24.46
0 | 0 | 277.6 | 55.71 | 339103.8 | 24.62
0 | 0 | 276.4 | 56.14 | 341692.6 | 24.83
0 | 0 | 275.2 | 56.5 | 343882.1 | 25.04
0 | 0 | 274 | 56.9 | 346360.9 | 25.15
0 | 0 | 272.8 | 57.1 | 347563.9 | 25.31
0 | 0 | 271.6 | 57.65 | 350904.1 | 25.49
0 | 0 | 270.4 | 58.03 | 353201.3 | 25.65
0 | 0 | 269.2 | 58.32 | 355000.3 | 25.83
0 | 0 | 268 | 58.77 | 357707.7 | 26.04
0 | 0 | 266.8 | 59.11 | 359804.6 | 26.16
0 | 0 | 265.6 | 59.56 | 362536.8 | 26.36
0 | 0 | 264.4 | 59.88 | 364442.8 | 26.51
0 | 0 | 263.2 | 60.2 | 366426.3 | 26.66
0 | 0 | 262 | 60.69 | 369376.8 | 26.86
0 | 0 | 260.8 | 61.13 | 372073.5 | 27.03
0 | 0 | 259.6 | 61.67 | 375359.7 | 27.27
0 | 0 | 258.4 | 62.05 | 377663.2 | 27.47
0 | 0 | 257.2 | 62.36 | 379542.3 | 27.62
0 | 0 | 256 | 62.86 | 382625.2 | 27.8
0 | 0 | 254.8 | 63.31 | 385322.2 | 28.01
0 | 0 | 253.6 | 63.68 | 387570.7 | 28.2
0 | 0 | 252.4 | 64.2 | 390789.1 | 28.39
0 | 0 | 251.2 | 64.51 | 392671.7 | 28.54
0 | 0 | 250 | 64.99 | 395579.1 | 28.76
0 | 0 | 248.8 | 65.34 | 397682.9 | 28.95
0 | 0 | 247.6 | 65.78 | 400364.7 | 29.15
0 | 0 | 246.3 | 66.41 | 404202.3 | 29.36
0 | 0 | 245.1 | 66.74 | 406200.9 | 29.55
0 | 0 | 243.9 | 67.27 | 409470.8 | 29.71
0 | 0 | 242.8 | 67.71 | 412125.9 | 29.98
0 | 0 | 241.5 | 68.16 | 414872 | 30.19
0 | 0 | 240.3 | 68.7 | 418185.4 | 30.36
0 | 0 | 239.1 | 69.09 | 420539.1 | 30.56
0 | 0 | 237.9 | 69.54 | 423272.9 | 30.8
0 | 0 | 236.7 | 70.01 | 426116.9 | 31
0 | 0 | 235.5 | 70.4 | 428494.6 | 31.21
0 | 0 | 234.3 | 70.99 | 432075.9 | 31.43
0 | 0 | 233.1 | 71.48 | 435080.9 | 31.65
0 | 0 | 231.9 | 72.05 | 438557.9 | 31.88
0 | 0 | 230.7 | 72.37 | 440521.4 | 32.05
0 | 0 | 229.5 | 72.78 | 443018.6 | 32.25
0 | 0 | 228.3 | 73.47 | 447185.1 | 32.54
0 | 0 | 227.1 | 73.87 | 449602.9 | 32.68
0 | 0 | 225.9 | 74.35 | 452540.3 | 32.93
0 | 0 | 224.7 | 74.85 | 455605.6 | 33.12
0 | 0 | 223.4 | 75.41 | 458970.6 | 33.4
0 | 0 | 222.3 | 75.98 | 462462 | 33.62
0 | 0 | 221.1 | 76.5 | 465609.1 | 33.88
0 | 0 | 219.8 | 77.11 | 469318.1 | 34.12
0 | 0 | 218.7 | 77.41 | 471166.7 | 34.33
0 | 0 | 217.4 | 77.97 | 474552.1 | 34.59
0 | 0 | 216.2 | 78.67 | 478815.8 | 34.81
0 | 0 | 215 | 79.13 | 481658.9 | 35.09
0 | 0 | 213.8 | 79.7 | 485108.1 | 35.31
0 | 0 | 212.6 | 80.07 | 487381.8 | 35.47
0 | 0 | 211.4 | 80.77 | 491639.2 | 35.76
0 | 0 | 210.2 | 81.26 | 494629.8 | 35.98
0 | 0 | 209 | 81.79 | 497834.6 | 36.24
0 | 0 | 207.8 | 82.42 | 501681.7 | 36.5
0 | 0 | 206.6 | 82.98 | 505089.4 | 36.73
0 | 0 | 205.4 | 83.44 | 507897.8 | 37
0 | 0 | 204.2 | 84.14 | 512135 | 37.23
0 | 0 | 203 | 84.6 | 514908.1 | 37.4
0 | 0 | 201.8 | 85.3 | 519175.8 | 37.73
0 | 0 | 200.6 | 85.8 | 522247.4 | 38.04
0 | 0 | 199.4 | 86.37 | 525698.5 | 38.23
0 | 0 | 198.2 | 87.13 | 530342.8 | 38.52
0 | 0 | 197 | 87.51 | 532665.2 | 38.74
0 | 0 | 195.7 | 88.11 | 536274.9 | 38.96
0 | 0 | 194.5 | 88.64 | 539543.3 | 39.25
0 | 0 | 193.3 | 89.37 | 543941 | 39.56
0 | 0 | 192.1 | 89.77 | 546397.5 | 39.8
0 | 0 | 191 | 90.63 | 551611.4 | 40.1
0 | 0 | 189.7 | 91.11 | 554558.5 | 40.36
0 | 0 | 188.6 | 91.73 | 558350 | 40.64
0 | 0 | 187.3 | 92.35 | 562088.8 | 40.87
0 | 0 | 186.1 | 93.06 | 566424.1 | 41.16
0 | 0 | 185 | 93.72 | 570434 | 41.51
0 | 0 | 183.7 | 94.36 | 574323.1 | 41.77
0 | 0 | 182.5 | 94.76 | 576802.8 | 41.99
0 | 0 | 181.3 | 95.57 | 581677.8 | 42.36
0 | 0 | 180.1 | 96.4 | 586727.3 | 42.66
0 | 0 | 179 | 97.03 | 590563.3 | 42.92
0 | 0 | 177.7 | 97.73 | 594877.4 | 43.3
0 | 0 | 176.5 | 98.22 | 597823.2 | 43.49
0 | 0 | 175.3 | 98.9 | 601987.3 | 43.78
0 | 0 | 174.1 | 99.63 | 606385.8 | 44.09
0 | 0 | 172.9 | 100.25 | 610211.4 | 44.41
0 | 0 | 171.7 | 100.87 | 613985.9 | 44.64
0 | 0 | 170.5 | 101.6 | 618384.1 | 45.04
0 | 0 | 169.3 | 102.39 | 623232.5 | 45.36
0 | 0 | 168.1 | 103.2 | 628156.5 | 45.61
0 | 0 | 166.9 | 103.95 | 632728.3 | 45.93
0 | 0 | 165.7 | 104.56 | 636424.6 | 46.31
0 | 0 | 164.5 | 104.97 | 638934.2 | 46.45
0 | 0 | 163.3 | 105.83 | 644135.2 | 46.87
0 | 0 | 162.1 | 106.66 | 649182.4 | 47.22
0 | 0 | 160.9 | 107.27 | 652889.8 | 47.47
0 | 0 | 159.7 | 107.83 | 656357.3 | 47.71
0 | 0 | 158.5 | 108.39 | 659706.2 | 48.09
0 | 0 | 157.3 | 109.18 | 664555.5 | 48.42
0 | 0 | 156.1 | 110.26 | 671091.5 | 48.81
0 | 0 | 154.9 | 111.15 | 676509.4 | 49.14
0 | 0 | 153.7 | 112.11 | 682347.8 | 49.54
0 | 0 | 152.5 | 112.63 | 685568.3 | 49.85
0 | 0 | 151.2 | 113.43 | 690412.1 | 50.21
0 | 0 | 150.1 | 114.19 | 695034.2 | 50.54
0 | 0 | 148.8 | 115.03 | 700162.1 | 50.89
0 | 0 | 147.6 | 115.87 | 705241.7 | 51.25
0 | 0 | 146.4 | 116.5 | 709100.8 | 51.52
0 | 0 | 145.2 | 117.25 | 713693.2 | 51.95
0 | 0 | 144 | 118.2 | 719458.5 | 52.29
0 | 0 | 142.8 | 118.85 | 723392.8 | 52.6
0 | 0 | 141.6 | 119.63 | 728131.5 | 52.94
0 | 0 | 140.4 | 120.42 | 732954.9 | 53.33
0 | 0 | 139.2 | 121.23 | 737903.4 | 53.7
0 | 0 | 138 | 122.19 | 743743.3 | 54.06
0 | 0 | 136.8 | 122.85 | 747759.7 | 54.35
0 | 0 | 135.6 | 123.6 | 752308.3 | 54.74
0 | 0 | 134.4 | 124.14 | 755626.9 | 55.01
0 | 0 | 133.2 | 125.34 | 762925.1 | 55.48
0 | 0 | 132 | 126.3 | 768756.2 | 55.88
0 | 0 | 130.8 | 127.21 | 774257.1 | 56.3
0 | 0 | 129.6 | 128.11 | 779773.1 | 56.72
0 | 0 | 128.4 | 128.92 | 784724.3 | 57.14
0 | 0 | 127.2 | 129.9 | 790666.1 | 57.54
0 | 0 | 126 | 130.57 | 794747.8 | 57.87
0 | 0 | 124.8 | 131.69 | 801528.1 | 58.31
0 | 0 | 123.6 | 132.39 | 805821.7 | 58.57
0 | 0 | 122.4 | 133.16 | 810503.7 | 58.98
0 | 0 | 121.2 | 134.13 | 816409.9 | 59.37
0 | 0 | 120 | 134.77 | 820315.7 | 59.7
0 | 0 | 118.8 | 135.55 | 825068.6 | 60.1
0 | 0 | 117.6 | 136.8 | 832644.1 | 60.71
0 | 0 | 116.4 | 137.78 | 838600.8 | 60.96
0 | 0 | 115.2 | 138.85 | 845116.9 | 61.41
0 | 0 | 113.9 | 139.71 | 850393.4 | 61.82
0 | 0 | 112.7 | 140.56 | 855546.7 | 62.21
0 | 0 | 111.5 | 141.52 | 861374.6 | 62.71
0 | 0 | 110.3 | 142.58 | 867834.4 | 63.13
0 | 0 | 109.1 | 143.24 | 871831.6 | 63.46
0 | 0 | 107.9 | 143.86 | 875657.7 | 63.67
0 | 0 | 106.7 | 145.1 | 883173.5 | 64.25
0 | 0 | 105.5 | 146.01 | 888735.8 | 64.68
0 | 0 | 104.3 | 147.08 | 895199.8 | 65.06
0 | 0 | 103.1 | 148.39 | 903186.1 | 65.64
0 | 0 | 101.9 | 149.16 | 907915.1 | 66
0 | 0 | 100.7 | 150.39 | 915393.9 | 66.54
0 | 0 | 99.5 | 151 | 919093.5 | 66.87
0 | 0 | 98.3 | 152.03 | 925375.3 | 67.29
0 | 0 | 97.1 | 153.23 | 932658 | 67.91
0 | 0 | 95.9 | 154.37 | 939631.6 | 68.36
0 | 0 | 94.7 | 155.31 | 945296.6 | 68.66
0 | 0 | 93.5 | 156.24 | 950970.5 | 69.17
0 | 0 | 92.3 | 157.28 | 957316 | 69.57
0 | 0 | 91.1 | 158.18 | 962795.8 | 70.12
0 | 0 | 89.9 | 159.39 | 970132.5 | 70.58
0 | 0 | 88.7 | 160.39 | 976266.9 | 70.85
0 | 0 | 87.4 | 161.48 | 982880.4 | 71.5
0 | 0 | 86.3 | 162.4 | 988460.1 | 71.89
0 | 0 | 85 | 163.35 | 994233.2 | 72.34
0 | 0 | 83.8 | 164.5 | 1001285.7 | 72.78
0 | 0 | 82.6 | 165.65 | 1008229.6 | 73.34
0 | 0 | 81.4 | 166.68 | 1014555 | 73.7
0 | 0 | 80.2 | 167.6 | 1020146.2 | 74.24
0 | 0 | 79 | 168.65 | 1026545.3 | 74.64
0 | 0 | 77.8 | 169.76 | 1033298.2 | 75.15
0 | 0 | 76.6 | 170.89 | 1040147.8 | 75.6
0 | 0 | 75.4 | 171.87 | 1046127.7 | 76.08
0 | 0 | 74.2 | 173.17 | 1054041.5 | 76.69
0 | 0 | 73 | 173.99 | 1059003.5 | 77.09
0 | 0 | 71.8 | 175.39 | 1067549.7 | 77.64
0 | 0 | 70.6 | 176.29 | 1073036.3 | 78.03
0 | 0 | 69.3 | 177.33 | 1079331.1 | 78.58
0 | 0 | 68.2 | 178.44 | 1086087.9 | 78.97
0 | 0 | 66.9 | 179.76 | 1094156.4 | 79.67
0 | 0 | 65.7 | 180.96 | 1101471.5 | 80.11
0 | 0 | 64.5 | 181.99 | 1107730.9 | 80.45
0 | 0 | 63.3 | 183.04 | 1114095.3 | 80.83
0 | 0 | 62.1 | 184.14 | 1120778.3 | 81.41
0 | 0 | 60.9 | 185.37 | 1128278.6 | 82.03
0 | 0 | 59.7 | 186.39 | 1134491.5 | 82.41
0 | 0 | 58.5 | 187.66 | 1142202 | 82.99
0 | 0 | 57.3 | 188.81 | 1149245.5 | 83.56
0 | 0 | 56.1 | 189.86 | 1155647.3 | 84.06
0 | 0 | 54.8 | 190.96 | 1162286.4 | 84.56
0 | 0 | 53.7 | 192.23 | 1170064.8 | 85.07
0 | 0 | 52.4 | 193.35 | 1176836.8 | 85.67
0 | 0 | 51.2 | 194.6 | 1184451.7 | 86.2
0 | 0 | 50 | 195.55 | 1190267.4 | 86.62
0 | 0 | 48.8 | 196.99 | 1198986.7 | 87.26
0 | 0 | 47.6 | 198.22 | 1206520.7 | 87.67
0 | 0 | 46.4 | 199.17 | 1212262.1 | 88.18
0 | 0 | 45.2 | 200.52 | 1220507.8 | 88.65
0 | 0 | 44 | 201.4 | 1225855.2 | 89.05
0 | 0 | 42.8 | 202.81 | 1234442.8 | 89.64
0 | 0 | 41.6 | 203.84 | 1240715.3 | 90.15
0 | 0 | 40.3 | 204.81 | 1246621 | 90.58
0 | 0 | 39.1 | 206.11 | 1254525.9 | 91.14
0 | 0 | 37.9 | 207.04 | 1260159.4 | 91.61
0 | 0 | 36.7 | 208.36 | 1268228.6 | 92.35
0 | 0 | 35.5 | 209.61 | 1275840.2 | 92.75
0 | 0 | 34.3 | 210.67 | 1282267.9 | 93.21
0 | 0 | 33.1 | 211.98 | 1290269.3 | 93.72
0 | 0 | 31.9 | 213.02 | 1296559.7 | 94.28
0 | 0 | 30.7 | 214.18 | 1303645.8 | 94.78
0 | 0 | 29.5 | 215.16 | 1309604.7 | 95.2
0 | 0 | 28.3 | 216.54 | 1317996.7 | 95.73
0 | 0 | 27 | 217.61 | 1324502.9 | 96.32
0 | 0 | 25.9 | 218.68 | 1331020.6 | 96.93
0 | 0 | 24.6 | 220.18 | 1340191.7 | 97.32
0 | 0 | 23.4 | 221.26 | 1346741.6 | 97.82
0 | 0 | 22.3 | 222.11 | 1351914.2 | 98.37
0 | 0 | 21 | 222.97 | 1357177.6 | 98.69
0 | 0 | 19.8 | 223.94 | 1363073.6 | 99.24
0 | 0 | 18.6 | 224.79 | 1368201.8 | 99.45
0 | 0 | 17.4 | 225.63 | 1373348.1 | 99.9
0 | 0 | 16.2 | 226.15 | 1376494.6 | 99.97
0 | 0 | 15 | 225.96 | 1375323.3 | 100
0 | 0 | 13.8 | 225.15 | 1370431.4 | 99.63
0 | 0 | 12.6 | 223.21 | 1358636.9 | 98.91
0 | 0 | 11.4 | 221.67 | 1349205.4 | 98.08
0 | 0 | 10.2 | 217.51 | 1323928 | 96.32
0 | 0 | 9 | 211.99 | 1290304.4 | 93.75
0 | 0 | 7.8 | 203.3 | 1237430 | 89.96
0 | 0 | 6.5 | 190.19 | 1157604.4 | 84.14
0 | 0 | 5.4 | 171.35 | 1042958.7 | 75.83
0 | 0 | 4.2 | 144.24 | 877952.6 | 63.81
0 | 0 | 2.9 | 123.51 | 751769.5 | 54.68
0 | 0 | 1.7 | 113.35 | 689949.4 | 50.24
0 | 0 | 0.5 | 106.61 | 648896.4 | 47.15
0 | 0 | -0.5 | 102.77 | 625541.5 | 45.53