## 1_Rendements énergie
Measurement time: | 12/08/2022 18:00:16
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 19.24 | 47561.7 | 20.48
0 | 0 | 309.1 | 20.74 | 51265.8 | 20.6
0 | 0 | 307.6 | 20.83 | 51504.6 | 20.67
0 | 0 | 306.6 | 21.02 | 51974.8 | 20.89
0 | 0 | 305.3 | 21.15 | 52292.6 | 21.01
0 | 0 | 304.1 | 21.35 | 52782.9 | 21.19
0 | 0 | 302.9 | 21.46 | 53066 | 21.33
0 | 0 | 301.7 | 21.62 | 53451.6 | 21.46
0 | 0 | 300.5 | 21.73 | 53713.5 | 21.55
0 | 0 | 299.3 | 21.83 | 53983.6 | 21.68
0 | 0 | 298.1 | 22.08 | 54597.4 | 21.97
0 | 0 | 296.9 | 22.18 | 54836.5 | 22.02
0 | 0 | 295.7 | 22.34 | 55238 | 22.21
0 | 0 | 294.5 | 22.52 | 55683.7 | 22.35
0 | 0 | 293.3 | 22.64 | 55983.6 | 22.46
0 | 0 | 292.1 | 22.78 | 56332 | 22.67
0 | 0 | 290.9 | 22.95 | 56744.8 | 22.82
0 | 0 | 289.7 | 23.07 | 57049 | 22.94
0 | 0 | 288.5 | 23.29 | 57574 | 23.13
0 | 0 | 287.3 | 23.45 | 57968 | 23.31
0 | 0 | 286.1 | 23.53 | 58174.1 | 23.42
0 | 0 | 284.9 | 23.7 | 58600 | 23.54
0 | 0 | 283.7 | 23.93 | 59151.6 | 23.76
0 | 0 | 282.5 | 24.01 | 59365.5 | 23.88
0 | 0 | 281.2 | 24.25 | 59960.5 | 24.07
0 | 0 | 280.1 | 24.39 | 60289 | 24.23
0 | 0 | 278.8 | 24.46 | 60477.1 | 24.31
0 | 0 | 277.7 | 24.68 | 61005.5 | 24.5
0 | 0 | 276.4 | 24.86 | 61469.1 | 24.72
0 | 0 | 275.2 | 25.02 | 61854 | 24.83
0 | 0 | 274 | 25.22 | 62360 | 25.07
0 | 0 | 272.8 | 25.32 | 62610.7 | 25.19
0 | 0 | 271.6 | 25.54 | 63143.9 | 25.37
0 | 0 | 270.4 | 25.66 | 63445.9 | 25.55
0 | 0 | 269.2 | 25.84 | 63885.9 | 25.73
0 | 0 | 268 | 26.05 | 64400.7 | 25.89
0 | 0 | 266.8 | 26.3 | 65024.5 | 26.08
0 | 0 | 265.6 | 26.42 | 65327.7 | 26.21
0 | 0 | 264.4 | 26.53 | 65584.6 | 26.36
0 | 0 | 263.2 | 26.75 | 66146.4 | 26.61
0 | 0 | 262 | 26.97 | 66670.8 | 26.79
0 | 0 | 260.8 | 27.05 | 66881.8 | 26.88
0 | 0 | 259.6 | 27.34 | 67602.7 | 27.15
0 | 0 | 258.4 | 27.47 | 67924.8 | 27.26
0 | 0 | 257.2 | 27.65 | 68368.3 | 27.45
0 | 0 | 256 | 27.85 | 68849.4 | 27.69
0 | 0 | 254.8 | 28.04 | 69317.7 | 27.86
0 | 0 | 253.6 | 28.19 | 69703.5 | 28.04
0 | 0 | 252.4 | 28.45 | 70348.8 | 28.25
0 | 0 | 251.2 | 28.64 | 70815.2 | 28.46
0 | 0 | 250 | 28.82 | 71244.8 | 28.62
0 | 0 | 248.8 | 29.06 | 71852.1 | 28.86
0 | 0 | 247.6 | 29.19 | 72175.6 | 29.03
0 | 0 | 246.4 | 29.43 | 72757.2 | 29.3
0 | 0 | 245.1 | 29.65 | 73312.1 | 29.45
0 | 0 | 244 | 29.9 | 73917.3 | 29.71
0 | 0 | 242.7 | 30.03 | 74247.9 | 29.82
0 | 0 | 241.6 | 30.24 | 74771.8 | 30.06
0 | 0 | 240.3 | 30.32 | 74956.5 | 30.08
0 | 0 | 239.1 | 30.57 | 75581.7 | 30.35
0 | 0 | 237.9 | 30.83 | 76219.8 | 30.63
0 | 0 | 236.7 | 31.01 | 76674 | 30.79
0 | 0 | 235.5 | 31.17 | 77051.9 | 30.99
0 | 0 | 234.3 | 31.43 | 77707.8 | 31.23
0 | 0 | 233.1 | 31.52 | 77936.1 | 31.35
0 | 0 | 231.9 | 31.82 | 78671.4 | 31.59
0 | 0 | 230.7 | 32.12 | 79421.2 | 31.9
0 | 0 | 229.5 | 32.33 | 79929.1 | 32.1
0 | 0 | 228.3 | 32.5 | 80362.2 | 32.32
0 | 0 | 227.1 | 32.75 | 80967.7 | 32.54
0 | 0 | 225.9 | 32.97 | 81509.4 | 32.75
0 | 0 | 224.7 | 33.21 | 82105.1 | 32.97
0 | 0 | 223.5 | 33.45 | 82702.7 | 33.24
0 | 0 | 222.2 | 33.64 | 83165.5 | 33.38
0 | 0 | 221 | 33.91 | 83846.6 | 33.69
0 | 0 | 219.9 | 34.1 | 84313.7 | 33.85
0 | 0 | 218.6 | 34.31 | 84814.5 | 34.1
0 | 0 | 217.5 | 34.65 | 85665 | 34.4
0 | 0 | 216.2 | 34.78 | 85995 | 34.58
0 | 0 | 215 | 35.02 | 86580.8 | 34.76
0 | 0 | 213.8 | 35.24 | 87132.5 | 34.98
0 | 0 | 212.6 | 35.45 | 87656.8 | 35.26
0 | 0 | 211.4 | 35.79 | 88497.1 | 35.59
0 | 0 | 210.2 | 35.96 | 88911.9 | 35.76
0 | 0 | 209 | 36.22 | 89556.4 | 35.98
0 | 0 | 207.8 | 36.44 | 90090.1 | 36.21
0 | 0 | 206.6 | 36.83 | 91054.7 | 36.49
0 | 0 | 205.4 | 37 | 91482 | 36.81
0 | 0 | 204.2 | 37.19 | 91958.6 | 36.95
0 | 0 | 203 | 37.47 | 92627.4 | 37.21
0 | 0 | 201.8 | 37.77 | 93378.5 | 37.5
0 | 0 | 200.6 | 38.03 | 94030.3 | 37.81
0 | 0 | 199.4 | 38.27 | 94618.1 | 38.03
0 | 0 | 198.2 | 38.55 | 95320.8 | 38.28
0 | 0 | 197 | 38.79 | 95913.9 | 38.59
0 | 0 | 195.8 | 39.1 | 96666.5 | 38.8
0 | 0 | 194.6 | 39.25 | 97033 | 38.94
0 | 0 | 193.4 | 39.64 | 98008.2 | 39.35
0 | 0 | 192.2 | 39.93 | 98723.7 | 39.68
0 | 0 | 191 | 40.15 | 99276.6 | 39.92
0 | 0 | 189.8 | 40.52 | 100176.2 | 40.22
0 | 0 | 188.5 | 40.73 | 100696.1 | 40.48
0 | 0 | 187.4 | 41.03 | 101429 | 40.76
0 | 0 | 186.1 | 41.28 | 102046.6 | 40.98
0 | 0 | 184.9 | 41.5 | 102609.4 | 41.24
0 | 0 | 183.7 | 41.87 | 103520.2 | 41.53
0 | 0 | 182.5 | 42.14 | 104174.2 | 41.87
0 | 0 | 181.3 | 42.38 | 104788.9 | 42.1
0 | 0 | 180.1 | 42.66 | 105475.2 | 42.38
0 | 0 | 178.9 | 42.95 | 106175.7 | 42.69
0 | 0 | 177.7 | 43.3 | 107043.9 | 43.02
0 | 0 | 176.5 | 43.56 | 107694.7 | 43.3
0 | 0 | 175.3 | 43.95 | 108666.8 | 43.64
0 | 0 | 174.1 | 44.18 | 109239.1 | 43.86
0 | 0 | 172.9 | 44.42 | 109830.9 | 44.13
0 | 0 | 171.7 | 44.82 | 110804.1 | 44.5
0 | 0 | 170.5 | 45.09 | 111469.7 | 44.85
0 | 0 | 169.3 | 45.38 | 112194.3 | 45.07
0 | 0 | 168.1 | 45.79 | 113207.6 | 45.41
0 | 0 | 166.9 | 46.02 | 113773.1 | 45.71
0 | 0 | 165.7 | 46.26 | 114381.3 | 45.98
0 | 0 | 164.4 | 46.66 | 115368.9 | 46.4
0 | 0 | 163.3 | 46.97 | 116120.2 | 46.7
0 | 0 | 162 | 47.22 | 116749.8 | 46.99
0 | 0 | 160.9 | 47.61 | 117715.4 | 47.29
0 | 0 | 159.6 | 47.99 | 118644.5 | 47.77
0 | 0 | 158.4 | 48.28 | 119363.4 | 47.94
0 | 0 | 157.2 | 48.55 | 120029.6 | 48.28
0 | 0 | 156 | 48.92 | 120939.2 | 48.58
0 | 0 | 154.8 | 49.25 | 121772.5 | 48.95
0 | 0 | 153.6 | 49.62 | 122682.4 | 49.35
0 | 0 | 152.4 | 50.03 | 123703.9 | 49.6
0 | 0 | 151.2 | 50.22 | 124168.6 | 49.91
0 | 0 | 150 | 50.46 | 124759.1 | 50.18
0 | 0 | 148.8 | 50.87 | 125765.6 | 50.56
0 | 0 | 147.6 | 51.25 | 126705.6 | 50.86
0 | 0 | 146.4 | 51.59 | 127539 | 51.15
0 | 0 | 145.2 | 51.97 | 128493.7 | 51.62
0 | 0 | 144 | 52.33 | 129378.5 | 52.01
0 | 0 | 142.8 | 52.72 | 130336.2 | 52.47
0 | 0 | 141.6 | 53.01 | 131055.7 | 52.71
0 | 0 | 140.4 | 53.4 | 132017.8 | 52.97
0 | 0 | 139.2 | 53.66 | 132656.7 | 53.35
0 | 0 | 138 | 54.05 | 133641 | 53.65
0 | 0 | 136.8 | 54.43 | 134578.9 | 54.14
0 | 0 | 135.6 | 54.75 | 135369.7 | 54.45
0 | 0 | 134.4 | 55.14 | 136318.4 | 54.76
0 | 0 | 133.2 | 55.49 | 137188.5 | 55.1
0 | 0 | 131.9 | 55.85 | 138072 | 55.48
0 | 0 | 130.8 | 56.25 | 139075.6 | 55.89
0 | 0 | 129.6 | 56.71 | 140213 | 56.4
0 | 0 | 128.3 | 57.02 | 140961.3 | 56.57
0 | 0 | 127.1 | 57.39 | 141899.8 | 56.96
0 | 0 | 126 | 57.87 | 143076.8 | 57.42
0 | 0 | 124.7 | 58.23 | 143964.5 | 57.84
0 | 0 | 123.5 | 58.57 | 144802.3 | 58.19
0 | 0 | 122.3 | 59.13 | 146193.7 | 58.73
0 | 0 | 121.1 | 59.48 | 147066 | 59.09
0 | 0 | 119.9 | 59.85 | 147964.3 | 59.48
0 | 0 | 118.7 | 60.24 | 148934.6 | 59.89
0 | 0 | 117.5 | 60.6 | 149836.2 | 60.29
0 | 0 | 116.3 | 60.88 | 150522.4 | 60.55
0 | 0 | 115.1 | 61.27 | 151475.4 | 60.89
0 | 0 | 113.9 | 61.79 | 152758.1 | 61.33
0 | 0 | 112.7 | 62.09 | 153506.7 | 61.68
0 | 0 | 111.5 | 62.54 | 154614.8 | 62.13
0 | 0 | 110.3 | 63.01 | 155792.1 | 62.56
0 | 0 | 109.1 | 63.51 | 157019.6 | 63
0 | 0 | 107.9 | 64.02 | 158289.7 | 63.56
0 | 0 | 106.7 | 64.51 | 159496.6 | 63.84
0 | 0 | 105.5 | 64.81 | 160233.8 | 64.36
0 | 0 | 104.2 | 65.18 | 161158.2 | 64.88
0 | 0 | 103.1 | 65.66 | 162345.9 | 65.19
0 | 0 | 101.9 | 66.12 | 163470.3 | 65.63
0 | 0 | 100.7 | 66.56 | 164553.7 | 66.23
0 | 0 | 99.5 | 66.8 | 165151.1 | 66.43
0 | 0 | 98.3 | 67.23 | 166216.7 | 66.84
0 | 0 | 97.1 | 67.74 | 167469.8 | 67.24
0 | 0 | 95.8 | 68.06 | 168266.5 | 67.73
0 | 0 | 94.7 | 68.39 | 169076.6 | 68.15
0 | 0 | 93.5 | 68.97 | 170521.4 | 68.6
0 | 0 | 92.2 | 69.61 | 172091.4 | 68.98
0 | 0 | 91.1 | 69.94 | 172923.4 | 69.54
0 | 0 | 89.8 | 70.5 | 174294.8 | 70.01
0 | 0 | 88.6 | 71.05 | 175650.2 | 70.44
0 | 0 | 87.4 | 71.54 | 176880.2 | 71.06
0 | 0 | 86.2 | 71.94 | 177867.6 | 71.55
0 | 0 | 85 | 72.45 | 179111.9 | 71.91
0 | 0 | 83.8 | 72.91 | 180252.2 | 72.35
0 | 0 | 82.6 | 73.37 | 181404 | 72.9
0 | 0 | 81.4 | 73.82 | 182515.6 | 73.36
0 | 0 | 80.2 | 74.32 | 183742 | 73.79
0 | 0 | 79 | 74.68 | 184629.3 | 74.18
0 | 0 | 77.8 | 75.26 | 186064.7 | 74.7
0 | 0 | 76.6 | 75.69 | 187144.5 | 75.09
0 | 0 | 75.4 | 75.94 | 187756.6 | 75.41
0 | 0 | 74.1 | 76.42 | 188938 | 75.96
0 | 0 | 73 | 76.9 | 190125.5 | 76.55
0 | 0 | 71.7 | 77.54 | 191710 | 77.05
0 | 0 | 70.5 | 78.1 | 193090.9 | 77.55
0 | 0 | 69.3 | 78.72 | 194624 | 77.98
0 | 0 | 68.1 | 79.04 | 195410 | 78.39
0 | 0 | 66.9 | 79.59 | 196786.3 | 79.02
0 | 0 | 65.7 | 80.07 | 197953.7 | 79.48
0 | 0 | 64.5 | 80.58 | 199215.5 | 79.89
0 | 0 | 63.3 | 81.05 | 200396.4 | 80.58
0 | 0 | 62 | 81.7 | 201999.9 | 81.11
0 | 0 | 60.9 | 82.02 | 202779.4 | 81.45
0 | 0 | 59.6 | 82.47 | 203895.1 | 82.01
0 | 0 | 58.5 | 83.17 | 205630.1 | 82.51
0 | 0 | 57.2 | 83.78 | 207122.8 | 83.14
0 | 0 | 56 | 84.05 | 207801.8 | 83.42
0 | 0 | 54.8 | 84.53 | 208993.7 | 83.93
0 | 0 | 53.6 | 84.95 | 210038.5 | 84.43
0 | 0 | 52.4 | 85.63 | 211702.6 | 85.05
0 | 0 | 51.2 | 86.21 | 213142.3 | 85.4
0 | 0 | 50 | 86.73 | 214435.1 | 86.13
0 | 0 | 48.8 | 87.06 | 215249.2 | 86.55
0 | 0 | 47.6 | 87.57 | 216492.8 | 86.98
0 | 0 | 46.4 | 88.17 | 217993.2 | 87.38
0 | 0 | 45.2 | 88.51 | 218825 | 87.97
0 | 0 | 44 | 89.08 | 220237.2 | 88.49
0 | 0 | 42.8 | 89.42 | 221077.5 | 88.88
0 | 0 | 41.6 | 89.98 | 222465.5 | 89.46
0 | 0 | 40.4 | 90.66 | 224139.6 | 89.96
0 | 0 | 39.1 | 91.34 | 225820.5 | 90.54
0 | 0 | 38 | 91.63 | 226542.2 | 90.99
0 | 0 | 36.7 | 92.09 | 227684.5 | 91.53
0 | 0 | 35.5 | 92.62 | 228982 | 91.93
0 | 0 | 34.3 | 93.2 | 230429.1 | 92.41
0 | 0 | 33.1 | 93.86 | 232066.6 | 93
0 | 0 | 31.9 | 94.45 | 233507.6 | 93.77
0 | 0 | 30.7 | 94.91 | 234651 | 94.29
0 | 0 | 29.5 | 95.33 | 235691.5 | 94.68
0 | 0 | 28.3 | 95.98 | 237296.9 | 95.29
0 | 0 | 27.1 | 96.25 | 237958.1 | 95.63
0 | 0 | 25.9 | 96.81 | 239336 | 96.15
0 | 0 | 24.7 | 97.32 | 240618.6 | 96.6
0 | 0 | 23.5 | 98.01 | 242304.1 | 97.24
0 | 0 | 22.3 | 98.33 | 243106.8 | 97.82
0 | 0 | 21.1 | 99 | 244773.2 | 98.39
0 | 0 | 19.8 | 99.28 | 245457.3 | 98.79
0 | 0 | 18.6 | 99.83 | 246816.3 | 99.35
0 | 0 | 17.4 | 99.99 | 247202.3 | 99.27
0 | 0 | 16.2 | 100.74 | 249072.3 | 99.76
0 | 0 | 15 | 100.58 | 248663 | 99.79
0 | 0 | 13.8 | 100.78 | 249172.3 | 99.99
0 | 0 | 12.6 | 100.54 | 248559 | 100
0 | 0 | 11.4 | 100.04 | 247344.5 | 99.33
0 | 0 | 10.2 | 99.02 | 244807.6 | 98.17
0 | 0 | 9 | 97.79 | 241773.5 | 97.12
0 | 0 | 7.8 | 95.33 | 235679.2 | 94.65
0 | 0 | 6.6 | 91.96 | 227360.3 | 91.27
0 | 0 | 5.4 | 86.89 | 214830.9 | 86.3
0 | 0 | 4.2 | 79.72 | 197103.6 | 79.19
0 | 0 | 3 | 69.42 | 171637.6 | 68.95
0 | 0 | 1.7 | 59.81 | 147860.3 | 59.41
0 | 0 | 0.6 | 54.3 | 134255.9 | 53.92
0 | 0 | -0.5 | 50.98 | 126046.2 | 50.65
Measurement time: | 12/08/2022 18:08:37
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 33.1 | 81833.9 | 33.6
0 | 0 | 309.1 | 34.77 | 85965.8 | 33.86
0 | 0 | 307.6 | 35.14 | 86867.2 | 34.19
0 | 0 | 306.6 | 35.13 | 86861.2 | 34.24
0 | 0 | 305.3 | 35.24 | 87125.6 | 34.33
0 | 0 | 304.2 | 35.38 | 87477.2 | 34.52
0 | 0 | 302.9 | 35.77 | 88431.4 | 34.67
0 | 0 | 301.7 | 35.69 | 88236.2 | 34.83
0 | 0 | 300.5 | 35.81 | 88538.9 | 35.03
0 | 0 | 299.3 | 36.29 | 89714.5 | 35.26
0 | 0 | 298.2 | 36.3 | 89741.2 | 35.37
0 | 0 | 296.9 | 36.37 | 89909.5 | 35.49
0 | 0 | 295.7 | 36.65 | 90610.7 | 35.73
0 | 0 | 294.5 | 36.81 | 91002.6 | 35.89
0 | 0 | 293.3 | 37.02 | 91532 | 36.11
0 | 0 | 292.1 | 37.24 | 92073.8 | 36.22
0 | 0 | 290.9 | 37.22 | 92019.1 | 36.38
0 | 0 | 289.7 | 37.58 | 92909.5 | 36.55
0 | 0 | 288.5 | 37.62 | 93011.1 | 36.73
0 | 0 | 287.3 | 37.97 | 93882.8 | 36.92
0 | 0 | 286.1 | 37.97 | 93884.1 | 37.04
0 | 0 | 284.9 | 38.3 | 94692.1 | 37.31
0 | 0 | 283.7 | 38.39 | 94907.2 | 37.48
0 | 0 | 282.5 | 38.67 | 95602.5 | 37.49
0 | 0 | 281.3 | 38.79 | 95903.7 | 37.79
0 | 0 | 280.1 | 39.1 | 96663.6 | 38.05
0 | 0 | 278.9 | 39.2 | 96917.8 | 38.26
0 | 0 | 277.7 | 39.33 | 97229.6 | 38.41
0 | 0 | 276.5 | 39.57 | 97820.6 | 38.54
0 | 0 | 275.2 | 39.79 | 98383.3 | 38.67
0 | 0 | 274.1 | 39.92 | 98698.7 | 38.95
0 | 0 | 272.8 | 40.23 | 99458.6 | 39.24
0 | 0 | 271.6 | 40.47 | 100065.5 | 39.41
0 | 0 | 270.4 | 40.53 | 100198 | 39.53
0 | 0 | 269.2 | 40.75 | 100740.1 | 39.78
0 | 0 | 268 | 40.83 | 100949.8 | 39.95
0 | 0 | 266.8 | 41.2 | 101872.6 | 40.18
0 | 0 | 265.6 | 41.44 | 102466.2 | 40.27
0 | 0 | 264.4 | 41.48 | 102555.3 | 40.55
0 | 0 | 263.2 | 41.73 | 103173.7 | 40.79
0 | 0 | 262 | 41.8 | 103346.9 | 40.88
0 | 0 | 260.8 | 42.34 | 104690.9 | 41.15
0 | 0 | 259.6 | 42.28 | 104524.2 | 41.23
0 | 0 | 258.4 | 42.52 | 105128.6 | 41.43
0 | 0 | 257.2 | 42.7 | 105580.4 | 41.63
0 | 0 | 256 | 42.91 | 106100.2 | 41.91
0 | 0 | 254.8 | 43.13 | 106644.1 | 42.11
0 | 0 | 253.6 | 43.35 | 107175.1 | 42.24
0 | 0 | 252.4 | 43.66 | 107930.9 | 42.49
0 | 0 | 251.2 | 43.79 | 108254.8 | 42.69
0 | 0 | 250 | 44.01 | 108796.4 | 42.94
0 | 0 | 248.8 | 44.22 | 109320.4 | 43.14
0 | 0 | 247.6 | 44.34 | 109623.4 | 43.41
0 | 0 | 246.4 | 44.62 | 110311.2 | 43.65
0 | 0 | 245.2 | 45.02 | 111302.9 | 43.86
0 | 0 | 244 | 45.16 | 111655.1 | 44.01
0 | 0 | 242.7 | 45.19 | 111717 | 44.19
0 | 0 | 241.6 | 45.55 | 112607.3 | 44.5
0 | 0 | 240.4 | 45.78 | 113175.7 | 44.69
0 | 0 | 239.1 | 46.12 | 114034.2 | 44.98
0 | 0 | 238 | 46.3 | 114477.2 | 45.1
0 | 0 | 236.7 | 46.45 | 114836.5 | 45.43
0 | 0 | 235.5 | 46.76 | 115595.3 | 45.64
0 | 0 | 234.3 | 46.96 | 116109.8 | 45.78
0 | 0 | 233.1 | 47.15 | 116568.6 | 45.89
0 | 0 | 231.9 | 47.3 | 116937.3 | 46.25
0 | 0 | 230.7 | 47.51 | 117469.4 | 46.53
0 | 0 | 229.5 | 47.89 | 118410.4 | 46.69
0 | 0 | 228.3 | 48.21 | 119182.8 | 47.01
0 | 0 | 227.1 | 48.32 | 119459.3 | 47.15
0 | 0 | 225.9 | 48.46 | 119810.1 | 47.28
0 | 0 | 224.7 | 48.7 | 120409.2 | 47.73
0 | 0 | 223.5 | 49.19 | 121612.5 | 48.03
0 | 0 | 222.3 | 49.26 | 121778.5 | 48.13
0 | 0 | 221.1 | 49.43 | 122205.7 | 48.38
0 | 0 | 219.9 | 49.69 | 122855.5 | 48.69
0 | 0 | 218.7 | 49.99 | 123599.2 | 48.83
0 | 0 | 217.5 | 50.5 | 124851.1 | 49.18
0 | 0 | 216.3 | 50.43 | 124683.7 | 49.35
0 | 0 | 215.1 | 50.86 | 125741.9 | 49.67
0 | 0 | 213.9 | 51.04 | 126201.1 | 49.76
0 | 0 | 212.6 | 51.42 | 127121 | 50.07
0 | 0 | 211.5 | 51.59 | 127555.2 | 50.47
0 | 0 | 210.2 | 51.87 | 128241.4 | 50.55
0 | 0 | 209 | 51.9 | 128327.1 | 50.86
0 | 0 | 207.9 | 52.28 | 129256.6 | 50.93
0 | 0 | 206.6 | 52.55 | 129922.6 | 51.35
0 | 0 | 205.4 | 52.72 | 130336.3 | 51.58
0 | 0 | 204.2 | 52.84 | 130634.9 | 51.9
0 | 0 | 203 | 53.25 | 131647.5 | 52.15
0 | 0 | 201.8 | 53.68 | 132719 | 52.34
0 | 0 | 200.6 | 53.82 | 133070.9 | 52.63
0 | 0 | 199.4 | 54.28 | 134208.1 | 52.99
0 | 0 | 198.2 | 54.39 | 134472.9 | 53.13
0 | 0 | 197 | 54.53 | 134812.3 | 53.31
0 | 0 | 195.8 | 54.82 | 135523.3 | 53.4
0 | 0 | 194.6 | 55.28 | 136675.6 | 53.86
0 | 0 | 193.4 | 55.44 | 137059.9 | 54.27
0 | 0 | 192.2 | 55.94 | 138313.7 | 54.42
0 | 0 | 191 | 55.89 | 138173 | 54.63
0 | 0 | 189.8 | 56.37 | 139370.8 | 54.85
0 | 0 | 188.5 | 56.54 | 139782.9 | 55.28
0 | 0 | 187.4 | 56.76 | 140334.3 | 55.51
0 | 0 | 186.2 | 57.11 | 141202.1 | 55.66
0 | 0 | 185 | 57.36 | 141825.2 | 55.97
0 | 0 | 183.8 | 57.51 | 142187.8 | 56.27
0 | 0 | 182.6 | 57.95 | 143262.1 | 56.66
0 | 0 | 181.4 | 58.36 | 144276.7 | 57
0 | 0 | 180.1 | 58.67 | 145054.6 | 57.29
0 | 0 | 178.9 | 58.77 | 145298.2 | 57.48
0 | 0 | 177.8 | 59.17 | 146297.5 | 57.81
0 | 0 | 176.5 | 59.38 | 146817.3 | 58.12
0 | 0 | 175.3 | 59.84 | 147933.2 | 58.26
0 | 0 | 174.1 | 60.06 | 148496.6 | 58.65
0 | 0 | 172.9 | 60.06 | 148491.9 | 58.88
0 | 0 | 171.7 | 60.57 | 149746.6 | 59.32
0 | 0 | 170.5 | 60.97 | 150738.4 | 59.52
0 | 0 | 169.3 | 61.38 | 151744.6 | 59.9
0 | 0 | 168.1 | 61.46 | 151941.3 | 60.12
0 | 0 | 166.9 | 61.85 | 152921.9 | 60.38
0 | 0 | 165.7 | 61.84 | 152889.6 | 60.61
0 | 0 | 164.5 | 62.51 | 154552.7 | 60.86
0 | 0 | 163.3 | 62.51 | 154555.2 | 61.28
0 | 0 | 162.1 | 62.99 | 155721.9 | 61.58
0 | 0 | 160.9 | 63.37 | 156677.2 | 61.78
0 | 0 | 159.7 | 63.84 | 157839.8 | 62.2
0 | 0 | 158.5 | 63.96 | 158125.9 | 62.42
0 | 0 | 157.3 | 64.15 | 158597.7 | 62.75
0 | 0 | 156.1 | 64.65 | 159834.5 | 63.14
0 | 0 | 154.9 | 64.74 | 160059.3 | 63.38
0 | 0 | 153.6 | 65.28 | 161406 | 63.55
0 | 0 | 152.4 | 65.49 | 161908.2 | 63.89
0 | 0 | 151.3 | 65.75 | 162559.3 | 64.34
0 | 0 | 150 | 66.03 | 163240.5 | 64.61
0 | 0 | 148.8 | 66.62 | 164701.8 | 65.13
0 | 0 | 147.6 | 66.91 | 165421.5 | 65.34
0 | 0 | 146.4 | 67.3 | 166391.9 | 65.69
0 | 0 | 145.2 | 67.48 | 166823.8 | 66.01
0 | 0 | 144 | 67.74 | 167478.5 | 66.09
0 | 0 | 142.8 | 68.22 | 168672.6 | 66.45
0 | 0 | 141.6 | 68.57 | 169526.7 | 66.87
0 | 0 | 140.4 | 68.72 | 169893.8 | 67.4
0 | 0 | 139.2 | 69.14 | 170932.3 | 67.53
0 | 0 | 138 | 69.66 | 172223.8 | 67.83
0 | 0 | 136.8 | 69.83 | 172652.2 | 68.15
0 | 0 | 135.6 | 70.25 | 173688.4 | 68.49
0 | 0 | 134.4 | 70.56 | 174440.5 | 69.08
0 | 0 | 133.2 | 71.13 | 175854.2 | 69.38
0 | 0 | 132 | 71.28 | 176236.4 | 69.61
0 | 0 | 130.8 | 71.66 | 177173.7 | 70.09
0 | 0 | 129.6 | 71.98 | 177967.1 | 70.42
0 | 0 | 128.4 | 72.48 | 179197.9 | 70.78
0 | 0 | 127.2 | 73.11 | 180757.5 | 71.13
0 | 0 | 126 | 72.71 | 179764.5 | 71.32
0 | 0 | 124.8 | 73.56 | 181862.9 | 71.68
0 | 0 | 123.6 | 73.67 | 182138.4 | 72.07
0 | 0 | 122.3 | 74.12 | 183260.5 | 72.59
0 | 0 | 121.1 | 74.59 | 184417.5 | 72.88
0 | 0 | 119.9 | 74.86 | 185087.5 | 73.09
0 | 0 | 118.7 | 75.32 | 186226.4 | 73.51
0 | 0 | 117.6 | 75.32 | 186214.2 | 73.74
0 | 0 | 116.3 | 76.08 | 188090.1 | 74.31
0 | 0 | 115.1 | 76.29 | 188617.7 | 74.61
0 | 0 | 113.9 | 76.78 | 189822.8 | 74.99
0 | 0 | 112.7 | 77.11 | 190648.5 | 75.32
0 | 0 | 111.5 | 77.32 | 191166.2 | 75.6
0 | 0 | 110.3 | 77.99 | 192828.2 | 75.93
0 | 0 | 109.1 | 78.05 | 192974.6 | 76.55
0 | 0 | 107.9 | 78.49 | 194044.7 | 76.75
0 | 0 | 106.7 | 79.09 | 195531.7 | 77.14
0 | 0 | 105.5 | 79.45 | 196422.9 | 77.31
0 | 0 | 104.3 | 79.64 | 196899.2 | 77.78
0 | 0 | 103.1 | 79.85 | 197410 | 78.43
0 | 0 | 101.9 | 80.69 | 199491.6 | 78.52
0 | 0 | 100.7 | 80.88 | 199964.6 | 79.08
0 | 0 | 99.5 | 81.22 | 200802.8 | 79.39
0 | 0 | 98.3 | 81.66 | 201899.7 | 79.8
0 | 0 | 97.1 | 82.06 | 202874.7 | 80.12
0 | 0 | 95.9 | 82.34 | 203578.6 | 80.47
0 | 0 | 94.7 | 82.86 | 204863.3 | 81.08
0 | 0 | 93.5 | 83.28 | 205906.9 | 81.37
0 | 0 | 92.2 | 83.68 | 206889.1 | 82.04
0 | 0 | 91.1 | 84.36 | 208557.7 | 82.21
0 | 0 | 89.9 | 84.53 | 208986.4 | 82.73
0 | 0 | 88.7 | 84.97 | 210064.3 | 82.89
0 | 0 | 87.5 | 85.23 | 210707.1 | 83.32
0 | 0 | 86.2 | 85.86 | 212269.2 | 83.93
0 | 0 | 85.1 | 86.28 | 213309.7 | 84.21
0 | 0 | 83.8 | 86.51 | 213875.1 | 84.55
0 | 0 | 82.6 | 87.07 | 215264.3 | 85.1
0 | 0 | 81.4 | 87.27 | 215761.8 | 85.35
0 | 0 | 80.2 | 87.72 | 216866.4 | 85.79
0 | 0 | 79 | 88.27 | 218232.7 | 86.04
0 | 0 | 77.8 | 88.51 | 218832.3 | 86.46
0 | 0 | 76.6 | 88.76 | 219437.7 | 86.88
0 | 0 | 75.4 | 89.51 | 221307.9 | 87.31
0 | 0 | 74.2 | 89.81 | 222037 | 87.8
0 | 0 | 73 | 90.55 | 223876.1 | 88.32
0 | 0 | 71.8 | 90.45 | 223614.5 | 88.58
0 | 0 | 70.5 | 91.08 | 225173 | 89.17
0 | 0 | 69.3 | 91.66 | 226607.8 | 89.55
0 | 0 | 68.1 | 91.97 | 227391.5 | 89.73
0 | 0 | 66.9 | 92.47 | 228615.1 | 90.37
0 | 0 | 65.7 | 92.99 | 229906.7 | 91.01
0 | 0 | 64.5 | 93.38 | 230873.5 | 91.29
0 | 0 | 63.3 | 93.72 | 231716.6 | 91.52
0 | 0 | 62.1 | 94.17 | 232830 | 91.95
0 | 0 | 60.9 | 94.7 | 234124.4 | 92.46
0 | 0 | 59.7 | 94.88 | 234581.7 | 92.85
0 | 0 | 58.5 | 95.4 | 235855.4 | 93.28
0 | 0 | 57.3 | 96.07 | 237512.2 | 93.54
0 | 0 | 56.1 | 96.47 | 238510.3 | 94.11
0 | 0 | 54.9 | 96.46 | 238493.9 | 94.7
0 | 0 | 53.6 | 97.28 | 240511.1 | 94.85
0 | 0 | 52.4 | 97.52 | 241096.7 | 95.35
0 | 0 | 51.2 | 97.99 | 242256.7 | 95.91
0 | 0 | 50 | 98.06 | 242449.4 | 96.14
0 | 0 | 48.8 | 98.69 | 243986.8 | 96.36
0 | 0 | 47.6 | 99.07 | 244945.8 | 96.78
0 | 0 | 46.4 | 99.32 | 245552.9 | 96.96
0 | 0 | 45.2 | 99.71 | 246528.1 | 97.43
0 | 0 | 44 | 100.15 | 247596.8 | 97.82
0 | 0 | 42.8 | 100.36 | 248131.4 | 98.12
0 | 0 | 41.6 | 100.88 | 249405.3 | 98.3
0 | 0 | 40.4 | 101.24 | 250306.6 | 98.7
0 | 0 | 39.2 | 100.95 | 249574.9 | 99.27
0 | 0 | 38 | 101.5 | 250941.2 | 99.27
0 | 0 | 36.8 | 102.1 | 252438.1 | 99.52
0 | 0 | 35.5 | 101.93 | 252018.7 | 99.72
0 | 0 | 34.3 | 102.32 | 252970.9 | 99.96
0 | 0 | 33.1 | 102.06 | 252335.1 | 100
0 | 0 | 31.9 | 102.23 | 252756.5 | 99.91
0 | 0 | 30.7 | 102.3 | 252930.6 | 99.7
0 | 0 | 29.5 | 102.27 | 252848.2 | 99.79
0 | 0 | 28.3 | 101.94 | 252031.8 | 99.74
0 | 0 | 27.1 | 101.62 | 251250.9 | 99.46
0 | 0 | 25.9 | 101.66 | 251350.4 | 99.28
0 | 0 | 24.7 | 100.87 | 249397.5 | 98.66
0 | 0 | 23.5 | 100.65 | 248854.2 | 98.45
0 | 0 | 22.3 | 99.78 | 246694.8 | 97.51
0 | 0 | 21 | 98.69 | 243992 | 96.83
0 | 0 | 19.8 | 98.07 | 242464.2 | 95.77
0 | 0 | 18.7 | 97 | 239813.6 | 94.61
0 | 0 | 17.4 | 95.7 | 236611 | 93.43
0 | 0 | 16.3 | 94 | 232402.7 | 91.95
0 | 0 | 15 | 92.08 | 227646.9 | 90
0 | 0 | 13.8 | 90.06 | 222648.8 | 87.87
0 | 0 | 12.6 | 87.56 | 216481.9 | 85.42
0 | 0 | 11.4 | 84.36 | 208571.4 | 82.84
0 | 0 | 10.2 | 81.48 | 201459.3 | 79.5
0 | 0 | 9 | 77.7 | 192105.4 | 75.86
0 | 0 | 7.8 | 72.85 | 180115.4 | 71.02
0 | 0 | 6.6 | 67.39 | 166609.8 | 65.82
0 | 0 | 5.4 | 60.8 | 150329.2 | 59.23
0 | 0 | 4.2 | 53.13 | 131355.7 | 51.76
0 | 0 | 3 | 44.31 | 109548.9 | 43.46
0 | 0 | 1.8 | 37.36 | 92377 | 36.44
0 | 0 | 0.6 | 33.46 | 82728.8 | 32.77
0 | 0 | -0.5 | 31.45 | 77754.3 | 30.64