## 2_Profils énergie
Measurement time: | 12/08/2022 18:00:16 | [5 empty] | Measurement time: | 30/08/2022 16:29:49
Radiation device: | Clinac2 Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV | [5 empty] | Energy: | 6 MV
Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F | [5 empty] | Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie | [5 empty] | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm
SSD: | 1000 mm | [5 empty] | SSD: | 1000 mm
Field size: | 100 x 100 mm | [5 empty] | Field size: | 100 x 100 mm
Gantry angle: | 0 | [5 empty] | Gantry angle: | 0
Measurement medium: | Water | [5 empty] | Measurement medium: | Water
Scan type: | Beam | [5 empty] | Scan type: | Inline
Scan mode: | Continuous | [5 empty] | Scan mode: | Continuous
Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 19.24 | 47561.7 | 20.48 |  | -155 | 0 | 100 | 97.82 | 3347 | 1.98 | 1.9954
0 | 0 | 309.1 | 20.74 | 51265.8 | 20.6 |  | -154.1 | 0 | 100 | 104.92 | 3590.1 | 2.14 | 2.1402
0 | 0 | 307.6 | 20.83 | 51504.6 | 20.67 |  | -152.6 | 0 | 100 | 107.93 | 3692.8 | 2.21 | 2.2016
0 | 0 | 306.6 | 21.02 | 51974.8 | 20.89 |  | -151.5 | 0 | 100 | 107.71 | 3685.5 | 2.2 | 2.1971
0 | 0 | 305.3 | 21.15 | 52292.6 | 21.01 |  | -150.4 | 0 | 100 | 110.71 | 3787.9 | 2.26 | 2.2583
0 | 0 | 304.1 | 21.35 | 52782.9 | 21.19 |  | -149.1 | 0 | 100 | 113.19 | 3872.8 | 2.31 | 2.3089
0 | 0 | 302.9 | 21.46 | 53066 | 21.33 |  | -147.9 | 0 | 100 | 112.91 | 3863.2 | 2.31 | 2.3032
0 | 0 | 301.7 | 21.62 | 53451.6 | 21.46 |  | -146.7 | 0 | 100 | 115.21 | 3941.9 | 2.35 | 2.3501
0 | 0 | 300.5 | 21.73 | 53713.5 | 21.55 |  | -145.5 | 0 | 100 | 118.72 | 4062 | 2.43 | 2.4217
0 | 0 | 299.3 | 21.83 | 53983.6 | 21.68 |  | -144.3 | 0 | 100 | 121.1 | 4143.7 | 2.47 | 2.4703
0 | 0 | 298.1 | 22.08 | 54597.4 | 21.97 |  | -143.1 | 0 | 100 | 121.05 | 4142 | 2.47 | 2.4693
0 | 0 | 296.9 | 22.18 | 54836.5 | 22.02 |  | -141.9 | 0 | 100 | 121.73 | 4164.9 | 2.49 | 2.4831
0 | 0 | 295.7 | 22.34 | 55238 | 22.21 |  | -140.7 | 0 | 100 | 124.22 | 4250.4 | 2.53 | 2.5339
0 | 0 | 294.5 | 22.52 | 55683.7 | 22.35 |  | -139.5 | 0 | 100 | 125.3 | 4287.4 | 2.56 | 2.5559
0 | 0 | 293.3 | 22.64 | 55983.6 | 22.46 |  | -138.3 | 0 | 100 | 125.42 | 4291.3 | 2.56 | 2.5584
0 | 0 | 292.1 | 22.78 | 56332 | 22.67 |  | -137.1 | 0 | 100 | 130.8 | 4475.5 | 2.68 | 2.6681
0 | 0 | 290.9 | 22.95 | 56744.8 | 22.82 |  | -135.9 | 0 | 100 | 131.48 | 4498.7 | 2.68 | 2.682
0 | 0 | 289.7 | 23.07 | 57049 | 22.94 |  | -134.7 | 0 | 100 | 132.49 | 4533.1 | 2.7 | 2.7026
0 | 0 | 288.5 | 23.29 | 57574 | 23.13 |  | -133.5 | 0 | 100 | 135.51 | 4636.7 | 2.76 | 2.7642
0 | 0 | 287.3 | 23.45 | 57968 | 23.31 |  | -132.3 | 0 | 100 | 136.41 | 4667.3 | 2.78 | 2.7826
0 | 0 | 286.1 | 23.53 | 58174.1 | 23.42 |  | -131.1 | 0 | 100 | 141.39 | 4837.7 | 2.88 | 2.8842
0 | 0 | 284.9 | 23.7 | 58600 | 23.54 |  | -129.9 | 0 | 100 | 142.63 | 4880.2 | 2.91 | 2.9095
0 | 0 | 283.7 | 23.93 | 59151.6 | 23.76 |  | -128.7 | 0 | 100 | 145.56 | 4980.5 | 2.97 | 2.9692
0 | 0 | 282.5 | 24.01 | 59365.5 | 23.88 |  | -127.5 | 0 | 100 | 151.43 | 5181.4 | 3.09 | 3.089
0 | 0 | 281.2 | 24.25 | 59960.5 | 24.07 |  | -126.3 | 0 | 100 | 152.81 | 5228.6 | 3.12 | 3.1171
0 | 0 | 280.1 | 24.39 | 60289 | 24.23 |  | -125.1 | 0 | 100 | 152.26 | 5209.8 | 3.11 | 3.1059
0 | 0 | 278.8 | 24.46 | 60477.1 | 24.31 |  | -123.9 | 0 | 100 | 155.96 | 5336.3 | 3.18 | 3.1814
0 | 0 | 277.7 | 24.68 | 61005.5 | 24.5 |  | -122.7 | 0 | 100 | 161.76 | 5534.6 | 3.3 | 3.2997
0 | 0 | 276.4 | 24.86 | 61469.1 | 24.72 |  | -121.5 | 0 | 100 | 161.02 | 5509.4 | 3.29 | 3.2846
0 | 0 | 275.2 | 25.02 | 61854 | 24.83 |  | -120.3 | 0 | 100 | 166.78 | 5706.5 | 3.4 | 3.4021
0 | 0 | 274 | 25.22 | 62360 | 25.07 |  | -119.1 | 0 | 100 | 166.77 | 5706 | 3.39 | 3.4019
0 | 0 | 272.8 | 25.32 | 62610.7 | 25.19 |  | -117.8 | 0 | 100 | 170.75 | 5842.3 | 3.48 | 3.4831
0 | 0 | 271.6 | 25.54 | 63143.9 | 25.37 |  | -116.7 | 0 | 100 | 175.5 | 6004.7 | 3.58 | 3.58
0 | 0 | 270.4 | 25.66 | 63445.9 | 25.55 |  | -115.4 | 0 | 100 | 178.38 | 6103.3 | 3.64 | 3.6387
0 | 0 | 269.2 | 25.84 | 63885.9 | 25.73 |  | -114.2 | 0 | 100 | 185.21 | 6337.2 | 3.78 | 3.778
0 | 0 | 268 | 26.05 | 64400.7 | 25.89 |  | -113 | 0 | 100 | 189 | 6466.7 | 3.85 | 3.8553
0 | 0 | 266.8 | 26.3 | 65024.5 | 26.08 |  | -111.8 | 0 | 100 | 189.41 | 6480.9 | 3.86 | 3.8637
0 | 0 | 265.6 | 26.42 | 65327.7 | 26.21 |  | -110.6 | 0 | 100 | 195.32 | 6683 | 3.98 | 3.9843
0 | 0 | 264.4 | 26.53 | 65584.6 | 26.36 |  | -109.4 | 0 | 100 | 199.33 | 6820.1 | 4.06 | 4.0661
0 | 0 | 263.2 | 26.75 | 66146.4 | 26.61 |  | -108.2 | 0 | 100 | 203.52 | 6963.5 | 4.15 | 4.1515
0 | 0 | 262 | 26.97 | 66670.8 | 26.79 |  | -107 | 0 | 100 | 211.38 | 7232.4 | 4.31 | 4.3119
0 | 0 | 260.8 | 27.05 | 66881.8 | 26.88 |  | -105.8 | 0 | 100 | 215.4 | 7370.1 | 4.39 | 4.3939
0 | 0 | 259.6 | 27.34 | 67602.7 | 27.15 |  | -104.6 | 0 | 100 | 218.59 | 7479.2 | 4.46 | 4.4589
0 | 0 | 258.4 | 27.47 | 67924.8 | 27.26 |  | -103.4 | 0 | 100 | 225.64 | 7720.3 | 4.6 | 4.6027
0 | 0 | 257.2 | 27.65 | 68368.3 | 27.45 |  | -102.2 | 0 | 100 | 232.07 | 7940.5 | 4.74 | 4.7339
0 | 0 | 256 | 27.85 | 68849.4 | 27.69 |  | -101 | 0 | 100 | 236.14 | 8079.6 | 4.81 | 4.8169
0 | 0 | 254.8 | 28.04 | 69317.7 | 27.86 |  | -99.8 | 0 | 100 | 241.8 | 8273.3 | 4.92 | 4.9324
0 | 0 | 253.6 | 28.19 | 69703.5 | 28.04 |  | -98.6 | 0 | 100 | 246.67 | 8439.9 | 5.03 | 5.0317
0 | 0 | 252.4 | 28.45 | 70348.8 | 28.25 |  | -97.4 | 0 | 100 | 252.61 | 8643.3 | 5.15 | 5.1529
0 | 0 | 251.2 | 28.64 | 70815.2 | 28.46 |  | -96.2 | 0 | 100 | 258.81 | 8855.3 | 5.28 | 5.2794
0 | 0 | 250 | 28.82 | 71244.8 | 28.62 |  | -95 | 0 | 100 | 264.84 | 9061.8 | 5.4 | 5.4024
0 | 0 | 248.8 | 29.06 | 71852.1 | 28.86 |  | -93.8 | 0 | 100 | 273.5 | 9358.1 | 5.58 | 5.579
0 | 0 | 247.6 | 29.19 | 72175.6 | 29.03 |  | -92.6 | 0 | 100 | 280.57 | 9599.8 | 5.72 | 5.7232
0 | 0 | 246.4 | 29.43 | 72757.2 | 29.3 |  | -91.3 | 0 | 100 | 291.65 | 9979 | 5.95 | 5.9493
0 | 0 | 245.1 | 29.65 | 73312.1 | 29.45 |  | -90.2 | 0 | 100 | 299.22 | 10238.1 | 6.11 | 6.1037
0 | 0 | 244 | 29.9 | 73917.3 | 29.71 |  | -89 | 0 | 100 | 304.87 | 10431.5 | 6.21 | 6.2189
0 | 0 | 242.7 | 30.03 | 74247.9 | 29.82 |  | -87.7 | 0 | 100 | 311.46 | 10656.7 | 6.35 | 6.3534
0 | 0 | 241.6 | 30.24 | 74771.8 | 30.06 |  | -86.6 | 0 | 100 | 325.18 | 11126.2 | 6.62 | 6.6332
0 | 0 | 240.3 | 30.32 | 74956.5 | 30.08 |  | -85.3 | 0 | 100 | 333.67 | 11416.6 | 6.79 | 6.8064
0 | 0 | 239.1 | 30.57 | 75581.7 | 30.35 |  | -84.1 | 0 | 100 | 344.21 | 11777.3 | 7.01 | 7.0214
0 | 0 | 237.9 | 30.83 | 76219.8 | 30.63 |  | -82.9 | 0 | 100 | 356.11 | 12184.6 | 7.26 | 7.2642
0 | 0 | 236.7 | 31.01 | 76674 | 30.79 |  | -81.7 | 0 | 100 | 364.08 | 12457.1 | 7.42 | 7.4267
0 | 0 | 235.5 | 31.17 | 77051.9 | 30.99 |  | -80.5 | 0 | 100 | 375.78 | 12857.6 | 7.65 | 7.6654
0 | 0 | 234.3 | 31.43 | 77707.8 | 31.23 |  | -79.3 | 0 | 100 | 388.87 | 13305.5 | 7.92 | 7.9324
0 | 0 | 233.1 | 31.52 | 77936.1 | 31.35 |  | -78.1 | 0 | 100 | 401.24 | 13728.8 | 8.18 | 8.1847
0 | 0 | 231.9 | 31.82 | 78671.4 | 31.59 |  | -76.9 | 0 | 100 | 413.18 | 14137.2 | 8.42 | 8.4283
0 | 0 | 230.7 | 32.12 | 79421.2 | 31.9 |  | -75.7 | 0 | 100 | 426.2 | 14582.8 | 8.68 | 8.6939
0 | 0 | 229.5 | 32.33 | 79929.1 | 32.1 |  | -74.5 | 0 | 100 | 440.18 | 15061 | 8.97 | 8.9791
0 | 0 | 228.3 | 32.5 | 80362.2 | 32.32 |  | -73.3 | 0 | 100 | 453.4 | 15513.5 | 9.26 | 9.2487
0 | 0 | 227.1 | 32.75 | 80967.7 | 32.54 |  | -72.1 | 0 | 100 | 472.25 | 16158.4 | 9.63 | 9.6333
0 | 0 | 225.9 | 32.97 | 81509.4 | 32.75 |  | -70.9 | 0 | 100 | 486.25 | 16637.3 | 9.91 | 9.9188
0 | 0 | 224.7 | 33.21 | 82105.1 | 32.97 |  | -69.7 | 0 | 100 | 502.98 | 17209.7 | 10.26 | 10.2601
0 | 0 | 223.5 | 33.45 | 82702.7 | 33.24 |  | -68.5 | 0 | 100 | 525.28 | 17972.9 | 10.71 | 10.715
0 | 0 | 222.2 | 33.64 | 83165.5 | 33.38 |  | -67.3 | 0 | 100 | 546.75 | 18707.3 | 11.13 | 11.153
0 | 0 | 221 | 33.91 | 83846.6 | 33.69 |  | -66.1 | 0 | 100 | 576.98 | 19741.9 | 11.77 | 11.7696
0 | 0 | 219.9 | 34.1 | 84313.7 | 33.85 |  | -64.9 | 0 | 100 | 610.91 | 20902.7 | 12.45 | 12.4617
0 | 0 | 218.6 | 34.31 | 84814.5 | 34.1 |  | -63.7 | 0 | 100 | 657.46 | 22495.5 | 13.4 | 13.4113
0 | 0 | 217.5 | 34.65 | 85665 | 34.4 |  | -62.5 | 0 | 100 | 730.37 | 24990.2 | 14.86 | 14.8985
0 | 0 | 216.2 | 34.78 | 85995 | 34.58 |  | -61.3 | 0 | 100 | 844.53 | 28896.1 | 17.21 | 17.2273
0 | 0 | 215 | 35.02 | 86580.8 | 34.76 |  | -60.1 | 0 | 100 | 1036.64 | 35469.3 | 21.12 | 21.146
0 | 0 | 213.8 | 35.24 | 87132.5 | 34.98 |  | -58.9 | 0 | 100 | 1328.49 | 45455.1 | 27.05 | 27.0994
0 | 0 | 212.6 | 35.45 | 87656.8 | 35.26 |  | -57.6 | 0 | 100 | 1733.18 | 59301.8 | 35.27 | 35.3545
0 | 0 | 211.4 | 35.79 | 88497.1 | 35.59 |  | -56.4 | 0 | 100 | 2226.35 | 76176.1 | 45.43 | 45.4145
0 | 0 | 210.2 | 35.96 | 88911.9 | 35.76 |  | -55.2 | 0 | 100 | 2807.32 | 96054.1 | 57.29 | 57.2655
0 | 0 | 209 | 36.22 | 89556.4 | 35.98 |  | -54 | 0 | 100 | 3371.3 | 115351.3 | 68.71 | 68.7699
0 | 0 | 207.8 | 36.44 | 90090.1 | 36.21 |  | -52.8 | 0 | 100 | 3852.93 | 131830.5 | 78.55 | 78.5945
0 | 0 | 206.6 | 36.83 | 91054.7 | 36.49 |  | -51.6 | 0 | 100 | 4207.02 | 143945.7 | 85.82 | 85.8174
0 | 0 | 205.4 | 37 | 91482 | 36.81 |  | -50.4 | 0 | 100 | 4398.35 | 150492.5 | 89.65 | 89.7203
0 | 0 | 204.2 | 37.19 | 91958.6 | 36.95 |  | -49.2 | 0 | 100 | 4505.77 | 154167.7 | 91.81 | 91.9115
0 | 0 | 203 | 37.47 | 92627.4 | 37.21 |  | -48 | 0 | 100 | 4584.95 | 156877.1 | 93.34 | 93.5267
0 | 0 | 201.8 | 37.77 | 93378.5 | 37.5 |  | -46.8 | 0 | 100 | 4638.44 | 158707.2 | 94.36 | 94.6178
0 | 0 | 200.6 | 38.03 | 94030.3 | 37.81 |  | -45.6 | 0 | 100 | 4653.46 | 159220.9 | 94.94 | 94.9242
0 | 0 | 199.4 | 38.27 | 94618.1 | 38.03 |  | -44.4 | 0 | 100 | 4681.66 | 160186 | 95.43 | 95.4995
0 | 0 | 198.2 | 38.55 | 95320.8 | 38.28 |  | -43.2 | 0 | 100 | 4707.85 | 161082.2 | 95.95 | 96.0337
0 | 0 | 197 | 38.79 | 95913.9 | 38.59 |  | -42 | 0 | 100 | 4717.17 | 161401.1 | 96.16 | 96.2238
0 | 0 | 195.8 | 39.1 | 96666.5 | 38.8 |  | -40.8 | 0 | 100 | 4728.38 | 161784.4 | 96.33 | 96.4525
0 | 0 | 194.6 | 39.25 | 97033 | 38.94 |  | -39.6 | 0 | 100 | 4750.88 | 162554.4 | 96.79 | 96.9114
0 | 0 | 193.4 | 39.64 | 98008.2 | 39.35 |  | -38.4 | 0 | 100 | 4767 | 163105.9 | 97.2 | 97.2403
0 | 0 | 192.2 | 39.93 | 98723.7 | 39.68 |  | -37.2 | 0 | 100 | 4775.2 | 163386.5 | 97.29 | 97.4075
0 | 0 | 191 | 40.15 | 99276.6 | 39.92 |  | -36 | 0 | 100 | 4781.4 | 163598.6 | 97.47 | 97.534
0 | 0 | 189.8 | 40.52 | 100176.2 | 40.22 |  | -34.8 | 0 | 100 | 4798.42 | 164180.9 | 97.85 | 97.8812
0 | 0 | 188.5 | 40.73 | 100696.1 | 40.48 |  | -33.6 | 0 | 100 | 4809.75 | 164568.8 | 98.01 | 98.1123
0 | 0 | 187.4 | 41.03 | 101429 | 40.76 |  | -32.4 | 0 | 100 | 4813.32 | 164690.6 | 98.01 | 98.1851
0 | 0 | 186.1 | 41.28 | 102046.6 | 40.98 |  | -31.2 | 0 | 100 | 4826.27 | 165133.7 | 98.5 | 98.4493
0 | 0 | 184.9 | 41.5 | 102609.4 | 41.24 |  | -30 | 0 | 100 | 4835.45 | 165448 | 98.56 | 98.6366
0 | 0 | 183.7 | 41.87 | 103520.2 | 41.53 |  | -28.8 | 0 | 100 | 4843.08 | 165709.2 | 98.73 | 98.7922
0 | 0 | 182.5 | 42.14 | 104174.2 | 41.87 |  | -27.6 | 0 | 100 | 4856.81 | 166179 | 98.9 | 99.0723
0 | 0 | 181.3 | 42.38 | 104788.9 | 42.1 |  | -26.4 | 0 | 100 | 4860.91 | 166319.2 | 99.01 | 99.1559
0 | 0 | 180.1 | 42.66 | 105475.2 | 42.38 |  | -25.1 | 0 | 100 | 4864.16 | 166430.5 | 99.24 | 99.2222
0 | 0 | 178.9 | 42.95 | 106175.7 | 42.69 |  | -24 | 0 | 100 | 4871.4 | 166677.9 | 99.1 | 99.3699
0 | 0 | 177.7 | 43.3 | 107043.9 | 43.02 |  | -22.8 | 0 | 100 | 4874.44 | 166782.2 | 99.34 | 99.4319
0 | 0 | 176.5 | 43.56 | 107694.7 | 43.3 |  | -21.6 | 0 | 100 | 4878.3 | 166914.2 | 99.59 | 99.5106
0 | 0 | 175.3 | 43.95 | 108666.8 | 43.64 |  | -20.4 | 0 | 100 | 4893.58 | 167437.1 | 99.79 | 99.8223
0 | 0 | 174.1 | 44.18 | 109239.1 | 43.86 |  | -19.1 | 0 | 100 | 4889.49 | 167297.1 | 99.69 | 99.7389
0 | 0 | 172.9 | 44.42 | 109830.9 | 44.13 |  | -18 | 0 | 100 | 4892.96 | 167415.8 | 99.78 | 99.8097
0 | 0 | 171.7 | 44.82 | 110804.1 | 44.5 |  | -16.7 | 0 | 100 | 4891.16 | 167354 | 99.74 | 99.773
0 | 0 | 170.5 | 45.09 | 111469.7 | 44.85 |  | -15.5 | 0 | 100 | 4891.44 | 167363.9 | 99.71 | 99.7787
0 | 0 | 169.3 | 45.38 | 112194.3 | 45.07 |  | -14.3 | 0 | 100 | 4885.61 | 167164.2 | 99.7 | 99.6598
0 | 0 | 168.1 | 45.79 | 113207.6 | 45.41 |  | -13.1 | 0 | 100 | 4899.41 | 167636.4 | 99.72 | 99.9413
0 | 0 | 166.9 | 46.02 | 113773.1 | 45.71 |  | -11.9 | 0 | 100 | 4892.25 | 167391.3 | 99.67 | 99.7952
0 | 0 | 165.7 | 46.26 | 114381.3 | 45.98 |  | -10.7 | 0 | 100 | 4899.41 | 167636.6 | 100 | 99.9413
0 | 0 | 164.4 | 46.66 | 115368.9 | 46.4 |  | -9.5 | 0 | 100 | 4899.38 | 167635.5 | 99.77 | 99.9406
0 | 0 | 163.3 | 46.97 | 116120.2 | 46.7 |  | -8.3 | 0 | 100 | 4892.85 | 167411.9 | 99.61 | 99.8074
0 | 0 | 162 | 47.22 | 116749.8 | 46.99 |  | -7.1 | 0 | 100 | 4901.85 | 167720.1 | 99.78 | 99.991
0 | 0 | 160.9 | 47.61 | 117715.4 | 47.29 |  | -5.9 | 0 | 100 | 4900.64 | 167678.4 | 99.91 | 99.9663
0 | 0 | 159.6 | 47.99 | 118644.5 | 47.77 |  | -4.7 | 0 | 100 | 4895.56 | 167504.7 | 99.72 | 99.8627
0 | 0 | 158.4 | 48.28 | 119363.4 | 47.94 |  | -3.5 | 0 | 100 | 4892.79 | 167409.8 | 99.74 | 99.8062
0 | 0 | 157.2 | 48.55 | 120029.6 | 48.28 |  | -2.3 | 0 | 100 | 4894.91 | 167482.6 | 99.73 | 99.8495
0 | 0 | 156 | 48.92 | 120939.2 | 48.58 |  | -1.1 | 0 | 100 | 4896.02 | 167520.4 | 99.78 | 99.8721
0 | 0 | 154.8 | 49.25 | 121772.5 | 48.95 |  | 0.1 | 0 | 100 | 4888.64 | 167268 | 99.66 | 99.7216
0 | 0 | 153.6 | 49.62 | 122682.4 | 49.35 |  | 1.3 | 0 | 100 | 4891.42 | 167363.2 | 99.75 | 99.7783
0 | 0 | 152.4 | 50.03 | 123703.9 | 49.6 |  | 2.5 | 0 | 100 | 4891.53 | 167366.7 | 99.9 | 99.7805
0 | 0 | 151.2 | 50.22 | 124168.6 | 49.91 |  | 3.7 | 0 | 100 | 4887.86 | 167241.4 | 99.68 | 99.7056
0 | 0 | 150 | 50.46 | 124759.1 | 50.18 |  | 4.9 | 0 | 100 | 4902.29 | 167734.9 | 99.97 | 100
0 | 0 | 148.8 | 50.87 | 125765.6 | 50.56 |  | 6.1 | 0 | 100 | 4890.69 | 167338.1 | 99.63 | 99.7634
0 | 0 | 147.6 | 51.25 | 126705.6 | 50.86 |  | 7.3 | 0 | 100 | 4898.39 | 167601.5 | 99.95 | 99.9204
0 | 0 | 146.4 | 51.59 | 127539 | 51.15 |  | 8.5 | 0 | 100 | 4888.96 | 167279 | 99.81 | 99.7281
0 | 0 | 145.2 | 51.97 | 128493.7 | 51.62 |  | 9.7 | 0 | 100 | 4890.84 | 167343.1 | 99.78 | 99.7664
0 | 0 | 144 | 52.33 | 129378.5 | 52.01 |  | 10.9 | 0 | 100 | 4896.41 | 167533.7 | 99.77 | 99.8801
0 | 0 | 142.8 | 52.72 | 130336.2 | 52.47 |  | 12.1 | 0 | 100 | 4888.9 | 167276.8 | 99.6 | 99.7269
0 | 0 | 141.6 | 53.01 | 131055.7 | 52.71 |  | 13.3 | 0 | 100 | 4891.49 | 167365.5 | 99.76 | 99.7797
0 | 0 | 140.4 | 53.4 | 132017.8 | 52.97 |  | 14.5 | 0 | 100 | 4892.41 | 167396.9 | 99.8 | 99.7985
0 | 0 | 139.2 | 53.66 | 132656.7 | 53.35 |  | 15.7 | 0 | 100 | 4885.76 | 167169.6 | 99.75 | 99.6628
0 | 0 | 138 | 54.05 | 133641 | 53.65 |  | 16.9 | 0 | 100 | 4888.39 | 167259.3 | 99.71 | 99.7165
0 | 0 | 136.8 | 54.43 | 134578.9 | 54.14 |  | 18.1 | 0 | 100 | 4885.12 | 167147.4 | 99.59 | 99.6498
0 | 0 | 135.6 | 54.75 | 135369.7 | 54.45 |  | 19.3 | 0 | 100 | 4888.34 | 167257.6 | 99.54 | 99.7154
0 | 0 | 134.4 | 55.14 | 136318.4 | 54.76 |  | 20.5 | 0 | 100 | 4882.39 | 167054.2 | 99.43 | 99.5941
0 | 0 | 133.2 | 55.49 | 137188.5 | 55.1 |  | 21.7 | 0 | 100 | 4878.13 | 166908.3 | 99.49 | 99.5072
0 | 0 | 131.9 | 55.85 | 138072 | 55.48 |  | 22.9 | 0 | 100 | 4875.19 | 166807.8 | 99.52 | 99.4472
0 | 0 | 130.8 | 56.25 | 139075.6 | 55.89 |  | 24.1 | 0 | 100 | 4874.35 | 166779.1 | 99.38 | 99.4301
0 | 0 | 129.6 | 56.71 | 140213 | 56.4 |  | 25.4 | 0 | 100 | 4857.2 | 166192.3 | 99.04 | 99.0802
0 | 0 | 128.3 | 57.02 | 140961.3 | 56.57 |  | 26.6 | 0 | 100 | 4854.68 | 166105.8 | 99.05 | 99.0288
0 | 0 | 127.1 | 57.39 | 141899.8 | 56.96 |  | 27.7 | 0 | 100 | 4843.69 | 165730.1 | 98.67 | 98.8046
0 | 0 | 126 | 57.87 | 143076.8 | 57.42 |  | 29 | 0 | 100 | 4844.71 | 165764.9 | 98.92 | 98.8254
0 | 0 | 124.7 | 58.23 | 143964.5 | 57.84 |  | 30.2 | 0 | 100 | 4830.07 | 165264.1 | 98.36 | 98.5268
0 | 0 | 123.5 | 58.57 | 144802.3 | 58.19 |  | 31.4 | 0 | 100 | 4822.46 | 165003.7 | 98.22 | 98.3716
0 | 0 | 122.3 | 59.13 | 146193.7 | 58.73 |  | 32.6 | 0 | 100 | 4813.12 | 164684.1 | 98.18 | 98.1811
0 | 0 | 121.1 | 59.48 | 147066 | 59.09 |  | 33.7 | 0 | 100 | 4800.17 | 164240.9 | 97.72 | 97.9169
0 | 0 | 119.9 | 59.85 | 147964.3 | 59.48 |  | 35 | 0 | 100 | 4790.88 | 163923.1 | 97.79 | 97.7274
0 | 0 | 118.7 | 60.24 | 148934.6 | 59.89 |  | 36.2 | 0 | 100 | 4777.32 | 163459.1 | 97.34 | 97.4508
0 | 0 | 117.5 | 60.6 | 149836.2 | 60.29 |  | 37.4 | 0 | 100 | 4762.1 | 162938.3 | 97.03 | 97.1403
0 | 0 | 116.3 | 60.88 | 150522.4 | 60.55 |  | 38.6 | 0 | 100 | 4755.28 | 162704.8 | 96.88 | 97.0012
0 | 0 | 115.1 | 61.27 | 151475.4 | 60.89 |  | 39.8 | 0 | 100 | 4739.59 | 162167.9 | 96.49 | 96.6811
0 | 0 | 113.9 | 61.79 | 152758.1 | 61.33 |  | 41 | 0 | 100 | 4723.54 | 161618.9 | 96.31 | 96.3537
0 | 0 | 112.7 | 62.09 | 153506.7 | 61.68 |  | 42.2 | 0 | 100 | 4711.5 | 161206.8 | 96.07 | 96.1081
0 | 0 | 111.5 | 62.54 | 154614.8 | 62.13 |  | 43.4 | 0 | 100 | 4682.66 | 160220.1 | 95.36 | 95.5198
0 | 0 | 110.3 | 63.01 | 155792.1 | 62.56 |  | 44.6 | 0 | 100 | 4670.4 | 159800.8 | 95.27 | 95.2698
0 | 0 | 109.1 | 63.51 | 157019.6 | 63 |  | 45.8 | 0 | 100 | 4635.93 | 158621.3 | 94.52 | 94.5666
0 | 0 | 107.9 | 64.02 | 158289.7 | 63.56 |  | 47 | 0 | 100 | 4587.42 | 156961.4 | 93.53 | 93.5771
0 | 0 | 106.7 | 64.51 | 159496.6 | 63.84 |  | 48.2 | 0 | 100 | 4517.63 | 154573.7 | 92.06 | 92.1535
0 | 0 | 105.5 | 64.81 | 160233.8 | 64.36 |  | 49.4 | 0 | 100 | 4411.79 | 150952.1 | 90.01 | 89.9945
0 | 0 | 104.2 | 65.18 | 161158.2 | 64.88 |  | 50.6 | 0 | 100 | 4216.55 | 144271.9 | 86.09 | 86.0118
0 | 0 | 103.1 | 65.66 | 162345.9 | 65.19 |  | 51.8 | 0 | 100 | 3930.02 | 134468.1 | 80.06 | 80.167
0 | 0 | 101.9 | 66.12 | 163470.3 | 65.63 |  | 53 | 0 | 100 | 3538.58 | 121074.9 | 72.14 | 72.1822
0 | 0 | 100.7 | 66.56 | 164553.7 | 66.23 |  | 54.2 | 0 | 100 | 3036.19 | 103885 | 61.84 | 61.9341
0 | 0 | 99.5 | 66.8 | 165151.1 | 66.43 |  | 55.4 | 0 | 100 | 2468.48 | 84460.5 | 50.42 | 50.3536
0 | 0 | 98.3 | 67.23 | 166216.7 | 66.84 |  | 56.6 | 0 | 100 | 1899.72 | 65000 | 38.78 | 38.7517
0 | 0 | 97.1 | 67.74 | 167469.8 | 67.24 |  | 57.8 | 0 | 100 | 1423.45 | 48704.2 | 29.05 | 29.0364
0 | 0 | 95.8 | 68.06 | 168266.5 | 67.73 |  | 59.1 | 0 | 100 | 1067.12 | 36512.1 | 21.74 | 21.7678
0 | 0 | 94.7 | 68.39 | 169076.6 | 68.15 |  | 60.3 | 0 | 100 | 864.71 | 29586.5 | 17.63 | 17.6389
0 | 0 | 93.5 | 68.97 | 170521.4 | 68.6 |  | 61.5 | 0 | 100 | 744.18 | 25462.5 | 15.17 | 15.1803
0 | 0 | 92.2 | 69.61 | 172091.4 | 68.98 |  | 62.7 | 0 | 100 | 670.48 | 22940.8 | 13.69 | 13.6769
0 | 0 | 91.1 | 69.94 | 172923.4 | 69.54 |  | 63.9 | 0 | 100 | 624.53 | 21368.7 | 12.72 | 12.7396
0 | 0 | 89.8 | 70.5 | 174294.8 | 70.01 |  | 65.1 | 0 | 100 | 582.49 | 19930.4 | 11.89 | 11.882
0 | 0 | 88.6 | 71.05 | 175650.2 | 70.44 |  | 66.3 | 0 | 100 | 548.99 | 18784 | 11.2 | 11.1986
0 | 0 | 87.4 | 71.54 | 176880.2 | 71.06 |  | 67.5 | 0 | 100 | 532.69 | 18226.4 | 10.87 | 10.8661
0 | 0 | 86.2 | 71.94 | 177867.6 | 71.55 |  | 68.7 | 0 | 100 | 512.93 | 17550.2 | 10.46 | 10.4631
0 | 0 | 85 | 72.45 | 179111.9 | 71.91 |  | 69.9 | 0 | 100 | 491.43 | 16814.6 | 10.02 | 10.0245
0 | 0 | 83.8 | 72.91 | 180252.2 | 72.35 |  | 71.1 | 0 | 100 | 471.34 | 16127.2 | 9.61 | 9.6147
0 | 0 | 82.6 | 73.37 | 181404 | 72.9 |  | 72.3 | 0 | 100 | 454.71 | 15558.3 | 9.26 | 9.2755
0 | 0 | 81.4 | 73.82 | 182515.6 | 73.36 |  | 73.5 | 0 | 100 | 445.31 | 15236.5 | 9.07 | 9.0837
0 | 0 | 80.2 | 74.32 | 183742 | 73.79 |  | 74.7 | 0 | 100 | 426.83 | 14604.4 | 8.71 | 8.7067
0 | 0 | 79 | 74.68 | 184629.3 | 74.18 |  | 75.9 | 0 | 100 | 416.8 | 14261 | 8.5 | 8.5021
0 | 0 | 77.8 | 75.26 | 186064.7 | 74.7 |  | 77.1 | 0 | 100 | 405.53 | 13875.5 | 8.27 | 8.2723
0 | 0 | 76.6 | 75.69 | 187144.5 | 75.09 |  | 78.3 | 0 | 100 | 393.82 | 13474.9 | 8.02 | 8.0334
0 | 0 | 75.4 | 75.94 | 187756.6 | 75.41 |  | 79.5 | 0 | 100 | 380.79 | 13028.9 | 7.77 | 7.7676
0 | 0 | 74.1 | 76.42 | 188938 | 75.96 |  | 80.7 | 0 | 100 | 372.67 | 12751.3 | 7.61 | 7.602
0 | 0 | 73 | 76.9 | 190125.5 | 76.55 |  | 81.9 | 0 | 100 | 360.34 | 12329.2 | 7.36 | 7.3504
0 | 0 | 71.7 | 77.54 | 191710 | 77.05 |  | 83.1 | 0 | 100 | 350.11 | 11979.3 | 7.12 | 7.1418
0 | 0 | 70.5 | 78.1 | 193090.9 | 77.55 |  | 84.3 | 0 | 100 | 337.67 | 11553.6 | 6.88 | 6.888
0 | 0 | 69.3 | 78.72 | 194624 | 77.98 |  | 85.5 | 0 | 100 | 326.95 | 11186.7 | 6.67 | 6.6693
0 | 0 | 68.1 | 79.04 | 195410 | 78.39 |  | 86.7 | 0 | 100 | 317.76 | 10872.5 | 6.49 | 6.4819
0 | 0 | 66.9 | 79.59 | 196786.3 | 79.02 |  | 87.9 | 0 | 100 | 310.09 | 10610 | 6.32 | 6.3254
0 | 0 | 65.7 | 80.07 | 197953.7 | 79.48 |  | 89.1 | 0 | 100 | 303.02 | 10368 | 6.17 | 6.1812
0 | 0 | 64.5 | 80.58 | 199215.5 | 79.89 |  | 90.3 | 0 | 100 | 292.8 | 10018.4 | 5.97 | 5.9727
0 | 0 | 63.3 | 81.05 | 200396.4 | 80.58 |  | 91.5 | 0 | 100 | 288.14 | 9858.8 | 5.89 | 5.8777
0 | 0 | 62 | 81.7 | 201999.9 | 81.11 |  | 92.7 | 0 | 100 | 279.53 | 9564.4 | 5.7 | 5.702
0 | 0 | 60.9 | 82.02 | 202779.4 | 81.45 |  | 93.9 | 0 | 100 | 272.38 | 9319.8 | 5.55 | 5.5562
0 | 0 | 59.6 | 82.47 | 203895.1 | 82.01 |  | 95.2 | 0 | 100 | 264.57 | 9052.4 | 5.4 | 5.3969
0 | 0 | 58.5 | 83.17 | 205630.1 | 82.51 |  | 96.4 | 0 | 100 | 258.69 | 8851.4 | 5.28 | 5.2769
0 | 0 | 57.2 | 83.78 | 207122.8 | 83.14 |  | 97.6 | 0 | 100 | 250.42 | 8568.5 | 5.11 | 5.1082
0 | 0 | 56 | 84.05 | 207801.8 | 83.42 |  | 98.8 | 0 | 100 | 248.9 | 8516.3 | 5.07 | 5.0772
0 | 0 | 54.8 | 84.53 | 208993.7 | 83.93 |  | 100 | 0 | 100 | 241.16 | 8251.4 | 4.92 | 4.9193
0 | 0 | 53.6 | 84.95 | 210038.5 | 84.43 |  | 101.2 | 0 | 100 | 241.53 | 8264.1 | 4.93 | 4.9269
0 | 0 | 52.4 | 85.63 | 211702.6 | 85.05 |  | 102.4 | 0 | 100 | 232.7 | 7961.9 | 4.74 | 4.7468
0 | 0 | 51.2 | 86.21 | 213142.3 | 85.4 |  | 103.6 | 0 | 100 | 222.63 | 7617.6 | 4.55 | 4.5413
0 | 0 | 50 | 86.73 | 214435.1 | 86.13 |  | 104.8 | 0 | 100 | 222.61 | 7616.7 | 4.54 | 4.5409
0 | 0 | 48.8 | 87.06 | 215249.2 | 86.55 |  | 106 | 0 | 100 | 214.57 | 7341.8 | 4.38 | 4.3769
0 | 0 | 47.6 | 87.57 | 216492.8 | 86.98 |  | 107.2 | 0 | 100 | 212.39 | 7267.1 | 4.33 | 4.3325
0 | 0 | 46.4 | 88.17 | 217993.2 | 87.38 |  | 108.4 | 0 | 100 | 205.9 | 7045 | 4.2 | 4.2001
0 | 0 | 45.2 | 88.51 | 218825 | 87.97 |  | 109.6 | 0 | 100 | 205.64 | 7036.2 | 4.19 | 4.1948
0 | 0 | 44 | 89.08 | 220237.2 | 88.49 |  | 110.8 | 0 | 100 | 201.4 | 6890.9 | 4.11 | 4.1083
0 | 0 | 42.8 | 89.42 | 221077.5 | 88.88 |  | 112 | 0 | 100 | 194 | 6637.8 | 3.96 | 3.9573
0 | 0 | 41.6 | 89.98 | 222465.5 | 89.46 |  | 113.2 | 0 | 100 | 190.77 | 6527.3 | 3.89 | 3.8914
0 | 0 | 40.4 | 90.66 | 224139.6 | 89.96 |  | 114.4 | 0 | 100 | 187.99 | 6432.2 | 3.83 | 3.8347
0 | 0 | 39.1 | 91.34 | 225820.5 | 90.54 |  | 115.6 | 0 | 100 | 183.54 | 6279.8 | 3.74 | 3.744
0 | 0 | 38 | 91.63 | 226542.2 | 90.99 |  | 116.8 | 0 | 100 | 180.95 | 6191.3 | 3.7 | 3.6911
0 | 0 | 36.7 | 92.09 | 227684.5 | 91.53 |  | 118 | 0 | 100 | 177.82 | 6084.4 | 3.62 | 3.6273
0 | 0 | 35.5 | 92.62 | 228982 | 91.93 |  | 119.3 | 0 | 100 | 170.82 | 5844.8 | 3.48 | 3.4845
0 | 0 | 34.3 | 93.2 | 230429.1 | 92.41 |  | 120.4 | 0 | 100 | 169.29 | 5792.2 | 3.45 | 3.4533
0 | 0 | 33.1 | 93.86 | 232066.6 | 93 |  | 121.7 | 0 | 100 | 163.48 | 5593.6 | 3.33 | 3.3348
0 | 0 | 31.9 | 94.45 | 233507.6 | 93.77 |  | 122.8 | 0 | 100 | 162.66 | 5565.4 | 3.32 | 3.318
0 | 0 | 30.7 | 94.91 | 234651 | 94.29 |  | 124.1 | 0 | 100 | 158.79 | 5433.1 | 3.24 | 3.2391
0 | 0 | 29.5 | 95.33 | 235691.5 | 94.68 |  | 125.3 | 0 | 100 | 157.59 | 5392.1 | 3.22 | 3.2146
0 | 0 | 28.3 | 95.98 | 237296.9 | 95.29 |  | 126.5 | 0 | 100 | 156.67 | 5360.5 | 3.19 | 3.1959
0 | 0 | 27.1 | 96.25 | 237958.1 | 95.63 |  | 127.7 | 0 | 100 | 152.51 | 5218.1 | 3.11 | 3.111
0 | 0 | 25.9 | 96.81 | 239336 | 96.15 |  | 128.9 | 0 | 100 | 150.41 | 5146.3 | 3.07 | 3.0682
0 | 0 | 24.7 | 97.32 | 240618.6 | 96.6 |  | 130.1 | 0 | 100 | 150.94 | 5164.6 | 3.08 | 3.079
0 | 0 | 23.5 | 98.01 | 242304.1 | 97.24 |  | 131.3 | 0 | 100 | 147.29 | 5039.7 | 3 | 3.0045
0 | 0 | 22.3 | 98.33 | 243106.8 | 97.82 |  | 132.5 | 0 | 100 | 145.19 | 4967.8 | 2.96 | 2.9617
0 | 0 | 21.1 | 99 | 244773.2 | 98.39 |  | 133.7 | 0 | 100 | 142.24 | 4866.9 | 2.9 | 2.9015
0 | 0 | 19.8 | 99.28 | 245457.3 | 98.79 |  | 134.9 | 0 | 100 | 142.08 | 4861.4 | 2.9 | 2.8982
0 | 0 | 18.6 | 99.83 | 246816.3 | 99.35 |  | 136.1 | 0 | 100 | 140.65 | 4812.4 | 2.87 | 2.8691
0 | 0 | 17.4 | 99.99 | 247202.3 | 99.27 |  | 137.3 | 0 | 100 | 135.36 | 4631.4 | 2.76 | 2.7612
0 | 0 | 16.2 | 100.74 | 249072.3 | 99.76 |  | 138.5 | 0 | 100 | 134.92 | 4616.3 | 2.75 | 2.7522
0 | 0 | 15 | 100.58 | 248663 | 99.79 |  | 139.7 | 0 | 100 | 131.98 | 4515.9 | 2.69 | 2.6922
0 | 0 | 13.8 | 100.78 | 249172.3 | 99.99 |  | 140.9 | 0 | 100 | 131.19 | 4488.6 | 2.68 | 2.6761
0 | 0 | 12.6 | 100.54 | 248559 | 100 |  | 142.1 | 0 | 100 | 128.2 | 4386.6 | 2.61 | 2.6151
0 | 0 | 11.4 | 100.04 | 247344.5 | 99.33 |  | 143.3 | 0 | 100 | 126.95 | 4343.7 | 2.59 | 2.5896
0 | 0 | 10.2 | 99.02 | 244807.6 | 98.17 |  | 144.5 | 0 | 100 | 125.08 | 4279.6 | 2.55 | 2.5515
0 | 0 | 9 | 97.79 | 241773.5 | 97.12 |  | 145.7 | 0 | 100 | 124.4 | 4256.5 | 2.53 | 2.5376
0 | 0 | 7.8 | 95.33 | 235679.2 | 94.65 |  | 146.9 | 0 | 100 | 120.69 | 4129.3 | 2.46 | 2.4619
0 | 0 | 6.6 | 91.96 | 227360.3 | 91.27 |  | 148.1 | 0 | 100 | 120.16 | 4111.3 | 2.45 | 2.4511
0 | 0 | 5.4 | 86.89 | 214830.9 | 86.3 |  | 149.4 | 0 | 100 | 121.42 | 4154.5 | 2.48 | 2.4768
0 | 0 | 4.2 | 79.72 | 197103.6 | 79.19 |  | 150.6 | 0 | 100 | 120.12 | 4109.9 | 2.45 | 2.4503
0 | 0 | 3 | 69.42 | 171637.6 | 68.95 |  | 151.8 | 0 | 100 | 112.92 | 3863.6 | 2.3 | 2.3034
0 | 0 | 1.7 | 59.81 | 147860.3 | 59.41 |  | 153 | 0 | 100 | 112.6 | 3852.7 | 2.3 | 2.2969
0 | 0 | 0.6 | 54.3 | 134255.9 | 53.92 |  | 155 | 0 | 100 | 112.22 | 3839.8 | 2.29 | 2.2891
0 | 0 | -0.5 | 50.98 | 126046.2 | 50.65
[7 empty] | Measurement time: | 30/08/2022 16:31:44
Measurement time: | 12/08/2022 18:08:37 | [5 empty] | Radiation device: | Novalis TB Accelerator
Radiation device: | Clinac2 Accelerator | [5 empty] | Energy: | 6 MV (FFF)
Energy: | 23 MV | [5 empty] | Controller: | CCU
Controller: | CCU | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Field detector: | CC 13 - F
Field detector: | CC 13 - F | [5 empty] | Reference detector: | CC 13 - R pas offset energie
Reference detector: | CC 13 - R pas offset energie | [5 empty] | SAD: | 1000 mm
SAD: | 1000 mm | [5 empty] | SSD: | 1000 mm
SSD: | 1000 mm | [5 empty] | Field size: | 100 x 100 mm
Field size: | 100 x 100 mm | [5 empty] | Gantry angle: | 0
Gantry angle: | 0 | [5 empty] | Measurement medium: | Water
Measurement medium: | Water | [5 empty] | Scan type: | Inline
Scan type: | Beam | [5 empty] | Scan mode: | Continuous
Scan mode: | Continuous | [5 empty] | Wedge: | none
Wedge: | none
Points [mm]:
Points [mm]: | [6 empty] | Inline | Crossline | Depth | Normalized field | Current field | Ratio
Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | 154.6 | 0 | 100 | 318.6 | 10901.2 | 1.96 | 1.9489
0 | 0 | 310 | 33.1 | 81833.9 | 33.6 |  | 153.4 | 0 | 100 | 325.9 | 11151 | 2 | 1.9936
0 | 0 | 309.1 | 34.77 | 85965.8 | 33.86 |  | 152.3 | 0 | 100 | 324.35 | 11097.8 | 1.99 | 1.9841
0 | 0 | 307.6 | 35.14 | 86867.2 | 34.19 |  | 151.2 | 0 | 100 | 330.19 | 11297.7 | 2.03 | 2.0198
0 | 0 | 306.6 | 35.13 | 86861.2 | 34.24 |  | 149.9 | 0 | 100 | 334.69 | 11451.6 | 2.05 | 2.0473
0 | 0 | 305.3 | 35.24 | 87125.6 | 34.33 |  | 148.7 | 0 | 100 | 336.9 | 11527.3 | 2.07 | 2.0608
0 | 0 | 304.2 | 35.38 | 87477.2 | 34.52 |  | 147.5 | 0 | 100 | 343.7 | 11759.8 | 2.11 | 2.1024
0 | 0 | 302.9 | 35.77 | 88431.4 | 34.67 |  | 146.3 | 0 | 100 | 346.29 | 11848.6 | 2.12 | 2.1183
0 | 0 | 301.7 | 35.69 | 88236.2 | 34.83 |  | 145.1 | 0 | 100 | 349.31 | 11951.7 | 2.15 | 2.1368
0 | 0 | 300.5 | 35.81 | 88538.9 | 35.03 |  | 143.9 | 0 | 100 | 356.26 | 12189.8 | 2.18 | 2.1793
0 | 0 | 299.3 | 36.29 | 89714.5 | 35.26 |  | 142.7 | 0 | 100 | 361.99 | 12385.8 | 2.22 | 2.2143
0 | 0 | 298.2 | 36.3 | 89741.2 | 35.37 |  | 141.5 | 0 | 100 | 364.6 | 12475 | 2.24 | 2.2303
0 | 0 | 296.9 | 36.37 | 89909.5 | 35.49 |  | 140.3 | 0 | 100 | 368.28 | 12601.1 | 2.26 | 2.2528
0 | 0 | 295.7 | 36.65 | 90610.7 | 35.73 |  | 139.1 | 0 | 100 | 377.06 | 12901.3 | 2.31 | 2.3065
0 | 0 | 294.5 | 36.81 | 91002.6 | 35.89 |  | 137.9 | 0 | 100 | 380.36 | 13014.3 | 2.33 | 2.3267
0 | 0 | 293.3 | 37.02 | 91532 | 36.11 |  | 136.7 | 0 | 100 | 383.09 | 13107.5 | 2.35 | 2.3434
0 | 0 | 292.1 | 37.24 | 92073.8 | 36.22 |  | 135.5 | 0 | 100 | 392.21 | 13419.7 | 2.4 | 2.3992
0 | 0 | 290.9 | 37.22 | 92019.1 | 36.38 |  | 134.2 | 0 | 100 | 393.02 | 13447.4 | 2.42 | 2.4041
0 | 0 | 289.7 | 37.58 | 92909.5 | 36.55 |  | 133 | 0 | 100 | 399.97 | 13685.3 | 2.45 | 2.4466
0 | 0 | 288.5 | 37.62 | 93011.1 | 36.73 |  | 131.8 | 0 | 100 | 409.75 | 14019.8 | 2.51 | 2.5065
0 | 0 | 287.3 | 37.97 | 93882.8 | 36.92 |  | 130.6 | 0 | 100 | 413.91 | 14162.1 | 2.54 | 2.5319
0 | 0 | 286.1 | 37.97 | 93884.1 | 37.04 |  | 129.4 | 0 | 100 | 422.79 | 14466 | 2.59 | 2.5862
0 | 0 | 284.9 | 38.3 | 94692.1 | 37.31 |  | 128.2 | 0 | 100 | 429.82 | 14706.6 | 2.64 | 2.6292
0 | 0 | 283.7 | 38.39 | 94907.2 | 37.48 |  | 127 | 0 | 100 | 434.72 | 14874.1 | 2.67 | 2.6592
0 | 0 | 282.5 | 38.67 | 95602.5 | 37.49 |  | 125.8 | 0 | 100 | 443.69 | 15181 | 2.71 | 2.7141
0 | 0 | 281.3 | 38.79 | 95903.7 | 37.79 |  | 124.6 | 0 | 100 | 450 | 15397.1 | 2.76 | 2.7527
0 | 0 | 280.1 | 39.1 | 96663.6 | 38.05 |  | 123.4 | 0 | 100 | 463.15 | 15847 | 2.84 | 2.8331
0 | 0 | 278.9 | 39.2 | 96917.8 | 38.26 |  | 122.2 | 0 | 100 | 467.07 | 15981 | 2.86 | 2.8571
0 | 0 | 277.7 | 39.33 | 97229.6 | 38.41 |  | 121 | 0 | 100 | 474.89 | 16248.6 | 2.92 | 2.9049
0 | 0 | 276.5 | 39.57 | 97820.6 | 38.54 |  | 119.8 | 0 | 100 | 485.9 | 16625.4 | 2.97 | 2.9723
0 | 0 | 275.2 | 39.79 | 98383.3 | 38.67 |  | 118.6 | 0 | 100 | 491.14 | 16804.7 | 3.01 | 3.0043
0 | 0 | 274.1 | 39.92 | 98698.7 | 38.95 |  | 117.4 | 0 | 100 | 499.69 | 17097.4 | 3.06 | 3.0566
0 | 0 | 272.8 | 40.23 | 99458.6 | 39.24 |  | 116.2 | 0 | 100 | 506.04 | 17314.5 | 3.11 | 3.0955
0 | 0 | 271.6 | 40.47 | 100065.5 | 39.41 |  | 115 | 0 | 100 | 520.84 | 17820.9 | 3.19 | 3.186
0 | 0 | 270.4 | 40.53 | 100198 | 39.53 |  | 113.8 | 0 | 100 | 529.51 | 18117.5 | 3.25 | 3.2391
0 | 0 | 269.2 | 40.75 | 100740.1 | 39.78 |  | 112.6 | 0 | 100 | 539.16 | 18447.6 | 3.3 | 3.2981
0 | 0 | 268 | 40.83 | 100949.8 | 39.95 |  | 111.4 | 0 | 100 | 552.1 | 18890.5 | 3.39 | 3.3772
0 | 0 | 266.8 | 41.2 | 101872.6 | 40.18 |  | 110.2 | 0 | 100 | 560.15 | 19165.8 | 3.43 | 3.4265
0 | 0 | 265.6 | 41.44 | 102466.2 | 40.27 |  | 109 | 0 | 100 | 569.49 | 19485.5 | 3.5 | 3.4836
0 | 0 | 264.4 | 41.48 | 102555.3 | 40.55 |  | 107.8 | 0 | 100 | 584.9 | 20012.6 | 3.58 | 3.5779
0 | 0 | 263.2 | 41.73 | 103173.7 | 40.79 |  | 106.6 | 0 | 100 | 592.37 | 20268.3 | 3.64 | 3.6236
0 | 0 | 262 | 41.8 | 103346.9 | 40.88 |  | 105.3 | 0 | 100 | 613.34 | 20985.8 | 3.76 | 3.7518
0 | 0 | 260.8 | 42.34 | 104690.9 | 41.15 |  | 104.2 | 0 | 100 | 622.03 | 21283 | 3.81 | 3.805
0 | 0 | 259.6 | 42.28 | 104524.2 | 41.23 |  | 103 | 0 | 100 | 637.72 | 21820.1 | 3.91 | 3.901
0 | 0 | 258.4 | 42.52 | 105128.6 | 41.43 |  | 101.8 | 0 | 100 | 654.33 | 22388.4 | 4.01 | 4.0026
0 | 0 | 257.2 | 42.7 | 105580.4 | 41.63 |  | 100.6 | 0 | 100 | 668.65 | 22878.3 | 4.1 | 4.0902
0 | 0 | 256 | 42.91 | 106100.2 | 41.91 |  | 99.3 | 0 | 100 | 682.3 | 23345.2 | 4.18 | 4.1737
0 | 0 | 254.8 | 43.13 | 106644.1 | 42.11 |  | 98.2 | 0 | 100 | 699.87 | 23946.6 | 4.29 | 4.2812
0 | 0 | 253.6 | 43.35 | 107175.1 | 42.24 |  | 96.9 | 0 | 100 | 717.91 | 24563.6 | 4.41 | 4.3915
0 | 0 | 252.4 | 43.66 | 107930.9 | 42.49 |  | 95.8 | 0 | 100 | 735.97 | 25181.6 | 4.51 | 4.502
0 | 0 | 251.2 | 43.79 | 108254.8 | 42.69 |  | 94.6 | 0 | 100 | 751.6 | 25716.4 | 4.61 | 4.5976
0 | 0 | 250 | 44.01 | 108796.4 | 42.94 |  | 93.3 | 0 | 100 | 774.62 | 26504 | 4.74 | 4.7384
0 | 0 | 248.8 | 44.22 | 109320.4 | 43.14 |  | 92.2 | 0 | 100 | 793.06 | 27135.1 | 4.87 | 4.8512
0 | 0 | 247.6 | 44.34 | 109623.4 | 43.41 |  | 90.9 | 0 | 100 | 820.53 | 28075 | 5.02 | 5.0192
0 | 0 | 246.4 | 44.62 | 110311.2 | 43.65 |  | 89.7 | 0 | 100 | 843.28 | 28853.4 | 5.18 | 5.1584
0 | 0 | 245.2 | 45.02 | 111302.9 | 43.86 |  | 88.5 | 0 | 100 | 863.99 | 29561.9 | 5.3 | 5.2851
0 | 0 | 244 | 45.16 | 111655.1 | 44.01 |  | 87.3 | 0 | 100 | 892.75 | 30546.2 | 5.47 | 5.461
0 | 0 | 242.7 | 45.19 | 111717 | 44.19 |  | 86.1 | 0 | 100 | 916.22 | 31349.1 | 5.61 | 5.6046
0 | 0 | 241.6 | 45.55 | 112607.3 | 44.5 |  | 84.9 | 0 | 100 | 942.24 | 32239.5 | 5.79 | 5.7638
0 | 0 | 240.4 | 45.78 | 113175.7 | 44.69 |  | 83.7 | 0 | 100 | 975.59 | 33380.3 | 5.97 | 5.9678
0 | 0 | 239.1 | 46.12 | 114034.2 | 44.98 |  | 82.5 | 0 | 100 | 1000.41 | 34229.5 | 6.15 | 6.1196
0 | 0 | 238 | 46.3 | 114477.2 | 45.1 |  | 81.3 | 0 | 100 | 1036.77 | 35473.7 | 6.34 | 6.342
0 | 0 | 236.7 | 46.45 | 114836.5 | 45.43 |  | 80.1 | 0 | 100 | 1068.27 | 36551.7 | 6.55 | 6.5347
0 | 0 | 235.5 | 46.76 | 115595.3 | 45.64 |  | 78.9 | 0 | 100 | 1102.46 | 37721.5 | 6.78 | 6.7438
0 | 0 | 234.3 | 46.96 | 116109.8 | 45.78 |  | 77.7 | 0 | 100 | 1141.41 | 39054 | 7 | 6.9821
0 | 0 | 233.1 | 47.15 | 116568.6 | 45.89 |  | 76.5 | 0 | 100 | 1185.59 | 40565.6 | 7.24 | 7.2523
0 | 0 | 231.9 | 47.3 | 116937.3 | 46.25 |  | 75.3 | 0 | 100 | 1215.18 | 41578 | 7.46 | 7.4333
0 | 0 | 230.7 | 47.51 | 117469.4 | 46.53 |  | 74.1 | 0 | 100 | 1260.81 | 43139.5 | 7.73 | 7.7125
0 | 0 | 229.5 | 47.89 | 118410.4 | 46.69 |  | 72.9 | 0 | 100 | 1311.29 | 44866.7 | 8.02 | 8.0213
0 | 0 | 228.3 | 48.21 | 119182.8 | 47.01 |  | 71.7 | 0 | 100 | 1354.17 | 46333.8 | 8.31 | 8.2836
0 | 0 | 227.1 | 48.32 | 119459.3 | 47.15 |  | 70.5 | 0 | 100 | 1404.91 | 48069.8 | 8.61 | 8.5939
0 | 0 | 225.9 | 48.46 | 119810.1 | 47.28 |  | 69.2 | 0 | 100 | 1462.77 | 50049.7 | 8.97 | 8.9479
0 | 0 | 224.7 | 48.7 | 120409.2 | 47.73 |  | 68.1 | 0 | 100 | 1520.24 | 52015.9 | 9.32 | 9.2994
0 | 0 | 223.5 | 49.19 | 121612.5 | 48.03 |  | 66.8 | 0 | 100 | 1598.71 | 54700.8 | 9.79 | 9.7794
0 | 0 | 222.3 | 49.26 | 121778.5 | 48.13 |  | 65.7 | 0 | 100 | 1682.03 | 57551.8 | 10.32 | 10.2891
0 | 0 | 221.1 | 49.43 | 122205.7 | 48.38 |  | 64.5 | 0 | 100 | 1788.65 | 61199.7 | 10.92 | 10.9413
0 | 0 | 219.9 | 49.69 | 122855.5 | 48.69 |  | 63.2 | 0 | 100 | 1924.64 | 65852.6 | 11.82 | 11.7732
0 | 0 | 218.7 | 49.99 | 123599.2 | 48.83 |  | 62.1 | 0 | 100 | 2136.12 | 73088.6 | 13.07 | 13.0668
0 | 0 | 217.5 | 50.5 | 124851.1 | 49.18 |  | 60.8 | 0 | 100 | 2465.61 | 84362.3 | 15.11 | 15.0823
0 | 0 | 216.3 | 50.43 | 124683.7 | 49.35 |  | 59.6 | 0 | 100 | 3059.83 | 104694.2 | 18.69 | 18.7172
0 | 0 | 215.1 | 50.86 | 125741.9 | 49.67 |  | 58.4 | 0 | 100 | 4040.4 | 138244.9 | 24.83 | 24.7154
0 | 0 | 213.9 | 51.04 | 126201.1 | 49.76 |  | 57.2 | 0 | 100 | 5572.24 | 190657.8 | 34.12 | 34.0858
0 | 0 | 212.6 | 51.42 | 127121 | 50.07 |  | 56 | 0 | 100 | 7170.94 | 245358.2 | 43.95 | 43.8652
0 | 0 | 211.5 | 51.59 | 127555.2 | 50.47 |  | 54.8 | 0 | 100 | 8975.05 | 307087.2 | 54.88 | 54.9011
0 | 0 | 210.2 | 51.87 | 128241.4 | 50.55 |  | 53.6 | 0 | 100 | 10516.23 | 359819.5 | 64.65 | 64.3286
0 | 0 | 209 | 51.9 | 128327.1 | 50.86 |  | 52.4 | 0 | 100 | 11752.95 | 402134.8 | 71.86 | 71.8937
0 | 0 | 207.9 | 52.28 | 129256.6 | 50.93 |  | 51.1 | 0 | 100 | 12707.74 | 434803.5 | 77.87 | 77.7342
0 | 0 | 206.6 | 52.55 | 129922.6 | 51.35 |  | 50 | 0 | 100 | 13295.11 | 454900.6 | 81.5 | 81.3272
0 | 0 | 205.4 | 52.72 | 130336.3 | 51.58 |  | 48.8 | 0 | 100 | 13732.22 | 469856.6 | 84.02 | 84.001
0 | 0 | 204.2 | 52.84 | 130634.9 | 51.9 |  | 47.6 | 0 | 100 | 13938.16 | 476903.1 | 85.6 | 85.2608
0 | 0 | 203 | 53.25 | 131647.5 | 52.15 |  | 46.4 | 0 | 100 | 14158.81 | 484452.9 | 86.67 | 86.6105
0 | 0 | 201.8 | 53.68 | 132719 | 52.34 |  | 45.2 | 0 | 100 | 14294.72 | 489102.9 | 87.55 | 87.4419
0 | 0 | 200.6 | 53.82 | 133070.9 | 52.63 |  | 44 | 0 | 100 | 14429.67 | 493720.2 | 88.42 | 88.2674
0 | 0 | 199.4 | 54.28 | 134208.1 | 52.99 |  | 42.7 | 0 | 100 | 14565.42 | 498365.1 | 89.22 | 89.0978
0 | 0 | 198.2 | 54.39 | 134472.9 | 53.13 |  | 41.6 | 0 | 100 | 14638.04 | 500849.9 | 89.91 | 89.542
0 | 0 | 197 | 54.53 | 134812.3 | 53.31 |  | 40.4 | 0 | 100 | 14779.18 | 505679.1 | 90.42 | 90.4054
0 | 0 | 195.8 | 54.82 | 135523.3 | 53.4 |  | 39.1 | 0 | 100 | 14863.5 | 508564.1 | 90.96 | 90.9212
0 | 0 | 194.6 | 55.28 | 136675.6 | 53.86 |  | 38 | 0 | 100 | 14908.5 | 510103.9 | 91.5 | 91.1964
0 | 0 | 193.4 | 55.44 | 137059.9 | 54.27 |  | 36.7 | 0 | 100 | 15031.39 | 514308.6 | 92.17 | 91.9482
0 | 0 | 192.2 | 55.94 | 138313.7 | 54.42 |  | 35.6 | 0 | 100 | 15108.81 | 516957.7 | 92.49 | 92.4217
0 | 0 | 191 | 55.89 | 138173 | 54.63 |  | 34.4 | 0 | 100 | 15178.31 | 519335.4 | 92.94 | 92.8469
0 | 0 | 189.8 | 56.37 | 139370.8 | 54.85 |  | 33.1 | 0 | 100 | 15240.29 | 521456.3 | 93.45 | 93.226
0 | 0 | 188.5 | 56.54 | 139782.9 | 55.28 |  | 32 | 0 | 100 | 15382.43 | 526319.8 | 94.01 | 94.0955
0 | 0 | 187.4 | 56.76 | 140334.3 | 55.51 |  | 30.7 | 0 | 100 | 15400.77 | 526947 | 94.38 | 94.2077
0 | 0 | 186.2 | 57.11 | 141202.1 | 55.66 |  | 29.6 | 0 | 100 | 15477.42 | 529569.7 | 94.76 | 94.6766
0 | 0 | 185 | 57.36 | 141825.2 | 55.97 |  | 28.3 | 0 | 100 | 15546.43 | 531931 | 95.18 | 95.0987
0 | 0 | 183.8 | 57.51 | 142187.8 | 56.27 |  | 27.2 | 0 | 100 | 15603.82 | 533894.8 | 95.49 | 95.4498
0 | 0 | 182.6 | 57.95 | 143262.1 | 56.66 |  | 26 | 0 | 100 | 15624.31 | 534595.9 | 95.93 | 95.5751
0 | 0 | 181.4 | 58.36 | 144276.7 | 57 |  | 24.7 | 0 | 100 | 15710.7 | 537551.7 | 96.29 | 96.1035
0 | 0 | 180.1 | 58.67 | 145054.6 | 57.29 |  | 23.6 | 0 | 100 | 15777.01 | 539820.6 | 96.52 | 96.5092
0 | 0 | 178.9 | 58.77 | 145298.2 | 57.48 |  | 22.3 | 0 | 100 | 15833.97 | 541769.2 | 96.81 | 96.8576
0 | 0 | 177.8 | 59.17 | 146297.5 | 57.81 |  | 21.1 | 0 | 100 | 15879.01 | 543310.5 | 97.37 | 97.1331
0 | 0 | 176.5 | 59.38 | 146817.3 | 58.12 |  | 20 | 0 | 100 | 15903.42 | 544145.6 | 97.55 | 97.2824
0 | 0 | 175.3 | 59.84 | 147933.2 | 58.26 |  | 18.7 | 0 | 100 | 15947.48 | 545653.1 | 97.86 | 97.5519
0 | 0 | 174.1 | 60.06 | 148496.6 | 58.65 |  | 17.5 | 0 | 100 | 16044.85 | 548984.7 | 98.08 | 98.1476
0 | 0 | 172.9 | 60.06 | 148491.9 | 58.88 |  | 16.3 | 0 | 100 | 15999.03 | 547416.9 | 98.26 | 97.8673
0 | 0 | 171.7 | 60.57 | 149746.6 | 59.32 |  | 15.1 | 0 | 100 | 16100.96 | 550904.5 | 98.59 | 98.4908
0 | 0 | 170.5 | 60.97 | 150738.4 | 59.52 |  | 14 | 0 | 100 | 16142.05 | 552310.7 | 98.73 | 98.7421
0 | 0 | 169.3 | 61.38 | 151744.6 | 59.9 |  | 12.7 | 0 | 100 | 16142.54 | 552327.4 | 98.89 | 98.7451
0 | 0 | 168.1 | 61.46 | 151941.3 | 60.12 |  | 11.5 | 0 | 100 | 16157.22 | 552829.6 | 99.18 | 98.8349
0 | 0 | 166.9 | 61.85 | 152921.9 | 60.38 |  | 10.3 | 0 | 100 | 16190.09 | 553954.3 | 99.16 | 99.036
0 | 0 | 165.7 | 61.84 | 152889.6 | 60.61 |  | 9.1 | 0 | 100 | 16201.82 | 554355.6 | 99.32 | 99.1078
0 | 0 | 164.5 | 62.51 | 154552.7 | 60.86 |  | 7.9 | 0 | 100 | 16248.75 | 555961.3 | 99.5 | 99.3948
0 | 0 | 163.3 | 62.51 | 154555.2 | 61.28 |  | 6.7 | 0 | 100 | 16253.24 | 556115 | 99.59 | 99.4223
0 | 0 | 162.1 | 62.99 | 155721.9 | 61.58 |  | 5.5 | 0 | 100 | 16255.89 | 556205.8 | 99.5 | 99.4385
0 | 0 | 160.9 | 63.37 | 156677.2 | 61.78 |  | 4.3 | 0 | 100 | 16275.05 | 556861.2 | 99.83 | 99.5557
0 | 0 | 159.7 | 63.84 | 157839.8 | 62.2 |  | 3.1 | 0 | 100 | 16301.21 | 557756.3 | 99.9 | 99.7157
0 | 0 | 158.5 | 63.96 | 158125.9 | 62.42 |  | 1.9 | 0 | 100 | 16347.68 | 559346.4 | 100 | 100
0 | 0 | 157.3 | 64.15 | 158597.7 | 62.75 |  | 0.7 | 0 | 100 | 16317.23 | 558304.4 | 99.82 | 99.8137
0 | 0 | 156.1 | 64.65 | 159834.5 | 63.14 |  | -0.5 | 0 | 100 | 16247.81 | 555929.2 | 99.9 | 99.3891
0 | 0 | 154.9 | 64.74 | 160059.3 | 63.38 |  | -1.7 | 0 | 100 | 16311.4 | 558104.9 | 99.94 | 99.7781
0 | 0 | 153.6 | 65.28 | 161406 | 63.55 |  | -2.9 | 0 | 100 | 16313.9 | 558190.3 | 99.93 | 99.7934
0 | 0 | 152.4 | 65.49 | 161908.2 | 63.89 |  | -4.1 | 0 | 100 | 16337.91 | 559012.1 | 99.92 | 99.9402
0 | 0 | 151.3 | 65.75 | 162559.3 | 64.34 |  | -5.4 | 0 | 100 | 16286.6 | 557256.6 | 99.78 | 99.6264
0 | 0 | 150 | 66.03 | 163240.5 | 64.61 |  | -6.5 | 0 | 100 | 16232.15 | 555393.5 | 99.62 | 99.2933
0 | 0 | 148.8 | 66.62 | 164701.8 | 65.13 |  | -7.8 | 0 | 100 | 16281.12 | 557069.1 | 99.78 | 99.5928
0 | 0 | 147.6 | 66.91 | 165421.5 | 65.34 |  | -9 | 0 | 100 | 16247.96 | 555934.4 | 99.59 | 99.39
0 | 0 | 146.4 | 67.3 | 166391.9 | 65.69 |  | -10.2 | 0 | 100 | 16253.41 | 556120.8 | 99.49 | 99.4233
0 | 0 | 145.2 | 67.48 | 166823.8 | 66.01 |  | -11.4 | 0 | 100 | 16210.69 | 554658.9 | 99.37 | 99.162
0 | 0 | 144 | 67.74 | 167478.5 | 66.09 |  | -12.6 | 0 | 100 | 16177.14 | 553511.3 | 98.97 | 98.9568
0 | 0 | 142.8 | 68.22 | 168672.6 | 66.45 |  | -13.8 | 0 | 100 | 16155.12 | 552757.6 | 98.87 | 98.8221
0 | 0 | 141.6 | 68.57 | 169526.7 | 66.87 |  | -15 | 0 | 100 | 16144.95 | 552409.8 | 98.75 | 98.7599
0 | 0 | 140.4 | 68.72 | 169893.8 | 67.4 |  | -16.2 | 0 | 100 | 16037.16 | 548721.5 | 98.48 | 98.1005
0 | 0 | 139.2 | 69.14 | 170932.3 | 67.53 |  | -17.4 | 0 | 100 | 16070.39 | 549858.8 | 98.24 | 98.3038
0 | 0 | 138 | 69.66 | 172223.8 | 67.83 |  | -18.6 | 0 | 100 | 15965 | 546252.7 | 97.94 | 97.6591
0 | 0 | 136.8 | 69.83 | 172652.2 | 68.15 |  | -19.8 | 0 | 100 | 15964.64 | 546240.3 | 97.79 | 97.6569
0 | 0 | 135.6 | 70.25 | 173688.4 | 68.49 |  | -21 | 0 | 100 | 15946.61 | 545623.3 | 97.56 | 97.5466
0 | 0 | 134.4 | 70.56 | 174440.5 | 69.08 |  | -22.2 | 0 | 100 | 15878.08 | 543278.7 | 97.2 | 97.1274
0 | 0 | 133.2 | 71.13 | 175854.2 | 69.38 |  | -23.4 | 0 | 100 | 15839.56 | 541960.7 | 96.89 | 96.8918
0 | 0 | 132 | 71.28 | 176236.4 | 69.61 |  | -24.6 | 0 | 100 | 15747.08 | 538796.5 | 96.66 | 96.3261
0 | 0 | 130.8 | 71.66 | 177173.7 | 70.09 |  | -25.8 | 0 | 100 | 15715.06 | 537700.7 | 96.27 | 96.1302
0 | 0 | 129.6 | 71.98 | 177967.1 | 70.42 |  | -27.1 | 0 | 100 | 15679.03 | 536467.9 | 95.97 | 95.9098
0 | 0 | 128.4 | 72.48 | 179197.9 | 70.78 |  | -28.2 | 0 | 100 | 15638.82 | 535092.2 | 95.56 | 95.6638
0 | 0 | 127.2 | 73.11 | 180757.5 | 71.13 |  | -29.5 | 0 | 100 | 15521.55 | 531079.6 | 95.13 | 94.9465
0 | 0 | 126 | 72.71 | 179764.5 | 71.32 |  | -30.6 | 0 | 100 | 15442.4 | 528371.5 | 94.76 | 94.4623
0 | 0 | 124.8 | 73.56 | 181862.9 | 71.68 |  | -31.9 | 0 | 100 | 15403.92 | 527055 | 94.28 | 94.2269
0 | 0 | 123.6 | 73.67 | 182138.4 | 72.07 |  | -33.1 | 0 | 100 | 15315.79 | 524039.6 | 93.85 | 93.6879
0 | 0 | 122.3 | 74.12 | 183260.5 | 72.59 |  | -34.2 | 0 | 100 | 15281.26 | 522858.1 | 93.41 | 93.4766
0 | 0 | 121.1 | 74.59 | 184417.5 | 72.88 |  | -35.5 | 0 | 100 | 15157.85 | 518635.7 | 93.01 | 92.7217
0 | 0 | 119.9 | 74.86 | 185087.5 | 73.09 |  | -36.6 | 0 | 100 | 15082.27 | 516049.6 | 92.52 | 92.2594
0 | 0 | 118.7 | 75.32 | 186226.4 | 73.51 |  | -37.9 | 0 | 100 | 15005.39 | 513418.9 | 91.9 | 91.7891
0 | 0 | 117.6 | 75.32 | 186214.2 | 73.74 |  | -39.1 | 0 | 100 | 14915.58 | 510346.3 | 91.34 | 91.2397
0 | 0 | 116.3 | 76.08 | 188090.1 | 74.31 |  | -40.2 | 0 | 100 | 14826.03 | 507282.1 | 90.8 | 90.692
0 | 0 | 115.1 | 76.29 | 188617.7 | 74.61 |  | -41.5 | 0 | 100 | 14717.14 | 503556.5 | 90.15 | 90.0259
0 | 0 | 113.9 | 76.78 | 189822.8 | 74.99 |  | -42.6 | 0 | 100 | 14638.11 | 500852.1 | 89.5 | 89.5424
0 | 0 | 112.7 | 77.11 | 190648.5 | 75.32 |  | -43.9 | 0 | 100 | 14535.6 | 497344.8 | 88.9 | 88.9154
0 | 0 | 111.5 | 77.32 | 191166.2 | 75.6 |  | -45.1 | 0 | 100 | 14382.32 | 492100.1 | 88.07 | 87.9777
0 | 0 | 110.3 | 77.99 | 192828.2 | 75.93 |  | -46.3 | 0 | 100 | 14207.96 | 486134.6 | 87.23 | 86.9112
0 | 0 | 109.1 | 78.05 | 192974.6 | 76.55 |  | -47.5 | 0 | 100 | 14082.64 | 481846.4 | 86.09 | 86.1446
0 | 0 | 107.9 | 78.49 | 194044.7 | 76.75 |  | -48.7 | 0 | 100 | 13792.53 | 471920.1 | 84.61 | 84.37
0 | 0 | 106.7 | 79.09 | 195531.7 | 77.14 |  | -49.9 | 0 | 100 | 13476.46 | 461105.6 | 82.47 | 82.4365
0 | 0 | 105.5 | 79.45 | 196422.9 | 77.31 |  | -51.1 | 0 | 100 | 12809.57 | 438287.5 | 78.45 | 78.3571
0 | 0 | 104.3 | 79.64 | 196899.2 | 77.78 |  | -52.3 | 0 | 100 | 11859.62 | 405784.6 | 72.5 | 72.5462
0 | 0 | 103.1 | 79.85 | 197410 | 78.43 |  | -53.5 | 0 | 100 | 10296.58 | 352304.2 | 63.12 | 62.985
0 | 0 | 101.9 | 80.69 | 199491.6 | 78.52 |  | -54.7 | 0 | 100 | 8651.14 | 296004.4 | 52.81 | 52.9197
0 | 0 | 100.7 | 80.88 | 199964.6 | 79.08 |  | -55.9 | 0 | 100 | 6837.29 | 233942.4 | 41.9 | 41.8242
0 | 0 | 99.5 | 81.22 | 200802.8 | 79.39 |  | -57.1 | 0 | 100 | 5273.87 | 180448.9 | 32.33 | 32.2607
0 | 0 | 98.3 | 81.66 | 201899.7 | 79.8 |  | -58.3 | 0 | 100 | 4092.54 | 140029 | 25.03 | 25.0344
0 | 0 | 97.1 | 82.06 | 202874.7 | 80.12 |  | -59.6 | 0 | 100 | 3113.59 | 106533.3 | 19.09 | 19.0461
0 | 0 | 95.9 | 82.34 | 203578.6 | 80.47 |  | -60.7 | 0 | 100 | 2548.94 | 87213.6 | 15.56 | 15.5921
0 | 0 | 94.7 | 82.86 | 204863.3 | 81.08 |  | -62 | 0 | 100 | 2179.06 | 74557.8 | 13.33 | 13.3295
0 | 0 | 93.5 | 83.28 | 205906.9 | 81.37 |  | -63.1 | 0 | 100 | 1955.63 | 66913.2 | 11.98 | 11.9627
0 | 0 | 92.2 | 83.68 | 206889.1 | 82.04 |  | -64.3 | 0 | 100 | 1809.15 | 61901.3 | 11.09 | 11.0667
0 | 0 | 91.1 | 84.36 | 208557.7 | 82.21 |  | -65.6 | 0 | 100 | 1698.94 | 58130.4 | 10.39 | 10.3925
0 | 0 | 89.9 | 84.53 | 208986.4 | 82.73 |  | -66.7 | 0 | 100 | 1612.44 | 55170.6 | 9.87 | 9.8634
0 | 0 | 88.7 | 84.97 | 210064.3 | 82.89 |  | -68 | 0 | 100 | 1529.93 | 52347.7 | 9.37 | 9.3587
0 | 0 | 87.5 | 85.23 | 210707.1 | 83.32 |  | -69.2 | 0 | 100 | 1470.42 | 50311.5 | 9 | 8.9947
0 | 0 | 86.2 | 85.86 | 212269.2 | 83.93 |  | -70.4 | 0 | 100 | 1413.59 | 48367 | 8.65 | 8.647
0 | 0 | 85.1 | 86.28 | 213309.7 | 84.21 |  | -71.6 | 0 | 100 | 1359.85 | 46528.1 | 8.34 | 8.3183
0 | 0 | 83.8 | 86.51 | 213875.1 | 84.55 |  | -72.8 | 0 | 100 | 1310.3 | 44832.8 | 8.02 | 8.0152
0 | 0 | 82.6 | 87.07 | 215264.3 | 85.1 |  | -74 | 0 | 100 | 1260.38 | 43124.8 | 7.72 | 7.7098
0 | 0 | 81.4 | 87.27 | 215761.8 | 85.35 |  | -75.2 | 0 | 100 | 1222.61 | 41832.2 | 7.48 | 7.4788
0 | 0 | 80.2 | 87.72 | 216866.4 | 85.79 |  | -76.4 | 0 | 100 | 1178.37 | 40318.7 | 7.2 | 7.2082
0 | 0 | 79 | 88.27 | 218232.7 | 86.04 |  | -77.6 | 0 | 100 | 1140.85 | 39035 | 6.99 | 6.9787
0 | 0 | 77.8 | 88.51 | 218832.3 | 86.46 |  | -78.8 | 0 | 100 | 1098.92 | 37600.2 | 6.73 | 6.7222
0 | 0 | 76.6 | 88.76 | 219437.7 | 86.88 |  | -80 | 0 | 100 | 1069.16 | 36582 | 6.54 | 6.5401
0 | 0 | 75.4 | 89.51 | 221307.9 | 87.31 |  | -81.2 | 0 | 100 | 1025.01 | 35071.3 | 6.28 | 6.2701
0 | 0 | 74.2 | 89.81 | 222037 | 87.8 |  | -82.4 | 0 | 100 | 997.54 | 34131.4 | 6.12 | 6.102
0 | 0 | 73 | 90.55 | 223876.1 | 88.32 |  | -83.6 | 0 | 100 | 970.54 | 33207.5 | 5.94 | 5.9369
0 | 0 | 71.8 | 90.45 | 223614.5 | 88.58 |  | -84.8 | 0 | 100 | 940.85 | 32191.9 | 5.75 | 5.7553
0 | 0 | 70.5 | 91.08 | 225173 | 89.17 |  | -86.1 | 0 | 100 | 904.14 | 30935.7 | 5.54 | 5.5307
0 | 0 | 69.3 | 91.66 | 226607.8 | 89.55 |  | -87.2 | 0 | 100 | 877.59 | 30027.4 | 5.37 | 5.3683
0 | 0 | 68.1 | 91.97 | 227391.5 | 89.73 |  | -88.4 | 0 | 100 | 854.43 | 29235 | 5.23 | 5.2266
0 | 0 | 66.9 | 92.47 | 228615.1 | 90.37 |  | -89.7 | 0 | 100 | 830.2 | 28406 | 5.07 | 5.0784
0 | 0 | 65.7 | 92.99 | 229906.7 | 91.01 |  | -90.8 | 0 | 100 | 805.57 | 27562.9 | 4.94 | 4.9277
0 | 0 | 64.5 | 93.38 | 230873.5 | 91.29 |  | -92.1 | 0 | 100 | 786.22 | 26900.9 | 4.82 | 4.8094
0 | 0 | 63.3 | 93.72 | 231716.6 | 91.52 |  | -93.2 | 0 | 100 | 758.41 | 25949.3 | 4.64 | 4.6393
0 | 0 | 62.1 | 94.17 | 232830 | 91.95 |  | -94.4 | 0 | 100 | 739.58 | 25305.2 | 4.53 | 4.5241
0 | 0 | 60.9 | 94.7 | 234124.4 | 92.46 |  | -95.7 | 0 | 100 | 720.66 | 24657.7 | 4.41 | 4.4083
0 | 0 | 59.7 | 94.88 | 234581.7 | 92.85 |  | -96.8 | 0 | 100 | 703.41 | 24067.5 | 4.3 | 4.3028
0 | 0 | 58.5 | 95.4 | 235855.4 | 93.28 |  | -98.1 | 0 | 100 | 680.56 | 23285.9 | 4.17 | 4.163
0 | 0 | 57.3 | 96.07 | 237512.2 | 93.54 |  | -99.3 | 0 | 100 | 665.03 | 22754.5 | 4.07 | 4.068
0 | 0 | 56.1 | 96.47 | 238510.3 | 94.11 |  | -100.5 | 0 | 100 | 649.76 | 22231.8 | 3.98 | 3.9746
0 | 0 | 54.9 | 96.46 | 238493.9 | 94.7 |  | -101.7 | 0 | 100 | 633.62 | 21679.8 | 3.88 | 3.8759
0 | 0 | 53.6 | 97.28 | 240511.1 | 94.85 |  | -102.9 | 0 | 100 | 616.93 | 21108.5 | 3.79 | 3.7738
0 | 0 | 52.4 | 97.52 | 241096.7 | 95.35 |  | -104.1 | 0 | 100 | 603.1 | 20635.4 | 3.69 | 3.6892
0 | 0 | 51.2 | 97.99 | 242256.7 | 95.91 |  | -105.3 | 0 | 100 | 595.13 | 20362.9 | 3.64 | 3.6405
0 | 0 | 50 | 98.06 | 242449.4 | 96.14 |  | -106.5 | 0 | 100 | 578.73 | 19801.6 | 3.54 | 3.5401
0 | 0 | 48.8 | 98.69 | 243986.8 | 96.36 |  | -107.7 | 0 | 100 | 564.96 | 19330.6 | 3.46 | 3.4559
0 | 0 | 47.6 | 99.07 | 244945.8 | 96.78 |  | -108.9 | 0 | 100 | 553.64 | 18943 | 3.38 | 3.3867
0 | 0 | 46.4 | 99.32 | 245552.9 | 96.96 |  | -110.1 | 0 | 100 | 539.36 | 18454.7 | 3.3 | 3.2993
0 | 0 | 45.2 | 99.71 | 246528.1 | 97.43 |  | -111.3 | 0 | 100 | 528.4 | 18079.4 | 3.23 | 3.2323
0 | 0 | 44 | 100.15 | 247596.8 | 97.82 |  | -112.5 | 0 | 100 | 517.08 | 17692.3 | 3.17 | 3.163
0 | 0 | 42.8 | 100.36 | 248131.4 | 98.12 |  | -113.7 | 0 | 100 | 507.09 | 17350.5 | 3.09 | 3.1019
0 | 0 | 41.6 | 100.88 | 249405.3 | 98.3 |  | -114.9 | 0 | 100 | 496.71 | 16995.1 | 3.05 | 3.0384
0 | 0 | 40.4 | 101.24 | 250306.6 | 98.7 |  | -116.1 | 0 | 100 | 486.39 | 16642.2 | 2.98 | 2.9753
0 | 0 | 39.2 | 100.95 | 249574.9 | 99.27 |  | -117.3 | 0 | 100 | 479.62 | 16410.5 | 2.93 | 2.9339
0 | 0 | 38 | 101.5 | 250941.2 | 99.27 |  | -118.5 | 0 | 100 | 466.56 | 15963.7 | 2.85 | 2.854
0 | 0 | 36.8 | 102.1 | 252438.1 | 99.52 |  | -119.7 | 0 | 100 | 457.37 | 15649.1 | 2.8 | 2.7978
0 | 0 | 35.5 | 101.93 | 252018.7 | 99.72 |  | -120.9 | 0 | 100 | 452.04 | 15467 | 2.76 | 2.7652
0 | 0 | 34.3 | 102.32 | 252970.9 | 99.96 |  | -122.1 | 0 | 100 | 442.1 | 15126.8 | 2.71 | 2.7044
0 | 0 | 33.1 | 102.06 | 252335.1 | 100 |  | -123.3 | 0 | 100 | 437.67 | 14975.2 | 2.67 | 2.6773
0 | 0 | 31.9 | 102.23 | 252756.5 | 99.91 |  | -124.5 | 0 | 100 | 425.01 | 14541.8 | 2.61 | 2.5998
0 | 0 | 30.7 | 102.3 | 252930.6 | 99.7 |  | -125.7 | 0 | 100 | 419.57 | 14355.9 | 2.56 | 2.5665
0 | 0 | 29.5 | 102.27 | 252848.2 | 99.79 |  | -126.9 | 0 | 100 | 407.93 | 13957.5 | 2.5 | 2.4953
0 | 0 | 28.3 | 101.94 | 252031.8 | 99.74 |  | -128.1 | 0 | 100 | 404.67 | 13845.9 | 2.48 | 2.4754
0 | 0 | 27.1 | 101.62 | 251250.9 | 99.46 |  | -129.3 | 0 | 100 | 399.59 | 13672.2 | 2.45 | 2.4443
0 | 0 | 25.9 | 101.66 | 251350.4 | 99.28 |  | -130.5 | 0 | 100 | 394.17 | 13486.7 | 2.41 | 2.4112
0 | 0 | 24.7 | 100.87 | 249397.5 | 98.66 |  | -131.8 | 0 | 100 | 381.27 | 13045.6 | 2.33 | 2.3323
0 | 0 | 23.5 | 100.65 | 248854.2 | 98.45 |  | -132.9 | 0 | 100 | 381.01 | 13036.4 | 2.32 | 2.3307
0 | 0 | 22.3 | 99.78 | 246694.8 | 97.51 |  | -134.1 | 0 | 100 | 371.7 | 12717.8 | 2.27 | 2.2737
0 | 0 | 21 | 98.69 | 243992 | 96.83 |  | -135.3 | 0 | 100 | 367.79 | 12584.3 | 2.25 | 2.2498
0 | 0 | 19.8 | 98.07 | 242464.2 | 95.77 |  | -136.5 | 0 | 100 | 361.42 | 12366.2 | 2.21 | 2.2108
0 | 0 | 18.7 | 97 | 239813.6 | 94.61 |  | -137.8 | 0 | 100 | 355.74 | 12171.9 | 2.18 | 2.1761
0 | 0 | 17.4 | 95.7 | 236611 | 93.43 |  | -138.9 | 0 | 100 | 348.14 | 11912 | 2.13 | 2.1296
0 | 0 | 16.3 | 94 | 232402.7 | 91.95 |  | -140.1 | 0 | 100 | 346.14 | 11843.2 | 2.12 | 2.1174
0 | 0 | 15 | 92.08 | 227646.9 | 90 |  | -141.3 | 0 | 100 | 343.16 | 11741.6 | 2.1 | 2.0991
0 | 0 | 13.8 | 90.06 | 222648.8 | 87.87 |  | -142.5 | 0 | 100 | 336.93 | 11528.1 | 2.06 | 2.061
0 | 0 | 12.6 | 87.56 | 216481.9 | 85.42 |  | -143.8 | 0 | 100 | 331.94 | 11357.5 | 2.03 | 2.0305
0 | 0 | 11.4 | 84.36 | 208571.4 | 82.84 |  | -144.9 | 0 | 100 | 326.51 | 11171.7 | 2 | 1.9973
0 | 0 | 10.2 | 81.48 | 201459.3 | 79.5 |  | -146.2 | 0 | 100 | 319.41 | 10928.8 | 1.95 | 1.9539
0 | 0 | 9 | 77.7 | 192105.4 | 75.86 |  | -147.3 | 0 | 100 | 317.62 | 10867.6 | 1.94 | 1.9429
0 | 0 | 7.8 | 72.85 | 180115.4 | 71.02 |  | -148.5 | 0 | 100 | 311.15 | 10646.2 | 1.9 | 1.9033
0 | 0 | 6.6 | 67.39 | 166609.8 | 65.82 |  | -149.8 | 0 | 100 | 308.99 | 10572.2 | 1.89 | 1.8901
0 | 0 | 5.4 | 60.8 | 150329.2 | 59.23 |  | -150.9 | 0 | 100 | 308.36 | 10550.8 | 1.88 | 1.8863
0 | 0 | 4.2 | 53.13 | 131355.7 | 51.76 |  | -152.2 | 0 | 100 | 299.73 | 10255.4 | 1.84 | 1.8335
0 | 0 | 3 | 44.31 | 109548.9 | 43.46 |  | -153.3 | 0 | 100 | 302.09 | 10336.2 | 1.85 | 1.8479
0 | 0 | 1.8 | 37.36 | 92377 | 36.44 |  | -155 | 0 | 100 | 295.69 | 10117.3 | 1.81 | 1.8088
0 | 0 | 0.6 | 33.46 | 82728.8 | 32.77
0 | 0 | -0.5 | 31.45 | 77754.3 | 30.64
Measurement time: | 12/08/2022 18:12:50
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
-100 | 0 | 100 | 2.43 | 6017.8 | 3.78 | 3.6215
-99.1 | 0 | 100 | 2.98 | 7363.3 | 4.43 | 4.4411
-97.7 | 0 | 100 | 3.03 | 7502 | 4.51 | 4.5156
-96.6 | 0 | 100 | 3.16 | 7801.4 | 4.69 | 4.7094
-95.4 | 0 | 100 | 3.24 | 8009.7 | 4.81 | 4.8286
-94.1 | 0 | 100 | 3.32 | 8198.1 | 4.92 | 4.9478
-93 | 0 | 100 | 3.43 | 8467.9 | 5.08 | 5.1118
-91.8 | 0 | 100 | 3.57 | 8821.7 | 5.3 | 5.3204
-90.6 | 0 | 100 | 3.68 | 9097.9 | 5.47 | 5.4844
-89.4 | 0 | 100 | 3.76 | 9301.1 | 5.59 | 5.6036
-88.1 | 0 | 100 | 3.92 | 9693.8 | 5.82 | 5.842
-86.9 | 0 | 100 | 4.02 | 9930.5 | 5.97 | 5.9911
-85.7 | 0 | 100 | 4.18 | 10341.6 | 6.21 | 6.2295
-84.5 | 0 | 100 | 4.32 | 10689.8 | 6.44 | 6.4382
-83.3 | 0 | 100 | 4.43 | 10940.6 | 6.6 | 6.6021
-82.1 | 0 | 100 | 4.62 | 11423.2 | 6.87 | 6.8852
-80.9 | 0 | 100 | 4.72 | 11681.8 | 7.02 | 7.0343
-79.7 | 0 | 100 | 4.92 | 12167.6 | 7.31 | 7.3323
-78.5 | 0 | 100 | 5.09 | 12589.6 | 7.56 | 7.5857
-77.3 | 0 | 100 | 5.25 | 12972.6 | 7.79 | 7.8241
-76.1 | 0 | 100 | 5.45 | 13463.5 | 8.1 | 8.1222
-74.9 | 0 | 100 | 5.65 | 13957.7 | 8.39 | 8.4203
-73.7 | 0 | 100 | 5.8 | 14339.6 | 8.63 | 8.6438
-72.5 | 0 | 100 | 6.06 | 14991.6 | 9.01 | 9.0313
-71.3 | 0 | 100 | 6.33 | 15644.6 | 9.4 | 9.4337
-70.1 | 0 | 100 | 6.54 | 16178.4 | 9.72 | 9.7466
-68.9 | 0 | 100 | 6.88 | 17008.1 | 10.22 | 10.2534
-67.7 | 0 | 100 | 7.14 | 17659.6 | 10.6 | 10.6408
-66.5 | 0 | 100 | 7.49 | 18510.3 | 11.14 | 11.1624
-65.3 | 0 | 100 | 7.92 | 19573.7 | 11.78 | 11.8033
-64.1 | 0 | 100 | 8.44 | 20854.5 | 12.54 | 12.5782
-62.9 | 0 | 100 | 9.15 | 22615.7 | 13.61 | 13.6364
-61.7 | 0 | 100 | 10.29 | 25446.7 | 15.29 | 15.3353
-60.5 | 0 | 100 | 12.2 | 30162.1 | 18.14 | 18.1818
-59.3 | 0 | 100 | 15.3 | 37836.1 | 22.75 | 22.8018
-58 | 0 | 100 | 20.1 | 49689.6 | 29.9 | 29.9553
-56.8 | 0 | 100 | 26.33 | 65108.2 | 39.15 | 39.2399
-55.6 | 0 | 100 | 33.92 | 83864.5 | 50.51 | 50.5514
-54.4 | 0 | 100 | 42.21 | 104351.1 | 62.76 | 62.9061
-53.2 | 0 | 100 | 49.84 | 123212 | 74.1 | 74.2772
-52 | 0 | 100 | 55.97 | 138366.6 | 83.2 | 83.4128
-50.8 | 0 | 100 | 59.72 | 147654.1 | 88.93 | 89.0015
-49.6 | 0 | 100 | 61.87 | 152956.1 | 92.23 | 92.2057
-48.4 | 0 | 100 | 63.12 | 156064.7 | 93.91 | 94.0686
-47.2 | 0 | 100 | 63.78 | 157694.1 | 94.91 | 95.0522
-46 | 0 | 100 | 64.21 | 158744.5 | 95.5 | 95.693
-44.8 | 0 | 100 | 64.44 | 159316.2 | 96.1 | 96.0358
-43.6 | 0 | 100 | 64.63 | 159781.4 | 96.33 | 96.3189
-42.4 | 0 | 100 | 64.97 | 160634.4 | 96.73 | 96.8256
-41.2 | 0 | 100 | 65.12 | 161010.6 | 96.94 | 97.0492
-40 | 0 | 100 | 65.34 | 161543.3 | 97.23 | 97.377
-38.8 | 0 | 100 | 65.32 | 161498.9 | 97.26 | 97.3472
-37.6 | 0 | 100 | 65.44 | 161799.3 | 97.57 | 97.5261
-36.4 | 0 | 100 | 65.73 | 162515.8 | 97.83 | 97.9583
-35.2 | 0 | 100 | 65.95 | 163057.7 | 97.99 | 98.2861
-34 | 0 | 100 | 66.13 | 163495.7 | 98.34 | 98.5544
-32.8 | 0 | 100 | 66.15 | 163542.7 | 98.48 | 98.5842
-31.6 | 0 | 100 | 66.36 | 164057.9 | 98.66 | 98.8972
-30.4 | 0 | 100 | 66.53 | 164473.8 | 98.97 | 99.1505
-29.2 | 0 | 100 | 66.57 | 164588.7 | 99.17 | 99.2101
-28 | 0 | 100 | 66.74 | 165015.9 | 99.27 | 99.4635
-26.8 | 0 | 100 | 66.82 | 165191.2 | 99.49 | 99.5827
-25.5 | 0 | 100 | 66.71 | 164924.1 | 99.31 | 99.4188
-24.4 | 0 | 100 | 66.96 | 165550.5 | 99.67 | 99.7914
-23.1 | 0 | 100 | 66.93 | 165476.1 | 99.79 | 99.7466
-21.9 | 0 | 100 | 66.96 | 165553 | 99.61 | 99.7914
-20.7 | 0 | 100 | 66.98 | 165608.6 | 99.68 | 99.8212
-19.5 | 0 | 100 | 67 | 165652 | 99.66 | 99.851
-18.3 | 0 | 100 | 67.08 | 165847.8 | 99.88 | 99.9702
-17.1 | 0 | 100 | 67.04 | 165748.2 | 99.74 | 99.9106
-15.9 | 0 | 100 | 66.99 | 165617.2 | 99.87 | 99.8361
-14.7 | 0 | 100 | 67.1 | 165890.4 | 99.93 | 100
-13.5 | 0 | 100 | 66.98 | 165599.5 | 99.8 | 99.8212
-12.3 | 0 | 100 | 66.93 | 165483 | 99.86 | 99.7466
-11.1 | 0 | 100 | 66.99 | 165627.2 | 100 | 99.8361
-9.9 | 0 | 100 | 66.89 | 165373 | 99.72 | 99.687
-8.7 | 0 | 100 | 66.99 | 165628.2 | 99.66 | 99.8361
-7.5 | 0 | 100 | 66.89 | 165363.8 | 99.56 | 99.687
-6.3 | 0 | 100 | 66.8 | 165157.8 | 99.51 | 99.5529
-5.1 | 0 | 100 | 66.98 | 165588.2 | 99.75 | 99.8212
-3.8 | 0 | 100 | 66.83 | 165232.3 | 99.55 | 99.5976
-2.7 | 0 | 100 | 66.71 | 164931.2 | 99.43 | 99.4188
-1.4 | 0 | 100 | 66.65 | 164781.7 | 99.11 | 99.3294
-0.3 | 0 | 100 | 66.62 | 164718.3 | 99.28 | 99.2846
1 | 0 | 100 | 66.68 | 164850.7 | 99.24 | 99.3741
2.2 | 0 | 100 | 66.68 | 164859.5 | 99.29 | 99.3741
3.4 | 0 | 100 | 66.53 | 164476.8 | 99.08 | 99.1505
4.6 | 0 | 100 | 66.48 | 164354.3 | 99.1 | 99.076
5.8 | 0 | 100 | 66.62 | 164704.4 | 99.24 | 99.2846
7 | 0 | 100 | 66.46 | 164323.9 | 99.09 | 99.0462
8.2 | 0 | 100 | 66.64 | 164766.1 | 99.27 | 99.3145
9.4 | 0 | 100 | 66.65 | 164776.5 | 99.38 | 99.3294
10.6 | 0 | 100 | 66.52 | 164451.7 | 99.15 | 99.1356
11.8 | 0 | 100 | 66.58 | 164614.9 | 99.24 | 99.225
13 | 0 | 100 | 66.58 | 164600.8 | 99.29 | 99.225
14.2 | 0 | 100 | 66.55 | 164545.8 | 99.3 | 99.1803
15.4 | 0 | 100 | 66.52 | 164468 | 99.29 | 99.1356
16.6 | 0 | 100 | 66.5 | 164414.9 | 99.05 | 99.1058
17.8 | 0 | 100 | 66.44 | 164259.1 | 98.88 | 99.0164
19 | 0 | 100 | 66.51 | 164446 | 98.84 | 99.1207
20.2 | 0 | 100 | 66.36 | 164075.3 | 98.83 | 98.8972
21.4 | 0 | 100 | 66.28 | 163856.4 | 98.77 | 98.7779
22.6 | 0 | 100 | 66.1 | 163433.9 | 98.45 | 98.5097
23.8 | 0 | 100 | 66.01 | 163189.2 | 98.38 | 98.3756
25 | 0 | 100 | 65.99 | 163152.9 | 98.27 | 98.3458
26.2 | 0 | 100 | 66.05 | 163310.6 | 98.48 | 98.4352
27.4 | 0 | 100 | 65.85 | 162796.9 | 98.16 | 98.1371
28.7 | 0 | 100 | 65.75 | 162550.3 | 97.94 | 97.9881
29.8 | 0 | 100 | 65.63 | 162264.2 | 97.81 | 97.8092
31 | 0 | 100 | 65.44 | 161779.4 | 97.41 | 97.5261
32.3 | 0 | 100 | 65.28 | 161400.1 | 97.26 | 97.2876
33.4 | 0 | 100 | 65.22 | 161236.3 | 97.24 | 97.1982
34.7 | 0 | 100 | 65.01 | 160737.9 | 97.07 | 96.8852
35.9 | 0 | 100 | 64.78 | 160150.6 | 96.5 | 96.5425
37.1 | 0 | 100 | 64.71 | 159980.3 | 96.4 | 96.4382
38.3 | 0 | 100 | 64.61 | 159749.8 | 96.24 | 96.2891
39.5 | 0 | 100 | 64.35 | 159099 | 96 | 95.9016
40.7 | 0 | 100 | 64.18 | 158674.8 | 95.71 | 95.6483
41.9 | 0 | 100 | 63.85 | 157854.8 | 95.26 | 95.1565
43.1 | 0 | 100 | 63.7 | 157493.9 | 94.9 | 94.9329
44.3 | 0 | 100 | 63.36 | 156657.5 | 94.59 | 94.4262
45.5 | 0 | 100 | 62.85 | 155375 | 93.58 | 93.6662
46.7 | 0 | 100 | 62.32 | 154069.3 | 92.74 | 92.8763
47.9 | 0 | 100 | 61.3 | 151555.5 | 91.37 | 91.3562
49.1 | 0 | 100 | 59.78 | 147802.8 | 89.09 | 89.0909
50.3 | 0 | 100 | 56.9 | 140672.5 | 84.87 | 84.7988
51.5 | 0 | 100 | 52.65 | 130164.1 | 78.45 | 78.465
52.7 | 0 | 100 | 47.05 | 116319.4 | 70.12 | 70.1192
53.9 | 0 | 100 | 39.68 | 98110 | 59.23 | 59.1356
55.1 | 0 | 100 | 31.62 | 78172.9 | 47.16 | 47.1237
56.3 | 0 | 100 | 23.99 | 59312.8 | 35.79 | 35.7526
57.6 | 0 | 100 | 17.64 | 43602.6 | 26.32 | 26.2891
58.8 | 0 | 100 | 13.28 | 32838.7 | 19.77 | 19.7914
60 | 0 | 100 | 10.87 | 26864.8 | 16.21 | 16.1997
61.2 | 0 | 100 | 9.49 | 23454.1 | 14.14 | 14.1431
62.4 | 0 | 100 | 8.66 | 21415.8 | 12.9 | 12.9061
63.6 | 0 | 100 | 8.03 | 19856.9 | 11.96 | 11.9672
64.8 | 0 | 100 | 7.56 | 18702.2 | 11.26 | 11.2668
66 | 0 | 100 | 7.19 | 17781 | 10.73 | 10.7154
67.2 | 0 | 100 | 6.86 | 16957.3 | 10.22 | 10.2235
68.4 | 0 | 100 | 6.54 | 16175 | 9.75 | 9.7466
69.6 | 0 | 100 | 6.37 | 15741.7 | 9.49 | 9.4933
70.8 | 0 | 100 | 6.06 | 14971 | 9.02 | 9.0313
72 | 0 | 100 | 5.86 | 14493.3 | 8.74 | 8.7332
73.2 | 0 | 100 | 5.65 | 13979.3 | 8.44 | 8.4203
74.4 | 0 | 100 | 5.48 | 13539.2 | 8.17 | 8.1669
75.6 | 0 | 100 | 5.28 | 13061.8 | 7.87 | 7.8689
76.8 | 0 | 100 | 5.08 | 12551.8 | 7.55 | 7.5708
78.1 | 0 | 100 | 4.88 | 12069.3 | 7.28 | 7.2727
79.3 | 0 | 100 | 4.79 | 11833 | 7.14 | 7.1386
80.5 | 0 | 100 | 4.57 | 11288.3 | 6.81 | 6.8107
81.7 | 0 | 100 | 4.41 | 10906.6 | 6.58 | 6.5723
82.9 | 0 | 100 | 4.26 | 10542.5 | 6.37 | 6.3487
84.1 | 0 | 100 | 4.18 | 10343.8 | 6.24 | 6.2295
85.3 | 0 | 100 | 4.04 | 9992.8 | 6.03 | 6.0209
86.5 | 0 | 100 | 3.92 | 9685.3 | 5.83 | 5.842
87.7 | 0 | 100 | 3.84 | 9498.3 | 5.73 | 5.7228
88.9 | 0 | 100 | 3.67 | 9064.5 | 5.47 | 5.4694
90.1 | 0 | 100 | 3.51 | 8684.6 | 5.24 | 5.231
91.3 | 0 | 100 | 3.43 | 8491.5 | 5.11 | 5.1118
92.5 | 0 | 100 | 3.34 | 8265.1 | 4.99 | 4.9776
93.7 | 0 | 100 | 3.25 | 8032.2 | 4.85 | 4.8435
94.9 | 0 | 100 | 3.18 | 7874.3 | 4.75 | 4.7392
96.1 | 0 | 100 | 3.04 | 7527.2 | 4.55 | 4.5306
97.3 | 0 | 100 | 2.93 | 7240.6 | 4.37 | 4.3666
98.5 | 0 | 100 | 2.93 | 7255.9 | 4.37 | 4.3666
100 | 0 | 100 | 2.81 | 6937.7 | 4.18 | 4.1878
Measurement time: | 12/08/2022 18:18:10
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
-100 | 0 | 100 | 2.2 | 5430.7 | 3.3 | 2.6793
-99.1 | 0 | 100 | 2.41 | 5946.4 | 2.93 | 2.9351
-97.6 | 0 | 100 | 2.56 | 6333.4 | 3.13 | 3.1178
-96.6 | 0 | 100 | 2.61 | 6453.4 | 3.18 | 3.1787
-95.4 | 0 | 100 | 2.69 | 6643.4 | 3.28 | 3.2761
-94.1 | 0 | 100 | 2.81 | 6954.2 | 3.43 | 3.4222
-93 | 0 | 100 | 2.91 | 7200.5 | 3.55 | 3.544
-91.8 | 0 | 100 | 3.01 | 7447.9 | 3.68 | 3.6658
-90.5 | 0 | 100 | 3.17 | 7830.4 | 3.85 | 3.8607
-89.4 | 0 | 100 | 3.22 | 7952 | 3.93 | 3.9216
-88.1 | 0 | 100 | 3.33 | 8220.9 | 4.06 | 4.0555
-86.9 | 0 | 100 | 3.46 | 8551.1 | 4.21 | 4.2139
-85.7 | 0 | 100 | 3.56 | 8809.9 | 4.34 | 4.3356
-84.5 | 0 | 100 | 3.73 | 9220.1 | 4.53 | 4.5427
-83.3 | 0 | 100 | 3.84 | 9502.5 | 4.7 | 4.6767
-82.1 | 0 | 100 | 3.99 | 9866.7 | 4.85 | 4.8593
-80.9 | 0 | 100 | 4.2 | 10381.3 | 5.12 | 5.1151
-79.7 | 0 | 100 | 4.38 | 10837.3 | 5.34 | 5.3343
-78.5 | 0 | 100 | 4.51 | 11161.1 | 5.51 | 5.4926
-77.3 | 0 | 100 | 4.72 | 11681.7 | 5.77 | 5.7484
-76.1 | 0 | 100 | 4.94 | 12215.4 | 5.99 | 6.0163
-74.9 | 0 | 100 | 5.14 | 12716.9 | 6.29 | 6.2599
-73.7 | 0 | 100 | 5.48 | 13546.9 | 6.69 | 6.674
-72.5 | 0 | 100 | 5.86 | 14482.4 | 7.14 | 7.1368
-71.3 | 0 | 100 | 6.2 | 15326.8 | 7.58 | 7.5508
-70.1 | 0 | 100 | 6.72 | 16608.4 | 8.18 | 8.1841
-68.9 | 0 | 100 | 7.19 | 17770.2 | 8.8 | 8.7565
-67.7 | 0 | 100 | 7.79 | 19263.2 | 9.53 | 9.4873
-66.5 | 0 | 100 | 8.61 | 21285.7 | 10.52 | 10.4859
-65.3 | 0 | 100 | 9.61 | 23758.6 | 11.74 | 11.7038
-64.1 | 0 | 100 | 10.78 | 26660.5 | 13.15 | 13.1287
-62.9 | 0 | 100 | 12.18 | 30104.8 | 14.9 | 14.8338
-61.6 | 0 | 100 | 14.27 | 35286.2 | 17.37 | 17.3791
-60.4 | 0 | 100 | 17.01 | 42052.5 | 20.75 | 20.7161
-59.2 | 0 | 100 | 20.85 | 51543 | 25.45 | 25.3928
-58 | 0 | 100 | 26.21 | 64789.7 | 32.18 | 31.9206
-56.8 | 0 | 100 | 33.33 | 82399.2 | 40.65 | 40.5919
-55.6 | 0 | 100 | 42.07 | 103999.5 | 51.38 | 51.2361
-54.4 | 0 | 100 | 51.26 | 126735.7 | 62.67 | 62.4284
-53.2 | 0 | 100 | 59.56 | 147250.7 | 72.67 | 72.5368
-52 | 0 | 100 | 65.97 | 163103.7 | 80.69 | 80.3434
-50.8 | 0 | 100 | 70.08 | 173256 | 85.8 | 85.3489
-49.6 | 0 | 100 | 72.99 | 180449.1 | 89.21 | 88.8929
-48.4 | 0 | 100 | 75.22 | 185961.1 | 91.61 | 91.6088
-47.2 | 0 | 100 | 76.57 | 189298.5 | 93.52 | 93.253
-46 | 0 | 100 | 77.55 | 191721 | 95.01 | 94.4465
-44.8 | 0 | 100 | 78.64 | 194433.8 | 96.15 | 95.774
-43.6 | 0 | 100 | 79.19 | 195785 | 96.72 | 96.4438
-42.4 | 0 | 100 | 79.87 | 197473.5 | 97.35 | 97.272
-41.2 | 0 | 100 | 80.2 | 198293.9 | 98.11 | 97.6739
-40 | 0 | 100 | 80.53 | 199093.8 | 98.35 | 98.0758
-38.8 | 0 | 100 | 81 | 200260.2 | 99.04 | 98.6482
-37.6 | 0 | 100 | 81.15 | 200625.7 | 98.97 | 98.8308
-36.4 | 0 | 100 | 81.37 | 201183.7 | 99.44 | 99.0988
-35.2 | 0 | 100 | 81.67 | 201910.8 | 99.66 | 99.4641
-34 | 0 | 100 | 81.64 | 201841.1 | 99.78 | 99.4276
-32.8 | 0 | 100 | 81.51 | 201517.3 | 99.73 | 99.2693
-31.6 | 0 | 100 | 81.77 | 202155.6 | 99.89 | 99.5859
-30.4 | 0 | 100 | 81.89 | 202452.6 | 99.99 | 99.7321
-29.2 | 0 | 100 | 81.92 | 202526.1 | 99.93 | 99.7686
-28 | 0 | 100 | 81.66 | 201893.5 | 99.8 | 99.452
-26.7 | 0 | 100 | 82.11 | 203005.2 | 99.9 | 100
-25.5 | 0 | 100 | 81.57 | 201675.9 | 100 | 99.3423
-24.3 | 0 | 100 | 81.79 | 202201.9 | 99.73 | 99.6103
-23.1 | 0 | 100 | 81.48 | 201453.6 | 99.83 | 99.2327
-21.9 | 0 | 100 | 81.67 | 201916.3 | 99.5 | 99.4641
-20.7 | 0 | 100 | 81.68 | 201931.8 | 99.78 | 99.4763
-19.5 | 0 | 100 | 81.37 | 201185.1 | 99.54 | 99.0988
-18.3 | 0 | 100 | 81.63 | 201806.3 | 99.81 | 99.4154
-17.1 | 0 | 100 | 81.53 | 201580.1 | 99.49 | 99.2936
-15.9 | 0 | 100 | 81.4 | 201253.9 | 99.78 | 99.1353
-14.7 | 0 | 100 | 81.46 | 201388 | 99.63 | 99.2084
-13.5 | 0 | 100 | 81.34 | 201106.2 | 99.53 | 99.0622
-12.3 | 0 | 100 | 81.51 | 201525.7 | 99.64 | 99.2693
-11.1 | 0 | 100 | 81.55 | 201627.6 | 99.38 | 99.318
-9.9 | 0 | 100 | 81.5 | 201489.5 | 99.31 | 99.2571
-8.7 | 0 | 100 | 81.16 | 200650.3 | 99.37 | 98.843
-7.5 | 0 | 100 | 81.33 | 201087.4 | 99.33 | 99.0501
-6.3 | 0 | 100 | 81.01 | 200293 | 98.99 | 98.6603
-5.1 | 0 | 100 | 81.35 | 201130.6 | 99.08 | 99.0744
-3.9 | 0 | 100 | 80.91 | 200032.2 | 99.22 | 98.5385
-2.7 | 0 | 100 | 81.24 | 200842.8 | 99.17 | 98.9404
-1.4 | 0 | 100 | 80.88 | 199967.2 | 99.16 | 98.502
-0.3 | 0 | 100 | 81.26 | 200908.9 | 99.11 | 98.9648
0.9 | 0 | 100 | 80.9 | 200007.8 | 98.95 | 98.5264
2.1 | 0 | 100 | 81.11 | 200521 | 98.89 | 98.7821
3.4 | 0 | 100 | 80.82 | 199812.9 | 99.13 | 98.4289
4.6 | 0 | 100 | 81.2 | 200754.2 | 99.16 | 98.8917
5.8 | 0 | 100 | 80.9 | 200001.6 | 99.01 | 98.5264
7 | 0 | 100 | 81.17 | 200687.6 | 99.02 | 98.8552
8.2 | 0 | 100 | 80.82 | 199814.1 | 98.98 | 98.4289
9.4 | 0 | 100 | 81.38 | 201211.6 | 99.29 | 99.1109
10.6 | 0 | 100 | 80.95 | 200146 | 99.28 | 98.5873
11.8 | 0 | 100 | 81.25 | 200877 | 99.22 | 98.9526
13 | 0 | 100 | 81.05 | 200396.2 | 99.24 | 98.709
14.2 | 0 | 100 | 80.94 | 200118.9 | 99.06 | 98.5751
15.4 | 0 | 100 | 81.23 | 200820 | 99.12 | 98.9283
16.6 | 0 | 100 | 80.68 | 199464.6 | 98.83 | 98.2584
17.8 | 0 | 100 | 81.11 | 200543 | 98.87 | 98.7821
19 | 0 | 100 | 80.67 | 199441.8 | 98.98 | 98.2463
20.2 | 0 | 100 | 81.02 | 200311.7 | 98.95 | 98.6725
21.4 | 0 | 100 | 81.17 | 200672.9 | 98.87 | 98.8552
22.6 | 0 | 100 | 80.69 | 199505.9 | 99.09 | 98.2706
23.8 | 0 | 100 | 80.95 | 200144.4 | 98.98 | 98.5873
25 | 0 | 100 | 80.95 | 200139.5 | 98.93 | 98.5873
26.2 | 0 | 100 | 80.76 | 199659 | 98.86 | 98.3559
27.4 | 0 | 100 | 81.04 | 200352.6 | 98.91 | 98.6969
28.6 | 0 | 100 | 80.83 | 199850.4 | 99.33 | 98.4411
29.8 | 0 | 100 | 80.91 | 200045.4 | 99.15 | 98.5385
31 | 0 | 100 | 80.88 | 199957 | 99.23 | 98.502
32.2 | 0 | 100 | 81.06 | 200416.4 | 99.34 | 98.7212
33.4 | 0 | 100 | 80.95 | 200139.7 | 99 | 98.5873
34.6 | 0 | 100 | 80.58 | 199232.6 | 98.91 | 98.1366
35.8 | 0 | 100 | 80.79 | 199737.2 | 98.69 | 98.3924
37 | 0 | 100 | 80.34 | 198628.2 | 98.52 | 97.8444
38.2 | 0 | 100 | 80.24 | 198383.3 | 98.22 | 97.7226
39.4 | 0 | 100 | 79.57 | 196717.2 | 97.87 | 96.9066
40.6 | 0 | 100 | 79.38 | 196254 | 97.38 | 96.6752
41.8 | 0 | 100 | 78.87 | 194995.2 | 96.49 | 96.0541
43.1 | 0 | 100 | 78.62 | 194385.9 | 95.85 | 95.7496
44.3 | 0 | 100 | 77.26 | 191019.5 | 94.79 | 94.0933
45.5 | 0 | 100 | 76.38 | 188840.1 | 93.79 | 93.0216
46.7 | 0 | 100 | 75.28 | 186117.3 | 92.07 | 91.6819
47.9 | 0 | 100 | 73.47 | 181640.9 | 90.1 | 89.4775
49.1 | 0 | 100 | 71.25 | 176145.3 | 87.04 | 86.7738
50.3 | 0 | 100 | 67.56 | 167034.2 | 83.01 | 82.2799
51.5 | 0 | 100 | 63.09 | 155990.8 | 77.22 | 76.836
52.7 | 0 | 100 | 56.87 | 140614.6 | 69.64 | 69.2607
53.9 | 0 | 100 | 49.1 | 121403.2 | 59.97 | 59.7978
55.1 | 0 | 100 | 39.99 | 98870.1 | 49.05 | 48.703
56.3 | 0 | 100 | 31.56 | 78035.8 | 38.53 | 38.4362
57.5 | 0 | 100 | 24.16 | 59728 | 29.62 | 29.4239
58.7 | 0 | 100 | 18.83 | 46544.6 | 23.06 | 22.9327
59.9 | 0 | 100 | 15.51 | 38343.7 | 18.99 | 18.8893
61.1 | 0 | 100 | 13.09 | 32354.3 | 15.99 | 15.942
62.3 | 0 | 100 | 11.32 | 27991.9 | 13.84 | 13.7864
63.5 | 0 | 100 | 10.03 | 24787 | 12.28 | 12.2153
64.7 | 0 | 100 | 8.95 | 22134.5 | 10.95 | 10.9
65.9 | 0 | 100 | 8.04 | 19871.6 | 9.87 | 9.7917
67.1 | 0 | 100 | 7.41 | 18316.6 | 9.06 | 9.0245
68.3 | 0 | 100 | 6.85 | 16932.1 | 8.4 | 8.3425
69.5 | 0 | 100 | 6.39 | 15792.6 | 7.81 | 7.7822
70.8 | 0 | 100 | 5.9 | 14583.4 | 7.22 | 7.1855
72 | 0 | 100 | 5.58 | 13788.3 | 6.82 | 6.7958
73.2 | 0 | 100 | 5.3 | 13104.8 | 6.49 | 6.4548
74.3 | 0 | 100 | 4.99 | 12335.8 | 6.11 | 6.0772
75.5 | 0 | 100 | 4.75 | 11742.4 | 5.8 | 5.7849
76.8 | 0 | 100 | 4.51 | 11162.1 | 5.55 | 5.4926
78 | 0 | 100 | 4.35 | 10745.7 | 5.33 | 5.2978
79.2 | 0 | 100 | 4.18 | 10331.6 | 5.12 | 5.0907
80.4 | 0 | 100 | 3.99 | 9855.1 | 4.87 | 4.8593
81.6 | 0 | 100 | 3.88 | 9597.3 | 4.75 | 4.7254
82.8 | 0 | 100 | 3.69 | 9118.8 | 4.53 | 4.494
84 | 0 | 100 | 3.56 | 8794.5 | 4.35 | 4.3356
85.2 | 0 | 100 | 3.47 | 8577.2 | 4.25 | 4.226
86.4 | 0 | 100 | 3.32 | 8213.5 | 4.06 | 4.0434
87.6 | 0 | 100 | 3.25 | 8046 | 3.99 | 3.9581
88.8 | 0 | 100 | 3.1 | 7665.1 | 3.8 | 3.7754
90 | 0 | 100 | 2.99 | 7388.1 | 3.66 | 3.6415
91.2 | 0 | 100 | 2.9 | 7170.3 | 3.54 | 3.5318
92.4 | 0 | 100 | 2.78 | 6868.8 | 3.41 | 3.3857
93.6 | 0 | 100 | 2.71 | 6696.2 | 3.32 | 3.3005
94.8 | 0 | 100 | 2.62 | 6487.5 | 3.22 | 3.1908
96 | 0 | 100 | 2.53 | 6260.5 | 3.1 | 3.0812
97.2 | 0 | 100 | 2.43 | 6016.2 | 2.99 | 2.9594
98.4 | 0 | 100 | 2.4 | 5923.3 | 2.94 | 2.9229
100 | 0 | 100 | 2.25 | 5564.8 | 2.76 | 2.7402