## 3_ Profils orientation
Measurement time: | 12/08/2022 18:20:49
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
99.1 | 0 | 0 | 6.34 | 15669.3 | 19.96 | 19.9811
97.6 | 0 | 0 | 6.45 | 15936.1 | 20.26 | 20.3278
96.5 | 0 | 0 | 6.61 | 16336.9 | 20.78 | 20.832
95.3 | 0 | 0 | 6.7 | 16567.4 | 21.21 | 21.1157
94.1 | 0 | 0 | 6.79 | 16777.1 | 21.4 | 21.3993
92.9 | 0 | 0 | 7.01 | 17331.5 | 22.05 | 22.0927
91.7 | 0 | 0 | 7.06 | 17447.9 | 22.29 | 22.2502
90.5 | 0 | 0 | 7.22 | 17861.2 | 22.81 | 22.7545
89.3 | 0 | 0 | 7.34 | 18153.5 | 23.15 | 23.1327
88 | 0 | 0 | 7.51 | 18556.5 | 23.6 | 23.6685
86.8 | 0 | 0 | 7.58 | 18730.6 | 23.92 | 23.8891
85.6 | 0 | 0 | 7.78 | 19226.9 | 24.45 | 24.5194
84.4 | 0 | 0 | 7.86 | 19425.1 | 24.88 | 24.7715
83.2 | 0 | 0 | 7.92 | 19586.1 | 24.95 | 24.9606
82 | 0 | 0 | 8.21 | 20300.7 | 25.83 | 25.8746
80.8 | 0 | 0 | 8.24 | 20374 | 26.07 | 25.9691
79.6 | 0 | 0 | 8.46 | 20903.8 | 26.6 | 26.6625
78.4 | 0 | 0 | 8.59 | 21246.2 | 27.07 | 27.0722
77.2 | 0 | 0 | 8.74 | 21618.7 | 27.56 | 27.5449
76 | 0 | 0 | 8.82 | 21805.2 | 27.81 | 27.797
74.8 | 0 | 0 | 9 | 22259.6 | 28.47 | 28.3643
73.6 | 0 | 0 | 9.16 | 22639.9 | 28.94 | 28.8686
72.4 | 0 | 0 | 9.37 | 23178.2 | 29.45 | 29.5304
71.2 | 0 | 0 | 9.45 | 23357.1 | 29.79 | 29.7825
70 | 0 | 0 | 9.66 | 23878.1 | 30.44 | 30.4444
68.8 | 0 | 0 | 9.83 | 24295.6 | 30.86 | 30.9801
67.6 | 0 | 0 | 9.89 | 24441.9 | 31.16 | 31.1692
66.4 | 0 | 0 | 10.21 | 25236.5 | 32.09 | 32.1777
65.2 | 0 | 0 | 10.38 | 25674.8 | 32.64 | 32.7135
63.9 | 0 | 0 | 10.52 | 26001.3 | 33.21 | 33.1547
62.8 | 0 | 0 | 10.72 | 26495.3 | 33.72 | 33.7851
61.5 | 0 | 0 | 10.93 | 27034.4 | 34.51 | 34.4469
60.3 | 0 | 0 | 11.05 | 27326.9 | 34.9 | 34.8251
59.2 | 0 | 0 | 11.33 | 28005.5 | 35.61 | 35.7075
57.9 | 0 | 0 | 11.57 | 28615.8 | 36.5 | 36.4639
56.7 | 0 | 0 | 11.82 | 29233 | 37.19 | 37.2518
55.5 | 0 | 0 | 12.15 | 30034 | 38.28 | 38.2918
54.3 | 0 | 0 | 12.49 | 30884.3 | 39.31 | 39.3634
53.1 | 0 | 0 | 13.49 | 33352.4 | 42.49 | 42.515
51.9 | 0 | 0 | 15.42 | 38125.3 | 48.68 | 48.5975
50.7 | 0 | 0 | 18.18 | 44951.1 | 57.16 | 57.2959
49.5 | 0 | 0 | 20.94 | 51766.9 | 65.97 | 65.9943
48.3 | 0 | 0 | 23.94 | 59187.2 | 75.43 | 75.4491
47.1 | 0 | 0 | 25.76 | 63689.3 | 81.34 | 81.185
45.9 | 0 | 0 | 27.63 | 68303.9 | 87.07 | 87.0785
44.7 | 0 | 0 | 28.75 | 71088.2 | 90.66 | 90.6083
43.5 | 0 | 0 | 29.35 | 72552.5 | 92.59 | 92.4992
42.2 | 0 | 0 | 29.85 | 73789.2 | 94.03 | 94.075
41.1 | 0 | 0 | 29.93 | 73995.4 | 94.75 | 94.3271
39.8 | 0 | 0 | 30.34 | 75022 | 95.61 | 95.6193
38.7 | 0 | 0 | 30.49 | 75388.3 | 95.95 | 96.092
37.5 | 0 | 0 | 30.42 | 75207.9 | 96.18 | 95.8714
36.3 | 0 | 0 | 30.69 | 75885.4 | 96.79 | 96.7223
35.1 | 0 | 0 | 30.66 | 75812.1 | 96.75 | 96.6278
33.8 | 0 | 0 | 30.96 | 76551.8 | 97.32 | 97.5733
32.7 | 0 | 0 | 30.96 | 76546 | 97.66 | 97.5733
31.4 | 0 | 0 | 30.95 | 76526.4 | 97.97 | 97.5418
30.2 | 0 | 0 | 30.99 | 76627 | 98.09 | 97.6678
29.1 | 0 | 0 | 31.04 | 76730.1 | 97.91 | 97.8254
27.8 | 0 | 0 | 31.24 | 77236.3 | 98.17 | 98.4557
26.7 | 0 | 0 | 31.24 | 77230.6 | 98.84 | 98.4557
25.4 | 0 | 0 | 31.09 | 76875.9 | 98.25 | 97.983
24.2 | 0 | 0 | 31.13 | 76964.4 | 98.09 | 98.109
23.1 | 0 | 0 | 31.28 | 77342.1 | 98.34 | 98.5818
21.8 | 0 | 0 | 31.29 | 77371.3 | 98.51 | 98.6133
20.6 | 0 | 0 | 31.2 | 77138.7 | 98.79 | 98.3297
19.4 | 0 | 0 | 31.24 | 77243.9 | 98.63 | 98.4557
18.2 | 0 | 0 | 31.3 | 77375.6 | 98.6 | 98.6448
17 | 0 | 0 | 31.42 | 77677.2 | 98.85 | 99.023
15.8 | 0 | 0 | 31.33 | 77470.5 | 98.76 | 98.7394
14.6 | 0 | 0 | 31.3 | 77383.3 | 99 | 98.6448
13.4 | 0 | 0 | 31.46 | 77791.6 | 99.31 | 99.1491
12.2 | 0 | 0 | 31.46 | 77773.1 | 99.27 | 99.1491
11 | 0 | 0 | 31.4 | 77626.4 | 98.9 | 98.96
9.7 | 0 | 0 | 31.58 | 78069.2 | 99.55 | 99.5273
8.6 | 0 | 0 | 31.44 | 77726.6 | 99.17 | 99.086
7.4 | 0 | 0 | 31.58 | 78082.4 | 99.67 | 99.5273
6.2 | 0 | 0 | 31.65 | 78239.6 | 99.24 | 99.7479
5 | 0 | 0 | 31.5 | 77887.5 | 99.67 | 99.2751
3.7 | 0 | 0 | 31.52 | 77916.9 | 99.47 | 99.3382
2.6 | 0 | 0 | 31.7 | 78368.6 | 99.65 | 99.9055
1.3 | 0 | 0 | 31.44 | 77739.5 | 99.46 | 99.086
0.2 | 0 | 0 | 31.52 | 77937.9 | 99.49 | 99.3382
-1 | 0 | 0 | 31.63 | 78197.8 | 99.66 | 99.6848
-2.3 | 0 | 0 | 31.54 | 77972.3 | 99.85 | 99.4012
-3.4 | 0 | 0 | 31.73 | 78450.8 | 99.69 | 100
-4.7 | 0 | 0 | 31.47 | 77806.4 | 99.5 | 99.1806
-5.9 | 0 | 0 | 31.69 | 78352.4 | 99.66 | 99.8739
-7.1 | 0 | 0 | 31.61 | 78143.6 | 99.74 | 99.6218
-8.3 | 0 | 0 | 31.6 | 78130.8 | 99.63 | 99.5903
-9.5 | 0 | 0 | 31.55 | 77992 | 99.77 | 99.4327
-10.7 | 0 | 0 | 31.58 | 78072.1 | 99.31 | 99.5273
-11.9 | 0 | 0 | 31.6 | 78118.9 | 100 | 99.5903
-13.1 | 0 | 0 | 31.69 | 78351.2 | 99.73 | 99.8739
-14.3 | 0 | 0 | 31.55 | 78007.6 | 99.72 | 99.4327
-15.5 | 0 | 0 | 31.53 | 77949.1 | 99.56 | 99.3697
-16.7 | 0 | 0 | 31.43 | 77706.2 | 99.56 | 99.0545
-17.9 | 0 | 0 | 31.47 | 77803.1 | 99.17 | 99.1806
-19.1 | 0 | 0 | 31.5 | 77876.9 | 99.51 | 99.2751
-20.4 | 0 | 0 | 31.61 | 78155.6 | 99.43 | 99.6218
-21.5 | 0 | 0 | 31.5 | 77866.8 | 99.4 | 99.2751
-22.7 | 0 | 0 | 31.43 | 77706 | 99.27 | 99.0545
-24 | 0 | 0 | 31.42 | 77670.1 | 99.05 | 99.023
-25.1 | 0 | 0 | 31.5 | 77880.7 | 99.43 | 99.2751
-26.4 | 0 | 0 | 31.38 | 77571.8 | 99 | 98.8969
-27.5 | 0 | 0 | 31.46 | 77771.4 | 98.95 | 99.1491
-28.8 | 0 | 0 | 31.38 | 77574.9 | 98.86 | 98.8969
-30 | 0 | 0 | 31.39 | 77609.2 | 99.01 | 98.9285
-31.2 | 0 | 0 | 31.3 | 77376.7 | 98.7 | 98.6448
-32.4 | 0 | 0 | 31.21 | 77151.1 | 98.63 | 98.3612
-33.5 | 0 | 0 | 31.26 | 77284.8 | 98.54 | 98.5188
-34.8 | 0 | 0 | 31.16 | 77039.3 | 98.18 | 98.2036
-36 | 0 | 0 | 31.14 | 76977.3 | 98.01 | 98.1406
-37.1 | 0 | 0 | 30.8 | 76138.7 | 97.38 | 97.069
-38.4 | 0 | 0 | 30.89 | 76377.9 | 97.1 | 97.3527
-39.6 | 0 | 0 | 30.68 | 75863.6 | 96.8 | 96.6908
-40.8 | 0 | 0 | 30.63 | 75719.9 | 96.31 | 96.5332
-42 | 0 | 0 | 30.51 | 75425.5 | 96.04 | 96.1551
-43.2 | 0 | 0 | 30.2 | 74658.6 | 95.66 | 95.1781
-44.4 | 0 | 0 | 30.14 | 74522 | 94.88 | 94.989
-45.6 | 0 | 0 | 29.81 | 73698.5 | 94.03 | 93.9489
-46.8 | 0 | 0 | 29.03 | 71774.5 | 91.48 | 91.4907
-48 | 0 | 0 | 27.24 | 67341.4 | 85.63 | 85.8494
-49.2 | 0 | 0 | 24.65 | 60938.8 | 78 | 77.6867
-50.4 | 0 | 0 | 21.9 | 54154.4 | 68.82 | 69.0199
-51.6 | 0 | 0 | 19.08 | 47171.1 | 60.32 | 60.1324
-52.8 | 0 | 0 | 16.78 | 41497.1 | 53.02 | 52.8837
-54 | 0 | 0 | 14.95 | 36952.2 | 47.08 | 47.1163
-55.2 | 0 | 0 | 13.5 | 33375.8 | 42.66 | 42.5465
-56.4 | 0 | 0 | 12.61 | 31176.7 | 39.75 | 39.7416
-57.6 | 0 | 0 | 12.2 | 30151.4 | 38.4 | 38.4494
-58.8 | 0 | 0 | 11.81 | 29204.7 | 37.32 | 37.2203
-60 | 0 | 0 | 11.55 | 28546.6 | 36.39 | 36.4009
-61.2 | 0 | 0 | 11.21 | 27708.3 | 35.37 | 35.3293
-62.5 | 0 | 0 | 11.03 | 27267.7 | 34.64 | 34.7621
-63.6 | 0 | 0 | 10.78 | 26652.5 | 34.03 | 33.9742
-64.9 | 0 | 0 | 10.63 | 26274.5 | 33.46 | 33.5014
-66.1 | 0 | 0 | 10.39 | 25687.1 | 32.72 | 32.745
-67.3 | 0 | 0 | 10.25 | 25332.9 | 32.25 | 32.3038
-68.5 | 0 | 0 | 10.03 | 24788.4 | 31.54 | 31.6105
-69.7 | 0 | 0 | 9.82 | 24279 | 31.11 | 30.9486
-70.9 | 0 | 0 | 9.68 | 23941.3 | 30.47 | 30.5074
-72.1 | 0 | 0 | 9.48 | 23441.9 | 29.95 | 29.8771
-73.3 | 0 | 0 | 9.35 | 23125 | 29.51 | 29.4674
-74.5 | 0 | 0 | 9.15 | 22612.5 | 28.95 | 28.8371
-75.7 | 0 | 0 | 9.03 | 22333.8 | 28.52 | 28.4589
-76.9 | 0 | 0 | 8.9 | 21992.3 | 27.96 | 28.0492
-78.1 | 0 | 0 | 8.74 | 21613.4 | 27.58 | 27.5449
-79.3 | 0 | 0 | 8.58 | 21206 | 27 | 27.0407
-80.5 | 0 | 0 | 8.42 | 20826.9 | 26.66 | 26.5364
-81.7 | 0 | 0 | 8.31 | 20548.2 | 26.18 | 26.1897
-82.9 | 0 | 0 | 8.13 | 20091.6 | 25.68 | 25.6224
-84.1 | 0 | 0 | 7.96 | 19670.5 | 25.07 | 25.0867
-85.3 | 0 | 0 | 7.88 | 19480.2 | 24.86 | 24.8345
-86.5 | 0 | 0 | 7.76 | 19181.6 | 24.51 | 24.4564
-87.7 | 0 | 0 | 7.6 | 18792 | 23.9 | 23.9521
-88.9 | 0 | 0 | 7.4 | 18300.2 | 23.37 | 23.3218
-90.1 | 0 | 0 | 7.32 | 18105.3 | 23.05 | 23.0697
-91.3 | 0 | 0 | 7.18 | 17747.1 | 22.62 | 22.6284
-92.6 | 0 | 0 | 6.99 | 17280.5 | 22.05 | 22.0296
-93.7 | 0 | 0 | 6.94 | 17161.2 | 21.91 | 21.872
-94.9 | 0 | 0 | 6.77 | 16725.5 | 21.32 | 21.3363
-96.2 | 0 | 0 | 6.74 | 16675 | 21.25 | 21.2417
-97.3 | 0 | 0 | 6.58 | 16267.6 | 20.8 | 20.7375
-98.6 | 0 | 0 | 6.45 | 15949.7 | 20.38 | 20.3278
-100 | 0 | 0 | 6.36 | 15723.5 | 20.08 | 20.0441
Measurement time: | 12/08/2022 18:21:40
Radiation device: | Clinac2 Accelerator
Energy: | 23 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 100 | 0 | 4.38 | 10837.7 | 13.84 | 13.8739
0 | 99.1 | 0 | 4.44 | 10976.7 | 14.11 | 14.064
0 | 97.6 | 0 | 4.62 | 11414.4 | 14.67 | 14.6341
0 | 96.5 | 0 | 4.69 | 11601.5 | 14.88 | 14.8559
0 | 95.3 | 0 | 4.79 | 11833.9 | 15.21 | 15.1726
0 | 94.1 | 0 | 4.9 | 12109.9 | 15.57 | 15.5211
0 | 92.9 | 0 | 5.01 | 12389 | 15.89 | 15.8695
0 | 91.7 | 0 | 5.15 | 12728.5 | 16.32 | 16.313
0 | 90.5 | 0 | 5.29 | 13083.5 | 16.74 | 16.7564
0 | 89.3 | 0 | 5.35 | 13224.5 | 16.98 | 16.9465
0 | 88.1 | 0 | 5.46 | 13506.9 | 17.37 | 17.2949
0 | 86.8 | 0 | 5.61 | 13864.7 | 17.75 | 17.77
0 | 85.7 | 0 | 5.79 | 14308.5 | 18.43 | 18.3402
0 | 84.5 | 0 | 5.85 | 14466.8 | 18.59 | 18.5303
0 | 83.2 | 0 | 6.07 | 15004.7 | 19.28 | 19.2271
0 | 82.1 | 0 | 6.18 | 15283 | 19.61 | 19.5755
0 | 80.8 | 0 | 6.29 | 15556.8 | 20 | 19.924
0 | 79.6 | 0 | 6.44 | 15920.6 | 20.45 | 20.3991
0 | 78.4 | 0 | 6.62 | 16354.7 | 20.96 | 20.9693
0 | 77.2 | 0 | 6.74 | 16660.1 | 21.39 | 21.3494
0 | 76 | 0 | 6.94 | 17169.2 | 22.02 | 21.9829
0 | 74.8 | 0 | 7.09 | 17522.1 | 22.51 | 22.458
0 | 73.6 | 0 | 7.31 | 18076.6 | 23.12 | 23.1549
0 | 72.4 | 0 | 7.44 | 18398.7 | 23.71 | 23.5667
0 | 71.2 | 0 | 7.64 | 18880.5 | 24.16 | 24.2002
0 | 70 | 0 | 7.82 | 19329.9 | 24.79 | 24.7704
0 | 68.8 | 0 | 8 | 19786 | 25.39 | 25.3405
0 | 67.6 | 0 | 8.18 | 20233.2 | 25.96 | 25.9107
0 | 66.4 | 0 | 8.36 | 20679 | 26.61 | 26.4808
0 | 65.2 | 0 | 8.62 | 21318.6 | 27.29 | 27.3044
0 | 64 | 0 | 8.8 | 21749.1 | 27.89 | 27.8746
0 | 62.8 | 0 | 9.03 | 22328.7 | 28.61 | 28.6031
0 | 61.6 | 0 | 9.25 | 22864.9 | 29.47 | 29.3
0 | 60.4 | 0 | 9.51 | 23510.3 | 30.16 | 30.1235
0 | 59.2 | 0 | 9.69 | 23946.2 | 30.7 | 30.6937
0 | 58 | 0 | 9.93 | 24541 | 31.58 | 31.4539
0 | 56.8 | 0 | 10.23 | 25285 | 32.55 | 32.4042
0 | 55.5 | 0 | 10.56 | 26116.4 | 33.43 | 33.4495
0 | 54.3 | 0 | 10.87 | 26865 | 34.4 | 34.4314
0 | 53.1 | 0 | 11.41 | 28213.2 | 36.12 | 36.1419
0 | 51.9 | 0 | 12.53 | 30975.4 | 39.75 | 39.6896
0 | 50.7 | 0 | 15.21 | 37603.5 | 48.35 | 48.1787
0 | 49.5 | 0 | 18 | 44501.3 | 57.17 | 57.0162
0 | 48.3 | 0 | 21.09 | 52138.4 | 66.76 | 66.8039
0 | 47.1 | 0 | 24.28 | 60033.6 | 76.85 | 76.9085
0 | 45.9 | 0 | 26.7 | 66008.9 | 85.02 | 84.574
0 | 44.7 | 0 | 28.31 | 70002.9 | 90.07 | 89.6737
0 | 43.5 | 0 | 28.88 | 71410.5 | 91.67 | 91.4793
0 | 42.3 | 0 | 29.18 | 72140 | 92.62 | 92.4295
0 | 41.1 | 0 | 29.39 | 72662.4 | 93.31 | 93.0947
0 | 39.9 | 0 | 29.57 | 73118.6 | 93.75 | 93.6649
0 | 38.7 | 0 | 29.66 | 73325.5 | 94.04 | 93.95
0 | 37.5 | 0 | 29.91 | 73939.4 | 94.71 | 94.7418
0 | 36.3 | 0 | 30.08 | 74368.8 | 95.47 | 95.2803
0 | 35.1 | 0 | 30.12 | 74463.6 | 95.82 | 95.407
0 | 33.9 | 0 | 30.36 | 75064.9 | 96.21 | 96.1672
0 | 32.7 | 0 | 30.41 | 75182.3 | 96.42 | 96.3256
0 | 31.5 | 0 | 30.45 | 75285.3 | 96.8 | 96.4523
0 | 30.3 | 0 | 30.68 | 75844.2 | 97.4 | 97.1809
0 | 29.1 | 0 | 30.68 | 75841.2 | 97.4 | 97.1809
0 | 27.9 | 0 | 30.63 | 75736.8 | 97.29 | 97.0225
0 | 26.7 | 0 | 30.84 | 76249 | 97.64 | 97.6877
0 | 25.5 | 0 | 30.78 | 76092.6 | 97.75 | 97.4976
0 | 24.2 | 0 | 30.94 | 76484.8 | 98.34 | 98.0044
0 | 23.1 | 0 | 30.99 | 76629 | 97.86 | 98.1628
0 | 21.9 | 0 | 30.97 | 76569.8 | 98.37 | 98.0995
0 | 20.6 | 0 | 31.04 | 76748.2 | 98.48 | 98.3212
0 | 19.4 | 0 | 31.03 | 76705.5 | 98.34 | 98.2895
0 | 18.2 | 0 | 31 | 76641.7 | 98.4 | 98.1945
0 | 17 | 0 | 31.15 | 77025.8 | 98.72 | 98.6696
0 | 15.8 | 0 | 31.1 | 76896.1 | 98.79 | 98.5112
0 | 14.6 | 0 | 31.23 | 77221.7 | 98.96 | 98.923
0 | 13.4 | 0 | 31.17 | 77064.4 | 99.26 | 98.733
0 | 12.2 | 0 | 31.27 | 77298.7 | 99.13 | 99.0497
0 | 11 | 0 | 31.45 | 77764.1 | 99.61 | 99.6199
0 | 9.8 | 0 | 31.18 | 77083.8 | 99.13 | 98.7646
0 | 8.6 | 0 | 31.36 | 77539.3 | 99.7 | 99.3348
0 | 7.4 | 0 | 31.45 | 77755.2 | 99.82 | 99.6199
0 | 6.2 | 0 | 31.37 | 77546.8 | 99.53 | 99.3665
0 | 5 | 0 | 31.5 | 77891 | 99.71 | 99.7783
0 | 3.7 | 0 | 31.25 | 77271 | 99.52 | 98.9864
0 | 2.6 | 0 | 31.39 | 77609.2 | 99.76 | 99.4298
0 | 1.4 | 0 | 31.46 | 77778.3 | 99.77 | 99.6516
0 | 0.2 | 0 | 31.57 | 78039.9 | 100 | 100
0 | -1 | 0 | 31.34 | 77472.3 | 99.83 | 99.2715
0 | -2.3 | 0 | 31.39 | 77613.2 | 99.62 | 99.4298
0 | -3.5 | 0 | 31.4 | 77638.3 | 99.54 | 99.4615
0 | -4.6 | 0 | 31.54 | 77976 | 99.79 | 99.905
0 | -5.9 | 0 | 31.48 | 77828.3 | 99.82 | 99.7149
0 | -7.1 | 0 | 31.34 | 77474.3 | 99.75 | 99.2715
0 | -8.3 | 0 | 31.43 | 77712 | 99.77 | 99.5565
0 | -9.5 | 0 | 31.34 | 77482.3 | 99.65 | 99.2715
0 | -10.7 | 0 | 31.42 | 77671.1 | 99.61 | 99.5249
0 | -11.9 | 0 | 31.41 | 77656.1 | 99.82 | 99.4932
0 | -13.1 | 0 | 31.33 | 77466.5 | 99.57 | 99.2398
0 | -14.3 | 0 | 31.32 | 77435.7 | 99.39 | 99.2081
0 | -15.5 | 0 | 31.34 | 77487.6 | 99.71 | 99.2715
0 | -16.7 | 0 | 31.27 | 77303.4 | 99.22 | 99.0497
0 | -17.9 | 0 | 31.3 | 77391.3 | 99.58 | 99.1448
0 | -19.1 | 0 | 31.17 | 77051.4 | 98.95 | 98.733
0 | -20.3 | 0 | 31.15 | 77006.1 | 98.91 | 98.6696
0 | -21.5 | 0 | 31.24 | 77243.1 | 99.14 | 98.9547
0 | -22.8 | 0 | 30.99 | 76630 | 98.67 | 98.1628
0 | -23.9 | 0 | 31.08 | 76837.2 | 98.71 | 98.4479
0 | -25.1 | 0 | 30.96 | 76547.9 | 98.73 | 98.0678
0 | -26.4 | 0 | 31.13 | 76957.9 | 98.4 | 98.6063
0 | -27.6 | 0 | 30.94 | 76499.1 | 98.09 | 98.0044
0 | -28.7 | 0 | 30.86 | 76308.3 | 98.14 | 97.751
0 | -30 | 0 | 30.95 | 76517.8 | 98.11 | 98.0361
0 | -31.2 | 0 | 30.87 | 76324 | 97.99 | 97.7827
0 | -32.3 | 0 | 30.85 | 76273.2 | 97.47 | 97.7194
0 | -33.6 | 0 | 30.66 | 75806.9 | 97.08 | 97.1175
0 | -34.8 | 0 | 30.42 | 75197.9 | 96.79 | 96.3573
0 | -36 | 0 | 30.46 | 75319.1 | 96.65 | 96.484
0 | -37.2 | 0 | 30.35 | 75029.9 | 96.17 | 96.1356
0 | -38.4 | 0 | 30.08 | 74362.4 | 95.48 | 95.2803
0 | -39.6 | 0 | 30.05 | 74299 | 95.4 | 95.1853
0 | -40.8 | 0 | 29.87 | 73844.2 | 95.02 | 94.6151
0 | -42 | 0 | 29.79 | 73653.9 | 94.5 | 94.3617
0 | -43.2 | 0 | 29.53 | 73001.9 | 93.62 | 93.5382
0 | -44.4 | 0 | 29.39 | 72654.5 | 93.43 | 93.0947
0 | -45.6 | 0 | 29.09 | 71930.6 | 92.42 | 92.1444
0 | -46.8 | 0 | 28.81 | 71218 | 91.56 | 91.2575
0 | -48 | 0 | 28.14 | 69581.8 | 89.23 | 89.1353
0 | -49.2 | 0 | 26.11 | 64543.9 | 83.15 | 82.7051
0 | -50.4 | 0 | 23.44 | 57953 | 74.42 | 74.2477
0 | -51.6 | 0 | 20.45 | 50553.4 | 64.84 | 64.7767
0 | -52.8 | 0 | 17.23 | 42610.9 | 54.65 | 54.5771
0 | -54 | 0 | 14.38 | 35545.1 | 45.68 | 45.5496
0 | -55.2 | 0 | 12.17 | 30100.3 | 38.77 | 38.5493
0 | -56.5 | 0 | 11.31 | 27971.2 | 35.83 | 35.8252
0 | -57.7 | 0 | 10.86 | 26858.1 | 34.54 | 34.3997
0 | -58.9 | 0 | 10.51 | 25992 | 33.3 | 33.2911
0 | -60.1 | 0 | 10.2 | 25212.5 | 32.51 | 32.3092
0 | -61.3 | 0 | 9.96 | 24636.3 | 31.63 | 31.5489
0 | -62.5 | 0 | 9.7 | 23972.1 | 30.79 | 30.7254
0 | -63.7 | 0 | 9.48 | 23440.9 | 30.12 | 30.0285
0 | -64.9 | 0 | 9.2 | 22739.1 | 29.24 | 29.1416
0 | -66.1 | 0 | 9.02 | 22299.4 | 28.65 | 28.5714
0 | -67.3 | 0 | 8.8 | 21752.2 | 27.97 | 27.8746
0 | -68.5 | 0 | 8.61 | 21274.8 | 27.27 | 27.2727
0 | -69.7 | 0 | 8.41 | 20788.8 | 26.68 | 26.6392
0 | -70.9 | 0 | 8.17 | 20196 | 25.93 | 25.879
0 | -72.1 | 0 | 8.03 | 19863.8 | 25.51 | 25.4355
0 | -73.3 | 0 | 7.83 | 19355 | 24.82 | 24.802
0 | -74.5 | 0 | 7.63 | 18856.6 | 24.2 | 24.1685
0 | -75.7 | 0 | 7.46 | 18442.8 | 23.73 | 23.63
0 | -76.9 | 0 | 7.28 | 18007.6 | 23.16 | 23.0599
0 | -78.1 | 0 | 7.14 | 17640.8 | 22.62 | 22.6164
0 | -79.3 | 0 | 6.93 | 17134.1 | 22 | 21.9512
0 | -80.5 | 0 | 6.84 | 16905.4 | 21.72 | 21.6661
0 | -81.7 | 0 | 6.63 | 16382.9 | 21.11 | 21.001
0 | -82.9 | 0 | 6.5 | 16076.2 | 20.64 | 20.5892
0 | -84.2 | 0 | 6.32 | 15618.2 | 20.07 | 20.019
0 | -85.4 | 0 | 6.2 | 15330.4 | 19.74 | 19.6389
0 | -86.5 | 0 | 6.1 | 15082 | 19.38 | 19.3221
0 | -87.8 | 0 | 5.93 | 14654 | 18.84 | 18.7837
0 | -89 | 0 | 5.8 | 14331.4 | 18.42 | 18.3719
0 | -90.1 | 0 | 5.7 | 14098.8 | 18.07 | 18.0551
0 | -91.4 | 0 | 5.52 | 13640.7 | 17.49 | 17.485
0 | -92.6 | 0 | 5.43 | 13415.9 | 17.22 | 17.1999
0 | -93.8 | 0 | 5.26 | 13009 | 16.71 | 16.6614
0 | -95 | 0 | 5.15 | 12733 | 16.41 | 16.313
0 | -96.2 | 0 | 5.06 | 12515.1 | 16.06 | 16.0279
0 | -97.4 | 0 | 4.97 | 12290.9 | 15.73 | 15.7428
0 | -98.6 | 0 | 4.81 | 11886.1 | 15.27 | 15.236
0 | -100 | 0 | 4.71 | 11639.2 | 14.91 | 14.9192
Measurement time: | 12/08/2022 18:23:00
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-100 | 0 | 0 | 3.75 | 9283.7 | 7.25 | 7.2506
-99.1 | 0 | 0 | 3.86 | 9553 | 7.45 | 7.4633
-97.6 | 0 | 0 | 4 | 9895.8 | 7.72 | 7.734
-96.6 | 0 | 0 | 4.15 | 10252.7 | 8.01 | 8.024
-95.4 | 0 | 0 | 4.17 | 10297.6 | 8.05 | 8.0626
-94.1 | 0 | 0 | 4.23 | 10465.5 | 8.18 | 8.1787
-92.9 | 0 | 0 | 4.33 | 10702.1 | 8.35 | 8.372
-91.7 | 0 | 0 | 4.38 | 10830.9 | 8.47 | 8.4687
-90.5 | 0 | 0 | 4.46 | 11038.1 | 8.62 | 8.6234
-89.3 | 0 | 0 | 4.55 | 11255.5 | 8.78 | 8.7974
-88.1 | 0 | 0 | 4.62 | 11430.3 | 8.94 | 8.9327
-86.9 | 0 | 0 | 4.77 | 11803.1 | 9.24 | 9.2227
-85.7 | 0 | 0 | 4.85 | 11984.5 | 9.36 | 9.3774
-84.5 | 0 | 0 | 5.03 | 12432.9 | 9.71 | 9.7254
-83.3 | 0 | 0 | 5.05 | 12473.1 | 9.73 | 9.7641
-82.1 | 0 | 0 | 5.17 | 12781.1 | 9.98 | 9.9961
-80.9 | 0 | 0 | 5.28 | 13059.8 | 10.21 | 10.2088
-79.7 | 0 | 0 | 5.43 | 13430.1 | 10.5 | 10.4988
-78.5 | 0 | 0 | 5.52 | 13649.6 | 10.66 | 10.6729
-77.3 | 0 | 0 | 5.61 | 13870.9 | 10.85 | 10.8469
-76.1 | 0 | 0 | 5.7 | 14102.8 | 11.01 | 11.0209
-74.9 | 0 | 0 | 5.87 | 14503.5 | 11.36 | 11.3496
-73.7 | 0 | 0 | 5.96 | 14735.7 | 11.52 | 11.5236
-72.5 | 0 | 0 | 6.12 | 15142.1 | 11.84 | 11.8329
-71.3 | 0 | 0 | 6.28 | 15528.7 | 12.13 | 12.1423
-70 | 0 | 0 | 6.41 | 15855.1 | 12.38 | 12.3937
-68.9 | 0 | 0 | 6.59 | 16304.4 | 12.76 | 12.7417
-67.7 | 0 | 0 | 6.68 | 16517 | 12.92 | 12.9157
-66.4 | 0 | 0 | 6.9 | 17051.9 | 13.32 | 13.3411
-65.2 | 0 | 0 | 7.03 | 17390.2 | 13.61 | 13.5924
-64 | 0 | 0 | 7.21 | 17827.7 | 13.92 | 13.9404
-62.8 | 0 | 0 | 7.43 | 18364 | 14.35 | 14.3658
-61.6 | 0 | 0 | 7.66 | 18933 | 14.81 | 14.8105
-60.4 | 0 | 0 | 7.83 | 19353.1 | 15.16 | 15.1392
-59.2 | 0 | 0 | 8.19 | 20258.4 | 15.84 | 15.8353
-58 | 0 | 0 | 8.53 | 21079.3 | 16.47 | 16.4927
-56.8 | 0 | 0 | 9.14 | 22591.1 | 17.66 | 17.6721
-55.6 | 0 | 0 | 10.35 | 25582.8 | 20.01 | 20.0116
-54.4 | 0 | 0 | 12.99 | 32124.7 | 25.13 | 25.116
-53.2 | 0 | 0 | 17.29 | 42757.9 | 33.43 | 33.43
-52 | 0 | 0 | 22.46 | 55541.2 | 43.46 | 43.4261
-50.8 | 0 | 0 | 28.72 | 70997.9 | 55.58 | 55.5298
-49.6 | 0 | 0 | 35.68 | 88212.5 | 69 | 68.9869
-48.4 | 0 | 0 | 42.27 | 104502 | 81.64 | 81.7285
-47.1 | 0 | 0 | 47.36 | 117094 | 91.52 | 91.57
-46 | 0 | 0 | 50 | 123627.1 | 96.84 | 96.6744
-44.8 | 0 | 0 | 50.76 | 125505.2 | 98.1 | 98.1439
-43.5 | 0 | 0 | 51.12 | 126396.8 | 98.86 | 98.8399
-42.4 | 0 | 0 | 51.1 | 126330.7 | 98.94 | 98.8012
-41.2 | 0 | 0 | 51.16 | 126491 | 98.93 | 98.9172
-39.9 | 0 | 0 | 51.14 | 126446.3 | 98.86 | 98.8786
-38.8 | 0 | 0 | 51.29 | 126817.6 | 99.21 | 99.1686
-37.5 | 0 | 0 | 51.3 | 126819.8 | 99.18 | 99.1879
-36.4 | 0 | 0 | 51.42 | 127125.5 | 99.44 | 99.42
-35.1 | 0 | 0 | 51.31 | 126851.9 | 99.3 | 99.2073
-33.9 | 0 | 0 | 51.37 | 127016.7 | 99.38 | 99.3233
-32.7 | 0 | 0 | 51.36 | 126971.4 | 99.41 | 99.3039
-31.5 | 0 | 0 | 51.43 | 127148.7 | 99.58 | 99.4393
-30.3 | 0 | 0 | 51.46 | 127231.5 | 99.41 | 99.4973
-29.1 | 0 | 0 | 51.48 | 127268.4 | 99.45 | 99.536
-27.9 | 0 | 0 | 51.58 | 127524.2 | 99.83 | 99.7293
-26.7 | 0 | 0 | 51.72 | 127868.2 | 100 | 100
-25.5 | 0 | 0 | 51.57 | 127507.2 | 99.63 | 99.71
-24.3 | 0 | 0 | 51.61 | 127594.5 | 99.85 | 99.7873
-23.1 | 0 | 0 | 51.66 | 127723.1 | 99.8 | 99.884
-21.9 | 0 | 0 | 51.65 | 127706.2 | 99.8 | 99.8647
-20.7 | 0 | 0 | 51.67 | 127734.8 | 99.97 | 99.9033
-19.5 | 0 | 0 | 51.61 | 127601 | 99.53 | 99.7873
-18.3 | 0 | 0 | 51.54 | 127426.2 | 99.63 | 99.652
-17.1 | 0 | 0 | 51.61 | 127592.7 | 99.87 | 99.7873
-15.9 | 0 | 0 | 51.47 | 127239.8 | 99.66 | 99.5166
-14.7 | 0 | 0 | 51.42 | 127129.2 | 99.52 | 99.42
-13.5 | 0 | 0 | 51.5 | 127319.6 | 99.59 | 99.5746
-12.3 | 0 | 0 | 51.35 | 126951.3 | 99.57 | 99.2846
-11 | 0 | 0 | 51.21 | 126606 | 99.05 | 99.0139
-9.8 | 0 | 0 | 51.16 | 126493.3 | 99.01 | 98.9172
-8.6 | 0 | 0 | 51.17 | 126499.8 | 98.89 | 98.9366
-7.4 | 0 | 0 | 51.24 | 126679 | 99.13 | 99.0719
-6.2 | 0 | 0 | 50.95 | 125954.6 | 98.74 | 98.5112
-5 | 0 | 0 | 50.99 | 126075.7 | 98.69 | 98.5886
-3.8 | 0 | 0 | 50.97 | 126015.7 | 98.72 | 98.5499
-2.6 | 0 | 0 | 50.98 | 126039.2 | 98.77 | 98.5692
-1.4 | 0 | 0 | 51.01 | 126104.4 | 98.6 | 98.6272
-0.2 | 0 | 0 | 51.02 | 126136.3 | 98.62 | 98.6466
1 | 0 | 0 | 50.86 | 125752.1 | 98.45 | 98.3372
2.2 | 0 | 0 | 50.91 | 125870.5 | 98.65 | 98.4339
3.4 | 0 | 0 | 50.76 | 125505 | 98.26 | 98.1439
4.6 | 0 | 0 | 50.95 | 125968.4 | 98.51 | 98.5112
5.8 | 0 | 0 | 51 | 126081.4 | 98.75 | 98.6079
7 | 0 | 0 | 50.77 | 125517.7 | 98.23 | 98.1632
8.2 | 0 | 0 | 50.95 | 125957.3 | 98.54 | 98.5112
9.4 | 0 | 0 | 51.06 | 126235.4 | 98.84 | 98.7239
10.6 | 0 | 0 | 51.04 | 126180.2 | 98.81 | 98.6852
11.8 | 0 | 0 | 51.01 | 126122.2 | 98.83 | 98.6272
13 | 0 | 0 | 50.98 | 126030.5 | 98.63 | 98.5692
14.2 | 0 | 0 | 51 | 126093.1 | 98.81 | 98.6079
15.4 | 0 | 0 | 51.2 | 126578.8 | 99.05 | 98.9946
16.6 | 0 | 0 | 51.15 | 126461.3 | 98.96 | 98.8979
17.8 | 0 | 0 | 51.11 | 126361.7 | 98.91 | 98.8206
19.1 | 0 | 0 | 51.1 | 126332.8 | 98.7 | 98.8012
20.2 | 0 | 0 | 51.07 | 126257.5 | 98.67 | 98.7432
21.4 | 0 | 0 | 51.11 | 126350.5 | 98.89 | 98.8206
22.7 | 0 | 0 | 51.12 | 126398.2 | 98.95 | 98.8399
23.9 | 0 | 0 | 51.05 | 126216.1 | 98.62 | 98.7046
25.1 | 0 | 0 | 51.02 | 126141.6 | 98.79 | 98.6466
26.3 | 0 | 0 | 50.96 | 125984.2 | 98.54 | 98.5305
27.4 | 0 | 0 | 50.86 | 125751.9 | 98.37 | 98.3372
28.7 | 0 | 0 | 50.83 | 125664.6 | 98.33 | 98.2792
29.9 | 0 | 0 | 50.85 | 125707.4 | 98.38 | 98.3179
31.1 | 0 | 0 | 50.74 | 125438.6 | 98.26 | 98.1052
32.3 | 0 | 0 | 50.59 | 125080.8 | 97.83 | 97.8152
33.5 | 0 | 0 | 50.52 | 124909.8 | 97.86 | 97.6798
34.7 | 0 | 0 | 50.39 | 124582.8 | 97.76 | 97.4285
35.9 | 0 | 0 | 50.41 | 124643.3 | 97.74 | 97.4671
37.1 | 0 | 0 | 50.29 | 124345.1 | 97.44 | 97.2351
38.3 | 0 | 0 | 50.31 | 124384.1 | 97.3 | 97.2738
39.5 | 0 | 0 | 50.25 | 124234.8 | 97.2 | 97.1578
40.7 | 0 | 0 | 50.27 | 124282.7 | 97.16 | 97.1964
41.9 | 0 | 0 | 49.9 | 123361.9 | 96.73 | 96.4811
43.1 | 0 | 0 | 49.71 | 122897.2 | 95.98 | 96.1137
44.3 | 0 | 0 | 48.63 | 120229.6 | 94.06 | 94.0255
45.5 | 0 | 0 | 46.02 | 113781.9 | 89.09 | 88.9791
46.7 | 0 | 0 | 41.94 | 103686.1 | 81.17 | 81.0905
47.9 | 0 | 0 | 37.02 | 91517 | 71.58 | 71.5777
49.1 | 0 | 0 | 30.6 | 75665.6 | 59.18 | 59.1647
50.3 | 0 | 0 | 23.72 | 58644.8 | 45.81 | 45.8623
51.5 | 0 | 0 | 17.22 | 42565.6 | 33.27 | 33.2947
52.7 | 0 | 0 | 12.2 | 30174.7 | 23.62 | 23.5886
53.9 | 0 | 0 | 9.7 | 23975.6 | 18.72 | 18.7548
55.2 | 0 | 0 | 8.81 | 21770.4 | 17.05 | 17.034
56.4 | 0 | 0 | 8.39 | 20749.5 | 16.22 | 16.222
57.5 | 0 | 0 | 8.13 | 20093.8 | 15.73 | 15.7193
58.8 | 0 | 0 | 7.92 | 19575 | 15.35 | 15.3132
60 | 0 | 0 | 7.67 | 18965.4 | 14.84 | 14.8299
61.2 | 0 | 0 | 7.49 | 18527.9 | 14.5 | 14.4818
62.4 | 0 | 0 | 7.27 | 17974.6 | 14.08 | 14.0565
63.6 | 0 | 0 | 7.14 | 17658.5 | 13.81 | 13.8051
64.8 | 0 | 0 | 6.93 | 17129.2 | 13.41 | 13.3991
66 | 0 | 0 | 6.75 | 16677.8 | 13.06 | 13.051
67.2 | 0 | 0 | 6.65 | 16447 | 12.87 | 12.8577
68.4 | 0 | 0 | 6.45 | 15951.4 | 12.5 | 12.471
69.6 | 0 | 0 | 6.33 | 15657.6 | 12.26 | 12.239
70.8 | 0 | 0 | 6.21 | 15342.4 | 12.02 | 12.007
72 | 0 | 0 | 6.1 | 15071.9 | 11.8 | 11.7943
73.2 | 0 | 0 | 6.02 | 14877.3 | 11.62 | 11.6396
74.4 | 0 | 0 | 5.84 | 14440.5 | 11.3 | 11.2916
75.6 | 0 | 0 | 5.7 | 14090.7 | 11.02 | 11.0209
76.8 | 0 | 0 | 5.59 | 13828.3 | 10.8 | 10.8082
78 | 0 | 0 | 5.5 | 13589.3 | 10.64 | 10.6342
79.2 | 0 | 0 | 5.41 | 13375.7 | 10.48 | 10.4602
80.4 | 0 | 0 | 5.27 | 13031.7 | 10.2 | 10.1895
81.6 | 0 | 0 | 5.23 | 12940.7 | 10.12 | 10.1121
82.9 | 0 | 0 | 5.11 | 12628.6 | 9.9 | 9.8801
84 | 0 | 0 | 5.04 | 12465.6 | 9.77 | 9.7448
85.3 | 0 | 0 | 4.97 | 12278.5 | 9.61 | 9.6094
86.5 | 0 | 0 | 4.84 | 11973.5 | 9.36 | 9.3581
87.7 | 0 | 0 | 4.75 | 11741.1 | 9.18 | 9.1841
88.9 | 0 | 0 | 4.69 | 11591.7 | 9.07 | 9.0681
90.1 | 0 | 0 | 4.58 | 11311.8 | 8.84 | 8.8554
91.3 | 0 | 0 | 4.48 | 11086.3 | 8.65 | 8.662
92.5 | 0 | 0 | 4.37 | 10798.6 | 8.46 | 8.4493
93.7 | 0 | 0 | 4.33 | 10697.8 | 8.37 | 8.372
94.9 | 0 | 0 | 4.25 | 10516.9 | 8.24 | 8.2173
96.1 | 0 | 0 | 4.21 | 10400.3 | 8.13 | 8.14
97.3 | 0 | 0 | 4.11 | 10169.2 | 7.96 | 7.9466
98.5 | 0 | 0 | 4.05 | 10021.3 | 7.83 | 7.8306
100 | 0 | 0 | 3.97 | 9820.5 | 7.68 | 7.6759
Measurement time: | 12/08/2022 18:23:53
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Crossline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 100 | 0 | 3.85 | 9530.3 | 7.41 | 7.4801
0 | 99.1 | 0 | 3.2 | 7911.7 | 6.22 | 6.2172
0 | 97.7 | 0 | 3.32 | 8196.2 | 6.44 | 6.4504
0 | 96.6 | 0 | 3.33 | 8242.6 | 6.46 | 6.4698
0 | 95.4 | 0 | 3.41 | 8431.5 | 6.62 | 6.6252
0 | 94.2 | 0 | 3.49 | 8624.7 | 6.77 | 6.7806
0 | 93 | 0 | 3.56 | 8791.5 | 6.9 | 6.9167
0 | 91.8 | 0 | 3.6 | 8904.1 | 7 | 6.9944
0 | 90.6 | 0 | 3.7 | 9158.7 | 7.2 | 7.1887
0 | 89.4 | 0 | 3.73 | 9218.4 | 7.24 | 7.2469
0 | 88.2 | 0 | 3.84 | 9499.1 | 7.47 | 7.4607
0 | 86.9 | 0 | 3.9 | 9636.7 | 7.56 | 7.5772
0 | 85.8 | 0 | 4 | 9885.4 | 7.77 | 7.7715
0 | 84.6 | 0 | 4.07 | 10052.6 | 7.9 | 7.9075
0 | 83.4 | 0 | 4.19 | 10357.3 | 8.14 | 8.1407
0 | 82.1 | 0 | 4.27 | 10556.5 | 8.29 | 8.2961
0 | 80.9 | 0 | 4.33 | 10714.9 | 8.42 | 8.4127
0 | 79.7 | 0 | 4.45 | 11002.2 | 8.66 | 8.6458
0 | 78.5 | 0 | 4.54 | 11231.5 | 8.82 | 8.8207
0 | 77.3 | 0 | 4.64 | 11474.4 | 9 | 9.015
0 | 76.1 | 0 | 4.76 | 11773.1 | 9.25 | 9.2481
0 | 74.9 | 0 | 4.87 | 12033.5 | 9.44 | 9.4618
0 | 73.7 | 0 | 5 | 12360.3 | 9.7 | 9.7144
0 | 72.5 | 0 | 5.09 | 12585.5 | 9.89 | 9.8893
0 | 71.3 | 0 | 5.28 | 13046.4 | 10.25 | 10.2584
0 | 70.1 | 0 | 5.38 | 13313.4 | 10.45 | 10.4527
0 | 68.9 | 0 | 5.53 | 13663.6 | 10.73 | 10.7441
0 | 67.7 | 0 | 5.69 | 14077 | 11.07 | 11.055
0 | 66.5 | 0 | 5.81 | 14359.2 | 11.28 | 11.2881
0 | 65.3 | 0 | 5.95 | 14706.4 | 11.55 | 11.5601
0 | 64.1 | 0 | 6.14 | 15191.8 | 11.93 | 11.9293
0 | 62.9 | 0 | 6.32 | 15634.7 | 12.28 | 12.279
0 | 61.7 | 0 | 6.56 | 16209.3 | 12.72 | 12.7453
0 | 60.5 | 0 | 6.73 | 16631.5 | 13.05 | 13.0756
0 | 59.3 | 0 | 6.98 | 17266.5 | 13.56 | 13.5613
0 | 58.1 | 0 | 7.21 | 17830.8 | 13.99 | 14.0082
0 | 56.9 | 0 | 7.44 | 18406.3 | 14.46 | 14.455
0 | 55.6 | 0 | 7.72 | 19078.4 | 15 | 14.999
0 | 54.5 | 0 | 8.07 | 19942.5 | 15.69 | 15.679
0 | 53.3 | 0 | 8.8 | 21757.4 | 17.12 | 17.0973
0 | 52 | 0 | 10.89 | 26926.3 | 21.14 | 21.158
0 | 50.8 | 0 | 16.51 | 40827.5 | 32.09 | 32.0769
0 | 49.7 | 0 | 23.46 | 58005 | 45.58 | 45.5799
0 | 48.4 | 0 | 30.99 | 76606.6 | 60.16 | 60.2098
0 | 47.2 | 0 | 38.79 | 95891.2 | 75.23 | 75.3643
0 | 46.1 | 0 | 45.07 | 111435.3 | 87.61 | 87.5656
0 | 44.8 | 0 | 49.14 | 121494.1 | 95.38 | 95.4731
0 | 43.6 | 0 | 49.91 | 123391.7 | 96.83 | 96.9691
0 | 42.4 | 0 | 50.48 | 124800.5 | 98.04 | 98.0765
0 | 41.2 | 0 | 50.5 | 124863.2 | 98 | 98.1154
0 | 40 | 0 | 50.66 | 125256.3 | 98.41 | 98.4263
0 | 38.8 | 0 | 50.79 | 125575.4 | 98.64 | 98.6788
0 | 37.6 | 0 | 50.9 | 125844.9 | 98.77 | 98.8926
0 | 36.4 | 0 | 50.9 | 125830.8 | 98.79 | 98.8926
0 | 35.2 | 0 | 50.96 | 125979.8 | 99.2 | 99.0091
0 | 34 | 0 | 50.87 | 125773.8 | 98.74 | 98.8343
0 | 32.8 | 0 | 50.9 | 125851.5 | 98.95 | 98.8926
0 | 31.6 | 0 | 51.17 | 126504.6 | 99.44 | 99.4171
0 | 30.4 | 0 | 51.14 | 126426.6 | 99.34 | 99.3588
0 | 29.1 | 0 | 51.23 | 126660.3 | 99.42 | 99.5337
0 | 28 | 0 | 51.11 | 126368.4 | 99.16 | 99.3006
0 | 26.8 | 0 | 51.24 | 126679.3 | 99.44 | 99.5531
0 | 25.5 | 0 | 51.35 | 126966.2 | 99.66 | 99.7669
0 | 24.3 | 0 | 51.3 | 126822.2 | 99.6 | 99.6697
0 | 23.1 | 0 | 51.36 | 126987.2 | 99.73 | 99.7863
0 | 21.9 | 0 | 51.41 | 127091.7 | 100 | 99.8834
0 | 20.7 | 0 | 51.28 | 126775 | 99.58 | 99.6309
0 | 19.5 | 0 | 51.22 | 126641 | 99.4 | 99.5143
0 | 18.3 | 0 | 51.39 | 127064.7 | 99.99 | 99.8446
0 | 17.1 | 0 | 51.34 | 126923.7 | 99.63 | 99.7474
0 | 15.9 | 0 | 51.47 | 127259.5 | 99.89 | 100
0 | 14.7 | 0 | 51.19 | 126568.3 | 99.5 | 99.456
0 | 13.5 | 0 | 51.27 | 126753.9 | 99.48 | 99.6114
0 | 12.3 | 0 | 51.24 | 126693.6 | 99.7 | 99.5531
0 | 11.1 | 0 | 51.13 | 126422 | 99.32 | 99.3394
0 | 9.9 | 0 | 51.09 | 126323.7 | 99.23 | 99.2617
0 | 8.7 | 0 | 51.07 | 126274.1 | 99.16 | 99.2228
0 | 7.5 | 0 | 51.08 | 126289.1 | 99.19 | 99.2423
0 | 6.3 | 0 | 51.09 | 126314.1 | 99.32 | 99.2617
0 | 5.1 | 0 | 50.96 | 125993.9 | 99.01 | 99.0091
0 | 3.9 | 0 | 50.81 | 125617.6 | 98.71 | 98.7177
0 | 2.7 | 0 | 51.03 | 126169 | 99.17 | 99.1451
0 | 1.5 | 0 | 50.83 | 125673.5 | 98.7 | 98.7566
0 | 0.3 | 0 | 50.81 | 125631.1 | 98.72 | 98.7177
0 | -0.9 | 0 | 50.79 | 125559.5 | 98.63 | 98.6788
0 | -2.2 | 0 | 50.8 | 125605.1 | 98.75 | 98.6983
0 | -3.4 | 0 | 50.8 | 125605.8 | 98.59 | 98.6983
0 | -4.5 | 0 | 50.81 | 125608.8 | 98.59 | 98.7177
0 | -5.8 | 0 | 50.85 | 125715.1 | 98.71 | 98.7954
0 | -7 | 0 | 50.83 | 125678.8 | 98.77 | 98.7566
0 | -8.2 | 0 | 51 | 126079.8 | 99.04 | 99.0868
0 | -9.4 | 0 | 50.89 | 125815.1 | 98.73 | 98.8731
0 | -10.6 | 0 | 51.02 | 126148.5 | 99.07 | 99.1257
0 | -11.8 | 0 | 51.08 | 126279.2 | 99.15 | 99.2423
0 | -13 | 0 | 51.03 | 126166.5 | 99.22 | 99.1451
0 | -14.2 | 0 | 50.94 | 125952.8 | 98.85 | 98.9703
0 | -15.4 | 0 | 51.07 | 126270.1 | 99.16 | 99.2228
0 | -16.6 | 0 | 51.17 | 126498.1 | 99.24 | 99.4171
0 | -17.8 | 0 | 51.18 | 126537.5 | 99.42 | 99.4366
0 | -19 | 0 | 51.18 | 126526.5 | 99.38 | 99.4366
0 | -20.2 | 0 | 51.14 | 126428.1 | 99.46 | 99.3588
0 | -21.4 | 0 | 51.28 | 126775.7 | 99.52 | 99.6309
0 | -22.6 | 0 | 51.18 | 126547.2 | 99.29 | 99.4366
0 | -23.8 | 0 | 51.23 | 126663.7 | 99.33 | 99.5337
0 | -25 | 0 | 51.2 | 126593.6 | 99.33 | 99.4754
0 | -26.2 | 0 | 51.22 | 126634.4 | 99.52 | 99.5143
0 | -27.4 | 0 | 51.17 | 126513.9 | 99.46 | 99.4171
0 | -28.7 | 0 | 51.01 | 126121.4 | 99.1 | 99.1063
0 | -29.9 | 0 | 51.14 | 126424.9 | 99.39 | 99.3588
0 | -31.1 | 0 | 51.06 | 126248.2 | 99.1 | 99.2034
0 | -32.2 | 0 | 50.89 | 125826.7 | 98.84 | 98.8731
0 | -33.4 | 0 | 50.97 | 126020.6 | 98.93 | 99.0286
0 | -34.7 | 0 | 50.93 | 125918 | 98.88 | 98.9508
0 | -35.9 | 0 | 51 | 126079.6 | 98.85 | 99.0868
0 | -37.1 | 0 | 50.76 | 125484.7 | 98.58 | 98.6206
0 | -38.3 | 0 | 50.75 | 125469.1 | 98.47 | 98.6011
0 | -39.5 | 0 | 50.7 | 125338.3 | 98.47 | 98.504
0 | -40.7 | 0 | 50.58 | 125044.1 | 98.25 | 98.2708
0 | -41.9 | 0 | 50.45 | 124732.8 | 97.86 | 98.0183
0 | -43.1 | 0 | 50.38 | 124553.5 | 97.98 | 97.8823
0 | -44.3 | 0 | 50.2 | 124105.2 | 97.62 | 97.5325
0 | -45.5 | 0 | 50.11 | 123901.7 | 97.36 | 97.3577
0 | -46.7 | 0 | 49.91 | 123401.5 | 96.92 | 96.9691
0 | -47.9 | 0 | 48.64 | 120250.2 | 94.52 | 94.5017
0 | -49.2 | 0 | 44.28 | 109466.8 | 85.99 | 86.0307
0 | -50.3 | 0 | 37.99 | 93924.4 | 73.67 | 73.81
0 | -51.5 | 0 | 30.27 | 74836.5 | 58.85 | 58.811
0 | -52.7 | 0 | 22.64 | 55974.3 | 44.05 | 43.9868
0 | -53.9 | 0 | 15.87 | 39223.9 | 30.83 | 30.8335
0 | -55.1 | 0 | 10.7 | 26466.3 | 20.78 | 20.7888
0 | -56.3 | 0 | 8.71 | 21525.6 | 16.92 | 16.9225
0 | -57.5 | 0 | 8.05 | 19894.9 | 15.64 | 15.6402
0 | -58.8 | 0 | 7.64 | 18893.2 | 14.84 | 14.8436
0 | -60 | 0 | 7.43 | 18367 | 14.43 | 14.4356
0 | -61.2 | 0 | 7.2 | 17799.5 | 13.97 | 13.9887
0 | -62.3 | 0 | 6.95 | 17177.6 | 13.5 | 13.503
0 | -63.6 | 0 | 6.72 | 16625.5 | 13.07 | 13.0561
0 | -64.8 | 0 | 6.5 | 16077 | 12.64 | 12.6287
0 | -66 | 0 | 6.31 | 15597.6 | 12.26 | 12.2596
0 | -67.2 | 0 | 6.09 | 15062.4 | 11.84 | 11.8321
0 | -68.4 | 0 | 5.89 | 14573.3 | 11.46 | 11.4436
0 | -69.6 | 0 | 5.8 | 14331.4 | 11.26 | 11.2687
0 | -70.8 | 0 | 5.59 | 13816.7 | 10.86 | 10.8607
0 | -72 | 0 | 5.47 | 13527 | 10.65 | 10.6276
0 | -73.2 | 0 | 5.38 | 13306.3 | 10.45 | 10.4527
0 | -74.3 | 0 | 5.2 | 12865.5 | 10.12 | 10.103
0 | -75.6 | 0 | 5.14 | 12712.7 | 9.98 | 9.9864
0 | -76.8 | 0 | 5.04 | 12454.1 | 9.79 | 9.7921
0 | -78 | 0 | 4.89 | 12093.6 | 9.51 | 9.5007
0 | -79.2 | 0 | 4.76 | 11756.9 | 9.23 | 9.2481
0 | -80.4 | 0 | 4.65 | 11506.3 | 9.05 | 9.0344
0 | -81.6 | 0 | 4.55 | 11241.9 | 8.83 | 8.8401
0 | -82.8 | 0 | 4.46 | 11027.9 | 8.67 | 8.6652
0 | -84 | 0 | 4.35 | 10748.7 | 8.45 | 8.4515
0 | -85.2 | 0 | 4.22 | 10430.9 | 8.19 | 8.199
0 | -86.4 | 0 | 4.18 | 10334 | 8.12 | 8.1212
0 | -87.6 | 0 | 4.07 | 10055.3 | 7.91 | 7.9075
0 | -88.8 | 0 | 4 | 9896.3 | 7.77 | 7.7715
0 | -90 | 0 | 3.97 | 9813.2 | 7.7 | 7.7132
0 | -91.2 | 0 | 3.85 | 9529.7 | 7.48 | 7.4801
0 | -92.4 | 0 | 3.73 | 9233.7 | 7.26 | 7.2469
0 | -93.6 | 0 | 3.71 | 9165.3 | 7.19 | 7.2081
0 | -94.8 | 0 | 3.61 | 8913.3 | 7 | 7.0138
0 | -96 | 0 | 3.56 | 8810.7 | 6.93 | 6.9167
0 | -97.3 | 0 | 3.46 | 8546.2 | 6.71 | 6.7224
0 | -98.4 | 0 | 3.43 | 8474.6 | 6.67 | 6.6641
0 | -100 | 0 | 3.37 | 8334.6 | 6.54 | 6.5475