## 4_Profils profondeur mesure
Measurement time: | 12/08/2022 18:26:00
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
100 | 0 | 50 | 2.76 | 6826.9 | 2.99 | 3.1662
99.1 | 0 | 50 | 2.38 | 5876.7 | 2.73 | 2.7303
97.6 | 0 | 50 | 2.45 | 6064 | 2.81 | 2.8106
96.5 | 0 | 50 | 2.53 | 6263.9 | 2.91 | 2.9024
95.3 | 0 | 50 | 2.6 | 6416.5 | 2.97 | 2.9827
94.1 | 0 | 50 | 2.7 | 6681.5 | 3.1 | 3.0974
92.9 | 0 | 50 | 2.77 | 6858.3 | 3.18 | 3.1777
91.7 | 0 | 50 | 2.89 | 7145.8 | 3.31 | 3.3154
90.5 | 0 | 50 | 2.89 | 7140.2 | 3.31 | 3.3154
89.3 | 0 | 50 | 3.04 | 7511.2 | 3.48 | 3.4874
88.1 | 0 | 50 | 3.14 | 7774.3 | 3.61 | 3.6022
86.9 | 0 | 50 | 3.21 | 7941.8 | 3.69 | 3.6825
85.7 | 0 | 50 | 3.37 | 8321.5 | 3.86 | 3.866
84.5 | 0 | 50 | 3.47 | 8569.7 | 3.97 | 3.9807
83.2 | 0 | 50 | 3.62 | 8950.4 | 4.15 | 4.1528
82 | 0 | 50 | 3.7 | 9147.8 | 4.24 | 4.2446
80.8 | 0 | 50 | 3.8 | 9398.2 | 4.36 | 4.3593
79.6 | 0 | 50 | 3.95 | 9758.6 | 4.53 | 4.5314
78.4 | 0 | 50 | 4.05 | 10009.6 | 4.64 | 4.6461
77.3 | 0 | 50 | 4.23 | 10448.5 | 4.85 | 4.8526
76 | 0 | 50 | 4.4 | 10874.4 | 5.05 | 5.0476
74.8 | 0 | 50 | 4.53 | 11209.5 | 5.2 | 5.1967
73.6 | 0 | 50 | 4.72 | 11679.1 | 5.42 | 5.4147
72.4 | 0 | 50 | 4.93 | 12187.9 | 5.65 | 5.6556
71.2 | 0 | 50 | 5.09 | 12594.3 | 5.84 | 5.8392
70 | 0 | 50 | 5.36 | 13251 | 6.14 | 6.1489
68.8 | 0 | 50 | 5.53 | 13663.4 | 6.35 | 6.3439
67.6 | 0 | 50 | 5.75 | 14217.9 | 6.59 | 6.5963
66.4 | 0 | 50 | 6.02 | 14890.1 | 6.91 | 6.906
65.2 | 0 | 50 | 6.36 | 15712.6 | 7.3 | 7.2961
64 | 0 | 50 | 6.64 | 16408.9 | 7.61 | 7.6173
62.8 | 0 | 50 | 7.04 | 17406.7 | 8.09 | 8.0762
61.6 | 0 | 50 | 7.55 | 18668.2 | 8.66 | 8.6612
60.3 | 0 | 50 | 8.16 | 20173.2 | 9.36 | 9.361
59.2 | 0 | 50 | 9.07 | 22416.5 | 10.4 | 10.405
57.9 | 0 | 50 | 10.49 | 25924.1 | 12.02 | 12.034
56.8 | 0 | 50 | 13.05 | 32270.5 | 14.99 | 14.9707
55.5 | 0 | 50 | 17.81 | 44027.6 | 20.42 | 20.4313
54.3 | 0 | 50 | 26.32 | 65060.7 | 30.13 | 30.1939
53.2 | 0 | 50 | 36.16 | 89388.7 | 41.49 | 41.4822
51.9 | 0 | 50 | 48.06 | 118827.8 | 55.12 | 55.1336
50.7 | 0 | 50 | 58.69 | 145114.5 | 67.35 | 67.3282
49.5 | 0 | 50 | 67.28 | 166340.6 | 77.22 | 77.1825
48.3 | 0 | 50 | 74.35 | 183813.9 | 85.36 | 85.2931
47.1 | 0 | 50 | 78.66 | 194476.4 | 90.12 | 90.2375
45.9 | 0 | 50 | 81.2 | 200754.1 | 93.14 | 93.1513
44.7 | 0 | 50 | 82.5 | 203966 | 94.81 | 94.6427
43.5 | 0 | 50 | 83.26 | 205840.6 | 95.42 | 95.5145
42.3 | 0 | 50 | 83.74 | 207024.1 | 95.85 | 96.0652
41.1 | 0 | 50 | 84.12 | 207972.6 | 96.54 | 96.5011
39.8 | 0 | 50 | 84.38 | 208628.1 | 96.8 | 96.7994
38.7 | 0 | 50 | 84.51 | 208946.2 | 96.99 | 96.9485
37.4 | 0 | 50 | 84.59 | 209137.5 | 97.09 | 97.0403
36.2 | 0 | 50 | 85.03 | 210230.2 | 97.57 | 97.545
35.1 | 0 | 50 | 85.15 | 210513.3 | 97.55 | 97.6827
33.8 | 0 | 50 | 85.22 | 210690.7 | 97.74 | 97.763
32.7 | 0 | 50 | 85.44 | 211244.2 | 97.98 | 98.0154
31.4 | 0 | 50 | 85.63 | 211695.4 | 98.19 | 98.2333
30.2 | 0 | 50 | 85.76 | 212027.9 | 98.37 | 98.3825
29 | 0 | 50 | 85.76 | 212025.6 | 98.53 | 98.3825
27.8 | 0 | 50 | 85.88 | 212317.6 | 98.57 | 98.5201
26.6 | 0 | 50 | 86.1 | 212863.9 | 98.75 | 98.7725
25.4 | 0 | 50 | 86.15 | 212991.1 | 98.81 | 98.8299
24.2 | 0 | 50 | 86.29 | 213332.4 | 98.87 | 98.9905
23 | 0 | 50 | 86.24 | 213211.2 | 98.93 | 98.9331
21.8 | 0 | 50 | 86.36 | 213507.5 | 99.15 | 99.0708
20.6 | 0 | 50 | 86.36 | 213519.2 | 99.05 | 99.0708
19.4 | 0 | 50 | 86.27 | 213280.6 | 99.1 | 98.9675
18.2 | 0 | 50 | 86.33 | 213438.5 | 99.07 | 99.0364
17 | 0 | 50 | 86.34 | 213460.4 | 99.23 | 99.0478
15.8 | 0 | 50 | 86.59 | 214069.1 | 99.17 | 99.3346
14.6 | 0 | 50 | 86.24 | 213211.5 | 98.89 | 98.9331
13.4 | 0 | 50 | 86.38 | 213550.2 | 99.24 | 99.0937
12.2 | 0 | 50 | 86.27 | 213282.8 | 98.91 | 98.9675
11 | 0 | 50 | 86.22 | 213166.8 | 98.94 | 98.9102
9.8 | 0 | 50 | 86.07 | 212803.8 | 99 | 98.7381
8.6 | 0 | 50 | 86.01 | 212656.1 | 98.73 | 98.6693
7.4 | 0 | 50 | 86.24 | 213216.3 | 98.87 | 98.9331
6.2 | 0 | 50 | 86.29 | 213345.4 | 98.97 | 98.9905
5 | 0 | 50 | 86.35 | 213490.5 | 99.04 | 99.0593
3.8 | 0 | 50 | 86.27 | 213281.9 | 99.11 | 98.9675
2.6 | 0 | 50 | 86.4 | 213600.1 | 99.01 | 99.1167
1.4 | 0 | 50 | 86.27 | 213298.7 | 99.02 | 98.9675
0.2 | 0 | 50 | 86.44 | 213719.6 | 99.11 | 99.1626
-1 | 0 | 50 | 86.34 | 213463.1 | 99.02 | 99.0478
-2.3 | 0 | 50 | 86.46 | 213751 | 99.09 | 99.1855
-3.4 | 0 | 50 | 86.32 | 213410.4 | 99.18 | 99.0249
-4.6 | 0 | 50 | 86.45 | 213739.5 | 99.2 | 99.174
-5.9 | 0 | 50 | 86.46 | 213766.3 | 99.2 | 99.1855
-7 | 0 | 50 | 86.61 | 214121.9 | 99.31 | 99.3576
-8.3 | 0 | 50 | 86.59 | 214092.5 | 99.42 | 99.3346
-9.5 | 0 | 50 | 86.68 | 214295.8 | 99.36 | 99.4379
-10.7 | 0 | 50 | 86.83 | 214670.8 | 99.78 | 99.61
-11.9 | 0 | 50 | 86.92 | 214890.3 | 99.75 | 99.7132
-13.1 | 0 | 50 | 86.93 | 214931.1 | 99.65 | 99.7247
-14.3 | 0 | 50 | 86.97 | 215014 | 99.82 | 99.7706
-15.5 | 0 | 50 | 87.04 | 215183.6 | 99.94 | 99.8509
-16.7 | 0 | 50 | 87.17 | 215503.1 | 99.95 | 100
-17.9 | 0 | 50 | 86.91 | 214861.3 | 99.64 | 99.7017
-19.1 | 0 | 50 | 87.15 | 215465.1 | 99.95 | 99.9771
-20.3 | 0 | 50 | 87.06 | 215246.5 | 99.92 | 99.8738
-21.5 | 0 | 50 | 87.1 | 215334.3 | 99.9 | 99.9197
-22.7 | 0 | 50 | 87.1 | 215353.8 | 100 | 99.9197
-23.9 | 0 | 50 | 87.03 | 215159.4 | 99.89 | 99.8394
-25.1 | 0 | 50 | 87 | 215093.2 | 99.57 | 99.805
-26.3 | 0 | 50 | 87.04 | 215182.8 | 99.93 | 99.8509
-27.5 | 0 | 50 | 86.82 | 214650.7 | 99.76 | 99.5985
-28.7 | 0 | 50 | 86.82 | 214655.6 | 99.71 | 99.5985
-29.9 | 0 | 50 | 86.86 | 214759.5 | 99.64 | 99.6444
-31.1 | 0 | 50 | 86.73 | 214421 | 99.42 | 99.4952
-32.3 | 0 | 50 | 86.57 | 214025.1 | 99.37 | 99.3117
-33.5 | 0 | 50 | 86.34 | 213469.6 | 99.11 | 99.0478
-34.7 | 0 | 50 | 86.38 | 213570.3 | 99.05 | 99.0937
-36 | 0 | 50 | 86.25 | 213249.9 | 98.84 | 98.9446
-37.1 | 0 | 50 | 86.07 | 212802.6 | 98.78 | 98.7381
-38.4 | 0 | 50 | 86.08 | 212825.2 | 98.87 | 98.7496
-39.5 | 0 | 50 | 85.73 | 211942.4 | 98.34 | 98.3481
-40.8 | 0 | 50 | 85.45 | 211266.9 | 98.1 | 98.0268
-41.9 | 0 | 50 | 85.32 | 210928.7 | 97.98 | 97.8777
-43.1 | 0 | 50 | 85.05 | 210264.9 | 97.45 | 97.568
-44.4 | 0 | 50 | 84.57 | 209090 | 97.19 | 97.0173
-45.5 | 0 | 50 | 83.93 | 207497.3 | 96.33 | 96.2831
-46.7 | 0 | 50 | 82.73 | 204536.8 | 94.97 | 94.9065
-48 | 0 | 50 | 80.47 | 198953.2 | 92.18 | 92.3139
-49.1 | 0 | 50 | 76.11 | 188167.6 | 87.34 | 87.3121
-50.4 | 0 | 50 | 68.2 | 168624.9 | 78.39 | 78.2379
-51.6 | 0 | 50 | 57.93 | 143211 | 66.58 | 66.4563
-52.8 | 0 | 50 | 46.25 | 114351.8 | 53.06 | 53.0572
-54 | 0 | 50 | 35.06 | 86669.2 | 40.28 | 40.2203
-55.2 | 0 | 50 | 25.85 | 63914.4 | 29.68 | 29.6547
-56.4 | 0 | 50 | 18.51 | 45765.2 | 21.23 | 21.2344
-57.6 | 0 | 50 | 13.69 | 33850.6 | 15.71 | 15.7049
-58.8 | 0 | 50 | 10.99 | 27163.8 | 12.61 | 12.6075
-60 | 0 | 50 | 9.36 | 23138.4 | 10.73 | 10.7376
-61.2 | 0 | 50 | 8.46 | 20912.9 | 9.69 | 9.7052
-62.5 | 0 | 50 | 7.75 | 19171.4 | 8.91 | 8.8907
-63.6 | 0 | 50 | 7.22 | 17852.3 | 8.28 | 8.2827
-64.9 | 0 | 50 | 6.86 | 16959.1 | 7.87 | 7.8697
-66.1 | 0 | 50 | 6.56 | 16211.6 | 7.52 | 7.5255
-67.2 | 0 | 50 | 6.18 | 15281.5 | 7.09 | 7.0896
-68.5 | 0 | 50 | 5.94 | 14693.9 | 6.82 | 6.8143
-69.6 | 0 | 50 | 5.73 | 14176.9 | 6.58 | 6.5734
-70.9 | 0 | 50 | 5.48 | 13543.6 | 6.29 | 6.2866
-72.1 | 0 | 50 | 5.31 | 13131.8 | 6.1 | 6.0915
-73.3 | 0 | 50 | 5.06 | 12521 | 5.81 | 5.8047
-74.5 | 0 | 50 | 4.88 | 12066.8 | 5.61 | 5.5983
-75.6 | 0 | 50 | 4.73 | 11698.2 | 5.42 | 5.4262
-76.9 | 0 | 50 | 4.5 | 11130.4 | 5.16 | 5.1623
-78.1 | 0 | 50 | 4.34 | 10721.8 | 4.98 | 4.9788
-79.3 | 0 | 50 | 4.23 | 10455.5 | 4.86 | 4.8526
-80.5 | 0 | 50 | 4.05 | 10006 | 4.65 | 4.6461
-81.7 | 0 | 50 | 3.87 | 9568.9 | 4.44 | 4.4396
-82.9 | 0 | 50 | 3.79 | 9369 | 4.35 | 4.3478
-84.1 | 0 | 50 | 3.68 | 9109.3 | 4.23 | 4.2216
-85.3 | 0 | 50 | 3.53 | 8719.4 | 4.05 | 4.0496
-86.6 | 0 | 50 | 3.41 | 8433.2 | 3.92 | 3.9119
-87.7 | 0 | 50 | 3.27 | 8084.3 | 3.75 | 3.7513
-88.9 | 0 | 50 | 3.23 | 7982.3 | 3.71 | 3.7054
-90.1 | 0 | 50 | 3.18 | 7859.3 | 3.64 | 3.648
-91.3 | 0 | 50 | 3.02 | 7462.9 | 3.47 | 3.4645
-92.5 | 0 | 50 | 2.87 | 7085.1 | 3.29 | 3.2924
-93.7 | 0 | 50 | 2.8 | 6927.5 | 3.21 | 3.2121
-95 | 0 | 50 | 2.7 | 6674 | 3.1 | 3.0974
-96.1 | 0 | 50 | 2.65 | 6539.7 | 3.03 | 3.04
-97.3 | 0 | 50 | 2.52 | 6230.1 | 2.89 | 2.8909
-98.6 | 0 | 50 | 2.49 | 6156.2 | 2.86 | 2.8565
-100 | 0 | 50 | 2.41 | 5946.3 | 2.76 | 2.7647
Measurement time: | 12/08/2022 18:26:54
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-100 | 0 | 100 | 2.52 | 6239.8 | 3.92 | 3.7696
-99.1 | 0 | 100 | 2.81 | 6946.8 | 4.2 | 4.2034
-97.7 | 0 | 100 | 2.95 | 7281.2 | 4.42 | 4.4129
-96.6 | 0 | 100 | 3.07 | 7581 | 4.6 | 4.5924
-95.4 | 0 | 100 | 3.12 | 7705.2 | 4.65 | 4.6672
-94.2 | 0 | 100 | 3.25 | 8028.7 | 4.85 | 4.8616
-93 | 0 | 100 | 3.36 | 8301.7 | 5.02 | 5.0262
-91.7 | 0 | 100 | 3.47 | 8586.7 | 5.19 | 5.1907
-90.6 | 0 | 100 | 3.57 | 8827.8 | 5.34 | 5.3403
-89.4 | 0 | 100 | 3.66 | 9044.8 | 5.47 | 5.4749
-88.2 | 0 | 100 | 3.73 | 9216.1 | 5.58 | 5.5797
-86.9 | 0 | 100 | 3.86 | 9534.3 | 5.78 | 5.7741
-85.7 | 0 | 100 | 4.01 | 9910.5 | 5.99 | 5.9985
-84.5 | 0 | 100 | 4.22 | 10442.6 | 6.32 | 6.3126
-83.3 | 0 | 100 | 4.32 | 10692.7 | 6.47 | 6.4622
-82.1 | 0 | 100 | 4.47 | 11052.4 | 6.68 | 6.6866
-80.9 | 0 | 100 | 4.64 | 11461.2 | 6.93 | 6.9409
-79.7 | 0 | 100 | 4.8 | 11875.8 | 7.19 | 7.1803
-78.5 | 0 | 100 | 4.98 | 12307.4 | 7.44 | 7.4495
-77.3 | 0 | 100 | 5.19 | 12839.6 | 7.78 | 7.7636
-76.1 | 0 | 100 | 5.32 | 13141.4 | 7.96 | 7.9581
-74.9 | 0 | 100 | 5.51 | 13616.8 | 8.25 | 8.2423
-73.7 | 0 | 100 | 5.73 | 14169.9 | 8.58 | 8.5714
-72.5 | 0 | 100 | 5.97 | 14764.6 | 8.93 | 8.9304
-71.3 | 0 | 100 | 6.18 | 15282.4 | 9.26 | 9.2446
-70.1 | 0 | 100 | 6.43 | 15900.9 | 9.62 | 9.6185
-68.9 | 0 | 100 | 6.7 | 16564.6 | 10.02 | 10.0224
-67.7 | 0 | 100 | 6.97 | 17240.1 | 10.42 | 10.4263
-66.5 | 0 | 100 | 7.38 | 18247.9 | 11.03 | 11.0396
-65.3 | 0 | 100 | 7.79 | 19260.6 | 11.67 | 11.653
-64.1 | 0 | 100 | 8.28 | 20481 | 12.39 | 12.3859
-62.9 | 0 | 100 | 9.05 | 22382.7 | 13.54 | 13.5378
-61.7 | 0 | 100 | 10.13 | 25056.2 | 15.19 | 15.1533
-60.5 | 0 | 100 | 12.05 | 29783.7 | 18.02 | 18.0254
-59.3 | 0 | 100 | 15.07 | 37263 | 22.57 | 22.543
-58 | 0 | 100 | 19.89 | 49177.1 | 29.74 | 29.7532
-56.8 | 0 | 100 | 25.98 | 64220.8 | 38.85 | 38.8631
-55.6 | 0 | 100 | 33.52 | 82867.4 | 50.18 | 50.1421
-54.4 | 0 | 100 | 41.72 | 103140 | 62.45 | 62.4084
-53.2 | 0 | 100 | 49.39 | 122108.1 | 73.83 | 73.8818
-52 | 0 | 100 | 55.56 | 137362.9 | 83.23 | 83.1114
-50.8 | 0 | 100 | 59.42 | 146914.5 | 88.73 | 88.8856
-49.6 | 0 | 100 | 61.64 | 152406.4 | 92.04 | 92.2064
-48.4 | 0 | 100 | 62.88 | 155452 | 94.09 | 94.0613
-47.2 | 0 | 100 | 63.44 | 156851.8 | 94.79 | 94.899
-46 | 0 | 100 | 63.96 | 158135 | 95.52 | 95.6769
-44.8 | 0 | 100 | 64.28 | 158917.1 | 96.11 | 96.1556
-43.6 | 0 | 100 | 64.52 | 159505.2 | 96.59 | 96.5146
-42.4 | 0 | 100 | 64.84 | 160295.2 | 96.85 | 96.9933
-41.2 | 0 | 100 | 65.17 | 161116.2 | 97.38 | 97.4869
-40 | 0 | 100 | 65.14 | 161057.4 | 97.46 | 97.442
-38.8 | 0 | 100 | 65.23 | 161275.3 | 97.49 | 97.5767
-37.6 | 0 | 100 | 65.53 | 162016.9 | 97.96 | 98.0254
-36.4 | 0 | 100 | 65.51 | 161951.8 | 97.92 | 97.9955
-35.2 | 0 | 100 | 65.88 | 162872.9 | 98.42 | 98.549
-34 | 0 | 100 | 65.91 | 162957.8 | 98.75 | 98.5939
-32.8 | 0 | 100 | 65.88 | 162883.2 | 98.72 | 98.549
-31.6 | 0 | 100 | 66.2 | 163661.4 | 98.93 | 99.0277
-30.4 | 0 | 100 | 66.29 | 163898.8 | 99.03 | 99.1623
-29.2 | 0 | 100 | 66.3 | 163920.7 | 99.11 | 99.1773
-28 | 0 | 100 | 66.56 | 164556.5 | 99.43 | 99.5662
-26.8 | 0 | 100 | 66.52 | 164466.8 | 99.53 | 99.5064
-25.5 | 0 | 100 | 66.52 | 164450.4 | 99.52 | 99.5064
-24.4 | 0 | 100 | 66.63 | 164729 | 99.68 | 99.6709
-23.2 | 0 | 100 | 66.75 | 165041.4 | 99.7 | 99.8504
-21.9 | 0 | 100 | 66.75 | 165029.2 | 99.86 | 99.8504
-20.8 | 0 | 100 | 66.81 | 165181.4 | 99.94 | 99.9402
-19.5 | 0 | 100 | 66.8 | 165153.5 | 99.84 | 99.9252
-18.3 | 0 | 100 | 66.79 | 165119.9 | 99.78 | 99.9102
-17.1 | 0 | 100 | 66.8 | 165160.5 | 99.78 | 99.9252
-15.9 | 0 | 100 | 66.78 | 165095 | 100 | 99.8953
-14.7 | 0 | 100 | 66.66 | 164813.1 | 99.65 | 99.7158
-13.5 | 0 | 100 | 66.75 | 165026.1 | 99.79 | 99.8504
-12.3 | 0 | 100 | 66.62 | 164718.9 | 99.69 | 99.6559
-11.1 | 0 | 100 | 66.85 | 165269.6 | 99.99 | 100
-9.9 | 0 | 100 | 66.76 | 165044.5 | 99.81 | 99.8654
-8.7 | 0 | 100 | 66.71 | 164942.3 | 99.74 | 99.7906
-7.5 | 0 | 100 | 66.64 | 164764.8 | 99.69 | 99.6859
-6.3 | 0 | 100 | 66.52 | 164469.6 | 99.62 | 99.5064
-5.1 | 0 | 100 | 66.56 | 164571.5 | 99.55 | 99.5662
-3.9 | 0 | 100 | 66.59 | 164622.8 | 99.46 | 99.6111
-2.7 | 0 | 100 | 66.53 | 164491.8 | 99.53 | 99.5213
-1.5 | 0 | 100 | 66.52 | 164471.7 | 99.55 | 99.5064
-0.3 | 0 | 100 | 66.37 | 164086.1 | 99.4 | 99.282
0.9 | 0 | 100 | 66.4 | 164155.7 | 99.28 | 99.3269
2.2 | 0 | 100 | 66.38 | 164111.7 | 99.26 | 99.2969
3.4 | 0 | 100 | 66.3 | 163916.4 | 99.3 | 99.1773
4.6 | 0 | 100 | 66.28 | 163855.7 | 99.26 | 99.1473
5.8 | 0 | 100 | 66.49 | 164388.2 | 99.39 | 99.4615
7 | 0 | 100 | 66.51 | 164447.7 | 99.35 | 99.4914
8.2 | 0 | 100 | 66.34 | 164014.5 | 99.33 | 99.2371
9.4 | 0 | 100 | 66.43 | 164244.9 | 99.34 | 99.3717
10.6 | 0 | 100 | 66.44 | 164258.1 | 99.6 | 99.3867
11.8 | 0 | 100 | 66.32 | 163967.3 | 99.25 | 99.2072
13 | 0 | 100 | 66.42 | 164220.5 | 99.59 | 99.3568
14.2 | 0 | 100 | 66.4 | 164171.1 | 99.4 | 99.3269
15.4 | 0 | 100 | 66.44 | 164262.4 | 99.35 | 99.3867
16.6 | 0 | 100 | 66.37 | 164096.5 | 99.35 | 99.282
17.8 | 0 | 100 | 66.34 | 164005.3 | 99.22 | 99.2371
19 | 0 | 100 | 66.48 | 164350.8 | 99.25 | 99.4465
20.2 | 0 | 100 | 66.33 | 163995.3 | 99.14 | 99.2221
21.4 | 0 | 100 | 66.26 | 163808.9 | 99.12 | 99.1174
22.6 | 0 | 100 | 66.04 | 163267.1 | 98.87 | 98.7883
23.8 | 0 | 100 | 66 | 163174.7 | 98.65 | 98.7285
25 | 0 | 100 | 66 | 163174.6 | 98.69 | 98.7285
26.3 | 0 | 100 | 65.9 | 162934.8 | 98.64 | 98.5789
27.4 | 0 | 100 | 65.68 | 162396 | 98.23 | 98.2498
28.7 | 0 | 100 | 65.55 | 162065.8 | 98.04 | 98.0553
29.9 | 0 | 100 | 65.54 | 162028.1 | 97.99 | 98.0404
31 | 0 | 100 | 65.26 | 161347.7 | 97.75 | 97.6215
32.3 | 0 | 100 | 65.37 | 161621.1 | 97.7 | 97.7861
33.5 | 0 | 100 | 65 | 160713.3 | 97.35 | 97.2326
34.7 | 0 | 100 | 64.91 | 160484.4 | 97.11 | 97.098
35.9 | 0 | 100 | 64.72 | 160013.9 | 96.93 | 96.8138
37.1 | 0 | 100 | 64.64 | 159822.6 | 96.63 | 96.6941
38.3 | 0 | 100 | 64.53 | 159537.6 | 96.58 | 96.5295
39.5 | 0 | 100 | 64.24 | 158828.3 | 96.04 | 96.0957
40.7 | 0 | 100 | 63.83 | 157807 | 95.59 | 95.4824
41.9 | 0 | 100 | 63.72 | 157547.5 | 95.45 | 95.3179
43.1 | 0 | 100 | 63.51 | 157010.4 | 94.96 | 95.0037
44.3 | 0 | 100 | 63.16 | 156159.4 | 94.44 | 94.4802
45.5 | 0 | 100 | 62.82 | 155304 | 94.12 | 93.9716
46.7 | 0 | 100 | 62.21 | 153807.9 | 93 | 93.0591
47.9 | 0 | 100 | 61.3 | 151550.9 | 91.67 | 91.6978
49.1 | 0 | 100 | 59.7 | 147592 | 89.18 | 89.3044
50.3 | 0 | 100 | 56.88 | 140619.3 | 85.25 | 85.086
51.5 | 0 | 100 | 52.71 | 130307.7 | 78.81 | 78.8482
52.7 | 0 | 100 | 46.88 | 115903.8 | 70.22 | 70.1272
53.9 | 0 | 100 | 39.7 | 98160 | 59.32 | 59.3867
55.1 | 0 | 100 | 31.84 | 78710.8 | 47.6 | 47.629
56.3 | 0 | 100 | 24.09 | 59560.8 | 36.1 | 36.0359
57.6 | 0 | 100 | 17.58 | 43469.9 | 26.28 | 26.2977
58.8 | 0 | 100 | 13.3 | 32893.5 | 19.89 | 19.8953
59.9 | 0 | 100 | 10.83 | 26779 | 16.21 | 16.2004
61.2 | 0 | 100 | 9.37 | 23176 | 14.04 | 14.0165
62.4 | 0 | 100 | 8.51 | 21042.3 | 12.74 | 12.73
63.6 | 0 | 100 | 7.93 | 19602.4 | 11.86 | 11.8624
64.8 | 0 | 100 | 7.51 | 18559.2 | 11.22 | 11.2341
66 | 0 | 100 | 7.03 | 17371.3 | 10.52 | 10.5161
67.2 | 0 | 100 | 6.8 | 16814.8 | 10.17 | 10.172
68.4 | 0 | 100 | 6.53 | 16133.7 | 9.76 | 9.7681
69.6 | 0 | 100 | 6.25 | 15455.3 | 9.34 | 9.3493
70.8 | 0 | 100 | 5.99 | 14798.2 | 8.97 | 8.9604
72 | 0 | 100 | 5.77 | 14265.7 | 8.64 | 8.6313
73.2 | 0 | 100 | 5.53 | 13675.3 | 8.27 | 8.2723
74.4 | 0 | 100 | 5.34 | 13212 | 8 | 7.988
75.6 | 0 | 100 | 5.14 | 12711.2 | 7.71 | 7.6889
76.8 | 0 | 100 | 4.95 | 12247 | 7.42 | 7.4046
78 | 0 | 100 | 4.79 | 11853.7 | 7.17 | 7.1653
79.2 | 0 | 100 | 4.6 | 11360.6 | 6.87 | 6.8811
80.4 | 0 | 100 | 4.45 | 11010 | 6.65 | 6.6567
81.6 | 0 | 100 | 4.28 | 10585.6 | 6.41 | 6.4024
82.8 | 0 | 100 | 4.15 | 10252.6 | 6.2 | 6.2079
84 | 0 | 100 | 4 | 9889.2 | 5.99 | 5.9835
85.3 | 0 | 100 | 3.89 | 9626.7 | 5.83 | 5.819
86.4 | 0 | 100 | 3.81 | 9423.9 | 5.71 | 5.6993
87.6 | 0 | 100 | 3.7 | 9154 | 5.54 | 5.5348
88.9 | 0 | 100 | 3.51 | 8680.7 | 5.27 | 5.2506
90 | 0 | 100 | 3.45 | 8535.6 | 5.17 | 5.1608
91.3 | 0 | 100 | 3.29 | 8142.6 | 4.94 | 4.9215
92.5 | 0 | 100 | 3.22 | 7958 | 4.82 | 4.8168
93.7 | 0 | 100 | 3.11 | 7696.2 | 4.67 | 4.6522
94.9 | 0 | 100 | 3.04 | 7504.1 | 4.55 | 4.5475
96.1 | 0 | 100 | 2.97 | 7330.5 | 4.44 | 4.4428
97.3 | 0 | 100 | 2.88 | 7119.2 | 4.31 | 4.3082
98.5 | 0 | 100 | 2.76 | 6819.9 | 4.13 | 4.1286
100 | 0 | 100 | 2.66 | 6580.2 | 3.98 | 3.9791
Measurement time: | 12/08/2022 18:27:42
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
100 | 0 | 150 | 3.14 | 7759.3 | 6.55 | 6.2092
99.1 | 0 | 150 | 2.83 | 6986.3 | 5.59 | 5.5962
97.6 | 0 | 150 | 2.95 | 7285.4 | 5.84 | 5.8335
96.5 | 0 | 150 | 3.04 | 7506.1 | 6.01 | 6.0115
95.4 | 0 | 150 | 3.15 | 7784 | 6.23 | 6.229
94.1 | 0 | 150 | 3.22 | 7956.5 | 6.37 | 6.3674
92.9 | 0 | 150 | 3.37 | 8326.8 | 6.67 | 6.664
91.7 | 0 | 150 | 3.4 | 8416.7 | 6.73 | 6.7234
90.5 | 0 | 150 | 3.5 | 8649.6 | 6.9 | 6.9211
89.3 | 0 | 150 | 3.62 | 8954.6 | 7.18 | 7.1584
88.1 | 0 | 150 | 3.74 | 9256.1 | 7.42 | 7.3957
86.9 | 0 | 150 | 3.82 | 9453.7 | 7.56 | 7.5539
85.7 | 0 | 150 | 3.98 | 9835.8 | 7.86 | 7.8703
84.5 | 0 | 150 | 4.11 | 10159.3 | 8.13 | 8.1273
83.3 | 0 | 150 | 4.23 | 10457 | 8.38 | 8.3646
82.1 | 0 | 150 | 4.38 | 10835.2 | 8.68 | 8.6613
80.9 | 0 | 150 | 4.47 | 11062.9 | 8.86 | 8.8392
79.7 | 0 | 150 | 4.6 | 11364 | 9.1 | 9.0963
78.4 | 0 | 150 | 4.85 | 11985.9 | 9.6 | 9.5907
77.3 | 0 | 150 | 4.98 | 12318.6 | 9.87 | 9.8477
76 | 0 | 150 | 5.14 | 12699.1 | 10.16 | 10.1641
74.8 | 0 | 150 | 5.3 | 13114.4 | 10.49 | 10.4805
73.7 | 0 | 150 | 5.46 | 13487.7 | 10.8 | 10.7969
72.4 | 0 | 150 | 5.69 | 14070.3 | 11.26 | 11.2517
71.3 | 0 | 150 | 5.89 | 14568.8 | 11.65 | 11.6472
70 | 0 | 150 | 6.18 | 15279.1 | 12.21 | 12.2207
68.8 | 0 | 150 | 6.43 | 15891.2 | 12.7 | 12.715
67.6 | 0 | 150 | 6.69 | 16545.2 | 13.25 | 13.2292
66.4 | 0 | 150 | 7.06 | 17449.4 | 13.98 | 13.9608
65.2 | 0 | 150 | 7.5 | 18543.5 | 14.82 | 14.8309
64 | 0 | 150 | 8.15 | 20147.3 | 16.13 | 16.1163
62.8 | 0 | 150 | 9.05 | 22374.1 | 17.89 | 17.896
61.6 | 0 | 150 | 10.74 | 26543 | 21.26 | 21.2379
60.4 | 0 | 150 | 13.78 | 34071.4 | 27.29 | 27.2494
59.2 | 0 | 150 | 18.06 | 44648.1 | 35.73 | 35.7129
58 | 0 | 150 | 23.64 | 58444.8 | 46.76 | 46.7471
56.8 | 0 | 150 | 29.43 | 72763.1 | 58.22 | 58.1966
55.6 | 0 | 150 | 34.43 | 85114 | 68.15 | 68.0838
54.3 | 0 | 150 | 38.96 | 96319 | 76.98 | 77.0417
53.2 | 0 | 150 | 41.96 | 103744.8 | 82.91 | 82.9741
51.9 | 0 | 150 | 44.28 | 109466.3 | 87.69 | 87.5618
50.8 | 0 | 150 | 45.57 | 112655.6 | 90.05 | 90.1127
49.6 | 0 | 150 | 46.21 | 114255.1 | 91.29 | 91.3783
48.3 | 0 | 150 | 46.95 | 116067.5 | 92.74 | 92.8416
47.2 | 0 | 150 | 47.11 | 116464.9 | 93.37 | 93.158
45.9 | 0 | 150 | 47.33 | 117014.1 | 93.78 | 93.593
44.7 | 0 | 150 | 47.54 | 117532.5 | 94.04 | 94.0083
43.5 | 0 | 150 | 47.77 | 118097.4 | 94.51 | 94.4631
42.3 | 0 | 150 | 48.05 | 118803.7 | 95.22 | 95.0168
41.2 | 0 | 150 | 48.21 | 119186.2 | 95.41 | 95.3332
39.9 | 0 | 150 | 48.34 | 119503.2 | 95.72 | 95.5903
38.7 | 0 | 150 | 48.53 | 119988.8 | 96.11 | 95.966
37.5 | 0 | 150 | 48.67 | 120320.3 | 96.17 | 96.2428
36.3 | 0 | 150 | 48.72 | 120448 | 96.44 | 96.3417
35.1 | 0 | 150 | 48.94 | 120999.1 | 96.82 | 96.7767
33.9 | 0 | 150 | 49 | 121154.1 | 96.76 | 96.8954
32.7 | 0 | 150 | 49.14 | 121485.1 | 97.24 | 97.1722
31.5 | 0 | 150 | 49.3 | 121884.9 | 97.67 | 97.4886
30.3 | 0 | 150 | 49.42 | 122184.2 | 97.85 | 97.7259
29.1 | 0 | 150 | 49.59 | 122601.9 | 98.18 | 98.0621
27.9 | 0 | 150 | 49.64 | 122732.8 | 98.25 | 98.161
26.7 | 0 | 150 | 49.76 | 123015.3 | 98.49 | 98.3983
25.5 | 0 | 150 | 49.73 | 122942.4 | 98.5 | 98.3389
24.2 | 0 | 150 | 49.83 | 123196.6 | 99.01 | 98.5367
23.1 | 0 | 150 | 49.99 | 123603.6 | 98.91 | 98.8531
21.8 | 0 | 150 | 50.05 | 123740.7 | 98.98 | 98.9717
20.7 | 0 | 150 | 50.1 | 123853.2 | 99.08 | 99.0706
19.5 | 0 | 150 | 50.17 | 124046.3 | 99.04 | 99.209
18.2 | 0 | 150 | 50.19 | 124082.1 | 99.09 | 99.2486
17.1 | 0 | 150 | 50.12 | 123914 | 99.2 | 99.1101
15.8 | 0 | 150 | 50.24 | 124206.9 | 99.17 | 99.3474
14.7 | 0 | 150 | 50.05 | 123750.4 | 99.23 | 98.9717
13.4 | 0 | 150 | 50.17 | 124048.4 | 99.29 | 99.209
12.2 | 0 | 150 | 50.22 | 124152.1 | 99.32 | 99.3079
11 | 0 | 150 | 50.16 | 124006 | 99.23 | 99.1892
9.8 | 0 | 150 | 50.29 | 124331.2 | 99.53 | 99.4463
8.6 | 0 | 150 | 50.25 | 124226.1 | 99.65 | 99.3672
7.4 | 0 | 150 | 50.18 | 124074.3 | 99.26 | 99.2288
6.2 | 0 | 150 | 50.34 | 124450.1 | 99.43 | 99.5452
5 | 0 | 150 | 50.26 | 124272 | 99.46 | 99.387
3.8 | 0 | 150 | 50.34 | 124452.9 | 99.55 | 99.5452
2.6 | 0 | 150 | 50.22 | 124161.4 | 99.64 | 99.3079
1.4 | 0 | 150 | 50.35 | 124473.6 | 99.71 | 99.565
0.2 | 0 | 150 | 50.33 | 124435.4 | 99.62 | 99.5254
-1 | 0 | 150 | 50.41 | 124642.1 | 99.81 | 99.6836
-2.2 | 0 | 150 | 50.3 | 124353.5 | 99.56 | 99.4661
-3.4 | 0 | 150 | 50.37 | 124531.1 | 99.6 | 99.6045
-4.6 | 0 | 150 | 50.41 | 124642.5 | 99.63 | 99.6836
-5.8 | 0 | 150 | 50.46 | 124745 | 100 | 99.7825
-7 | 0 | 150 | 50.43 | 124690.7 | 99.76 | 99.7232
-8.3 | 0 | 150 | 50.4 | 124613.8 | 99.66 | 99.6638
-9.4 | 0 | 150 | 50.36 | 124515.6 | 99.79 | 99.5847
-10.7 | 0 | 150 | 50.52 | 124911.3 | 99.93 | 99.9011
-11.8 | 0 | 150 | 50.47 | 124770.8 | 100 | 99.8023
-13.1 | 0 | 150 | 50.44 | 124703.2 | 99.74 | 99.7429
-14.3 | 0 | 150 | 50.49 | 124825.9 | 99.87 | 99.8418
-15.5 | 0 | 150 | 50.57 | 125017 | 99.9 | 100
-16.7 | 0 | 150 | 50.46 | 124767.2 | 99.9 | 99.7825
-17.9 | 0 | 150 | 50.56 | 125001.4 | 99.91 | 99.9802
-19.1 | 0 | 150 | 50.55 | 124981.3 | 99.92 | 99.9605
-20.3 | 0 | 150 | 50.46 | 124754.3 | 99.7 | 99.7825
-21.5 | 0 | 150 | 50.39 | 124575.5 | 99.74 | 99.6441
-22.7 | 0 | 150 | 50.32 | 124399.3 | 99.55 | 99.5056
-23.9 | 0 | 150 | 50.34 | 124468.2 | 99.5 | 99.5452
-25.1 | 0 | 150 | 50.36 | 124518.9 | 99.48 | 99.5847
-26.3 | 0 | 150 | 50.11 | 123901.8 | 99.4 | 99.0904
-27.5 | 0 | 150 | 50.16 | 124021.7 | 99.26 | 99.1892
-28.7 | 0 | 150 | 50.05 | 123734 | 99.09 | 98.9717
-29.9 | 0 | 150 | 49.97 | 123538.2 | 98.91 | 98.8135
-31.1 | 0 | 150 | 49.94 | 123474.6 | 98.76 | 98.7542
-32.4 | 0 | 150 | 49.85 | 123257.6 | 98.57 | 98.5762
-33.5 | 0 | 150 | 49.61 | 122647.5 | 98.24 | 98.1016
-34.7 | 0 | 150 | 49.56 | 122523.9 | 97.87 | 98.0028
-36 | 0 | 150 | 49.55 | 122512.5 | 97.92 | 97.983
-37.1 | 0 | 150 | 49.29 | 121864.8 | 97.49 | 97.4689
-38.4 | 0 | 150 | 49.22 | 121700.8 | 97.27 | 97.3304
-39.5 | 0 | 150 | 48.99 | 121112.3 | 96.97 | 96.8756
-40.8 | 0 | 150 | 48.98 | 121085 | 96.85 | 96.8558
-42 | 0 | 150 | 48.82 | 120697.8 | 96.68 | 96.5395
-43.1 | 0 | 150 | 48.7 | 120402.6 | 96.17 | 96.3022
-44.4 | 0 | 150 | 48.51 | 119937.7 | 95.93 | 95.9264
-45.6 | 0 | 150 | 48.29 | 119380.7 | 95.54 | 95.4914
-46.8 | 0 | 150 | 48.03 | 118737.6 | 95.17 | 94.9773
-48 | 0 | 150 | 47.75 | 118062.9 | 94.44 | 94.4236
-49.2 | 0 | 150 | 47.5 | 117448.5 | 93.94 | 93.9292
-50.4 | 0 | 150 | 47 | 116201.1 | 92.89 | 92.9405
-51.6 | 0 | 150 | 46.18 | 114166.4 | 91.33 | 91.319
-52.8 | 0 | 150 | 44.88 | 110967.7 | 88.61 | 88.7483
-54 | 0 | 150 | 42.29 | 104548 | 83.62 | 83.6267
-55.2 | 0 | 150 | 38.35 | 94814.9 | 75.83 | 75.8355
-56.5 | 0 | 150 | 33 | 81588.1 | 65.2 | 65.2561
-57.6 | 0 | 150 | 27.38 | 67694.5 | 54.16 | 54.1428
-58.8 | 0 | 150 | 21.92 | 54183.5 | 43.42 | 43.3459
-60 | 0 | 150 | 17.08 | 42223.4 | 33.77 | 33.775
-61.2 | 0 | 150 | 13.49 | 33343.4 | 26.64 | 26.6759
-62.5 | 0 | 150 | 10.79 | 26684.5 | 21.36 | 21.3368
-63.7 | 0 | 150 | 9.25 | 22857.1 | 18.3 | 18.2915
-64.9 | 0 | 150 | 8.18 | 20233.3 | 16.18 | 16.1756
-66.1 | 0 | 150 | 7.64 | 18898.7 | 15.1 | 15.1078
-67.3 | 0 | 150 | 7.16 | 17713.2 | 14.17 | 14.1586
-68.5 | 0 | 150 | 6.8 | 16803.5 | 13.45 | 13.4467
-69.7 | 0 | 150 | 6.49 | 16042.2 | 12.83 | 12.8337
-70.9 | 0 | 150 | 6.24 | 15423.3 | 12.36 | 12.3393
-72.1 | 0 | 150 | 5.99 | 14820 | 11.86 | 11.845
-73.3 | 0 | 150 | 5.82 | 14392 | 11.5 | 11.5088
-74.5 | 0 | 150 | 5.57 | 13780.2 | 11.02 | 11.0144
-75.7 | 0 | 150 | 5.41 | 13374.3 | 10.69 | 10.698
-76.9 | 0 | 150 | 5.19 | 12828.9 | 10.26 | 10.263
-78.1 | 0 | 150 | 5.11 | 12630.8 | 10.12 | 10.1048
-79.3 | 0 | 150 | 4.88 | 12064.3 | 9.64 | 9.65
-80.5 | 0 | 150 | 4.74 | 11727.1 | 9.39 | 9.3731
-81.7 | 0 | 150 | 4.59 | 11347.7 | 9.08 | 9.0765
-83 | 0 | 150 | 4.45 | 10998.9 | 8.81 | 8.7997
-84.2 | 0 | 150 | 4.3 | 10640.7 | 8.5 | 8.5031
-85.3 | 0 | 150 | 4.21 | 10410.8 | 8.31 | 8.3251
-86.6 | 0 | 150 | 4.03 | 9956.8 | 7.94 | 7.9692
-87.7 | 0 | 150 | 3.96 | 9794.7 | 7.84 | 7.8307
-89 | 0 | 150 | 3.83 | 9480.1 | 7.59 | 7.5737
-90.2 | 0 | 150 | 3.69 | 9132 | 7.3 | 7.2968
-91.3 | 0 | 150 | 3.61 | 8920 | 7.13 | 7.1386
-92.6 | 0 | 150 | 3.45 | 8521.9 | 6.82 | 6.8222
-93.7 | 0 | 150 | 3.35 | 8289.6 | 6.62 | 6.6245
-95 | 0 | 150 | 3.28 | 8119.1 | 6.5 | 6.4861
-96.2 | 0 | 150 | 3.15 | 7786.4 | 6.23 | 6.229
-97.4 | 0 | 150 | 3.09 | 7634.2 | 6.12 | 6.1103
-98.6 | 0 | 150 | 3 | 7429.3 | 5.94 | 5.9324
-100 | 0 | 150 | 2.91 | 7193 | 5.75 | 5.7544