## 7_Rendements DSP
Measurement time: | 12/08/2022 18:43:49
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1100 mm
Field size: | 90 x 90 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | dose relative (%)
0 | 0 | 310 | 17.36 | 42929.3 | 21.02 | 21.1218
0 | 0 | 309.1 | 17.32 | 42809.3 | 21.08 | 21.0731
0 | 0 | 307.6 | 17.53 | 43336.9 | 21.33 | 21.3286
0 | 0 | 306.6 | 17.54 | 43377.3 | 21.31 | 21.3408
0 | 0 | 305.3 | 17.72 | 43805.1 | 21.56 | 21.5598
0 | 0 | 304.1 | 17.81 | 44030.7 | 21.65 | 21.6693
0 | 0 | 302.9 | 17.85 | 44128.6 | 21.75 | 21.718
0 | 0 | 301.7 | 17.93 | 44328.1 | 21.86 | 21.8153
0 | 0 | 300.5 | 18.1 | 44753.5 | 22.01 | 22.0221
0 | 0 | 299.3 | 18.26 | 45141.2 | 22.22 | 22.2168
0 | 0 | 298.1 | 18.37 | 45422.1 | 22.33 | 22.3507
0 | 0 | 296.9 | 18.48 | 45690.5 | 22.51 | 22.4845
0 | 0 | 295.7 | 18.62 | 46023.3 | 22.68 | 22.6548
0 | 0 | 294.5 | 18.69 | 46219.5 | 22.78 | 22.74
0 | 0 | 293.3 | 18.88 | 46670.4 | 22.95 | 22.9712
0 | 0 | 292.1 | 18.95 | 46848.1 | 23.05 | 23.0563
0 | 0 | 290.9 | 19.14 | 47320.7 | 23.31 | 23.2875
0 | 0 | 289.7 | 19.3 | 47707.1 | 23.49 | 23.4822
0 | 0 | 288.5 | 19.39 | 47947.8 | 23.65 | 23.5917
0 | 0 | 287.3 | 19.56 | 48368.3 | 23.77 | 23.7985
0 | 0 | 286.1 | 19.56 | 48356.3 | 23.79 | 23.7985
0 | 0 | 284.9 | 19.74 | 48801.6 | 24.01 | 24.0175
0 | 0 | 283.7 | 19.93 | 49261.6 | 24.24 | 24.2487
0 | 0 | 282.5 | 19.97 | 49364.5 | 24.33 | 24.2974
0 | 0 | 281.2 | 20.18 | 49884.1 | 24.56 | 24.5529
0 | 0 | 280.1 | 20.26 | 50087.2 | 24.72 | 24.6502
0 | 0 | 278.8 | 20.49 | 50647.2 | 24.91 | 24.93
0 | 0 | 277.7 | 20.61 | 50962.4 | 25.08 | 25.076
0 | 0 | 276.4 | 20.78 | 51370.3 | 25.27 | 25.2829
0 | 0 | 275.2 | 20.89 | 51637.6 | 25.37 | 25.4167
0 | 0 | 274.1 | 21.01 | 51951.4 | 25.56 | 25.5627
0 | 0 | 272.8 | 21.13 | 52232.8 | 25.73 | 25.7087
0 | 0 | 271.6 | 21.29 | 52641.1 | 25.93 | 25.9034
0 | 0 | 270.4 | 21.42 | 52949.5 | 26.06 | 26.0616
0 | 0 | 269.2 | 21.63 | 53484.5 | 26.32 | 26.3171
0 | 0 | 268.1 | 21.68 | 53605.1 | 26.39 | 26.3779
0 | 0 | 266.8 | 21.88 | 54099.8 | 26.62 | 26.6212
0 | 0 | 265.6 | 22 | 54383.4 | 26.81 | 26.7672
0 | 0 | 264.4 | 22.16 | 54784.3 | 26.92 | 26.9619
0 | 0 | 263.2 | 22.3 | 55129.8 | 27.12 | 27.1323
0 | 0 | 262 | 22.44 | 55479.5 | 27.24 | 27.3026
0 | 0 | 260.8 | 22.62 | 55918.5 | 27.47 | 27.5216
0 | 0 | 259.6 | 22.8 | 56371.4 | 27.74 | 27.7406
0 | 0 | 258.4 | 22.88 | 56579.6 | 27.85 | 27.8379
0 | 0 | 257.2 | 23.08 | 57050.9 | 28.08 | 28.0813
0 | 0 | 256 | 23.22 | 57403.8 | 28.26 | 28.2516
0 | 0 | 254.8 | 23.35 | 57731.9 | 28.44 | 28.4098
0 | 0 | 253.6 | 23.51 | 58135.6 | 28.65 | 28.6045
0 | 0 | 252.4 | 23.72 | 58654.5 | 28.86 | 28.86
0 | 0 | 251.2 | 23.9 | 59085.2 | 29.09 | 29.079
0 | 0 | 250 | 24.03 | 59407.9 | 29.24 | 29.2371
0 | 0 | 248.7 | 24.17 | 59757.6 | 29.36 | 29.4075
0 | 0 | 247.6 | 24.25 | 59955.7 | 29.52 | 29.5048
0 | 0 | 246.4 | 24.54 | 60677.8 | 29.89 | 29.8576
0 | 0 | 245.1 | 24.66 | 60968.1 | 30 | 30.0037
0 | 0 | 244 | 24.8 | 61308.2 | 30.24 | 30.174
0 | 0 | 242.7 | 25 | 61813.2 | 30.46 | 30.4173
0 | 0 | 241.6 | 25.17 | 62233 | 30.62 | 30.6242
0 | 0 | 240.3 | 25.36 | 62696.3 | 30.86 | 30.8553
0 | 0 | 239.1 | 25.52 | 63100 | 31.04 | 31.05
0 | 0 | 238 | 25.69 | 63505.3 | 31.29 | 31.2568
0 | 0 | 236.7 | 25.85 | 63906.1 | 31.5 | 31.4515
0 | 0 | 235.5 | 26 | 64275.6 | 31.67 | 31.634
0 | 0 | 234.3 | 26.26 | 64933.2 | 31.98 | 31.9504
0 | 0 | 233.1 | 26.41 | 65294 | 32.15 | 32.1329
0 | 0 | 231.9 | 26.55 | 65645.1 | 32.33 | 32.3032
0 | 0 | 230.7 | 26.71 | 66026.2 | 32.54 | 32.4979
0 | 0 | 229.5 | 26.91 | 66536.1 | 32.7 | 32.7412
0 | 0 | 228.3 | 27.2 | 67248 | 33.12 | 33.0941
0 | 0 | 227.1 | 27.33 | 67562.7 | 33.28 | 33.2522
0 | 0 | 225.9 | 27.47 | 67906.9 | 33.44 | 33.4226
0 | 0 | 224.7 | 27.66 | 68378 | 33.69 | 33.6537
0 | 0 | 223.5 | 27.91 | 68992.8 | 34.01 | 33.9579
0 | 0 | 222.3 | 28.1 | 69474.6 | 34.23 | 34.1891
0 | 0 | 221.1 | 28.28 | 69907.1 | 34.4 | 34.4081
0 | 0 | 219.9 | 28.48 | 70409.6 | 34.71 | 34.6514
0 | 0 | 218.6 | 28.68 | 70902.9 | 34.94 | 34.8948
0 | 0 | 217.5 | 28.91 | 71476.9 | 35.18 | 35.1746
0 | 0 | 216.3 | 29.04 | 71802 | 35.31 | 35.3328
0 | 0 | 215.1 | 29.29 | 72415.4 | 35.64 | 35.6369
0 | 0 | 213.9 | 29.34 | 72535.8 | 35.7 | 35.6978
0 | 0 | 212.6 | 29.58 | 73142 | 36.04 | 35.9898
0 | 0 | 211.4 | 29.81 | 73689.5 | 36.26 | 36.2696
0 | 0 | 210.2 | 30.01 | 74202.9 | 36.55 | 36.513
0 | 0 | 209 | 30.26 | 74811.3 | 36.82 | 36.8171
0 | 0 | 207.8 | 30.36 | 75055.1 | 36.88 | 36.9388
0 | 0 | 206.6 | 30.53 | 75490.4 | 37.19 | 37.1456
0 | 0 | 205.4 | 30.8 | 76145.5 | 37.47 | 37.4741
0 | 0 | 204.2 | 30.99 | 76613.5 | 37.72 | 37.7053
0 | 0 | 203 | 31.22 | 77196.1 | 37.97 | 37.9852
0 | 0 | 201.8 | 31.44 | 77728.5 | 38.23 | 38.2528
0 | 0 | 200.6 | 31.68 | 78322.8 | 38.55 | 38.5448
0 | 0 | 199.4 | 31.9 | 78855.8 | 38.88 | 38.8125
0 | 0 | 198.2 | 32.12 | 79407.9 | 39.03 | 39.0802
0 | 0 | 197 | 32.24 | 79700.9 | 39.2 | 39.2262
0 | 0 | 195.8 | 32.5 | 80346.8 | 39.66 | 39.5425
0 | 0 | 194.6 | 32.74 | 80953.9 | 39.76 | 39.8345
0 | 0 | 193.4 | 32.93 | 81406 | 40.18 | 40.0657
0 | 0 | 192.2 | 33.15 | 81961.7 | 40.34 | 40.3334
0 | 0 | 191 | 33.44 | 82671.4 | 40.74 | 40.6862
0 | 0 | 189.8 | 33.61 | 83086.9 | 40.88 | 40.8931
0 | 0 | 188.6 | 33.9 | 83808.4 | 41.25 | 41.2459
0 | 0 | 187.4 | 34.16 | 84465.6 | 41.57 | 41.5622
0 | 0 | 186.2 | 34.33 | 84886.5 | 41.72 | 41.7691
0 | 0 | 185 | 34.52 | 85350.6 | 42.04 | 42.0002
0 | 0 | 183.8 | 34.84 | 86124.8 | 42.32 | 42.3896
0 | 0 | 182.5 | 35.07 | 86707 | 42.56 | 42.6694
0 | 0 | 181.3 | 35.31 | 87300.4 | 43.01 | 42.9614
0 | 0 | 180.1 | 35.43 | 87605.4 | 43.04 | 43.1074
0 | 0 | 178.9 | 35.66 | 88174 | 43.46 | 43.3873
0 | 0 | 177.7 | 36.01 | 89020.2 | 43.81 | 43.8131
0 | 0 | 176.5 | 36.21 | 89514.9 | 44.04 | 44.0565
0 | 0 | 175.3 | 36.44 | 90084.4 | 44.38 | 44.3363
0 | 0 | 174.1 | 36.7 | 90736.4 | 44.69 | 44.6526
0 | 0 | 172.9 | 36.89 | 91203.6 | 44.94 | 44.8838
0 | 0 | 171.7 | 37.13 | 91805.2 | 45.29 | 45.1758
0 | 0 | 170.5 | 37.36 | 92374.8 | 45.52 | 45.4557
0 | 0 | 169.3 | 37.55 | 92842.5 | 45.8 | 45.6868
0 | 0 | 168.1 | 37.88 | 93644.4 | 46.11 | 46.0883
0 | 0 | 166.9 | 38.16 | 94332.7 | 46.46 | 46.429
0 | 0 | 165.7 | 38.52 | 95223.1 | 46.89 | 46.867
0 | 0 | 164.5 | 38.83 | 95998.6 | 47.13 | 47.2442
0 | 0 | 163.3 | 38.97 | 96340.2 | 47.41 | 47.4145
0 | 0 | 162.1 | 39.17 | 96832.9 | 47.66 | 47.6579
0 | 0 | 160.9 | 39.59 | 97887.3 | 48.12 | 48.1689
0 | 0 | 159.7 | 39.79 | 98378.1 | 48.4 | 48.4122
0 | 0 | 158.4 | 40 | 98903.2 | 48.66 | 48.6677
0 | 0 | 157.2 | 40.3 | 99645.9 | 49.1 | 49.0327
0 | 0 | 156.1 | 40.66 | 100521.4 | 49.48 | 49.4707
0 | 0 | 154.8 | 40.79 | 100836.9 | 49.72 | 49.6289
0 | 0 | 153.6 | 41.09 | 101586.9 | 50.01 | 49.9939
0 | 0 | 152.4 | 41.46 | 102493.6 | 50.55 | 50.4441
0 | 0 | 151.2 | 41.74 | 103189.8 | 50.78 | 50.7848
0 | 0 | 150 | 41.96 | 103731.9 | 51.1 | 51.0524
0 | 0 | 148.8 | 42.2 | 104328.1 | 51.37 | 51.3444
0 | 0 | 147.6 | 42.43 | 104903.9 | 51.66 | 51.6243
0 | 0 | 146.4 | 42.79 | 105794.6 | 52.14 | 52.0623
0 | 0 | 145.2 | 43.01 | 106339.5 | 52.39 | 52.33
0 | 0 | 144 | 43.27 | 106970.7 | 52.64 | 52.6463
0 | 0 | 142.8 | 43.65 | 107907.8 | 53.06 | 53.1087
0 | 0 | 141.6 | 43.88 | 108481.2 | 53.35 | 53.3885
0 | 0 | 140.4 | 44.19 | 109259.8 | 53.8 | 53.7657
0 | 0 | 139.2 | 44.5 | 110024.4 | 54.17 | 54.1428
0 | 0 | 138 | 44.84 | 110859.6 | 54.58 | 54.5565
0 | 0 | 136.8 | 45.06 | 111411.7 | 54.88 | 54.8242
0 | 0 | 135.6 | 45.36 | 112145.1 | 55.19 | 55.1892
0 | 0 | 134.4 | 45.8 | 113236.7 | 55.66 | 55.7245
0 | 0 | 133.2 | 45.98 | 113674.6 | 56.07 | 55.9435
0 | 0 | 132 | 46.29 | 114454.7 | 56.44 | 56.3207
0 | 0 | 130.8 | 46.68 | 115407.9 | 56.76 | 56.7952
0 | 0 | 129.6 | 46.94 | 116049.4 | 57.09 | 57.1116
0 | 0 | 128.4 | 47.36 | 117082.9 | 57.57 | 57.6226
0 | 0 | 127.2 | 47.56 | 117589.5 | 57.95 | 57.8659
0 | 0 | 126 | 47.9 | 118430.6 | 58.22 | 58.2796
0 | 0 | 124.8 | 48.2 | 119161.4 | 58.72 | 58.6446
0 | 0 | 123.6 | 48.53 | 119993.1 | 59.05 | 59.0461
0 | 0 | 122.3 | 48.91 | 120933.1 | 59.54 | 59.5085
0 | 0 | 121.2 | 49.26 | 121783.4 | 60.02 | 59.9343
0 | 0 | 120 | 49.52 | 122438 | 60.25 | 60.2506
0 | 0 | 118.8 | 49.83 | 123197.8 | 60.7 | 60.6278
0 | 0 | 117.6 | 50.21 | 124128.3 | 61.15 | 61.0902
0 | 0 | 116.3 | 50.43 | 124675.7 | 61.4 | 61.3578
0 | 0 | 115.2 | 50.84 | 125684.7 | 61.96 | 61.8567
0 | 0 | 113.9 | 51.15 | 126454.2 | 62.31 | 62.2338
0 | 0 | 112.7 | 51.52 | 127379.3 | 62.78 | 62.684
0 | 0 | 111.5 | 51.84 | 128161.1 | 63.1 | 63.0734
0 | 0 | 110.3 | 52.2 | 129046 | 63.53 | 63.5114
0 | 0 | 109.1 | 52.52 | 129841.5 | 63.94 | 63.9007
0 | 0 | 107.9 | 52.83 | 130616.4 | 64.26 | 64.2779
0 | 0 | 106.7 | 53.18 | 131489 | 64.71 | 64.7037
0 | 0 | 105.5 | 53.52 | 132309.7 | 65.21 | 65.1174
0 | 0 | 104.3 | 53.85 | 133146.9 | 65.54 | 65.5189
0 | 0 | 103.1 | 54.16 | 133906.2 | 65.89 | 65.8961
0 | 0 | 101.9 | 54.6 | 134978.3 | 66.38 | 66.4314
0 | 0 | 100.7 | 54.92 | 135788.7 | 66.9 | 66.8208
0 | 0 | 99.5 | 55.3 | 136709.3 | 67.27 | 67.2831
0 | 0 | 98.3 | 55.57 | 137376.7 | 67.5 | 67.6116
0 | 0 | 97.1 | 56.06 | 138607 | 68.2 | 68.2078
0 | 0 | 95.9 | 56.39 | 139413 | 68.57 | 68.6093
0 | 0 | 94.7 | 56.76 | 140322.5 | 69 | 69.0595
0 | 0 | 93.5 | 57.12 | 141214.9 | 69.37 | 69.4975
0 | 0 | 92.3 | 57.44 | 142019.2 | 69.81 | 69.8868
0 | 0 | 91.1 | 57.84 | 143001.2 | 70.38 | 70.3735
0 | 0 | 89.9 | 58.24 | 143997.4 | 70.94 | 70.8602
0 | 0 | 88.7 | 58.52 | 144684.7 | 71.17 | 71.2009
0 | 0 | 87.5 | 58.99 | 145838.9 | 71.77 | 71.7727
0 | 0 | 86.3 | 59.29 | 146583.9 | 72.19 | 72.1377
0 | 0 | 85.1 | 59.82 | 147900.2 | 72.79 | 72.7826
0 | 0 | 83.8 | 60.11 | 148604.9 | 73.23 | 73.1354
0 | 0 | 82.6 | 60.47 | 149498 | 73.6 | 73.5734
0 | 0 | 81.5 | 60.84 | 150427.7 | 73.99 | 74.0236
0 | 0 | 80.2 | 61.22 | 151353.6 | 74.55 | 74.4859
0 | 0 | 79 | 61.56 | 152196.8 | 74.92 | 74.8996
0 | 0 | 77.8 | 62.1 | 153525.3 | 75.84 | 75.5566
0 | 0 | 76.6 | 62.37 | 154207 | 75.92 | 75.8851
0 | 0 | 75.4 | 62.76 | 155164.7 | 76.39 | 76.3597
0 | 0 | 74.2 | 63.2 | 156240.7 | 76.96 | 76.895
0 | 0 | 73 | 63.66 | 157386.1 | 77.46 | 77.4547
0 | 0 | 71.8 | 64 | 158236.6 | 77.79 | 77.8684
0 | 0 | 70.6 | 64.26 | 158879.5 | 78.16 | 78.1847
0 | 0 | 69.3 | 64.81 | 160243 | 78.82 | 78.8539
0 | 0 | 68.2 | 65.13 | 161021.6 | 79.26 | 79.2432
0 | 0 | 66.9 | 65.5 | 161950.7 | 79.75 | 79.6934
0 | 0 | 65.7 | 65.86 | 162825.2 | 80.3 | 80.1314
0 | 0 | 64.5 | 66.28 | 163863.7 | 80.64 | 80.6424
0 | 0 | 63.3 | 66.7 | 164914.8 | 81.08 | 81.1534
0 | 0 | 62.1 | 67.2 | 166143.7 | 81.68 | 81.7618
0 | 0 | 60.9 | 67.5 | 166894.5 | 82.09 | 82.1268
0 | 0 | 59.7 | 67.97 | 168051.7 | 82.79 | 82.6986
0 | 0 | 58.5 | 68.24 | 168712.4 | 83.09 | 83.0271
0 | 0 | 57.3 | 68.76 | 170007.6 | 83.74 | 83.6598
0 | 0 | 56 | 69.12 | 170893.1 | 84.29 | 84.0978
0 | 0 | 54.9 | 69.45 | 171709.7 | 84.48 | 84.4993
0 | 0 | 53.6 | 69.97 | 172995.8 | 85.23 | 85.132
0 | 0 | 52.5 | 70.24 | 173649.1 | 85.47 | 85.4605
0 | 0 | 51.2 | 70.62 | 174590.9 | 86.09 | 85.9229
0 | 0 | 50 | 71.12 | 175829.8 | 86.55 | 86.5312
0 | 0 | 48.8 | 71.6 | 177020.8 | 87.22 | 87.1152
0 | 0 | 47.6 | 72.1 | 178248.7 | 87.69 | 87.7236
0 | 0 | 46.4 | 72.35 | 178881.1 | 88.12 | 88.0277
0 | 0 | 45.2 | 72.77 | 179924.3 | 88.58 | 88.5388
0 | 0 | 44 | 73.17 | 180897.8 | 89.23 | 89.0254
0 | 0 | 42.8 | 73.58 | 181908.7 | 89.84 | 89.5243
0 | 0 | 41.6 | 74.07 | 183131.5 | 90.17 | 90.1205
0 | 0 | 40.4 | 74.54 | 184296.9 | 90.63 | 90.6923
0 | 0 | 39.1 | 74.96 | 185323.2 | 91.33 | 91.2033
0 | 0 | 38 | 75.39 | 186402.6 | 91.89 | 91.7265
0 | 0 | 36.8 | 75.82 | 187444.6 | 92.06 | 92.2497
0 | 0 | 35.5 | 76.14 | 188246.9 | 92.66 | 92.639
0 | 0 | 34.3 | 76.57 | 189307.4 | 93.17 | 93.1622
0 | 0 | 33.1 | 76.96 | 190271.2 | 93.57 | 93.6367
0 | 0 | 31.9 | 77.24 | 190964.1 | 93.9 | 93.9774
0 | 0 | 30.7 | 77.76 | 192259.8 | 94.68 | 94.61
0 | 0 | 29.5 | 78.2 | 193334.9 | 95.17 | 95.1454
0 | 0 | 28.3 | 78.34 | 193693.4 | 95.32 | 95.3157
0 | 0 | 27.1 | 78.94 | 195162.7 | 95.98 | 96.0457
0 | 0 | 25.9 | 79.35 | 196181.3 | 96.48 | 96.5446
0 | 0 | 24.7 | 79.77 | 197211 | 97.01 | 97.0556
0 | 0 | 23.5 | 80.23 | 198346.5 | 97.36 | 97.6153
0 | 0 | 22.3 | 80.53 | 199108.9 | 98.07 | 97.9803
0 | 0 | 21 | 80.91 | 200047.1 | 98.43 | 98.4426
0 | 0 | 19.8 | 81.27 | 200924.2 | 99.12 | 98.8806
0 | 0 | 18.6 | 81.75 | 202112 | 99.48 | 99.4647
0 | 0 | 17.4 | 81.94 | 202581.9 | 99.42 | 99.6958
0 | 0 | 16.2 | 82.15 | 203108.7 | 100 | 99.9513
0 | 0 | 15 | 82.05 | 202857.2 | 99.99 | 99.8297
0 | 0 | 13.8 | 82.19 | 203203.7 | 99.78 | 100
0 | 0 | 12.6 | 82.07 | 202894.8 | 99.75 | 99.854
0 | 0 | 11.4 | 81.6 | 201755.2 | 99.21 | 99.2822
0 | 0 | 10.2 | 80.69 | 199483 | 98.06 | 98.175
0 | 0 | 9 | 79.42 | 196359.9 | 96.71 | 96.6298
0 | 0 | 7.8 | 77.44 | 191447.4 | 94.4 | 94.2207
0 | 0 | 6.6 | 74.59 | 184418.7 | 90.98 | 90.7531
0 | 0 | 5.3 | 70.25 | 173674.9 | 85.46 | 85.4727
0 | 0 | 4.1 | 63.95 | 158105.3 | 77.87 | 77.8075
0 | 0 | 2.9 | 55.58 | 137411 | 67.6 | 67.6238
0 | 0 | 1.7 | 47.68 | 117874.3 | 58.05 | 58.0119
0 | 0 | 0.5 | 43.35 | 107179.3 | 52.8 | 52.7436
0 | 0 | -0.5 | 40.75 | 100742.4 | 49.54 | 49.5802
Measurement time: | 12/08/2022 18:48:09
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 900 mm
Field size: | 110 x 110 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 22.31 | 55169 | 19.31 | 17.8694
0 | 0 | 309.1 | 24.25 | 59966.2 | 19.4 | 19.4233
0 | 0 | 307.6 | 24.45 | 60441.3 | 19.57 | 19.5835
0 | 0 | 306.6 | 24.68 | 61014.7 | 19.74 | 19.7677
0 | 0 | 305.3 | 24.83 | 61387.3 | 19.88 | 19.8879
0 | 0 | 304.1 | 24.97 | 61741.2 | 19.98 | 20
0 | 0 | 302.9 | 25.17 | 62233.7 | 20.16 | 20.1602
0 | 0 | 301.7 | 25.32 | 62604.4 | 20.26 | 20.2803
0 | 0 | 300.5 | 25.43 | 62864.9 | 20.34 | 20.3684
0 | 0 | 299.3 | 25.7 | 63551.6 | 20.56 | 20.5847
0 | 0 | 298.1 | 25.9 | 64037.7 | 20.73 | 20.7449
0 | 0 | 296.9 | 26.02 | 64332 | 20.82 | 20.841
0 | 0 | 295.7 | 26.21 | 64791.3 | 20.99 | 20.9932
0 | 0 | 294.5 | 26.4 | 65267.4 | 21.11 | 21.1454
0 | 0 | 293.3 | 26.53 | 65588.2 | 21.2 | 21.2495
0 | 0 | 292.1 | 26.75 | 66132 | 21.41 | 21.4257
0 | 0 | 290.9 | 26.96 | 66645.9 | 21.59 | 21.5939
0 | 0 | 289.7 | 27.13 | 67073.2 | 21.69 | 21.7301
0 | 0 | 288.5 | 27.24 | 67355 | 21.8 | 21.8182
0 | 0 | 287.2 | 27.43 | 67821.2 | 22.04 | 21.9704
0 | 0 | 286.1 | 27.69 | 68455.2 | 22.18 | 22.1786
0 | 0 | 284.8 | 27.9 | 68986.2 | 22.33 | 22.3468
0 | 0 | 283.6 | 28.08 | 69419 | 22.46 | 22.491
0 | 0 | 282.5 | 28.28 | 69912.5 | 22.61 | 22.6512
0 | 0 | 281.2 | 28.41 | 70230.1 | 22.71 | 22.7553
0 | 0 | 280.1 | 28.67 | 70888.2 | 22.95 | 22.9636
0 | 0 | 278.8 | 28.83 | 71266.5 | 23.08 | 23.0917
0 | 0 | 277.7 | 28.99 | 71684.1 | 23.2 | 23.2199
0 | 0 | 276.4 | 29.21 | 72205.2 | 23.39 | 23.3961
0 | 0 | 275.2 | 29.44 | 72780.6 | 23.58 | 23.5803
0 | 0 | 274.1 | 29.63 | 73252.6 | 23.75 | 23.7325
0 | 0 | 272.8 | 29.8 | 73682.3 | 23.87 | 23.8686
0 | 0 | 271.6 | 30.06 | 74312.5 | 24.07 | 24.0769
0 | 0 | 270.4 | 30.25 | 74787.9 | 24.21 | 24.2291
0 | 0 | 269.2 | 30.48 | 75361.4 | 24.39 | 24.4133
0 | 0 | 268 | 30.74 | 75989.8 | 24.62 | 24.6215
0 | 0 | 266.8 | 30.87 | 76329.5 | 24.72 | 24.7257
0 | 0 | 265.6 | 31.12 | 76937.6 | 24.92 | 24.9259
0 | 0 | 264.4 | 31.29 | 77368.7 | 25.06 | 25.0621
0 | 0 | 263.2 | 31.54 | 77977.6 | 25.24 | 25.2623
0 | 0 | 262 | 31.71 | 78389 | 25.36 | 25.3985
0 | 0 | 260.8 | 31.97 | 79038.5 | 25.59 | 25.6067
0 | 0 | 259.6 | 32.15 | 79494.3 | 25.73 | 25.7509
0 | 0 | 258.4 | 32.44 | 80194.7 | 25.95 | 25.9832
0 | 0 | 257.2 | 32.7 | 80852.5 | 26.16 | 26.1914
0 | 0 | 256 | 32.91 | 81356.6 | 26.37 | 26.3596
0 | 0 | 254.8 | 33.07 | 81763.3 | 26.5 | 26.4878
0 | 0 | 253.6 | 33.24 | 82177.2 | 26.65 | 26.6239
0 | 0 | 252.4 | 33.56 | 82965 | 26.89 | 26.8803
0 | 0 | 251.2 | 33.92 | 83853.5 | 27.16 | 27.1686
0 | 0 | 250 | 34.06 | 84205.3 | 27.27 | 27.2807
0 | 0 | 248.7 | 34.33 | 84876.7 | 27.49 | 27.497
0 | 0 | 247.6 | 34.51 | 85318.4 | 27.63 | 27.6412
0 | 0 | 246.3 | 34.76 | 85932.9 | 27.86 | 27.8414
0 | 0 | 245.1 | 35.09 | 86766.2 | 28.13 | 28.1057
0 | 0 | 243.9 | 35.35 | 87403.3 | 28.28 | 28.314
0 | 0 | 242.7 | 35.55 | 87903.2 | 28.49 | 28.4742
0 | 0 | 241.6 | 35.76 | 88416.2 | 28.64 | 28.6424
0 | 0 | 240.3 | 36.05 | 89128 | 28.81 | 28.8746
0 | 0 | 239.1 | 36.3 | 89741.5 | 29.1 | 29.0749
0 | 0 | 238 | 36.59 | 90475.3 | 29.34 | 29.3072
0 | 0 | 236.7 | 36.81 | 91006 | 29.45 | 29.4834
0 | 0 | 235.5 | 37.04 | 91574.8 | 29.66 | 29.6676
0 | 0 | 234.3 | 37.33 | 92282.7 | 29.91 | 29.8999
0 | 0 | 233.1 | 37.57 | 92878.1 | 30.07 | 30.0921
0 | 0 | 231.9 | 37.82 | 93493.7 | 30.3 | 30.2924
0 | 0 | 230.7 | 38.06 | 94087.4 | 30.43 | 30.4846
0 | 0 | 229.5 | 38.41 | 94956.6 | 30.79 | 30.7649
0 | 0 | 228.3 | 38.56 | 95337.6 | 30.91 | 30.8851
0 | 0 | 227.1 | 38.91 | 96206.6 | 31.2 | 31.1654
0 | 0 | 225.9 | 39.2 | 96928.2 | 31.39 | 31.3977
0 | 0 | 224.6 | 39.56 | 97799.9 | 31.67 | 31.686
0 | 0 | 223.5 | 39.64 | 97992.7 | 31.78 | 31.7501
0 | 0 | 222.2 | 40.03 | 98961.4 | 32.08 | 32.0625
0 | 0 | 221.1 | 40.39 | 99852.4 | 32.33 | 32.3508
0 | 0 | 219.9 | 40.56 | 100286 | 32.49 | 32.487
0 | 0 | 218.6 | 40.9 | 101115.6 | 32.77 | 32.7593
0 | 0 | 217.5 | 41.12 | 101663.9 | 32.97 | 32.9355
0 | 0 | 216.2 | 41.28 | 102050.3 | 33.06 | 33.0637
0 | 0 | 215.1 | 41.65 | 102973.7 | 33.4 | 33.36
0 | 0 | 213.8 | 41.93 | 103667.3 | 33.61 | 33.5843
0 | 0 | 212.6 | 42.26 | 104480.2 | 33.87 | 33.8486
0 | 0 | 211.4 | 42.52 | 105115.3 | 34.11 | 34.0569
0 | 0 | 210.2 | 42.96 | 106209.3 | 34.37 | 34.4093
0 | 0 | 209 | 43.25 | 106917.6 | 34.66 | 34.6416
0 | 0 | 207.9 | 43.51 | 107578.4 | 34.81 | 34.8498
0 | 0 | 206.6 | 43.85 | 108416.8 | 35.09 | 35.1221
0 | 0 | 205.4 | 44.05 | 108895.6 | 35.23 | 35.2823
0 | 0 | 204.2 | 44.39 | 109738 | 35.54 | 35.5547
0 | 0 | 203 | 44.7 | 110516.3 | 35.81 | 35.803
0 | 0 | 201.8 | 45.01 | 111272 | 36 | 36.0513
0 | 0 | 200.6 | 45.24 | 111839.1 | 36.24 | 36.2355
0 | 0 | 199.4 | 45.65 | 112866.6 | 36.52 | 36.5639
0 | 0 | 198.2 | 45.91 | 113516.7 | 36.85 | 36.7721
0 | 0 | 197 | 46.21 | 114239 | 37.03 | 37.0124
0 | 0 | 195.8 | 46.5 | 114955.9 | 37.32 | 37.2447
0 | 0 | 194.6 | 46.75 | 115594.7 | 37.45 | 37.4449
0 | 0 | 193.4 | 47.19 | 116661.1 | 37.86 | 37.7974
0 | 0 | 192.2 | 47.56 | 117576.2 | 38.15 | 38.0937
0 | 0 | 190.9 | 47.96 | 118580.9 | 38.42 | 38.4141
0 | 0 | 189.8 | 48.21 | 119195.3 | 38.63 | 38.6143
0 | 0 | 188.5 | 48.55 | 120037.1 | 38.98 | 38.8867
0 | 0 | 187.3 | 49.02 | 121193.7 | 39.29 | 39.2631
0 | 0 | 186.1 | 49.28 | 121833.4 | 39.47 | 39.4714
0 | 0 | 184.9 | 49.68 | 122831.6 | 39.86 | 39.7918
0 | 0 | 183.7 | 50.07 | 123784.6 | 40.11 | 40.1041
0 | 0 | 182.5 | 50.33 | 124438.8 | 40.36 | 40.3124
0 | 0 | 181.3 | 50.63 | 125185 | 40.58 | 40.5527
0 | 0 | 180.1 | 51.09 | 126321.3 | 40.94 | 40.9211
0 | 0 | 178.9 | 51.52 | 127374.3 | 41.27 | 41.2655
0 | 0 | 177.7 | 51.88 | 128262.1 | 41.59 | 41.5539
0 | 0 | 176.5 | 52.18 | 128995 | 41.85 | 41.7942
0 | 0 | 175.3 | 52.48 | 129748.4 | 42.07 | 42.0344
0 | 0 | 174.1 | 52.86 | 130677.4 | 42.4 | 42.3388
0 | 0 | 172.9 | 53.22 | 131572.9 | 42.62 | 42.6272
0 | 0 | 171.7 | 53.71 | 132786.2 | 43.02 | 43.0196
0 | 0 | 170.5 | 53.96 | 133400.1 | 43.23 | 43.2199
0 | 0 | 169.3 | 54.4 | 134495.8 | 43.6 | 43.5723
0 | 0 | 168.1 | 54.71 | 135264.1 | 43.83 | 43.8206
0 | 0 | 166.9 | 55.26 | 136625.1 | 44.3 | 44.2611
0 | 0 | 165.7 | 55.56 | 137355 | 44.54 | 44.5014
0 | 0 | 164.5 | 56.01 | 138484.8 | 44.88 | 44.8618
0 | 0 | 163.3 | 56.41 | 139464.7 | 45.21 | 45.1822
0 | 0 | 162.1 | 56.73 | 140249.3 | 45.54 | 45.4385
0 | 0 | 160.9 | 57.25 | 141531.5 | 45.89 | 45.855
0 | 0 | 159.7 | 57.6 | 142397.6 | 46.13 | 46.1354
0 | 0 | 158.4 | 58 | 143384.6 | 46.52 | 46.4557
0 | 0 | 157.3 | 58.34 | 144228 | 46.72 | 46.7281
0 | 0 | 156 | 58.8 | 145386.5 | 47.13 | 47.0965
0 | 0 | 154.8 | 59.22 | 146423.6 | 47.38 | 47.4329
0 | 0 | 153.6 | 59.69 | 147584.9 | 47.82 | 47.8094
0 | 0 | 152.4 | 59.93 | 148174.2 | 48.01 | 48.0016
0 | 0 | 151.2 | 60.28 | 149028.1 | 48.35 | 48.2819
0 | 0 | 150 | 60.92 | 150612.2 | 48.8 | 48.7946
0 | 0 | 148.8 | 61.39 | 151785.1 | 49.12 | 49.171
0 | 0 | 147.6 | 61.79 | 152773.7 | 49.4 | 49.4914
0 | 0 | 146.4 | 62.19 | 153752 | 49.88 | 49.8118
0 | 0 | 145.2 | 62.57 | 154704.3 | 50.12 | 50.1161
0 | 0 | 144 | 63.06 | 155901.9 | 50.45 | 50.5086
0 | 0 | 142.8 | 63.46 | 156890.6 | 50.87 | 50.829
0 | 0 | 141.6 | 63.8 | 157731.4 | 51.16 | 51.1013
0 | 0 | 140.4 | 64.37 | 159136.9 | 51.63 | 51.5579
0 | 0 | 139.2 | 64.78 | 160150.9 | 51.94 | 51.8863
0 | 0 | 138 | 65.11 | 160970 | 52.24 | 52.1506
0 | 0 | 136.8 | 65.64 | 162287.4 | 52.61 | 52.5751
0 | 0 | 135.6 | 66.1 | 163413.4 | 52.93 | 52.9435
0 | 0 | 134.4 | 66.68 | 164860.5 | 53.41 | 53.4081
0 | 0 | 133.2 | 67.05 | 165776.9 | 53.69 | 53.7044
0 | 0 | 132 | 67.48 | 166838.2 | 54.14 | 54.0489
0 | 0 | 130.8 | 67.99 | 168101.4 | 54.58 | 54.4573
0 | 0 | 129.6 | 68.51 | 169385.6 | 55 | 54.8738
0 | 0 | 128.4 | 68.88 | 170289.4 | 55.26 | 55.1702
0 | 0 | 127.2 | 69.65 | 172205.6 | 55.74 | 55.7869
0 | 0 | 126 | 69.9 | 172826.5 | 55.98 | 55.9872
0 | 0 | 124.8 | 70.53 | 174365.7 | 56.41 | 56.4918
0 | 0 | 123.6 | 70.93 | 175373.8 | 56.75 | 56.8122
0 | 0 | 122.3 | 71.55 | 176899.3 | 57.28 | 57.3088
0 | 0 | 121.2 | 71.92 | 177810.5 | 57.62 | 57.6051
0 | 0 | 119.9 | 72.25 | 178633.8 | 57.87 | 57.8694
0 | 0 | 118.7 | 72.86 | 180129.8 | 58.37 | 58.358
0 | 0 | 117.5 | 73.33 | 181290.3 | 58.76 | 58.7345
0 | 0 | 116.3 | 73.74 | 182315.7 | 59.16 | 59.0629
0 | 0 | 115.1 | 74.25 | 183584 | 59.49 | 59.4714
0 | 0 | 113.9 | 74.84 | 185039 | 59.88 | 59.9439
0 | 0 | 112.7 | 75.31 | 186185 | 60.35 | 60.3204
0 | 0 | 111.5 | 76.03 | 187977.8 | 60.88 | 60.8971
0 | 0 | 110.3 | 76.2 | 188403.5 | 61.12 | 61.0332
0 | 0 | 109.1 | 76.92 | 190164 | 61.62 | 61.6099
0 | 0 | 107.9 | 77.35 | 191239.8 | 62.09 | 61.9543
0 | 0 | 106.7 | 77.73 | 192185.7 | 62.35 | 62.2587
0 | 0 | 105.5 | 78.5 | 194078.3 | 62.88 | 62.8755
0 | 0 | 104.3 | 79.12 | 195607.7 | 63.39 | 63.372
0 | 0 | 103.1 | 79.63 | 196868.2 | 63.78 | 63.7805
0 | 0 | 101.9 | 80.12 | 198077.1 | 64.19 | 64.173
0 | 0 | 100.7 | 80.69 | 199501.7 | 64.61 | 64.6296
0 | 0 | 99.5 | 81.32 | 201046.7 | 65.11 | 65.1342
0 | 0 | 98.3 | 81.77 | 202169.9 | 65.54 | 65.4946
0 | 0 | 97.1 | 82.51 | 203982.8 | 66.09 | 66.0873
0 | 0 | 95.9 | 82.94 | 205067.6 | 66.53 | 66.4317
0 | 0 | 94.7 | 83.54 | 206552.1 | 66.92 | 66.9123
0 | 0 | 93.5 | 84.1 | 207919.6 | 67.43 | 67.3608
0 | 0 | 92.3 | 84.57 | 209085.2 | 67.92 | 67.7373
0 | 0 | 91.1 | 85.09 | 210377.5 | 68.2 | 68.1538
0 | 0 | 89.9 | 85.61 | 211662.3 | 68.68 | 68.5703
0 | 0 | 88.7 | 86.65 | 214217.4 | 69.34 | 69.4033
0 | 0 | 87.5 | 87.13 | 215418.1 | 69.72 | 69.7877
0 | 0 | 86.2 | 87.55 | 216451.3 | 70.14 | 70.1241
0 | 0 | 85.1 | 88.2 | 218057.5 | 70.6 | 70.6448
0 | 0 | 83.8 | 88.78 | 219488.9 | 71.17 | 71.1093
0 | 0 | 82.6 | 89.31 | 220804.3 | 71.7 | 71.5338
0 | 0 | 81.4 | 89.93 | 222335.9 | 72.14 | 72.0304
0 | 0 | 80.2 | 90.56 | 223907.2 | 72.52 | 72.535
0 | 0 | 79 | 91.03 | 225059.5 | 72.84 | 72.9115
0 | 0 | 77.8 | 91.51 | 226240.3 | 73.4 | 73.296
0 | 0 | 76.6 | 92.09 | 227685 | 73.87 | 73.7605
0 | 0 | 75.4 | 92.77 | 229367.3 | 74.26 | 74.3052
0 | 0 | 74.2 | 93.43 | 230999.9 | 74.9 | 74.8338
0 | 0 | 73 | 93.95 | 232286.9 | 75.31 | 75.2503
0 | 0 | 71.8 | 94.7 | 234129.8 | 75.87 | 75.851
0 | 0 | 70.5 | 95.29 | 235581.7 | 76.24 | 76.3236
0 | 0 | 69.4 | 95.84 | 236952.7 | 76.72 | 76.7641
0 | 0 | 68.1 | 96.42 | 238380.6 | 77.32 | 77.2287
0 | 0 | 67 | 97.15 | 240185.8 | 77.79 | 77.8134
0 | 0 | 65.7 | 97.88 | 241988.5 | 78.37 | 78.3981
0 | 0 | 64.5 | 98.69 | 243992.2 | 78.96 | 79.0469
0 | 0 | 63.3 | 99.19 | 245237.7 | 79.44 | 79.4473
0 | 0 | 62.1 | 99.7 | 246487.9 | 79.96 | 79.8558
0 | 0 | 60.9 | 100.36 | 248121.5 | 80.32 | 80.3845
0 | 0 | 59.7 | 100.9 | 249472.3 | 80.85 | 80.817
0 | 0 | 58.5 | 101.7 | 251427.4 | 81.42 | 81.4577
0 | 0 | 57.3 | 102.26 | 252815.6 | 81.98 | 81.9063
0 | 0 | 56.1 | 102.96 | 254551.1 | 82.52 | 82.467
0 | 0 | 54.9 | 103.67 | 256296.3 | 83.07 | 83.0356
0 | 0 | 53.7 | 104.25 | 257744.5 | 83.49 | 83.5002
0 | 0 | 52.5 | 104.94 | 259459 | 84.05 | 84.0529
0 | 0 | 51.3 | 105.58 | 261030.3 | 84.55 | 84.5655
0 | 0 | 50 | 106.18 | 262510.5 | 85.03 | 85.0461
0 | 0 | 48.9 | 106.77 | 263980.9 | 85.61 | 85.5186
0 | 0 | 47.6 | 107.42 | 265579.2 | 86.1 | 86.0392
0 | 0 | 46.5 | 108.29 | 267720.3 | 86.69 | 86.7361
0 | 0 | 45.2 | 108.78 | 268937.4 | 87.14 | 87.1286
0 | 0 | 44 | 109.32 | 270289 | 87.64 | 87.5611
0 | 0 | 42.8 | 109.97 | 271875.4 | 88.08 | 88.0817
0 | 0 | 41.6 | 110.89 | 274168.8 | 88.76 | 88.8186
0 | 0 | 40.4 | 111.55 | 275789.5 | 89.25 | 89.3472
0 | 0 | 39.2 | 112.09 | 277115.7 | 89.79 | 89.7797
0 | 0 | 38 | 112.77 | 278814.9 | 90.42 | 90.3244
0 | 0 | 36.8 | 113.49 | 280577.2 | 90.92 | 90.9011
0 | 0 | 35.6 | 114.13 | 282162.1 | 91.42 | 91.4137
0 | 0 | 34.4 | 114.75 | 283709.5 | 91.97 | 91.9103
0 | 0 | 33.2 | 115.43 | 285375.6 | 92.46 | 92.4549
0 | 0 | 32 | 115.93 | 286611.7 | 92.9 | 92.8554
0 | 0 | 30.7 | 116.77 | 288704.7 | 93.5 | 93.5282
0 | 0 | 29.5 | 117.52 | 290542.9 | 94.11 | 94.129
0 | 0 | 28.3 | 118.19 | 292205.5 | 94.67 | 94.6656
0 | 0 | 27.1 | 118.78 | 293676.8 | 95.04 | 95.1382
0 | 0 | 25.9 | 119.43 | 295263.2 | 95.68 | 95.6588
0 | 0 | 24.7 | 119.99 | 296655.2 | 96.18 | 96.1073
0 | 0 | 23.5 | 120.44 | 297774.7 | 96.67 | 96.4678
0 | 0 | 22.3 | 121.39 | 300124.3 | 97.3 | 97.2287
0 | 0 | 21.1 | 121.98 | 301570.6 | 97.71 | 97.7012
0 | 0 | 19.9 | 122.82 | 303652.3 | 98.29 | 98.374
0 | 0 | 18.7 | 123.63 | 305666.2 | 98.85 | 99.0228
0 | 0 | 17.5 | 123.91 | 306337.3 | 99.21 | 99.2471
0 | 0 | 16.3 | 124.24 | 307157.8 | 99.56 | 99.5114
0 | 0 | 15 | 124.85 | 308684.9 | 99.89 | 100
0 | 0 | 13.8 | 124.8 | 308556.9 | 100 | 99.96
0 | 0 | 12.6 | 124.43 | 307635.8 | 99.85 | 99.6636
0 | 0 | 11.4 | 123.79 | 306059.1 | 99.19 | 99.151
0 | 0 | 10.2 | 122.97 | 304032.5 | 98.45 | 98.4942
0 | 0 | 9 | 121.05 | 299270.5 | 97.16 | 96.9563
0 | 0 | 7.8 | 118.54 | 293072.1 | 94.98 | 94.9459
0 | 0 | 6.6 | 114.7 | 283586.8 | 91.93 | 91.8702
0 | 0 | 5.4 | 108.63 | 268565.6 | 87.02 | 87.0084
0 | 0 | 4.2 | 100.12 | 247542.8 | 80.23 | 80.1922
0 | 0 | 3 | 87.27 | 215770.4 | 69.88 | 69.8999
0 | 0 | 1.8 | 75.13 | 185742.8 | 60.31 | 60.1762
0 | 0 | 0.6 | 68.88 | 170296.6 | 55.2 | 55.1702
0 | 0 | -0.5 | 64.7 | 159954 | 51.88 | 51.8222
Measurement time: | 12/08/2022 18:51:58
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 20.62 | 50985.5 | 20.43 | 20.5563
0 | 0 | 309.1 | 20.48 | 50623 | 20.37 | 20.4167
0 | 0 | 307.6 | 20.65 | 51042.1 | 20.54 | 20.5862
0 | 0 | 306.6 | 20.81 | 51445.5 | 20.66 | 20.7457
0 | 0 | 305.3 | 20.93 | 51753.1 | 20.82 | 20.8653
0 | 0 | 304.1 | 21.06 | 52070.9 | 20.93 | 20.9949
0 | 0 | 302.9 | 21.1 | 52163 | 20.96 | 21.0348
0 | 0 | 301.7 | 21.3 | 52659.5 | 21.18 | 21.2342
0 | 0 | 300.5 | 21.44 | 53019 | 21.33 | 21.3737
0 | 0 | 299.3 | 21.58 | 53350.2 | 21.45 | 21.5133
0 | 0 | 298.1 | 21.75 | 53779.2 | 21.62 | 21.6828
0 | 0 | 296.9 | 21.87 | 54060.2 | 21.75 | 21.8024
0 | 0 | 295.7 | 22.05 | 54514.3 | 21.92 | 21.9819
0 | 0 | 294.5 | 22.28 | 55077.1 | 22.17 | 22.2111
0 | 0 | 293.3 | 22.37 | 55299.1 | 22.22 | 22.3009
0 | 0 | 292.1 | 22.41 | 55400.5 | 22.32 | 22.3407
0 | 0 | 290.9 | 22.65 | 55993.4 | 22.5 | 22.58
0 | 0 | 289.7 | 22.78 | 56312.6 | 22.68 | 22.7096
0 | 0 | 288.5 | 22.98 | 56807.9 | 22.84 | 22.909
0 | 0 | 287.3 | 23.08 | 57072.6 | 22.95 | 23.0087
0 | 0 | 286.1 | 23.2 | 57352 | 23.1 | 23.1283
0 | 0 | 284.9 | 23.47 | 58024.3 | 23.34 | 23.3975
0 | 0 | 283.7 | 23.56 | 58247.1 | 23.45 | 23.4872
0 | 0 | 282.5 | 23.72 | 58656 | 23.63 | 23.6467
0 | 0 | 281.3 | 23.91 | 59113.3 | 23.79 | 23.8361
0 | 0 | 280 | 24.12 | 59636.3 | 23.97 | 24.0455
0 | 0 | 278.8 | 24.25 | 59949.3 | 24.1 | 24.1751
0 | 0 | 277.7 | 24.42 | 60376 | 24.28 | 24.3445
0 | 0 | 276.4 | 24.67 | 60986 | 24.55 | 24.5938
0 | 0 | 275.2 | 24.76 | 61221.9 | 24.63 | 24.6835
0 | 0 | 274.1 | 24.92 | 61615.9 | 24.81 | 24.843
0 | 0 | 272.8 | 25.08 | 61996.5 | 24.97 | 25.0025
0 | 0 | 271.6 | 25.26 | 62445.9 | 25.13 | 25.1819
0 | 0 | 270.4 | 25.43 | 62859.5 | 25.27 | 25.3514
0 | 0 | 269.2 | 25.6 | 63282.3 | 25.51 | 25.5209
0 | 0 | 268 | 25.75 | 63655.9 | 25.68 | 25.6704
0 | 0 | 266.8 | 25.97 | 64199.3 | 25.83 | 25.8897
0 | 0 | 265.6 | 26.09 | 64498.3 | 25.95 | 26.0094
0 | 0 | 264.4 | 26.28 | 64978.1 | 26.13 | 26.1988
0 | 0 | 263.2 | 26.42 | 65313.9 | 26.29 | 26.3384
0 | 0 | 262 | 26.71 | 66043.8 | 26.54 | 26.6275
0 | 0 | 260.8 | 26.79 | 66231.1 | 26.69 | 26.7072
0 | 0 | 259.6 | 26.99 | 66732.3 | 26.88 | 26.9066
0 | 0 | 258.4 | 27.18 | 67189.8 | 27.02 | 27.096
0 | 0 | 257.2 | 27.37 | 67671.8 | 27.26 | 27.2854
0 | 0 | 256 | 27.59 | 68224.1 | 27.47 | 27.5047
0 | 0 | 254.8 | 27.68 | 68429.5 | 27.57 | 27.5945
0 | 0 | 253.6 | 27.95 | 69091.6 | 27.85 | 27.8636
0 | 0 | 252.4 | 28.13 | 69543 | 28.01 | 28.0431
0 | 0 | 251.2 | 28.31 | 69983.9 | 28.15 | 28.2225
0 | 0 | 250 | 28.39 | 70193.5 | 28.27 | 28.3023
0 | 0 | 248.8 | 28.68 | 70906.1 | 28.55 | 28.5914
0 | 0 | 247.6 | 28.92 | 71494.6 | 28.82 | 28.8306
0 | 0 | 246.3 | 29.1 | 71945.2 | 28.95 | 29.0101
0 | 0 | 245.1 | 29.21 | 72222.7 | 29.11 | 29.1197
0 | 0 | 243.9 | 29.42 | 72737.5 | 29.31 | 29.3291
0 | 0 | 242.7 | 29.67 | 73346.5 | 29.56 | 29.5783
0 | 0 | 241.5 | 29.9 | 73920 | 29.79 | 29.8076
0 | 0 | 240.3 | 30.11 | 74438.1 | 29.96 | 30.0169
0 | 0 | 239.1 | 30.28 | 74850.9 | 30.11 | 30.1864
0 | 0 | 237.9 | 30.49 | 75372 | 30.33 | 30.3958
0 | 0 | 236.7 | 30.76 | 76055.8 | 30.63 | 30.6649
0 | 0 | 235.5 | 30.93 | 76473.8 | 30.8 | 30.8344
0 | 0 | 234.3 | 31.16 | 77026.5 | 31.04 | 31.0637
0 | 0 | 233.1 | 31.35 | 77505.4 | 31.23 | 31.2531
0 | 0 | 231.9 | 31.59 | 78112.6 | 31.48 | 31.4924
0 | 0 | 230.7 | 31.81 | 78638.5 | 31.69 | 31.7117
0 | 0 | 229.5 | 32.02 | 79172.4 | 31.89 | 31.921
0 | 0 | 228.3 | 32.2 | 79618.2 | 32.04 | 32.1005
0 | 0 | 227.1 | 32.5 | 80344.9 | 32.37 | 32.3996
0 | 0 | 225.9 | 32.68 | 80798.7 | 32.58 | 32.579
0 | 0 | 224.7 | 32.88 | 81296 | 32.77 | 32.7784
0 | 0 | 223.5 | 33.19 | 82052.6 | 33.05 | 33.0874
0 | 0 | 222.3 | 33.33 | 82396.1 | 33.14 | 33.227
0 | 0 | 221.1 | 33.6 | 83075.3 | 33.46 | 33.4962
0 | 0 | 219.9 | 33.81 | 83580.8 | 33.72 | 33.7055
0 | 0 | 218.7 | 34 | 84055.1 | 33.83 | 33.8949
0 | 0 | 217.5 | 34.29 | 84771.1 | 34.12 | 34.184
0 | 0 | 216.3 | 34.52 | 85351.2 | 34.42 | 34.4133
0 | 0 | 215 | 34.77 | 85967.4 | 34.69 | 34.6625
0 | 0 | 213.9 | 35.03 | 86597.9 | 34.84 | 34.9217
0 | 0 | 212.7 | 35.29 | 87251.1 | 35.14 | 35.1809
0 | 0 | 211.5 | 35.46 | 87661.4 | 35.31 | 35.3504
0 | 0 | 210.2 | 35.76 | 88414.8 | 35.59 | 35.6495
0 | 0 | 209 | 36.06 | 89161.7 | 35.88 | 35.9486
0 | 0 | 207.8 | 36.27 | 89675.2 | 36.1 | 36.1579
0 | 0 | 206.6 | 36.34 | 89838 | 36.31 | 36.2277
0 | 0 | 205.4 | 36.74 | 90842.4 | 36.6 | 36.6265
0 | 0 | 204.2 | 36.94 | 91322.4 | 36.8 | 36.8258
0 | 0 | 203 | 37.2 | 91969.9 | 37.07 | 37.085
0 | 0 | 201.8 | 37.42 | 92514.5 | 37.31 | 37.3044
0 | 0 | 200.6 | 37.61 | 92990.5 | 37.44 | 37.4938
0 | 0 | 199.4 | 37.83 | 93533.1 | 37.73 | 37.7131
0 | 0 | 198.2 | 38.18 | 94394.5 | 38.05 | 38.062
0 | 0 | 197 | 38.51 | 95207.4 | 38.32 | 38.391
0 | 0 | 195.8 | 38.72 | 95723.6 | 38.55 | 38.6003
0 | 0 | 194.6 | 39.04 | 96520.3 | 38.9 | 38.9194
0 | 0 | 193.4 | 39.28 | 97116.6 | 39.15 | 39.1586
0 | 0 | 192.2 | 39.5 | 97654.1 | 39.3 | 39.3779
0 | 0 | 191 | 39.83 | 98484.4 | 39.68 | 39.7069
0 | 0 | 189.8 | 40.17 | 99307.9 | 39.95 | 40.0459
0 | 0 | 188.5 | 40.47 | 100043.7 | 40.29 | 40.3449
0 | 0 | 187.3 | 40.62 | 100418.2 | 40.5 | 40.4945
0 | 0 | 186.1 | 40.98 | 101305 | 40.74 | 40.8534
0 | 0 | 184.9 | 41.28 | 102063.5 | 41.12 | 41.1524
0 | 0 | 183.8 | 41.68 | 103048.5 | 41.49 | 41.5512
0 | 0 | 182.6 | 41.84 | 103438.8 | 41.67 | 41.7107
0 | 0 | 181.3 | 42.07 | 104011.3 | 41.93 | 41.94
0 | 0 | 180.1 | 42.31 | 104611.2 | 42.17 | 42.1792
0 | 0 | 178.9 | 42.69 | 105556.9 | 42.55 | 42.5581
0 | 0 | 177.7 | 42.99 | 106274.2 | 42.83 | 42.8571
0 | 0 | 176.5 | 43.26 | 106944.6 | 43.11 | 43.1263
0 | 0 | 175.3 | 43.51 | 107565 | 43.35 | 43.3755
0 | 0 | 174.1 | 43.78 | 108230.2 | 43.59 | 43.6447
0 | 0 | 172.9 | 44.19 | 109247.9 | 43.96 | 44.0534
0 | 0 | 171.7 | 44.48 | 109959.8 | 44.33 | 44.3425
0 | 0 | 170.5 | 44.86 | 110910 | 44.65 | 44.7214
0 | 0 | 169.3 | 45.07 | 111421.9 | 44.83 | 44.9307
0 | 0 | 168.1 | 45.34 | 112092 | 45.2 | 45.1999
0 | 0 | 166.9 | 45.72 | 113027.8 | 45.52 | 45.5787
0 | 0 | 165.7 | 45.93 | 113556.4 | 45.71 | 45.7881
0 | 0 | 164.5 | 46.25 | 114345.2 | 46.08 | 46.1071
0 | 0 | 163.3 | 46.52 | 115004.8 | 46.38 | 46.3762
0 | 0 | 162.1 | 46.8 | 115705.2 | 46.61 | 46.6554
0 | 0 | 160.9 | 47.25 | 116827.8 | 47.11 | 47.104
0 | 0 | 159.7 | 47.66 | 117836 | 47.48 | 47.5127
0 | 0 | 158.4 | 47.98 | 118612.3 | 47.76 | 47.8317
0 | 0 | 157.3 | 48.34 | 119517 | 48.13 | 48.1906
0 | 0 | 156.1 | 48.68 | 120354.3 | 48.49 | 48.5296
0 | 0 | 154.8 | 48.98 | 121093 | 48.82 | 48.8286
0 | 0 | 153.7 | 49.29 | 121856.8 | 49.07 | 49.1377
0 | 0 | 152.4 | 49.59 | 122615.4 | 49.47 | 49.4367
0 | 0 | 151.3 | 49.9 | 123363.6 | 49.67 | 49.7458
0 | 0 | 150 | 50.22 | 124158 | 50.04 | 50.0648
0 | 0 | 148.8 | 50.51 | 124879.1 | 50.33 | 50.3539
0 | 0 | 147.6 | 50.9 | 125839 | 50.78 | 50.7427
0 | 0 | 146.4 | 51.35 | 126963.9 | 51.16 | 51.1913
0 | 0 | 145.2 | 51.54 | 127434.7 | 51.4 | 51.3807
0 | 0 | 144 | 52 | 128558.9 | 51.78 | 51.8393
0 | 0 | 142.8 | 52.32 | 129356.2 | 52.15 | 52.1583
0 | 0 | 141.6 | 52.76 | 130436.2 | 52.56 | 52.5969
0 | 0 | 140.4 | 53.06 | 131177.6 | 52.8 | 52.896
0 | 0 | 139.2 | 53.43 | 132091.3 | 53.23 | 53.2649
0 | 0 | 138 | 53.79 | 132981.6 | 53.64 | 53.6238
0 | 0 | 136.8 | 54.07 | 133684.8 | 53.97 | 53.9029
0 | 0 | 135.6 | 54.45 | 134625.9 | 54.31 | 54.2817
0 | 0 | 134.4 | 54.87 | 135656.5 | 54.5 | 54.7004
0 | 0 | 133.2 | 55.18 | 136424.4 | 54.95 | 55.0095
0 | 0 | 132 | 55.74 | 137809.9 | 55.43 | 55.5677
0 | 0 | 130.8 | 56.1 | 138705.1 | 55.84 | 55.9266
0 | 0 | 129.6 | 56.4 | 139450.8 | 56.19 | 56.2257
0 | 0 | 128.4 | 56.68 | 140124.9 | 56.5 | 56.5048
0 | 0 | 127.2 | 57.13 | 141250.6 | 56.97 | 56.9534
0 | 0 | 126 | 57.46 | 142067.4 | 57.3 | 57.2824
0 | 0 | 124.8 | 57.91 | 143163.6 | 57.73 | 57.731
0 | 0 | 123.6 | 58.38 | 144327.6 | 58.22 | 58.1996
0 | 0 | 122.4 | 58.68 | 145078.3 | 58.52 | 58.4987
0 | 0 | 121.2 | 59.16 | 146271.3 | 58.87 | 58.9772
0 | 0 | 120 | 59.4 | 146867.6 | 59.29 | 59.2164
0 | 0 | 118.7 | 59.92 | 148151.2 | 59.71 | 59.7348
0 | 0 | 117.5 | 60.37 | 149260.2 | 60.14 | 60.1834
0 | 0 | 116.3 | 60.68 | 150031.6 | 60.66 | 60.4925
0 | 0 | 115.1 | 61.03 | 150896.9 | 60.83 | 60.8414
0 | 0 | 113.9 | 61.55 | 152176.7 | 61.33 | 61.3598
0 | 0 | 112.7 | 62.14 | 153637.9 | 61.84 | 61.948
0 | 0 | 111.5 | 62.48 | 154471.5 | 62.09 | 62.2869
0 | 0 | 110.3 | 62.87 | 155444.5 | 62.51 | 62.6757
0 | 0 | 109.1 | 63.23 | 156317 | 62.93 | 63.0346
0 | 0 | 107.9 | 63.57 | 157160.8 | 63.39 | 63.3735
0 | 0 | 106.7 | 64.01 | 158263 | 63.82 | 63.8122
0 | 0 | 105.5 | 64.56 | 159604.6 | 64.18 | 64.3605
0 | 0 | 104.3 | 64.88 | 160406.1 | 64.54 | 64.6795
0 | 0 | 103.1 | 65.21 | 161210.3 | 65.01 | 65.0085
0 | 0 | 101.9 | 65.53 | 162001.2 | 65.31 | 65.3275
0 | 0 | 100.7 | 66.16 | 163563.9 | 65.98 | 65.9555
0 | 0 | 99.5 | 66.56 | 164558.7 | 66.38 | 66.3543
0 | 0 | 98.3 | 67.02 | 165703.1 | 66.89 | 66.8129
0 | 0 | 97.1 | 67.59 | 167117 | 67.42 | 67.3811
0 | 0 | 95.9 | 67.95 | 167994.1 | 67.68 | 67.74
0 | 0 | 94.7 | 68.47 | 169272.2 | 68.11 | 68.2584
0 | 0 | 93.5 | 68.88 | 170307.1 | 68.68 | 68.6671
0 | 0 | 92.3 | 69.27 | 171266.3 | 69 | 69.0559
0 | 0 | 91.1 | 69.67 | 172252.7 | 69.41 | 69.4547
0 | 0 | 89.9 | 70.16 | 173467.2 | 70.06 | 69.9432
0 | 0 | 88.7 | 70.62 | 174608.9 | 70.41 | 70.4018
0 | 0 | 87.5 | 71.16 | 175932.1 | 70.82 | 70.9401
0 | 0 | 86.3 | 71.58 | 176974.7 | 71.37 | 71.3588
0 | 0 | 85.1 | 72.11 | 178270.4 | 71.89 | 71.8871
0 | 0 | 83.8 | 72.48 | 179189.5 | 72.44 | 72.256
0 | 0 | 82.7 | 72.93 | 180304.5 | 72.79 | 72.7046
0 | 0 | 81.5 | 73.54 | 181815.1 | 73.21 | 73.3127
0 | 0 | 80.2 | 74.06 | 183103.7 | 73.81 | 73.8311
0 | 0 | 79 | 74.39 | 183907.3 | 74.15 | 74.1601
0 | 0 | 77.8 | 74.79 | 184898.2 | 74.63 | 74.5589
0 | 0 | 76.6 | 75.35 | 186298.8 | 75.18 | 75.1171
0 | 0 | 75.4 | 75.77 | 187324.2 | 75.45 | 75.5358
0 | 0 | 74.2 | 76.27 | 188573.3 | 76.01 | 76.0343
0 | 0 | 73 | 76.67 | 189562.3 | 76.47 | 76.4331
0 | 0 | 71.8 | 77.23 | 190927.7 | 77.01 | 76.9913
0 | 0 | 70.6 | 77.78 | 192310.3 | 77.58 | 77.5396
0 | 0 | 69.4 | 78.26 | 193494 | 78.02 | 78.0181
0 | 0 | 68.2 | 78.76 | 194711 | 78.44 | 78.5166
0 | 0 | 67 | 79.15 | 195677.2 | 78.83 | 78.9054
0 | 0 | 65.7 | 79.68 | 197005.4 | 79.42 | 79.4338
0 | 0 | 64.5 | 80.23 | 198366.7 | 79.94 | 79.9821
0 | 0 | 63.3 | 80.83 | 199833.7 | 80.5 | 80.5802
0 | 0 | 62.1 | 81.11 | 200531.3 | 80.84 | 80.8593
0 | 0 | 60.9 | 81.65 | 201873.7 | 81.4 | 81.3977
0 | 0 | 59.7 | 82.18 | 203177.8 | 82.03 | 81.926
0 | 0 | 58.5 | 82.67 | 204384.1 | 82.47 | 82.4145
0 | 0 | 57.3 | 83.2 | 205704.3 | 82.91 | 82.9429
0 | 0 | 56.1 | 83.68 | 206888.2 | 83.34 | 83.4214
0 | 0 | 54.9 | 84.06 | 207823.1 | 83.89 | 83.8002
0 | 0 | 53.7 | 84.56 | 209060.1 | 84.37 | 84.2987
0 | 0 | 52.5 | 85.33 | 210956.4 | 85.03 | 85.0663
0 | 0 | 51.3 | 85.74 | 211979.1 | 85.3 | 85.475
0 | 0 | 50 | 86.41 | 213634.3 | 86.09 | 86.143
0 | 0 | 48.8 | 86.83 | 214675.5 | 86.44 | 86.5617
0 | 0 | 47.6 | 87.08 | 215297.2 | 86.93 | 86.8109
0 | 0 | 46.4 | 87.56 | 216476.9 | 87.27 | 87.2894
0 | 0 | 45.2 | 88.15 | 217948.6 | 87.95 | 87.8776
0 | 0 | 44 | 88.63 | 219135.3 | 88.35 | 88.3561
0 | 0 | 42.8 | 89.4 | 221038.4 | 89.13 | 89.1237
0 | 0 | 41.6 | 89.76 | 221910.5 | 89.7 | 89.4826
0 | 0 | 40.4 | 90.36 | 223399.4 | 90.29 | 90.0807
0 | 0 | 39.2 | 91 | 224989.8 | 90.75 | 90.7188
0 | 0 | 38 | 91.57 | 226388.7 | 91.23 | 91.287
0 | 0 | 36.8 | 92.01 | 227488.6 | 91.62 | 91.7257
0 | 0 | 35.5 | 92.48 | 228639.2 | 92.25 | 92.1942
0 | 0 | 34.3 | 93.16 | 230328.8 | 92.84 | 92.8721
0 | 0 | 33.1 | 93.57 | 231346.5 | 93.3 | 93.2808
0 | 0 | 31.9 | 94.17 | 232821.8 | 93.79 | 93.879
0 | 0 | 30.7 | 94.49 | 233618.4 | 94.18 | 94.198
0 | 0 | 29.5 | 94.99 | 234843.9 | 94.72 | 94.6964
0 | 0 | 28.3 | 95.61 | 236391.8 | 95.39 | 95.3145
0 | 0 | 27.1 | 96.04 | 237437.3 | 95.74 | 95.7432
0 | 0 | 25.8 | 96.61 | 238864.8 | 96.2 | 96.3114
0 | 0 | 24.7 | 97.06 | 239974.3 | 96.73 | 96.76
0 | 0 | 23.4 | 97.54 | 241147 | 97.28 | 97.2386
0 | 0 | 22.2 | 98.2 | 242795.3 | 97.62 | 97.8965
0 | 0 | 21.1 | 98.61 | 243808.9 | 98.3 | 98.3053
0 | 0 | 19.8 | 99.21 | 245274.2 | 98.68 | 98.9034
0 | 0 | 18.6 | 99.59 | 246211.4 | 99.28 | 99.2822
0 | 0 | 17.4 | 99.63 | 246325 | 99.49 | 99.3221
0 | 0 | 16.2 | 99.94 | 247084.9 | 99.69 | 99.6311
0 | 0 | 15 | 100.28 | 247931.2 | 99.99 | 99.9701
0 | 0 | 13.8 | 100.31 | 247992.5 | 100 | 100
0 | 0 | 12.6 | 100.02 | 247282.4 | 99.86 | 99.7109
0 | 0 | 11.4 | 99.63 | 246311.4 | 99.28 | 99.3221
0 | 0 | 10.2 | 98.67 | 243937.9 | 98.3 | 98.3651
0 | 0 | 9 | 97.17 | 240231.4 | 96.82 | 96.8697
0 | 0 | 7.8 | 94.81 | 234398.9 | 94.52 | 94.517
0 | 0 | 6.6 | 91.55 | 226344.7 | 91.32 | 91.2671
0 | 0 | 5.4 | 86.21 | 213132.1 | 85.8 | 85.9436
0 | 0 | 4.2 | 78.68 | 194527.2 | 78.53 | 78.4368
0 | 0 | 2.9 | 68.76 | 170008.7 | 68.5 | 68.5475
0 | 0 | 1.7 | 58.88 | 145570.7 | 58.67 | 58.698
0 | 0 | 0.6 | 53.86 | 133167.5 | 53.65 | 53.6935
0 | 0 | -0.5 | 50.66 | 125253.9 | 50.52 | 50.5034