## 8_Rendements taille de champ
Measurement time: | 12/08/2022 18:54:12
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 310 | 20.25 | 50053.1 | 20.2 | 20.2035
0 | 0 | 309.1 | 20.35 | 50304.8 | 20.29 | 20.3033
0 | 0 | 307.6 | 20.68 | 51118.5 | 20.61 | 20.6325
0 | 0 | 306.6 | 20.74 | 51270.8 | 20.69 | 20.6924
0 | 0 | 305.3 | 20.87 | 51599.4 | 20.8 | 20.8221
0 | 0 | 304.1 | 20.95 | 51791.2 | 20.9 | 20.9019
0 | 0 | 302.9 | 21.03 | 51986.8 | 20.94 | 20.9817
0 | 0 | 301.7 | 21.24 | 52513.5 | 21.19 | 21.1913
0 | 0 | 300.5 | 21.44 | 53017.3 | 21.4 | 21.3908
0 | 0 | 299.3 | 21.58 | 53353.7 | 21.58 | 21.5305
0 | 0 | 298.1 | 21.73 | 53713.2 | 21.68 | 21.6801
0 | 0 | 296.9 | 21.8 | 53885.9 | 21.77 | 21.75
0 | 0 | 295.7 | 22.01 | 54426.3 | 21.98 | 21.9595
0 | 0 | 294.5 | 22.16 | 54780.9 | 22.14 | 22.1091
0 | 0 | 293.3 | 22.3 | 55137.8 | 22.27 | 22.2488
0 | 0 | 292.1 | 22.49 | 55597.2 | 22.46 | 22.4384
0 | 0 | 290.9 | 22.61 | 55905 | 22.58 | 22.5581
0 | 0 | 289.7 | 22.71 | 56146.2 | 22.72 | 22.6579
0 | 0 | 288.5 | 22.87 | 56546.2 | 22.84 | 22.8175
0 | 0 | 287.3 | 23.04 | 56963.4 | 22.96 | 22.9871
0 | 0 | 286.1 | 23.22 | 57407 | 23.2 | 23.1667
0 | 0 | 284.8 | 23.37 | 57783.1 | 23.35 | 23.3164
0 | 0 | 283.7 | 23.54 | 58196.3 | 23.49 | 23.486
0 | 0 | 282.5 | 23.68 | 58536.8 | 23.63 | 23.6257
0 | 0 | 281.2 | 23.92 | 59134.2 | 23.9 | 23.8651
0 | 0 | 280.1 | 24.05 | 59470.6 | 24.03 | 23.9948
0 | 0 | 278.8 | 24.12 | 59635.2 | 24.12 | 24.0647
0 | 0 | 277.6 | 24.29 | 60047.2 | 24.25 | 24.2343
0 | 0 | 276.4 | 24.51 | 60594.3 | 24.49 | 24.4538
0 | 0 | 275.2 | 24.69 | 61040.6 | 24.66 | 24.6333
0 | 0 | 274 | 24.91 | 61579.7 | 24.86 | 24.8528
0 | 0 | 272.8 | 24.94 | 61672.1 | 24.92 | 24.8828
0 | 0 | 271.6 | 25.15 | 62189.2 | 25.11 | 25.0923
0 | 0 | 270.4 | 25.34 | 62640 | 25.34 | 25.2819
0 | 0 | 269.2 | 25.5 | 63048 | 25.49 | 25.4415
0 | 0 | 268 | 25.67 | 63477.5 | 25.61 | 25.6111
0 | 0 | 266.8 | 25.8 | 63781.9 | 25.76 | 25.7408
0 | 0 | 265.6 | 26.05 | 64405.5 | 26.01 | 25.9902
0 | 0 | 264.4 | 26.25 | 64888.7 | 26.27 | 26.1898
0 | 0 | 263.2 | 26.41 | 65291.7 | 26.41 | 26.3494
0 | 0 | 262 | 26.54 | 65620.8 | 26.51 | 26.4791
0 | 0 | 260.8 | 26.74 | 66116.2 | 26.73 | 26.6786
0 | 0 | 259.6 | 27.02 | 66809.6 | 26.95 | 26.958
0 | 0 | 258.4 | 27.18 | 67191 | 27.17 | 27.1176
0 | 0 | 257.2 | 27.37 | 67672.3 | 27.34 | 27.3072
0 | 0 | 256 | 27.56 | 68148.9 | 27.55 | 27.4968
0 | 0 | 254.8 | 27.76 | 68631.2 | 27.71 | 27.6963
0 | 0 | 253.6 | 27.86 | 68872.9 | 27.81 | 27.7961
0 | 0 | 252.4 | 28.12 | 69534.2 | 28.11 | 28.0555
0 | 0 | 251.2 | 28.21 | 69740.5 | 28.19 | 28.1453
0 | 0 | 250 | 28.5 | 70452.4 | 28.49 | 28.4346
0 | 0 | 248.8 | 28.64 | 70803.9 | 28.63 | 28.5743
0 | 0 | 247.6 | 28.86 | 71345.3 | 28.82 | 28.7938
0 | 0 | 246.4 | 29.03 | 71784.4 | 29.03 | 28.9634
0 | 0 | 245.1 | 29.24 | 72295.9 | 29.2 | 29.1729
0 | 0 | 243.9 | 29.45 | 72801.8 | 29.39 | 29.3824
0 | 0 | 242.7 | 29.61 | 73211.8 | 29.59 | 29.5421
0 | 0 | 241.5 | 29.82 | 73719.6 | 29.81 | 29.7516
0 | 0 | 240.3 | 30.02 | 74228.3 | 29.95 | 29.9511
0 | 0 | 239.1 | 30.34 | 75020.5 | 30.26 | 30.2704
0 | 0 | 237.9 | 30.54 | 75516.3 | 30.55 | 30.4699
0 | 0 | 236.7 | 30.64 | 75754.1 | 30.7 | 30.5697
0 | 0 | 235.5 | 30.92 | 76443.9 | 30.88 | 30.849
0 | 0 | 234.3 | 31.16 | 77033.1 | 31.1 | 31.0885
0 | 0 | 233.1 | 31.31 | 77412.2 | 31.26 | 31.2382
0 | 0 | 231.9 | 31.51 | 77915.5 | 31.44 | 31.4377
0 | 0 | 230.7 | 31.84 | 78709.4 | 31.77 | 31.7669
0 | 0 | 229.5 | 31.89 | 78850.8 | 31.87 | 31.8168
0 | 0 | 228.3 | 32.21 | 79644.5 | 32.21 | 32.1361
0 | 0 | 227.1 | 32.46 | 80251.6 | 32.44 | 32.3855
0 | 0 | 225.9 | 32.68 | 80800 | 32.62 | 32.605
0 | 0 | 224.7 | 32.89 | 81313.3 | 32.84 | 32.8145
0 | 0 | 223.5 | 33.1 | 81845.3 | 33.08 | 33.024
0 | 0 | 222.2 | 33.41 | 82596 | 33.38 | 33.3333
0 | 0 | 221.1 | 33.63 | 83139.7 | 33.6 | 33.5528
0 | 0 | 219.9 | 33.75 | 83430.8 | 33.77 | 33.6726
0 | 0 | 218.6 | 34 | 84061.1 | 34.03 | 33.922
0 | 0 | 217.5 | 34.29 | 84770.7 | 34.24 | 34.2113
0 | 0 | 216.2 | 34.4 | 85052.4 | 34.38 | 34.3211
0 | 0 | 215 | 34.81 | 86067.1 | 34.75 | 34.7301
0 | 0 | 213.8 | 34.92 | 86331.2 | 34.89 | 34.8399
0 | 0 | 212.6 | 35.18 | 86989.1 | 35.17 | 35.0993
0 | 0 | 211.4 | 35.45 | 87644.1 | 35.42 | 35.3687
0 | 0 | 210.2 | 35.65 | 88137.4 | 35.71 | 35.5682
0 | 0 | 209 | 35.94 | 88851.9 | 35.87 | 35.8575
0 | 0 | 207.8 | 36.14 | 89357 | 36.08 | 36.0571
0 | 0 | 206.6 | 36.5 | 90236.1 | 36.43 | 36.4162
0 | 0 | 205.4 | 36.52 | 90282.9 | 36.5 | 36.4362
0 | 0 | 204.2 | 36.85 | 91117.6 | 36.85 | 36.7654
0 | 0 | 203 | 37.22 | 92015 | 37.17 | 37.1346
0 | 0 | 201.8 | 37.49 | 92681.5 | 37.41 | 37.404
0 | 0 | 200.6 | 37.6 | 92949.6 | 37.57 | 37.5137
0 | 0 | 199.4 | 37.9 | 93694.4 | 37.87 | 37.813
0 | 0 | 198.2 | 38.19 | 94409.6 | 38.12 | 38.1024
0 | 0 | 197 | 38.59 | 95419.5 | 38.56 | 38.5014
0 | 0 | 195.8 | 38.72 | 95724.9 | 38.65 | 38.6311
0 | 0 | 194.6 | 39.04 | 96520.9 | 38.94 | 38.9504
0 | 0 | 193.4 | 39.15 | 96800.1 | 39.13 | 39.0602
0 | 0 | 192.2 | 39.54 | 97754.6 | 39.52 | 39.4493
0 | 0 | 191 | 39.75 | 98263.7 | 39.67 | 39.6588
0 | 0 | 189.8 | 40.12 | 99189.1 | 40.09 | 40.0279
0 | 0 | 188.5 | 40.37 | 99801.8 | 40.34 | 40.2774
0 | 0 | 187.3 | 40.68 | 100572.4 | 40.63 | 40.5867
0 | 0 | 186.2 | 40.86 | 101012.7 | 40.86 | 40.7662
0 | 0 | 185 | 41.19 | 101848.2 | 41.17 | 41.0955
0 | 0 | 183.8 | 41.45 | 102486.4 | 41.44 | 41.3549
0 | 0 | 182.5 | 41.77 | 103281.2 | 41.74 | 41.6741
0 | 0 | 181.3 | 42.1 | 104074.1 | 41.98 | 42.0034
0 | 0 | 180.1 | 42.49 | 105048.9 | 42.37 | 42.3925
0 | 0 | 178.9 | 42.64 | 105416.5 | 42.61 | 42.5422
0 | 0 | 177.7 | 43.03 | 106378.1 | 42.97 | 42.9313
0 | 0 | 176.5 | 43.23 | 106883.2 | 43.22 | 43.1308
0 | 0 | 175.3 | 43.58 | 107754.4 | 43.63 | 43.48
0 | 0 | 174.1 | 43.94 | 108624.5 | 43.91 | 43.8392
0 | 0 | 172.9 | 44.25 | 109409 | 44.3 | 44.1485
0 | 0 | 171.7 | 44.48 | 109972.2 | 44.44 | 44.3779
0 | 0 | 170.5 | 44.74 | 110615 | 44.77 | 44.6373
0 | 0 | 169.3 | 45.02 | 111300.5 | 45.03 | 44.9167
0 | 0 | 168.1 | 45.32 | 112048.4 | 45.25 | 45.216
0 | 0 | 166.9 | 45.62 | 112799.9 | 45.61 | 45.5153
0 | 0 | 165.7 | 46.02 | 113775.5 | 45.95 | 45.9144
0 | 0 | 164.4 | 46.29 | 114432.8 | 46.25 | 46.1838
0 | 0 | 163.3 | 46.6 | 115217.1 | 46.59 | 46.4931
0 | 0 | 162 | 46.93 | 116017.9 | 46.86 | 46.8223
0 | 0 | 160.9 | 47.24 | 116799.6 | 47.15 | 47.1316
0 | 0 | 159.6 | 47.57 | 117610.8 | 47.55 | 47.4608
0 | 0 | 158.4 | 47.96 | 118579.4 | 47.89 | 47.8499
0 | 0 | 157.2 | 48.26 | 119304.3 | 48.18 | 48.1493
0 | 0 | 156 | 48.57 | 120089.1 | 48.51 | 48.4585
0 | 0 | 154.8 | 48.96 | 121053.6 | 48.92 | 48.8477
0 | 0 | 153.6 | 49.31 | 121902.6 | 49.25 | 49.1968
0 | 0 | 152.4 | 49.52 | 122437.6 | 49.57 | 49.4064
0 | 0 | 151.2 | 49.87 | 123289.8 | 49.8 | 49.7556
0 | 0 | 150 | 50.15 | 123998.1 | 50.09 | 50.0349
0 | 0 | 148.8 | 50.61 | 125129.7 | 50.54 | 50.4939
0 | 0 | 147.6 | 50.89 | 125806.2 | 50.84 | 50.7732
0 | 0 | 146.4 | 51.23 | 126658.6 | 51.21 | 51.1124
0 | 0 | 145.2 | 51.53 | 127404 | 51.43 | 51.4118
0 | 0 | 144 | 51.93 | 128392.8 | 51.85 | 51.8108
0 | 0 | 142.8 | 52.25 | 129185.3 | 52.2 | 52.1301
0 | 0 | 141.6 | 52.66 | 130204.7 | 52.56 | 52.5392
0 | 0 | 140.4 | 52.95 | 130915.7 | 52.85 | 52.8285
0 | 0 | 139.2 | 53.3 | 131771.7 | 53.24 | 53.1777
0 | 0 | 138 | 53.61 | 132550.8 | 53.61 | 53.487
0 | 0 | 136.8 | 54.07 | 133672.7 | 53.99 | 53.9459
0 | 0 | 135.6 | 54.62 | 135028.8 | 54.5 | 54.4947
0 | 0 | 134.4 | 54.92 | 135790.6 | 54.77 | 54.794
0 | 0 | 133.2 | 55.21 | 136504.2 | 55.13 | 55.0833
0 | 0 | 132 | 55.64 | 137554.6 | 55.55 | 55.5123
0 | 0 | 130.7 | 55.93 | 138289.3 | 55.87 | 55.8017
0 | 0 | 129.6 | 56.27 | 139109.8 | 56.25 | 56.1409
0 | 0 | 128.3 | 56.65 | 140062.1 | 56.64 | 56.52
0 | 0 | 127.1 | 56.89 | 140650.9 | 56.91 | 56.7595
0 | 0 | 125.9 | 57.31 | 141698.2 | 57.34 | 57.1785
0 | 0 | 124.7 | 57.72 | 142705.7 | 57.75 | 57.5875
0 | 0 | 123.5 | 58.15 | 143766.1 | 58.15 | 58.0166
0 | 0 | 122.3 | 58.53 | 144718.6 | 58.43 | 58.3957
0 | 0 | 121.1 | 58.99 | 145855.7 | 58.86 | 58.8546
0 | 0 | 119.9 | 59.48 | 147054 | 59.29 | 59.3435
0 | 0 | 118.7 | 59.88 | 148033.6 | 59.73 | 59.7426
0 | 0 | 117.5 | 60.2 | 148826.1 | 60.08 | 60.0619
0 | 0 | 116.3 | 60.67 | 150005 | 60.55 | 60.5308
0 | 0 | 115.1 | 60.96 | 150726.3 | 60.88 | 60.8201
0 | 0 | 113.9 | 61.53 | 152114.6 | 61.44 | 61.3888
0 | 0 | 112.7 | 61.98 | 153235.2 | 61.76 | 61.8378
0 | 0 | 111.5 | 62.25 | 153895.6 | 62.19 | 62.1072
0 | 0 | 110.3 | 62.68 | 154974.3 | 62.56 | 62.5362
0 | 0 | 109.1 | 62.96 | 155670.3 | 62.86 | 62.8155
0 | 0 | 107.9 | 63.47 | 156918.3 | 63.42 | 63.3244
0 | 0 | 106.7 | 64.01 | 158263.3 | 63.91 | 63.8631
0 | 0 | 105.5 | 64.45 | 159355.1 | 64.33 | 64.3021
0 | 0 | 104.2 | 64.89 | 160432.7 | 64.84 | 64.7411
0 | 0 | 103.1 | 65.28 | 161395.6 | 65.2 | 65.1302
0 | 0 | 101.8 | 65.63 | 162249.8 | 65.6 | 65.4794
0 | 0 | 100.6 | 66.2 | 163681.8 | 66.08 | 66.0481
0 | 0 | 99.5 | 66.52 | 164455.2 | 66.47 | 66.3674
0 | 0 | 98.3 | 67.02 | 165686.6 | 66.97 | 66.8662
0 | 0 | 97.1 | 67.49 | 166855.4 | 67.4 | 67.3351
0 | 0 | 95.9 | 67.83 | 167708.1 | 67.79 | 67.6743
0 | 0 | 94.6 | 68.31 | 168893.5 | 68.22 | 68.1532
0 | 0 | 93.5 | 68.85 | 170212.6 | 68.72 | 68.692
0 | 0 | 92.2 | 69.37 | 171504.3 | 69.18 | 69.2108
0 | 0 | 91.1 | 69.77 | 172487.8 | 69.6 | 69.6099
0 | 0 | 89.8 | 70.16 | 173467.3 | 70.15 | 69.999
0 | 0 | 88.6 | 70.49 | 174278.6 | 70.4 | 70.3282
0 | 0 | 87.4 | 71.11 | 175804.9 | 71 | 70.9468
0 | 0 | 86.2 | 71.62 | 177078.6 | 71.48 | 71.4557
0 | 0 | 85 | 71.89 | 177741.1 | 71.71 | 71.725
0 | 0 | 83.8 | 72.4 | 178986.2 | 72.42 | 72.2339
0 | 0 | 82.6 | 73.01 | 180507.3 | 73.04 | 72.8425
0 | 0 | 81.4 | 73.43 | 181549.3 | 73.35 | 73.2615
0 | 0 | 80.2 | 73.78 | 182410.2 | 73.81 | 73.6107
0 | 0 | 79 | 74.37 | 183866.7 | 74.22 | 74.1993
0 | 0 | 77.8 | 74.82 | 184989.5 | 74.7 | 74.6483
0 | 0 | 76.6 | 75.39 | 186402.5 | 75.26 | 75.217
0 | 0 | 75.4 | 75.72 | 187210.5 | 75.62 | 75.5462
0 | 0 | 74.1 | 76.35 | 188774.1 | 76.36 | 76.1748
0 | 0 | 73 | 76.67 | 189557.2 | 76.62 | 76.4941
0 | 0 | 71.7 | 77.22 | 190919.1 | 77.18 | 77.0428
0 | 0 | 70.5 | 77.64 | 191948.1 | 77.47 | 77.4618
0 | 0 | 69.3 | 78.36 | 193721.7 | 78.18 | 78.1802
0 | 0 | 68.1 | 78.76 | 194733 | 78.51 | 78.5793
0 | 0 | 66.9 | 79.21 | 195830.3 | 79.03 | 79.0282
0 | 0 | 65.7 | 79.55 | 196672.4 | 79.53 | 79.3675
0 | 0 | 64.5 | 79.98 | 197749.5 | 80.06 | 79.7965
0 | 0 | 63.3 | 80.66 | 199418 | 80.6 | 80.4749
0 | 0 | 62.1 | 81.34 | 201102.7 | 81.37 | 81.1533
0 | 0 | 60.9 | 81.6 | 201736.2 | 81.5 | 81.4128
0 | 0 | 59.7 | 82.29 | 203437.6 | 82.19 | 82.1012
0 | 0 | 58.5 | 82.71 | 204477.6 | 82.47 | 82.5202
0 | 0 | 57.3 | 83.37 | 206114.6 | 83.15 | 83.1787
0 | 0 | 56 | 83.81 | 207212.5 | 83.78 | 83.6177
0 | 0 | 54.8 | 84.26 | 208309.6 | 84.17 | 84.0666
0 | 0 | 53.6 | 84.78 | 209605 | 84.68 | 84.5855
0 | 0 | 52.4 | 85.32 | 210936.9 | 85.15 | 85.1242
0 | 0 | 51.2 | 85.77 | 212064.9 | 85.74 | 85.5732
0 | 0 | 50 | 86.37 | 213532.2 | 86.25 | 86.1718
0 | 0 | 48.8 | 86.78 | 214542.2 | 86.68 | 86.5809
0 | 0 | 47.6 | 87.38 | 216038.6 | 87.24 | 87.1795
0 | 0 | 46.4 | 87.7 | 216829.8 | 87.66 | 87.4988
0 | 0 | 45.2 | 88.28 | 218248.8 | 88.3 | 88.0774
0 | 0 | 44 | 88.7 | 219295.7 | 88.8 | 88.4965
0 | 0 | 42.8 | 89.1 | 220297.8 | 89.11 | 88.8955
0 | 0 | 41.6 | 89.75 | 221884.7 | 89.7 | 89.544
0 | 0 | 40.4 | 90.26 | 223146.9 | 90.18 | 90.0529
0 | 0 | 39.2 | 90.83 | 224551.5 | 90.71 | 90.6216
0 | 0 | 38 | 91.48 | 226160.8 | 91.33 | 91.2701
0 | 0 | 36.7 | 91.87 | 227127.9 | 91.71 | 91.6592
0 | 0 | 35.5 | 92.38 | 228403 | 92.43 | 92.168
0 | 0 | 34.3 | 92.87 | 229605.9 | 92.82 | 92.6569
0 | 0 | 33.1 | 93.42 | 230962.4 | 93.15 | 93.2056
0 | 0 | 31.9 | 94.01 | 232418.5 | 93.9 | 93.7943
0 | 0 | 30.7 | 94.39 | 233359.9 | 94.33 | 94.1734
0 | 0 | 29.5 | 94.92 | 234681 | 94.98 | 94.7022
0 | 0 | 28.3 | 95.54 | 236196.2 | 95.42 | 95.3208
0 | 0 | 27.1 | 96.24 | 237945.4 | 95.87 | 96.0192
0 | 0 | 25.9 | 96.58 | 238772.5 | 96.42 | 96.3584
0 | 0 | 24.7 | 96.9 | 239569.2 | 96.88 | 96.6776
0 | 0 | 23.5 | 97.44 | 240917.7 | 97.6 | 97.2164
0 | 0 | 22.3 | 97.9 | 242037.5 | 97.85 | 97.6753
0 | 0 | 21 | 98.4 | 243283.7 | 98.34 | 98.1742
0 | 0 | 19.8 | 98.76 | 244173.8 | 98.7 | 98.5334
0 | 0 | 18.6 | 99.23 | 245327.4 | 99.28 | 99.0023
0 | 0 | 17.4 | 99.51 | 246015.6 | 99.66 | 99.2817
0 | 0 | 16.2 | 99.83 | 246806.3 | 99.86 | 99.6009
0 | 0 | 15 | 100.05 | 247363.4 | 100 | 99.8204
0 | 0 | 13.8 | 100.23 | 247804.1 | 99.96 | 100
0 | 0 | 12.6 | 99.9 | 246994.4 | 99.72 | 99.6708
0 | 0 | 11.4 | 99.64 | 246347.1 | 99.41 | 99.4114
0 | 0 | 10.2 | 98.62 | 243834.6 | 98.58 | 98.3937
0 | 0 | 9 | 97.23 | 240376.9 | 97.06 | 97.0069
0 | 0 | 7.8 | 94.71 | 234144.2 | 94.75 | 94.4927
0 | 0 | 6.6 | 91.68 | 226660.7 | 91.52 | 91.4696
0 | 0 | 5.4 | 86.23 | 213188.9 | 86.14 | 86.0321
0 | 0 | 4.2 | 78.73 | 194657.8 | 78.67 | 78.5493
0 | 0 | 2.9 | 68.16 | 168516.7 | 68.12 | 68.0036
0 | 0 | 1.7 | 58.98 | 145816.8 | 59.04 | 58.8447
0 | 0 | 0.5 | 53.75 | 132881.1 | 53.64 | 53.6267
0 | 0 | -0.5 | 50.5 | 124865 | 50.46 | 50.3841
Measurement time: | 12/08/2022 18:59:41
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 16.29 | 40265.3 | 17.01 | 17.2326
0 | 0 | 309.1 | 16.54 | 40882.8 | 17.43 | 17.4971
0 | 0 | 307.6 | 16.76 | 41425.7 | 17.67 | 17.7298
0 | 0 | 306.6 | 16.78 | 41489.7 | 17.73 | 17.751
0 | 0 | 305.3 | 16.94 | 41887.5 | 17.87 | 17.9202
0 | 0 | 304.1 | 17.1 | 42273.9 | 18.03 | 18.0895
0 | 0 | 302.9 | 17.17 | 42450.3 | 18.12 | 18.1635
0 | 0 | 301.7 | 17.31 | 42805.8 | 18.27 | 18.3116
0 | 0 | 300.5 | 17.44 | 43124.6 | 18.39 | 18.4492
0 | 0 | 299.3 | 17.58 | 43474.5 | 18.57 | 18.5973
0 | 0 | 298.1 | 17.72 | 43814 | 18.68 | 18.7454
0 | 0 | 296.9 | 17.8 | 44017.9 | 18.77 | 18.83
0 | 0 | 295.7 | 17.89 | 44240.4 | 18.92 | 18.9252
0 | 0 | 294.5 | 18.14 | 44844.3 | 19.18 | 19.1897
0 | 0 | 293.3 | 18.17 | 44921.2 | 19.19 | 19.2214
0 | 0 | 292.1 | 18.32 | 45304.9 | 19.36 | 19.3801
0 | 0 | 290.9 | 18.5 | 45740.3 | 19.53 | 19.5705
0 | 0 | 289.7 | 18.52 | 45787.4 | 19.57 | 19.5917
0 | 0 | 288.5 | 18.7 | 46244.5 | 19.78 | 19.7821
0 | 0 | 287.3 | 18.89 | 46691.6 | 19.94 | 19.9831
0 | 0 | 286.1 | 19.02 | 47025.9 | 20.08 | 20.1206
0 | 0 | 284.9 | 19.13 | 47296.9 | 20.2 | 20.237
0 | 0 | 283.7 | 19.25 | 47580.8 | 20.33 | 20.3639
0 | 0 | 282.5 | 19.37 | 47897.8 | 20.45 | 20.4908
0 | 0 | 281.3 | 19.55 | 48327 | 20.63 | 20.6813
0 | 0 | 280.1 | 19.66 | 48612.1 | 20.75 | 20.7976
0 | 0 | 278.9 | 19.92 | 49236.9 | 21 | 21.0727
0 | 0 | 277.7 | 19.94 | 49290.9 | 21.03 | 21.0938
0 | 0 | 276.5 | 20.05 | 49564.9 | 21.17 | 21.2102
0 | 0 | 275.2 | 20.22 | 49999.6 | 21.34 | 21.39
0 | 0 | 274.1 | 20.41 | 50464.5 | 21.56 | 21.591
0 | 0 | 272.9 | 20.49 | 50670 | 21.65 | 21.6757
0 | 0 | 271.7 | 20.69 | 51163.2 | 21.85 | 21.8872
0 | 0 | 270.5 | 20.89 | 51642.5 | 22.07 | 22.0988
0 | 0 | 269.2 | 21 | 51931.3 | 22.19 | 22.2152
0 | 0 | 268.1 | 21.12 | 52207.6 | 22.31 | 22.3421
0 | 0 | 266.8 | 21.33 | 52726.8 | 22.58 | 22.5643
0 | 0 | 265.6 | 21.4 | 52897.2 | 22.59 | 22.6383
0 | 0 | 264.4 | 21.48 | 53114.3 | 22.68 | 22.7229
0 | 0 | 263.2 | 21.73 | 53713.7 | 22.93 | 22.9874
0 | 0 | 262 | 21.93 | 54213.2 | 23.17 | 23.199
0 | 0 | 260.8 | 22.09 | 54617.4 | 23.36 | 23.3682
0 | 0 | 259.6 | 22.21 | 54912.6 | 23.45 | 23.4952
0 | 0 | 258.4 | 22.37 | 55304.1 | 23.67 | 23.6644
0 | 0 | 257.2 | 22.56 | 55780.2 | 23.89 | 23.8654
0 | 0 | 256 | 22.76 | 56278.7 | 24.04 | 24.077
0 | 0 | 254.8 | 22.89 | 56585.7 | 24.16 | 24.2145
0 | 0 | 253.6 | 23.02 | 56921.6 | 24.35 | 24.3521
0 | 0 | 252.4 | 23.17 | 57285.8 | 24.5 | 24.5107
0 | 0 | 251.2 | 23.33 | 57689 | 24.65 | 24.68
0 | 0 | 250 | 23.47 | 58031.2 | 24.83 | 24.8281
0 | 0 | 248.8 | 23.71 | 58613.1 | 25.05 | 25.082
0 | 0 | 247.6 | 23.87 | 59026.9 | 25.22 | 25.2512
0 | 0 | 246.4 | 24.08 | 59527.7 | 25.45 | 25.4734
0 | 0 | 245.2 | 24.22 | 59890.1 | 25.59 | 25.6215
0 | 0 | 244 | 24.44 | 60427.2 | 25.8 | 25.8542
0 | 0 | 242.7 | 24.57 | 60740.2 | 25.98 | 25.9917
0 | 0 | 241.6 | 24.74 | 61170.1 | 26.13 | 26.1716
0 | 0 | 240.3 | 24.89 | 61541.2 | 26.37 | 26.3303
0 | 0 | 239.1 | 25.12 | 62097.5 | 26.53 | 26.5736
0 | 0 | 237.9 | 25.33 | 62619.1 | 26.8 | 26.7957
0 | 0 | 236.7 | 25.5 | 63035.9 | 26.95 | 26.9756
0 | 0 | 235.5 | 25.64 | 63403 | 27.09 | 27.1237
0 | 0 | 234.3 | 25.85 | 63910.4 | 27.32 | 27.3458
0 | 0 | 233.1 | 26.08 | 64478.4 | 27.54 | 27.5891
0 | 0 | 231.9 | 26.19 | 64759.1 | 27.69 | 27.7055
0 | 0 | 230.7 | 26.48 | 65460.5 | 28.05 | 28.0123
0 | 0 | 229.5 | 26.56 | 65675.7 | 28.16 | 28.0969
0 | 0 | 228.3 | 26.82 | 66312.4 | 28.42 | 28.3719
0 | 0 | 227.1 | 26.99 | 66724.4 | 28.56 | 28.5518
0 | 0 | 225.9 | 27.23 | 67311.9 | 28.81 | 28.8057
0 | 0 | 224.7 | 27.39 | 67716.6 | 28.93 | 28.9749
0 | 0 | 223.5 | 27.64 | 68335.4 | 29.19 | 29.2394
0 | 0 | 222.3 | 27.81 | 68745.4 | 29.43 | 29.4192
0 | 0 | 221.1 | 28.09 | 69460.2 | 29.72 | 29.7154
0 | 0 | 219.9 | 28.22 | 69779.5 | 29.85 | 29.853
0 | 0 | 218.6 | 28.43 | 70280.2 | 30.04 | 30.0751
0 | 0 | 217.5 | 28.66 | 70853.7 | 30.33 | 30.3184
0 | 0 | 216.3 | 28.85 | 71315.5 | 30.5 | 30.5194
0 | 0 | 215.1 | 29.03 | 71769.5 | 30.75 | 30.7098
0 | 0 | 213.9 | 29.23 | 72260.9 | 30.95 | 30.9214
0 | 0 | 212.6 | 29.48 | 72875.4 | 31.16 | 31.1859
0 | 0 | 211.5 | 29.72 | 73487.1 | 31.42 | 31.4398
0 | 0 | 210.2 | 29.84 | 73780.5 | 31.59 | 31.5667
0 | 0 | 209 | 30.18 | 74607.9 | 31.92 | 31.9264
0 | 0 | 207.8 | 30.37 | 75091.1 | 32.13 | 32.1274
0 | 0 | 206.6 | 30.62 | 75703 | 32.37 | 32.3918
0 | 0 | 205.4 | 30.81 | 76183.5 | 32.58 | 32.5928
0 | 0 | 204.2 | 31.03 | 76708 | 32.82 | 32.8256
0 | 0 | 203 | 31.2 | 77146.9 | 33.02 | 33.0054
0 | 0 | 201.8 | 31.56 | 78017.1 | 33.43 | 33.3862
0 | 0 | 200.6 | 31.76 | 78524.3 | 33.62 | 33.5978
0 | 0 | 199.4 | 31.95 | 78997.5 | 33.79 | 33.7988
0 | 0 | 198.2 | 32.21 | 79645 | 34.09 | 34.0738
0 | 0 | 197 | 32.48 | 80311.4 | 34.34 | 34.3595
0 | 0 | 195.8 | 32.71 | 80877.9 | 34.63 | 34.6028
0 | 0 | 194.6 | 32.9 | 81329.8 | 34.82 | 34.8038
0 | 0 | 193.4 | 33.18 | 82039.6 | 35.16 | 35.1
0 | 0 | 192.2 | 33.41 | 82589.6 | 35.32 | 35.3433
0 | 0 | 191 | 33.66 | 83215.4 | 35.59 | 35.6077
0 | 0 | 189.8 | 33.89 | 83780.9 | 35.89 | 35.8511
0 | 0 | 188.5 | 34.14 | 84409 | 36.15 | 36.1155
0 | 0 | 187.4 | 34.3 | 84798.5 | 36.34 | 36.2848
0 | 0 | 186.2 | 34.7 | 85789.2 | 36.69 | 36.7079
0 | 0 | 185 | 34.86 | 86188.9 | 36.97 | 36.8772
0 | 0 | 183.8 | 35.21 | 87049.5 | 37.25 | 37.2474
0 | 0 | 182.5 | 35.33 | 87356.6 | 37.4 | 37.3744
0 | 0 | 181.4 | 35.68 | 88202.8 | 37.8 | 37.7446
0 | 0 | 180.1 | 35.96 | 88899.2 | 38.14 | 38.0408
0 | 0 | 178.9 | 36.27 | 89660 | 38.38 | 38.3688
0 | 0 | 177.8 | 36.53 | 90326.8 | 38.69 | 38.6438
0 | 0 | 176.5 | 36.84 | 91088.2 | 39.01 | 38.9718
0 | 0 | 175.3 | 37 | 91472.9 | 39.23 | 39.141
0 | 0 | 174.1 | 37.26 | 92124.4 | 39.46 | 39.4161
0 | 0 | 172.9 | 37.58 | 92917.3 | 39.79 | 39.7546
0 | 0 | 171.7 | 37.84 | 93550.2 | 40.05 | 40.0296
0 | 0 | 170.5 | 38.22 | 94495.1 | 40.49 | 40.4316
0 | 0 | 169.3 | 38.42 | 94983.5 | 40.73 | 40.6432
0 | 0 | 168.1 | 38.74 | 95778.3 | 41.05 | 40.9817
0 | 0 | 166.9 | 39.04 | 96511.3 | 41.42 | 41.2991
0 | 0 | 165.7 | 39.27 | 97086.2 | 41.55 | 41.5424
0 | 0 | 164.5 | 39.55 | 97785.7 | 41.93 | 41.8386
0 | 0 | 163.3 | 39.91 | 98663.1 | 42.26 | 42.2194
0 | 0 | 162.1 | 40.12 | 99186.4 | 42.45 | 42.4416
0 | 0 | 160.9 | 40.52 | 100177.2 | 42.85 | 42.8647
0 | 0 | 159.7 | 40.75 | 100760.1 | 43.2 | 43.108
0 | 0 | 158.4 | 41.06 | 101518.2 | 43.47 | 43.4359
0 | 0 | 157.3 | 41.41 | 102382.7 | 43.85 | 43.8062
0 | 0 | 156 | 41.67 | 103032.9 | 44.17 | 44.0812
0 | 0 | 154.9 | 41.91 | 103624.3 | 44.45 | 44.3351
0 | 0 | 153.6 | 42.24 | 104437 | 44.74 | 44.6842
0 | 0 | 152.5 | 42.56 | 105234.1 | 45.06 | 45.0227
0 | 0 | 151.3 | 42.86 | 105970.3 | 45.44 | 45.3401
0 | 0 | 150 | 43.19 | 106771.1 | 45.77 | 45.6892
0 | 0 | 148.8 | 43.55 | 107681.4 | 46.14 | 46.07
0 | 0 | 147.6 | 43.88 | 108498.3 | 46.46 | 46.4191
0 | 0 | 146.4 | 44.18 | 109234.4 | 46.72 | 46.7365
0 | 0 | 145.2 | 44.6 | 110278.3 | 47.24 | 47.1808
0 | 0 | 144 | 44.88 | 110969.3 | 47.58 | 47.477
0 | 0 | 142.8 | 45.2 | 111751.4 | 47.86 | 47.8155
0 | 0 | 141.6 | 45.56 | 112642.6 | 48.25 | 48.1963
0 | 0 | 140.4 | 45.94 | 113575.6 | 48.74 | 48.5983
0 | 0 | 139.2 | 46.28 | 114424.2 | 49.09 | 48.958
0 | 0 | 138 | 46.54 | 115075.1 | 49.34 | 49.233
0 | 0 | 136.8 | 46.9 | 115949.3 | 49.65 | 49.6139
0 | 0 | 135.6 | 47.26 | 116845.4 | 50.09 | 49.9947
0 | 0 | 134.4 | 47.66 | 117821.8 | 50.36 | 50.4179
0 | 0 | 133.2 | 47.97 | 118607.3 | 50.87 | 50.7458
0 | 0 | 132 | 48.44 | 119761.2 | 51.34 | 51.243
0 | 0 | 130.8 | 48.71 | 120421 | 51.56 | 51.5286
0 | 0 | 129.6 | 49.06 | 121297.3 | 51.91 | 51.8989
0 | 0 | 128.4 | 49.45 | 122254.6 | 52.32 | 52.3114
0 | 0 | 127.2 | 49.76 | 123014.7 | 52.78 | 52.6394
0 | 0 | 125.9 | 50.19 | 124092.6 | 53.28 | 53.0943
0 | 0 | 124.7 | 50.54 | 124954 | 53.56 | 53.4645
0 | 0 | 123.5 | 50.89 | 125815.5 | 53.92 | 53.8348
0 | 0 | 122.3 | 51.34 | 126940.6 | 54.33 | 54.3108
0 | 0 | 121.2 | 51.62 | 127625.6 | 54.71 | 54.607
0 | 0 | 119.9 | 52.08 | 128747.7 | 55.16 | 55.0936
0 | 0 | 118.7 | 52.51 | 129815.6 | 55.57 | 55.5485
0 | 0 | 117.5 | 52.81 | 130554 | 55.91 | 55.8659
0 | 0 | 116.3 | 53.26 | 131680.8 | 56.48 | 56.3419
0 | 0 | 115.1 | 53.68 | 132715.6 | 56.85 | 56.7862
0 | 0 | 113.9 | 54.05 | 133619.8 | 57.26 | 57.1776
0 | 0 | 112.7 | 54.52 | 134791 | 57.87 | 57.6748
0 | 0 | 111.5 | 54.81 | 135502.6 | 58.11 | 57.9816
0 | 0 | 110.3 | 55.28 | 136676.6 | 58.6 | 58.4788
0 | 0 | 109.1 | 55.79 | 137939.7 | 59.06 | 59.0183
0 | 0 | 107.9 | 56.13 | 138779 | 59.38 | 59.378
0 | 0 | 106.7 | 56.54 | 139793.2 | 59.9 | 59.8117
0 | 0 | 105.5 | 56.87 | 140607.8 | 60.33 | 60.1608
0 | 0 | 104.3 | 57.29 | 141641 | 60.79 | 60.6051
0 | 0 | 103.1 | 57.81 | 142935.6 | 61.16 | 61.1552
0 | 0 | 101.9 | 58.13 | 143725.5 | 61.5 | 61.4937
0 | 0 | 100.7 | 58.63 | 144950.4 | 62.08 | 62.0226
0 | 0 | 99.5 | 59.11 | 146151.6 | 62.57 | 62.5304
0 | 0 | 98.3 | 59.59 | 147338 | 63.23 | 63.0382
0 | 0 | 97.1 | 59.95 | 148224.9 | 63.54 | 63.419
0 | 0 | 95.8 | 60.24 | 148932.2 | 63.84 | 63.7258
0 | 0 | 94.7 | 60.76 | 150212.6 | 64.4 | 64.2759
0 | 0 | 93.5 | 61.37 | 151727.4 | 65.02 | 64.9212
0 | 0 | 92.2 | 61.65 | 152428.8 | 65.33 | 65.2174
0 | 0 | 91.1 | 62.21 | 153798 | 65.84 | 65.8098
0 | 0 | 89.8 | 62.69 | 154989.4 | 66.34 | 66.3176
0 | 0 | 88.7 | 63.09 | 155981 | 66.89 | 66.7407
0 | 0 | 87.4 | 63.63 | 157310.3 | 67.47 | 67.312
0 | 0 | 86.2 | 64.1 | 158484.5 | 67.99 | 67.8092
0 | 0 | 85 | 64.48 | 159422.1 | 68.5 | 68.2111
0 | 0 | 83.8 | 64.93 | 160520.1 | 68.91 | 68.6872
0 | 0 | 82.6 | 65.5 | 161943.6 | 69.38 | 69.2902
0 | 0 | 81.4 | 65.9 | 162916.5 | 69.79 | 69.7133
0 | 0 | 80.2 | 66.56 | 164565.6 | 70.51 | 70.4115
0 | 0 | 79 | 66.92 | 165461.9 | 70.79 | 70.7923
0 | 0 | 77.8 | 67.39 | 166603.5 | 71.42 | 71.2895
0 | 0 | 76.6 | 67.94 | 167962.3 | 71.98 | 71.8714
0 | 0 | 75.4 | 68.43 | 169180.2 | 72.51 | 72.3897
0 | 0 | 74.2 | 68.76 | 169996 | 72.95 | 72.7388
0 | 0 | 73 | 69.25 | 171200.3 | 73.54 | 73.2572
0 | 0 | 71.8 | 69.79 | 172545.8 | 74.05 | 73.8284
0 | 0 | 70.6 | 70.39 | 174028.2 | 74.59 | 74.4631
0 | 0 | 69.4 | 71.01 | 175567.3 | 75.16 | 75.119
0 | 0 | 68.1 | 71.39 | 176512.5 | 75.7 | 75.521
0 | 0 | 66.9 | 71.93 | 177839.8 | 76.22 | 76.0922
0 | 0 | 65.7 | 72.27 | 178670.8 | 76.75 | 76.4519
0 | 0 | 64.5 | 72.94 | 180327.2 | 77.42 | 77.1607
0 | 0 | 63.3 | 73.56 | 181863.7 | 77.77 | 77.8166
0 | 0 | 62.1 | 74.16 | 183341 | 78.55 | 78.4513
0 | 0 | 60.9 | 74.48 | 184130.3 | 79.13 | 78.7898
0 | 0 | 59.7 | 75.04 | 185531.8 | 79.58 | 79.3822
0 | 0 | 58.5 | 75.65 | 187036 | 80.07 | 80.0275
0 | 0 | 57.3 | 76.06 | 188057.6 | 80.69 | 80.4612
0 | 0 | 56.1 | 76.66 | 189526.1 | 81.36 | 81.0959
0 | 0 | 54.9 | 77.24 | 190966 | 81.79 | 81.7095
0 | 0 | 53.6 | 77.8 | 192343 | 82.4 | 82.3019
0 | 0 | 52.5 | 78.32 | 193622.3 | 82.89 | 82.852
0 | 0 | 51.2 | 78.97 | 195238.3 | 83.61 | 83.5396
0 | 0 | 50 | 79.37 | 196238.4 | 84.23 | 83.9628
0 | 0 | 48.8 | 80.03 | 197862.5 | 84.68 | 84.661
0 | 0 | 47.6 | 80.43 | 198863.4 | 85.39 | 85.0841
0 | 0 | 46.4 | 81.27 | 200934.8 | 86.08 | 85.9727
0 | 0 | 45.2 | 81.66 | 201887.7 | 86.52 | 86.3853
0 | 0 | 44 | 82.11 | 202993.7 | 87.09 | 86.8613
0 | 0 | 42.8 | 82.68 | 204406.9 | 87.72 | 87.4643
0 | 0 | 41.6 | 83.24 | 205796.3 | 88.26 | 88.0567
0 | 0 | 40.4 | 83.9 | 207432.8 | 88.71 | 88.7549
0 | 0 | 39.2 | 84.41 | 208698.1 | 89.26 | 89.2944
0 | 0 | 38 | 84.99 | 210136.8 | 90.05 | 89.908
0 | 0 | 36.7 | 85.49 | 211359.1 | 90.52 | 90.4369
0 | 0 | 35.5 | 85.96 | 212513.6 | 91.09 | 90.9341
0 | 0 | 34.3 | 86.61 | 214136.3 | 91.86 | 91.6217
0 | 0 | 33.1 | 87.18 | 215548.5 | 92.44 | 92.2247
0 | 0 | 31.9 | 87.66 | 216732.5 | 93.16 | 92.7325
0 | 0 | 30.7 | 88.4 | 218557.7 | 93.63 | 93.5153
0 | 0 | 29.5 | 88.81 | 219558.1 | 94.2 | 93.949
0 | 0 | 28.3 | 89.47 | 221199 | 95.03 | 94.6472
0 | 0 | 27.1 | 90.07 | 222689.6 | 95.34 | 95.2819
0 | 0 | 25.8 | 90.53 | 223812.3 | 95.96 | 95.7685
0 | 0 | 24.7 | 91.08 | 225177.7 | 96.53 | 96.3504
0 | 0 | 23.4 | 91.64 | 226555.2 | 97.03 | 96.9428
0 | 0 | 22.2 | 92.09 | 227686.1 | 97.56 | 97.4188
0 | 0 | 21 | 92.62 | 228992 | 97.96 | 97.9795
0 | 0 | 19.8 | 92.94 | 229789.1 | 98.56 | 98.318
0 | 0 | 18.6 | 93.51 | 231184.1 | 99.26 | 98.921
0 | 0 | 17.4 | 93.84 | 231994.4 | 99.35 | 99.2701
0 | 0 | 16.2 | 94.3 | 233153.1 | 99.85 | 99.7567
0 | 0 | 15 | 94.53 | 233720.7 | 100 | 100
0 | 0 | 13.8 | 94.18 | 232851.6 | 99.82 | 99.6297
0 | 0 | 12.6 | 93.77 | 231837.6 | 99.55 | 99.196
0 | 0 | 11.4 | 93.24 | 230520.1 | 98.97 | 98.6354
0 | 0 | 10.2 | 92.32 | 228235.1 | 97.73 | 97.6621
0 | 0 | 9 | 90.6 | 223984.7 | 95.93 | 95.8426
0 | 0 | 7.8 | 87.97 | 217493.2 | 93.13 | 93.0604
0 | 0 | 6.6 | 84.21 | 208207.4 | 89.14 | 89.0828
0 | 0 | 5.3 | 78.78 | 194776 | 83.67 | 83.3386
0 | 0 | 4.1 | 71.18 | 175970.3 | 75.46 | 75.2988
0 | 0 | 2.9 | 60.49 | 149542.8 | 64.11 | 63.9903
0 | 0 | 1.7 | 51.04 | 126190.1 | 54.16 | 53.9934
0 | 0 | 0.5 | 46.29 | 114433.7 | 49.11 | 48.9686
0 | 0 | -0.5 | 43.11 | 106574.1 | 45.77 | 45.6046
Measurement time: | 12/08/2022 19:04:06
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 24.63 | 60901.6 | 22.75 | 23.6986
0 | 0 | 309.1 | 23.53 | 58183.6 | 22.6 | 22.6402
0 | 0 | 307.6 | 23.71 | 58631.3 | 22.79 | 22.8134
0 | 0 | 306.6 | 23.83 | 58906.6 | 22.87 | 22.9289
0 | 0 | 305.3 | 24.01 | 59350.3 | 23.03 | 23.1021
0 | 0 | 304.1 | 24.17 | 59756.9 | 23.19 | 23.256
0 | 0 | 302.9 | 24.39 | 60312 | 23.41 | 23.4677
0 | 0 | 301.7 | 24.48 | 60520.1 | 23.51 | 23.5543
0 | 0 | 300.5 | 24.62 | 60857.4 | 23.63 | 23.689
0 | 0 | 299.3 | 24.78 | 61276.3 | 23.8 | 23.843
0 | 0 | 298.1 | 24.94 | 61656.9 | 23.91 | 23.9969
0 | 0 | 296.9 | 25.14 | 62149.1 | 24.14 | 24.1894
0 | 0 | 295.7 | 25.28 | 62489.2 | 24.25 | 24.3241
0 | 0 | 294.5 | 25.44 | 62885.4 | 24.38 | 24.478
0 | 0 | 293.3 | 25.62 | 63336.9 | 24.62 | 24.6512
0 | 0 | 292.1 | 25.79 | 63755.7 | 24.77 | 24.8148
0 | 0 | 290.9 | 25.95 | 64163.2 | 24.9 | 24.9687
0 | 0 | 289.7 | 26.06 | 64422.8 | 25.01 | 25.0746
0 | 0 | 288.5 | 26.19 | 64753.6 | 25.15 | 25.1997
0 | 0 | 287.3 | 26.41 | 65298.7 | 25.34 | 25.4113
0 | 0 | 286.1 | 26.58 | 65710.8 | 25.53 | 25.5749
0 | 0 | 284.9 | 26.83 | 66326.3 | 25.76 | 25.8155
0 | 0 | 283.7 | 26.98 | 66714 | 25.92 | 25.9598
0 | 0 | 282.5 | 27.17 | 67185.2 | 26.09 | 26.1426
0 | 0 | 281.3 | 27.27 | 67408.9 | 26.22 | 26.2388
0 | 0 | 280.1 | 27.49 | 67960.4 | 26.4 | 26.4505
0 | 0 | 278.9 | 27.71 | 68510 | 26.66 | 26.6622
0 | 0 | 277.7 | 27.87 | 68914.3 | 26.77 | 26.8161
0 | 0 | 276.5 | 28.04 | 69329.5 | 26.9 | 26.9797
0 | 0 | 275.3 | 28.21 | 69740.7 | 27.09 | 27.1433
0 | 0 | 274.1 | 28.36 | 70126.3 | 27.24 | 27.2876
0 | 0 | 272.8 | 28.52 | 70504.2 | 27.4 | 27.4415
0 | 0 | 271.7 | 28.73 | 71028 | 27.62 | 27.6436
0 | 0 | 270.4 | 28.91 | 71470.6 | 27.74 | 27.8168
0 | 0 | 269.2 | 29.07 | 71863.7 | 27.92 | 27.9707
0 | 0 | 268.1 | 29.34 | 72533.4 | 28.21 | 28.2305
0 | 0 | 266.8 | 29.44 | 72797 | 28.29 | 28.3268
0 | 0 | 265.6 | 29.68 | 73390.6 | 28.48 | 28.5577
0 | 0 | 264.4 | 29.82 | 73737.4 | 28.69 | 28.6924
0 | 0 | 263.2 | 30.05 | 74290.1 | 28.85 | 28.9137
0 | 0 | 262 | 30.27 | 74844.8 | 29.11 | 29.1254
0 | 0 | 260.8 | 30.46 | 75311 | 29.23 | 29.3082
0 | 0 | 259.6 | 30.61 | 75678.7 | 29.41 | 29.4525
0 | 0 | 258.4 | 30.77 | 76086.4 | 29.59 | 29.6065
0 | 0 | 257.2 | 30.95 | 76524.6 | 29.75 | 29.7797
0 | 0 | 256 | 31.24 | 77225.4 | 30.05 | 30.0587
0 | 0 | 254.8 | 31.38 | 77579.2 | 30.19 | 30.1934
0 | 0 | 253.6 | 31.6 | 78125.4 | 30.37 | 30.4051
0 | 0 | 252.4 | 31.71 | 78389.1 | 30.49 | 30.5109
0 | 0 | 251.2 | 31.99 | 79081.5 | 30.73 | 30.7803
0 | 0 | 250 | 32.2 | 79620.4 | 30.94 | 30.9824
0 | 0 | 248.8 | 32.48 | 80299.4 | 31.19 | 31.2518
0 | 0 | 247.6 | 32.73 | 80910.1 | 31.4 | 31.4924
0 | 0 | 246.4 | 32.87 | 81278.1 | 31.58 | 31.6271
0 | 0 | 245.2 | 33.11 | 81853.5 | 31.81 | 31.858
0 | 0 | 244 | 33.28 | 82287.2 | 32.02 | 32.0216
0 | 0 | 242.8 | 33.54 | 82930.3 | 32.24 | 32.2717
0 | 0 | 241.6 | 33.72 | 83374.7 | 32.42 | 32.4449
0 | 0 | 240.3 | 33.91 | 83838.5 | 32.62 | 32.6277
0 | 0 | 239.1 | 34.12 | 84357.1 | 32.8 | 32.8298
0 | 0 | 238 | 34.4 | 85050.4 | 33.08 | 33.0992
0 | 0 | 236.7 | 34.59 | 85526.3 | 33.26 | 33.282
0 | 0 | 235.5 | 34.86 | 86191.7 | 33.48 | 33.5418
0 | 0 | 234.3 | 35.1 | 86770.3 | 33.69 | 33.7727
0 | 0 | 233.1 | 35.23 | 87102.9 | 33.81 | 33.8978
0 | 0 | 231.9 | 35.53 | 87840 | 34.13 | 34.1865
0 | 0 | 230.7 | 35.82 | 88566.5 | 34.4 | 34.4655
0 | 0 | 229.5 | 36.01 | 89022.8 | 34.62 | 34.6483
0 | 0 | 228.3 | 36.17 | 89423.2 | 34.73 | 34.8023
0 | 0 | 227.1 | 36.43 | 90079 | 35.01 | 35.0524
0 | 0 | 225.9 | 36.63 | 90571.7 | 35.24 | 35.2449
0 | 0 | 224.7 | 36.93 | 91314.3 | 35.54 | 35.5335
0 | 0 | 223.5 | 37.22 | 92018.5 | 35.78 | 35.8126
0 | 0 | 222.3 | 37.39 | 92444 | 35.95 | 35.9761
0 | 0 | 221.1 | 37.64 | 93065.2 | 36.17 | 36.2167
0 | 0 | 219.9 | 37.96 | 93852 | 36.46 | 36.5246
0 | 0 | 218.6 | 38.23 | 94520.6 | 36.79 | 36.7844
0 | 0 | 217.5 | 38.42 | 94988.6 | 36.97 | 36.9672
0 | 0 | 216.3 | 38.64 | 95528.8 | 37.14 | 37.1789
0 | 0 | 215.1 | 38.86 | 96079.8 | 37.36 | 37.3906
0 | 0 | 213.9 | 39.21 | 96931.7 | 37.68 | 37.7273
0 | 0 | 212.6 | 39.41 | 97439.7 | 37.87 | 37.9198
0 | 0 | 211.4 | 39.6 | 97908.2 | 38.09 | 38.1026
0 | 0 | 210.2 | 39.91 | 98666 | 38.32 | 38.4008
0 | 0 | 209 | 40.02 | 98939.8 | 38.56 | 38.5067
0 | 0 | 207.8 | 40.42 | 99923.6 | 38.88 | 38.8916
0 | 0 | 206.6 | 40.53 | 100211.1 | 38.99 | 38.9974
0 | 0 | 205.4 | 40.84 | 100982.5 | 39.23 | 39.2957
0 | 0 | 204.2 | 41.01 | 101396.3 | 39.44 | 39.4593
0 | 0 | 203 | 41.45 | 102466.7 | 39.82 | 39.8826
0 | 0 | 201.8 | 41.67 | 103019.1 | 40.1 | 40.0943
0 | 0 | 200.6 | 41.99 | 103820.5 | 40.38 | 40.4022
0 | 0 | 199.4 | 42.26 | 104474.6 | 40.64 | 40.662
0 | 0 | 198.2 | 42.47 | 105001.3 | 40.87 | 40.864
0 | 0 | 197 | 42.84 | 105904.2 | 41.21 | 41.2201
0 | 0 | 195.8 | 43.01 | 106325.2 | 41.39 | 41.3836
0 | 0 | 194.6 | 43.33 | 107125.8 | 41.72 | 41.6915
0 | 0 | 193.4 | 43.63 | 107878.2 | 41.95 | 41.9802
0 | 0 | 192.2 | 43.86 | 108425.8 | 42.2 | 42.2015
0 | 0 | 191 | 44.07 | 108961 | 42.39 | 42.4035
0 | 0 | 189.8 | 44.4 | 109760.9 | 42.74 | 42.7211
0 | 0 | 188.6 | 44.79 | 110741.9 | 43 | 43.0963
0 | 0 | 187.4 | 45.13 | 111565.7 | 43.4 | 43.4235
0 | 0 | 186.2 | 45.32 | 112036.4 | 43.64 | 43.6063
0 | 0 | 185 | 45.66 | 112898.1 | 43.89 | 43.9334
0 | 0 | 183.8 | 45.88 | 113434.2 | 44.13 | 44.1451
0 | 0 | 182.5 | 46.26 | 114379.7 | 44.57 | 44.5107
0 | 0 | 181.4 | 46.47 | 114890.2 | 44.67 | 44.7128
0 | 0 | 180.1 | 46.78 | 115645.5 | 45.01 | 45.0111
0 | 0 | 178.9 | 47.16 | 116592.6 | 45.29 | 45.3767
0 | 0 | 177.8 | 47.47 | 117364 | 45.67 | 45.675
0 | 0 | 176.5 | 47.68 | 117886.3 | 45.84 | 45.877
0 | 0 | 175.3 | 48.04 | 118773 | 46.18 | 46.2234
0 | 0 | 174.1 | 48.25 | 119293.9 | 46.43 | 46.4255
0 | 0 | 172.9 | 48.62 | 120209 | 46.77 | 46.7815
0 | 0 | 171.7 | 48.87 | 120828.7 | 47.07 | 47.022
0 | 0 | 170.5 | 49.23 | 121704.7 | 47.36 | 47.3684
0 | 0 | 169.3 | 49.49 | 122349.6 | 47.59 | 47.6186
0 | 0 | 168.1 | 49.8 | 123126.6 | 47.91 | 47.9169
0 | 0 | 166.9 | 50.24 | 124200.7 | 48.21 | 48.3402
0 | 0 | 165.7 | 50.57 | 125026.6 | 48.63 | 48.6578
0 | 0 | 164.5 | 50.73 | 125423.5 | 48.84 | 48.8117
0 | 0 | 163.3 | 51.07 | 126268.6 | 49.03 | 49.1388
0 | 0 | 162.1 | 51.37 | 127015.3 | 49.42 | 49.4275
0 | 0 | 160.9 | 51.81 | 128087.2 | 49.77 | 49.8509
0 | 0 | 159.7 | 52.13 | 128872 | 50.21 | 50.1588
0 | 0 | 158.5 | 52.47 | 129729.1 | 50.37 | 50.4859
0 | 0 | 157.3 | 52.72 | 130341.7 | 50.74 | 50.7265
0 | 0 | 156.1 | 53.02 | 131077.7 | 50.98 | 51.0151
0 | 0 | 154.9 | 53.51 | 132295.6 | 51.51 | 51.4866
0 | 0 | 153.7 | 53.78 | 132953.6 | 51.71 | 51.7464
0 | 0 | 152.5 | 54.13 | 133835.3 | 52.16 | 52.0831
0 | 0 | 151.3 | 54.52 | 134793.6 | 52.37 | 52.4584
0 | 0 | 150.1 | 54.91 | 135764 | 52.76 | 52.8336
0 | 0 | 148.9 | 55.22 | 136534.9 | 53.08 | 53.1319
0 | 0 | 147.7 | 55.6 | 137457.3 | 53.5 | 53.4975
0 | 0 | 146.4 | 55.84 | 138058.4 | 53.78 | 53.7285
0 | 0 | 145.3 | 56.2 | 138942.1 | 54.1 | 54.0749
0 | 0 | 144 | 56.6 | 139940.2 | 54.44 | 54.4597
0 | 0 | 142.9 | 56.85 | 140544.1 | 54.72 | 54.7003
0 | 0 | 141.7 | 57.29 | 141645.5 | 55.04 | 55.1236
0 | 0 | 140.4 | 57.81 | 142924.8 | 55.43 | 55.624
0 | 0 | 139.3 | 58.03 | 143473.2 | 55.75 | 55.8357
0 | 0 | 138 | 58.35 | 144271.1 | 56.08 | 56.1436
0 | 0 | 136.8 | 58.69 | 145112.7 | 56.44 | 56.4707
0 | 0 | 135.6 | 59.12 | 146171.4 | 56.92 | 56.8844
0 | 0 | 134.4 | 59.59 | 147322.8 | 57.34 | 57.3367
0 | 0 | 133.2 | 59.85 | 147960.3 | 57.48 | 57.5868
0 | 0 | 132 | 60.34 | 149174.7 | 58.05 | 58.0583
0 | 0 | 130.8 | 60.73 | 150143 | 58.44 | 58.4336
0 | 0 | 129.6 | 61.08 | 151017.5 | 58.7 | 58.7703
0 | 0 | 128.4 | 61.49 | 152028 | 59.03 | 59.1648
0 | 0 | 127.2 | 61.83 | 152855.7 | 59.5 | 59.492
0 | 0 | 126 | 62.07 | 153454.1 | 59.76 | 59.7229
0 | 0 | 124.8 | 62.57 | 154703.6 | 60.23 | 60.204
0 | 0 | 123.6 | 63.01 | 155772.2 | 60.55 | 60.6273
0 | 0 | 122.4 | 63.17 | 156178.9 | 60.8 | 60.7813
0 | 0 | 121.2 | 63.71 | 157520.7 | 61.3 | 61.3009
0 | 0 | 120 | 64.18 | 158683.1 | 61.71 | 61.7531
0 | 0 | 118.8 | 64.47 | 159396.4 | 61.95 | 62.0321
0 | 0 | 117.6 | 64.95 | 160589.6 | 62.41 | 62.494
0 | 0 | 116.4 | 65.26 | 161337.5 | 62.76 | 62.7923
0 | 0 | 115.2 | 65.57 | 162108.7 | 63.1 | 63.0905
0 | 0 | 114 | 65.91 | 162961.2 | 63.44 | 63.4177
0 | 0 | 112.8 | 66.41 | 164200.5 | 63.88 | 63.8988
0 | 0 | 111.6 | 66.75 | 165035.9 | 64.24 | 64.2259
0 | 0 | 110.3 | 67.16 | 166050.3 | 64.66 | 64.6204
0 | 0 | 109.2 | 67.61 | 167164.3 | 64.97 | 65.0534
0 | 0 | 107.9 | 67.95 | 167992.8 | 65.35 | 65.3805
0 | 0 | 106.7 | 68.46 | 169255.2 | 65.82 | 65.8713
0 | 0 | 105.5 | 69.05 | 170727.7 | 66.31 | 66.4389
0 | 0 | 104.3 | 69.44 | 171684.2 | 66.73 | 66.8142
0 | 0 | 103.2 | 69.84 | 172663.4 | 67.21 | 67.1991
0 | 0 | 101.9 | 70.26 | 173714.4 | 67.61 | 67.6032
0 | 0 | 100.7 | 70.69 | 174781.1 | 67.82 | 68.0169
0 | 0 | 99.5 | 71.1 | 175785.7 | 68.36 | 68.4114
0 | 0 | 98.3 | 71.47 | 176697.9 | 68.72 | 68.7674
0 | 0 | 97.1 | 71.78 | 177464.8 | 69.05 | 69.0657
0 | 0 | 95.9 | 72.29 | 178724.1 | 69.6 | 69.5564
0 | 0 | 94.7 | 72.71 | 179762.2 | 69.96 | 69.9606
0 | 0 | 93.5 | 73.28 | 181179.3 | 70.47 | 70.509
0 | 0 | 92.3 | 73.68 | 182168.3 | 70.87 | 70.8939
0 | 0 | 91.1 | 74 | 182963.7 | 71.22 | 71.2018
0 | 0 | 89.9 | 74.49 | 184177.2 | 71.71 | 71.6732
0 | 0 | 88.7 | 75.13 | 185744.6 | 72.2 | 72.289
0 | 0 | 87.5 | 75.67 | 187078.1 | 72.62 | 72.8086
0 | 0 | 86.3 | 76.14 | 188239.8 | 73.17 | 73.2608
0 | 0 | 85.1 | 76.52 | 189175.7 | 73.62 | 73.6265
0 | 0 | 83.8 | 76.92 | 190171.6 | 74.02 | 74.0114
0 | 0 | 82.6 | 77.37 | 191291.8 | 74.5 | 74.4443
0 | 0 | 81.4 | 77.68 | 192043.2 | 74.69 | 74.7426
0 | 0 | 80.2 | 78.22 | 193384.5 | 75.18 | 75.2622
0 | 0 | 79 | 78.69 | 194548 | 75.85 | 75.7144
0 | 0 | 77.8 | 79.2 | 195813.8 | 76.18 | 76.2051
0 | 0 | 76.6 | 79.75 | 197162.6 | 76.63 | 76.7343
0 | 0 | 75.4 | 80.11 | 198059.4 | 77.02 | 77.0807
0 | 0 | 74.2 | 80.72 | 199566.2 | 77.59 | 77.6677
0 | 0 | 73 | 81.05 | 200377.9 | 78.07 | 77.9852
0 | 0 | 71.8 | 81.59 | 201713.4 | 78.5 | 78.5048
0 | 0 | 70.6 | 82.02 | 202782.2 | 78.93 | 78.9185
0 | 0 | 69.4 | 82.3 | 203483.6 | 79.21 | 79.1879
0 | 0 | 68.1 | 82.7 | 204456.5 | 79.46 | 79.5728
0 | 0 | 66.9 | 83.14 | 205560.6 | 79.96 | 79.9962
0 | 0 | 65.7 | 83.61 | 206711.9 | 80.6 | 80.4484
0 | 0 | 64.6 | 83.9 | 207424.1 | 80.94 | 80.7274
0 | 0 | 63.3 | 84.55 | 209031.7 | 81.5 | 81.3528
0 | 0 | 62.1 | 85.31 | 210921.2 | 81.92 | 82.0841
0 | 0 | 60.9 | 85.88 | 212315.9 | 82.52 | 82.6325
0 | 0 | 59.7 | 86.3 | 213354.3 | 82.94 | 83.0367
0 | 0 | 58.5 | 86.74 | 214442.3 | 83.5 | 83.46
0 | 0 | 57.3 | 87.17 | 215509 | 83.92 | 83.8738
0 | 0 | 56.1 | 87.76 | 216971.1 | 84.45 | 84.4415
0 | 0 | 54.9 | 88.37 | 218475.2 | 84.94 | 85.0284
0 | 0 | 53.6 | 88.7 | 219296.6 | 85.49 | 85.3459
0 | 0 | 52.5 | 89.22 | 220590.3 | 85.84 | 85.8462
0 | 0 | 51.3 | 89.73 | 221837.4 | 86.35 | 86.337
0 | 0 | 50 | 90.13 | 222838.1 | 86.62 | 86.7218
0 | 0 | 48.8 | 90.73 | 224307.1 | 87.19 | 87.2991
0 | 0 | 47.6 | 91.27 | 225659.7 | 87.6 | 87.8187
0 | 0 | 46.4 | 91.62 | 226518.7 | 88.11 | 88.1555
0 | 0 | 45.2 | 91.89 | 227187.7 | 88.4 | 88.4153
0 | 0 | 44 | 92.68 | 229139.3 | 88.98 | 89.1754
0 | 0 | 42.8 | 93.14 | 230270.5 | 89.63 | 89.618
0 | 0 | 41.6 | 93.54 | 231252.6 | 90.1 | 90.0029
0 | 0 | 40.4 | 94.09 | 232618.2 | 90.42 | 90.5321
0 | 0 | 39.2 | 94.43 | 233469.2 | 90.85 | 90.8592
0 | 0 | 38 | 95.07 | 235039.3 | 91.35 | 91.475
0 | 0 | 36.8 | 95.49 | 236090.4 | 91.9 | 91.8791
0 | 0 | 35.5 | 96.17 | 237775.9 | 92.45 | 92.5334
0 | 0 | 34.3 | 96.72 | 239118.6 | 92.86 | 93.0626
0 | 0 | 33.1 | 97.03 | 239889.7 | 93.45 | 93.3609
0 | 0 | 31.9 | 97.7 | 241536.5 | 94.13 | 94.0056
0 | 0 | 30.7 | 98.16 | 242691.5 | 94.5 | 94.4482
0 | 0 | 29.5 | 98.56 | 243678.2 | 94.84 | 94.8331
0 | 0 | 28.3 | 99.04 | 244856 | 95.26 | 95.2949
0 | 0 | 27.1 | 99.6 | 246254.4 | 95.85 | 95.8337
0 | 0 | 25.9 | 100.21 | 247752.2 | 96.29 | 96.4207
0 | 0 | 24.7 | 100.64 | 248809.3 | 96.9 | 96.8344
0 | 0 | 23.5 | 101.17 | 250121.4 | 97.22 | 97.3444
0 | 0 | 22.3 | 101.43 | 250759.8 | 97.81 | 97.5945
0 | 0 | 21.1 | 102.12 | 252476.5 | 98.1 | 98.2584
0 | 0 | 19.8 | 102.61 | 253683.9 | 98.54 | 98.7299
0 | 0 | 18.7 | 102.88 | 254349.3 | 99.09 | 98.9897
0 | 0 | 17.4 | 102.98 | 254602.8 | 99.16 | 99.0859
0 | 0 | 16.2 | 103.32 | 255440.2 | 99.58 | 99.4131
0 | 0 | 15 | 103.55 | 256004.8 | 99.77 | 99.6344
0 | 0 | 13.8 | 103.93 | 256959.7 | 100 | 100
0 | 0 | 12.6 | 103.82 | 256679.2 | 99.86 | 99.8942
0 | 0 | 11.4 | 103.69 | 256360.9 | 99.66 | 99.7691
0 | 0 | 10.2 | 102.8 | 254163.4 | 99.15 | 98.9127
0 | 0 | 9 | 101.6 | 251179 | 97.85 | 97.7581
0 | 0 | 7.8 | 99.73 | 246571.7 | 95.94 | 95.9588
0 | 0 | 6.6 | 96.77 | 239249.5 | 92.88 | 93.1107
0 | 0 | 5.4 | 91.69 | 226690.8 | 88.25 | 88.2228
0 | 0 | 4.2 | 85.23 | 210719.2 | 82.04 | 82.0071
0 | 0 | 3 | 74.97 | 185360.2 | 72.17 | 72.1351
0 | 0 | 1.8 | 65.58 | 162129.5 | 63.11 | 63.1002
0 | 0 | 0.6 | 60.53 | 149659.8 | 58.25 | 58.2411
0 | 0 | -0.5 | 57.11 | 141201.1 | 54.97 | 54.9504
Measurement time: | 12/08/2022 19:07:57
Radiation device: | Clinac2 Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 200 x 200 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 23.97 | 59258.7 | 24.07 | 22.467
0 | 0 | 309.1 | 25.91 | 64060.4 | 24.28 | 24.2853
0 | 0 | 307.6 | 26.2 | 64772.1 | 24.52 | 24.5571
0 | 0 | 306.6 | 26.31 | 65057.1 | 24.62 | 24.6602
0 | 0 | 305.3 | 26.51 | 65543.8 | 24.84 | 24.8477
0 | 0 | 304.1 | 26.69 | 65996.7 | 25.01 | 25.0164
0 | 0 | 302.9 | 26.83 | 66333.9 | 25.08 | 25.1476
0 | 0 | 301.7 | 26.94 | 66594.4 | 25.25 | 25.2507
0 | 0 | 300.5 | 27.11 | 67030.4 | 25.42 | 25.4101
0 | 0 | 299.3 | 27.27 | 67408.7 | 25.54 | 25.56
0 | 0 | 298.1 | 27.48 | 67938.6 | 25.72 | 25.7569
0 | 0 | 296.9 | 27.62 | 68289.4 | 25.86 | 25.8881
0 | 0 | 295.7 | 27.82 | 68775.3 | 26.06 | 26.0755
0 | 0 | 294.5 | 27.92 | 69029.3 | 26.13 | 26.1693
0 | 0 | 293.3 | 28.18 | 69674.3 | 26.37 | 26.413
0 | 0 | 292.1 | 28.35 | 70086.6 | 26.49 | 26.5723
0 | 0 | 290.9 | 28.45 | 70326.4 | 26.64 | 26.666
0 | 0 | 289.7 | 28.67 | 70872.6 | 26.82 | 26.8722
0 | 0 | 288.5 | 28.88 | 71401.7 | 27.02 | 27.0691
0 | 0 | 287.3 | 29.04 | 71797.8 | 27.21 | 27.219
0 | 0 | 286.1 | 29.15 | 72075.4 | 27.3 | 27.3221
0 | 0 | 284.9 | 29.47 | 72859.3 | 27.58 | 27.6221
0 | 0 | 283.7 | 29.6 | 73179.8 | 27.67 | 27.7439
0 | 0 | 282.5 | 29.82 | 73727.8 | 27.95 | 27.9501
0 | 0 | 281.3 | 30.01 | 74193.8 | 28.05 | 28.1282
0 | 0 | 280.1 | 30.22 | 74714.8 | 28.3 | 28.3251
0 | 0 | 278.8 | 30.34 | 75010.5 | 28.41 | 28.4375
0 | 0 | 277.7 | 30.52 | 75451.3 | 28.58 | 28.6062
0 | 0 | 276.4 | 30.7 | 75908.7 | 28.78 | 28.775
0 | 0 | 275.2 | 31.01 | 76656 | 29.06 | 29.0655
0 | 0 | 274.1 | 31.08 | 76838.7 | 29.17 | 29.1311
0 | 0 | 272.8 | 31.35 | 77511.1 | 29.38 | 29.3842
0 | 0 | 271.6 | 31.5 | 77879 | 29.53 | 29.5248
0 | 0 | 270.4 | 31.69 | 78351.4 | 29.7 | 29.7029
0 | 0 | 269.2 | 31.88 | 78826.8 | 29.87 | 29.881
0 | 0 | 268 | 32.06 | 79274 | 29.99 | 30.0497
0 | 0 | 266.8 | 32.27 | 79779.3 | 30.23 | 30.2465
0 | 0 | 265.6 | 32.45 | 80235.1 | 30.4 | 30.4152
0 | 0 | 264.4 | 32.59 | 80584 | 30.5 | 30.5464
0 | 0 | 263.2 | 32.89 | 81305.4 | 30.74 | 30.8276
0 | 0 | 262 | 33.16 | 81981.1 | 31.05 | 31.0807
0 | 0 | 260.8 | 33.27 | 82248.7 | 31.15 | 31.1838
0 | 0 | 259.6 | 33.39 | 82552.1 | 31.28 | 31.2963
0 | 0 | 258.4 | 33.64 | 83181 | 31.55 | 31.5306
0 | 0 | 257.2 | 33.87 | 83741.2 | 31.77 | 31.7462
0 | 0 | 256 | 34.1 | 84296.1 | 31.95 | 31.9618
0 | 0 | 254.8 | 34.27 | 84739.3 | 32.11 | 32.1211
0 | 0 | 253.6 | 34.49 | 85269.4 | 32.28 | 32.3273
0 | 0 | 252.4 | 34.7 | 85780.6 | 32.5 | 32.5241
0 | 0 | 251.2 | 34.89 | 86252.7 | 32.68 | 32.7022
0 | 0 | 250 | 35.05 | 86654.3 | 32.9 | 32.8522
0 | 0 | 248.8 | 35.2 | 87032.8 | 32.98 | 32.9928
0 | 0 | 247.6 | 35.55 | 87888.4 | 33.27 | 33.3208
0 | 0 | 246.4 | 35.66 | 88154.4 | 33.36 | 33.4239
0 | 0 | 245.2 | 35.95 | 88887.4 | 33.65 | 33.6958
0 | 0 | 244 | 36.28 | 89698.7 | 33.95 | 34.0051
0 | 0 | 242.7 | 36.33 | 89820.5 | 34.08 | 34.0519
0 | 0 | 241.6 | 36.66 | 90640.7 | 34.33 | 34.3612
0 | 0 | 240.3 | 36.92 | 91281 | 34.55 | 34.6049
0 | 0 | 239.1 | 37.12 | 91763.4 | 34.74 | 34.7924
0 | 0 | 237.9 | 37.41 | 92501.5 | 35.02 | 35.0642
0 | 0 | 236.7 | 37.65 | 93083.1 | 35.2 | 35.2892
0 | 0 | 235.5 | 37.83 | 93518.6 | 35.44 | 35.4579
0 | 0 | 234.3 | 38.05 | 94081.8 | 35.67 | 35.6641
0 | 0 | 233.1 | 38.33 | 94766.2 | 35.91 | 35.9265
0 | 0 | 231.9 | 38.53 | 95253.5 | 36.07 | 36.114
0 | 0 | 230.7 | 38.82 | 95973.2 | 36.41 | 36.3858
0 | 0 | 229.5 | 39.03 | 96496 | 36.51 | 36.5826
0 | 0 | 228.3 | 39.23 | 96979.3 | 36.67 | 36.7701
0 | 0 | 227.1 | 39.49 | 97630.3 | 36.98 | 37.0138
0 | 0 | 225.9 | 39.72 | 98189.7 | 37.22 | 37.2294
0 | 0 | 224.7 | 39.95 | 98781.7 | 37.46 | 37.4449
0 | 0 | 223.4 | 40.27 | 99564 | 37.71 | 37.7449
0 | 0 | 222.2 | 40.35 | 99771.5 | 37.79 | 37.8199
0 | 0 | 221 | 40.66 | 100519.9 | 38.09 | 38.1104
0 | 0 | 219.9 | 40.97 | 101303.7 | 38.3 | 38.401
0 | 0 | 218.6 | 41.17 | 101775.3 | 38.59 | 38.5884
0 | 0 | 217.4 | 41.43 | 102424.9 | 38.82 | 38.8321
0 | 0 | 216.2 | 41.68 | 103039.7 | 39 | 39.0665
0 | 0 | 215 | 42.01 | 103868.3 | 39.34 | 39.3758
0 | 0 | 213.8 | 42.25 | 104452 | 39.6 | 39.6007
0 | 0 | 212.6 | 42.54 | 105167.2 | 39.8 | 39.8725
0 | 0 | 211.4 | 42.7 | 105577.8 | 39.97 | 40.0225
0 | 0 | 210.2 | 43 | 106322.8 | 40.24 | 40.3037
0 | 0 | 209 | 43.24 | 106899.3 | 40.49 | 40.5286
0 | 0 | 207.8 | 43.47 | 107482.5 | 40.78 | 40.7442
0 | 0 | 206.6 | 43.87 | 108452.3 | 41.08 | 41.1191
0 | 0 | 205.4 | 44.01 | 108811 | 41.16 | 41.2504
0 | 0 | 204.2 | 44.25 | 109392.8 | 41.43 | 41.4753
0 | 0 | 203 | 44.54 | 110106.5 | 41.72 | 41.7471
0 | 0 | 201.8 | 44.83 | 110836.2 | 42.04 | 42.0189
0 | 0 | 200.6 | 45.12 | 111551.8 | 42.23 | 42.2907
0 | 0 | 199.4 | 45.4 | 112243.7 | 42.47 | 42.5532
0 | 0 | 198.2 | 45.7 | 112995.1 | 42.71 | 42.8344
0 | 0 | 197 | 45.87 | 113408 | 42.95 | 42.9937
0 | 0 | 195.8 | 46.34 | 114558.6 | 43.4 | 43.4342
0 | 0 | 194.6 | 46.65 | 115330.8 | 43.63 | 43.7248
0 | 0 | 193.4 | 46.93 | 116019.2 | 43.88 | 43.9873
0 | 0 | 192.1 | 47.07 | 116366.6 | 44.1 | 44.1185
0 | 0 | 190.9 | 47.36 | 117086.3 | 44.4 | 44.3903
0 | 0 | 189.8 | 47.7 | 117943.4 | 44.7 | 44.709
0 | 0 | 188.5 | 48 | 118668.5 | 44.96 | 44.9902
0 | 0 | 187.3 | 48.36 | 119553.3 | 45.28 | 45.3276
0 | 0 | 186.1 | 48.57 | 120080.6 | 45.51 | 45.5244
0 | 0 | 184.9 | 48.85 | 120785.9 | 45.68 | 45.7869
0 | 0 | 183.7 | 49.14 | 121490.2 | 46.05 | 46.0587
0 | 0 | 182.5 | 49.57 | 122565.3 | 46.45 | 46.4617
0 | 0 | 181.3 | 49.79 | 123104.1 | 46.56 | 46.6679
0 | 0 | 180.1 | 50.03 | 123681.4 | 46.88 | 46.8929
0 | 0 | 178.9 | 50.33 | 124426.5 | 47.13 | 47.1741
0 | 0 | 177.7 | 50.72 | 125391.5 | 47.53 | 47.5396
0 | 0 | 176.5 | 51 | 126087.5 | 47.73 | 47.802
0 | 0 | 175.3 | 51.29 | 126809.9 | 48.02 | 48.0739
0 | 0 | 174.1 | 51.62 | 127632.2 | 48.37 | 48.3832
0 | 0 | 172.9 | 51.91 | 128332.3 | 48.68 | 48.655
0 | 0 | 171.7 | 52.2 | 129062.6 | 48.91 | 48.9268
0 | 0 | 170.4 | 52.53 | 129878 | 49.17 | 49.2361
0 | 0 | 169.3 | 52.86 | 130681.8 | 49.49 | 49.5454
0 | 0 | 168.1 | 53.04 | 131141.4 | 49.66 | 49.7141
0 | 0 | 166.9 | 53.52 | 132312.4 | 50.07 | 50.164
0 | 0 | 165.7 | 53.81 | 133045.8 | 50.46 | 50.4358
0 | 0 | 164.4 | 54.17 | 133930.8 | 50.77 | 50.7733
0 | 0 | 163.2 | 54.47 | 134669 | 51.07 | 51.0545
0 | 0 | 162 | 54.86 | 135633.4 | 51.37 | 51.42
0 | 0 | 160.8 | 55.15 | 136346.6 | 51.68 | 51.6918
0 | 0 | 159.6 | 55.51 | 137251.1 | 52.03 | 52.0292
0 | 0 | 158.4 | 55.88 | 138143.1 | 52.27 | 52.376
0 | 0 | 157.2 | 56.17 | 138881.3 | 52.63 | 52.6479
0 | 0 | 156 | 56.41 | 139458.2 | 52.86 | 52.8728
0 | 0 | 154.8 | 56.87 | 140605.1 | 53.25 | 53.304
0 | 0 | 153.6 | 57.19 | 141393.7 | 53.54 | 53.6039
0 | 0 | 152.4 | 57.43 | 141977.9 | 53.83 | 53.8288
0 | 0 | 151.2 | 57.95 | 143277.3 | 54.28 | 54.3162
0 | 0 | 150 | 58.2 | 143896.1 | 54.57 | 54.5506
0 | 0 | 148.8 | 58.53 | 144710.9 | 54.89 | 54.8599
0 | 0 | 147.6 | 58.93 | 145686.7 | 55.26 | 55.2348
0 | 0 | 146.4 | 59.29 | 146588.2 | 55.52 | 55.5722
0 | 0 | 145.2 | 59.66 | 147499.2 | 55.94 | 55.919
0 | 0 | 144 | 59.99 | 148318.4 | 56.22 | 56.2283
0 | 0 | 142.8 | 60.28 | 149023.6 | 56.52 | 56.5001
0 | 0 | 141.6 | 60.71 | 150105.2 | 56.79 | 56.9032
0 | 0 | 140.3 | 61.01 | 150850 | 57.17 | 57.1844
0 | 0 | 139.2 | 61.35 | 151681.4 | 57.62 | 57.503
0 | 0 | 137.9 | 61.7 | 152547.3 | 57.93 | 57.8311
0 | 0 | 136.8 | 62.11 | 153551.8 | 58.06 | 58.2154
0 | 0 | 135.6 | 62.51 | 154558.5 | 58.59 | 58.5903
0 | 0 | 134.3 | 62.91 | 155526.7 | 58.93 | 58.9652
0 | 0 | 133.1 | 63.41 | 156760.2 | 59.34 | 59.4339
0 | 0 | 131.9 | 63.72 | 157546.4 | 59.68 | 59.7244
0 | 0 | 130.7 | 64.04 | 158318.7 | 60.01 | 60.0244
0 | 0 | 129.5 | 64.46 | 159360.3 | 60.44 | 60.418
0 | 0 | 128.3 | 64.85 | 160343.8 | 60.77 | 60.7836
0 | 0 | 127.1 | 65.19 | 161183.1 | 61.09 | 61.1023
0 | 0 | 125.9 | 65.47 | 161860.6 | 61.34 | 61.3647
0 | 0 | 124.7 | 66.01 | 163199.3 | 61.85 | 61.8708
0 | 0 | 123.5 | 66.15 | 163539.1 | 62.07 | 62.0021
0 | 0 | 122.3 | 66.7 | 164898.8 | 62.53 | 62.5176
0 | 0 | 121.1 | 67.06 | 165796.8 | 62.85 | 62.855
0 | 0 | 119.9 | 67.37 | 166553.6 | 63.26 | 63.1456
0 | 0 | 118.7 | 67.79 | 167602.3 | 63.64 | 63.5392
0 | 0 | 117.5 | 68.17 | 168550 | 64.12 | 63.8954
0 | 0 | 116.3 | 68.69 | 169823 | 64.54 | 64.3828
0 | 0 | 115.1 | 68.98 | 170535.9 | 64.85 | 64.6546
0 | 0 | 113.9 | 69.47 | 171754.1 | 65.2 | 65.1139
0 | 0 | 112.7 | 70.11 | 173332.2 | 65.67 | 65.7138
0 | 0 | 111.5 | 70.4 | 174043.8 | 65.77 | 65.9856
0 | 0 | 110.2 | 70.91 | 175302.3 | 66.29 | 66.4636
0 | 0 | 109.1 | 71.33 | 176359.3 | 66.65 | 66.8572
0 | 0 | 107.8 | 71.45 | 176641.7 | 66.88 | 66.9697
0 | 0 | 106.7 | 71.68 | 177223.6 | 67.18 | 67.1853
0 | 0 | 105.5 | 72.31 | 178768.8 | 67.73 | 67.7758
0 | 0 | 104.2 | 72.72 | 179780.3 | 68.11 | 68.1601
0 | 0 | 103.1 | 73.1 | 180728.4 | 68.42 | 68.5163
0 | 0 | 101.8 | 73.43 | 181553 | 68.85 | 68.8256
0 | 0 | 100.6 | 73.79 | 182432.8 | 69.03 | 69.163
0 | 0 | 99.5 | 74.31 | 183714.5 | 69.71 | 69.6504
0 | 0 | 98.2 | 74.84 | 185029.9 | 70.08 | 70.1472
0 | 0 | 97 | 75.34 | 186260.4 | 70.54 | 70.6158
0 | 0 | 95.8 | 75.61 | 186944.4 | 70.88 | 70.8689
0 | 0 | 94.6 | 76.02 | 187957.8 | 71.32 | 71.2532
0 | 0 | 93.4 | 76.74 | 189715.9 | 71.8 | 71.928
0 | 0 | 92.2 | 77 | 190379.4 | 72.19 | 72.1717
0 | 0 | 91 | 77.37 | 191290.5 | 72.49 | 72.5185
0 | 0 | 89.8 | 77.82 | 192402.7 | 72.93 | 72.9403
0 | 0 | 88.6 | 78.26 | 193489.7 | 73.31 | 73.3527
0 | 0 | 87.4 | 78.78 | 194767.3 | 73.86 | 73.8401
0 | 0 | 86.2 | 79.07 | 195483.8 | 74.13 | 74.1119
0 | 0 | 85 | 79.69 | 197026.5 | 74.66 | 74.693
0 | 0 | 83.8 | 80.07 | 197952.5 | 75.09 | 75.0492
0 | 0 | 82.6 | 80.45 | 198905.2 | 75.56 | 75.4054
0 | 0 | 81.4 | 81.03 | 200323.8 | 75.79 | 75.949
0 | 0 | 80.2 | 81.4 | 201257.5 | 76.23 | 76.2958
0 | 0 | 79 | 81.74 | 202101.1 | 76.46 | 76.6145
0 | 0 | 77.7 | 82.42 | 203779.9 | 77.01 | 77.2519
0 | 0 | 76.5 | 82.6 | 204222.9 | 77.33 | 77.4206
0 | 0 | 75.3 | 83.15 | 205574 | 77.86 | 77.9361
0 | 0 | 74.1 | 83.63 | 206762.4 | 78.34 | 78.386
0 | 0 | 72.9 | 84.06 | 207835.3 | 78.79 | 78.789
0 | 0 | 71.7 | 84.47 | 208849.6 | 79.16 | 79.1733
0 | 0 | 70.5 | 84.93 | 209981.4 | 79.59 | 79.6045
0 | 0 | 69.3 | 85.68 | 211823.6 | 80.11 | 80.3074
0 | 0 | 68.1 | 85.96 | 212517.5 | 80.55 | 80.5699
0 | 0 | 66.9 | 86.25 | 213251.5 | 80.91 | 80.8417
0 | 0 | 65.7 | 86.72 | 214412.5 | 81.29 | 81.2822
0 | 0 | 64.5 | 87.3 | 215833.4 | 81.84 | 81.8259
0 | 0 | 63.3 | 87.81 | 217107.9 | 82.16 | 82.3039
0 | 0 | 62.1 | 88.33 | 218371.3 | 82.64 | 82.7913
0 | 0 | 60.9 | 88.83 | 219625.1 | 83.05 | 83.2599
0 | 0 | 59.6 | 89.08 | 220224.8 | 83.45 | 83.4942
0 | 0 | 58.4 | 89.51 | 221306.2 | 83.79 | 83.8973
0 | 0 | 57.2 | 90.25 | 223125.7 | 84.49 | 84.5909
0 | 0 | 56 | 90.64 | 224086.3 | 84.82 | 84.9564
0 | 0 | 54.8 | 91.11 | 225267.3 | 85.35 | 85.3969
0 | 0 | 53.6 | 91.59 | 226454 | 85.73 | 85.8468
0 | 0 | 52.4 | 92.07 | 227632.5 | 86.16 | 86.2967
0 | 0 | 51.2 | 92.5 | 228703.5 | 86.67 | 86.6998
0 | 0 | 50 | 93.04 | 230021.2 | 87.1 | 87.2059
0 | 0 | 48.8 | 93.34 | 230779.7 | 87.46 | 87.4871
0 | 0 | 47.5 | 94.07 | 232564.1 | 87.92 | 88.1713
0 | 0 | 46.4 | 94.4 | 233383.2 | 88.45 | 88.4806
0 | 0 | 45.1 | 95.17 | 235294.5 | 89.05 | 89.2024
0 | 0 | 43.9 | 95.39 | 235849.9 | 89.22 | 89.4086
0 | 0 | 42.7 | 95.89 | 237072.1 | 89.96 | 89.8772
0 | 0 | 41.5 | 96.36 | 238223.9 | 90.42 | 90.3177
0 | 0 | 40.3 | 96.87 | 239484.4 | 90.84 | 90.7958
0 | 0 | 39.1 | 97.32 | 240602.5 | 91.2 | 91.2175
0 | 0 | 37.9 | 97.98 | 242241.6 | 91.63 | 91.8362
0 | 0 | 36.7 | 98.45 | 243401.6 | 92.1 | 92.2767
0 | 0 | 35.5 | 98.55 | 243652.6 | 92.31 | 92.3704
0 | 0 | 34.3 | 98.98 | 244719.3 | 92.82 | 92.7735
0 | 0 | 33.1 | 99.5 | 245994.1 | 93.24 | 93.2608
0 | 0 | 31.8 | 100.23 | 247803.9 | 93.85 | 93.9451
0 | 0 | 30.7 | 100.71 | 248987.5 | 94.32 | 94.395
0 | 0 | 29.4 | 101.3 | 250444.9 | 94.83 | 94.948
0 | 0 | 28.2 | 101.63 | 251269.9 | 95.04 | 95.2573
0 | 0 | 27 | 101.97 | 252108.5 | 95.58 | 95.576
0 | 0 | 25.8 | 102.64 | 253750.4 | 96.04 | 96.204
0 | 0 | 24.6 | 103.01 | 254677.5 | 96.59 | 96.5508
0 | 0 | 23.4 | 103.55 | 256010.1 | 97.03 | 97.0569
0 | 0 | 22.2 | 104.06 | 257279 | 97.64 | 97.5349
0 | 0 | 21 | 104.63 | 258687.6 | 97.98 | 98.0692
0 | 0 | 19.8 | 105.03 | 259678.9 | 98.27 | 98.4441
0 | 0 | 18.5 | 105.33 | 260401.2 | 98.75 | 98.7253
0 | 0 | 17.3 | 105.59 | 261044.4 | 98.96 | 98.969
0 | 0 | 16.1 | 105.97 | 262005.8 | 99.27 | 99.3251
0 | 0 | 14.9 | 106.41 | 263081.8 | 99.7 | 99.7376
0 | 0 | 13.7 | 106.69 | 263782.3 | 99.92 | 100
0 | 0 | 12.5 | 106.67 | 263735.5 | 100 | 99.9813
0 | 0 | 11.3 | 106.29 | 262778.8 | 99.8 | 99.6251
0 | 0 | 10.1 | 105.6 | 261070.6 | 98.94 | 98.9783
0 | 0 | 8.9 | 104.6 | 258604.6 | 97.91 | 98.0411
0 | 0 | 7.7 | 103.05 | 254785.6 | 96.53 | 96.5882
0 | 0 | 6.5 | 100.01 | 247271.2 | 93.8 | 93.7389
0 | 0 | 5.3 | 95.75 | 236721.5 | 89.56 | 89.746
0 | 0 | 4.1 | 89.42 | 221088 | 83.67 | 83.8129
0 | 0 | 2.9 | 79.67 | 196961.5 | 74.58 | 74.6743
0 | 0 | 1.7 | 71.34 | 176389.8 | 66.77 | 66.8666
0 | 0 | 0.4 | 66.23 | 163736.6 | 62.19 | 62.077
0 | 0 | -0.5 | 63.02 | 155808.4 | 59.06 | 59.0683