## Rendements énergie et DSP
Measurement time: | 02/08/2022 19:51:44 | [5 empty] | Measurement time: | 12/08/2022 18:00:16
Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Clinac2 Accelerator
Energy: | 6 MV | [5 empty] | Energy: | 6 MV
Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F | [5 empty] | Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie | [5 empty] | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm
SSD: | 900 mm | [5 empty] | SSD: | 1000 mm
Field size: | 100 x 100 mm | [5 empty] | Field size: | 100 x 100 mm
Gantry angle: | 0 | [5 empty] | Gantry angle: | 0
Measurement medium: | Water | [5 empty] | Measurement medium: | Water
Scan type: | Beam | [5 empty] | Scan type: | Beam
Scan mode: | Continuous | [5 empty] | Scan mode: | Continuous
Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 310 | 40.35 | 58577.9 | 19.25 | 19.1223 | 0 | 0 | 310 | 19.24 | 47561.7 | 20.48 | 19.0911
0 | 0 | 309.1 | 40.78 | 59206.4 | 19.3 | 19.3261 | 0 | 0 | 309.1 | 20.74 | 51265.8 | 20.6 | 20.5795
0 | 0 | 307.6 | 41.23 | 59856.5 | 19.52 | 19.5394 | 0 | 0 | 307.6 | 20.83 | 51504.6 | 20.67 | 20.6688
0 | 0 | 306.6 | 41.54 | 60306.7 | 19.68 | 19.6863 | 0 | 0 | 306.6 | 21.02 | 51974.8 | 20.89 | 20.8573
0 | 0 | 305.3 | 41.82 | 60709.2 | 19.77 | 19.819 | 0 | 0 | 305.3 | 21.15 | 52292.6 | 21.01 | 20.9863
0 | 0 | 304.1 | 42.11 | 61133.4 | 19.92 | 19.9564 | 0 | 0 | 304.1 | 21.35 | 52782.9 | 21.19 | 21.1848
0 | 0 | 302.9 | 42.41 | 61564.9 | 20.12 | 20.0986 | 0 | 0 | 302.9 | 21.46 | 53066 | 21.33 | 21.2939
0 | 0 | 301.7 | 42.73 | 62033 | 20.2 | 20.2502 | 0 | 0 | 301.7 | 21.62 | 53451.6 | 21.46 | 21.4527
0 | 0 | 300.5 | 43.02 | 62454.6 | 20.36 | 20.3877 | 0 | 0 | 300.5 | 21.73 | 53713.5 | 21.55 | 21.5618
0 | 0 | 299.3 | 43.28 | 62838 | 20.48 | 20.5109 | 0 | 0 | 299.3 | 21.83 | 53983.6 | 21.68 | 21.661
0 | 0 | 298.1 | 43.59 | 63276.6 | 20.62 | 20.6578 | 0 | 0 | 298.1 | 22.08 | 54597.4 | 21.97 | 21.9091
0 | 0 | 296.9 | 43.8 | 63593.1 | 20.74 | 20.7573 | 0 | 0 | 296.9 | 22.18 | 54836.5 | 22.02 | 22.0083
0 | 0 | 295.7 | 44.13 | 64061.4 | 20.87 | 20.9137 | 0 | 0 | 295.7 | 22.34 | 55238 | 22.21 | 22.1671
0 | 0 | 294.5 | 44.44 | 64521.1 | 21.02 | 21.0606 | 0 | 0 | 294.5 | 22.52 | 55683.7 | 22.35 | 22.3457
0 | 0 | 293.3 | 44.72 | 64923.2 | 21.18 | 21.1933 | 0 | 0 | 293.3 | 22.64 | 55983.6 | 22.46 | 22.4648
0 | 0 | 292.1 | 45.06 | 65412.3 | 21.35 | 21.3544 | 0 | 0 | 292.1 | 22.78 | 56332 | 22.67 | 22.6037
0 | 0 | 290.9 | 45.42 | 65936.4 | 21.51 | 21.525 | 0 | 0 | 290.9 | 22.95 | 56744.8 | 22.82 | 22.7724
0 | 0 | 289.7 | 45.68 | 66314.5 | 21.63 | 21.6483 | 0 | 0 | 289.7 | 23.07 | 57049 | 22.94 | 22.8914
0 | 0 | 288.5 | 46.08 | 66898.4 | 21.84 | 21.8378 | 0 | 0 | 288.5 | 23.29 | 57574 | 23.13 | 23.1097
0 | 0 | 287.3 | 46.33 | 67264.7 | 21.96 | 21.9563 | 0 | 0 | 287.3 | 23.45 | 57968 | 23.31 | 23.2685
0 | 0 | 286.1 | 46.61 | 67665.8 | 22.09 | 22.089 | 0 | 0 | 286.1 | 23.53 | 58174.1 | 23.42 | 23.3479
0 | 0 | 284.9 | 46.86 | 68035.1 | 22.18 | 22.2075 | 0 | 0 | 284.9 | 23.7 | 58600 | 23.54 | 23.5166
0 | 0 | 283.7 | 47.17 | 68487.2 | 22.32 | 22.3544 | 0 | 0 | 283.7 | 23.93 | 59151.6 | 23.76 | 23.7448
0 | 0 | 282.5 | 47.61 | 69120.8 | 22.53 | 22.5629 | 0 | 0 | 282.5 | 24.01 | 59365.5 | 23.88 | 23.8242
0 | 0 | 281.3 | 47.87 | 69503.8 | 22.67 | 22.6861 | 0 | 0 | 281.2 | 24.25 | 59960.5 | 24.07 | 24.0623
0 | 0 | 280 | 48.23 | 70020 | 22.8 | 22.8567 | 0 | 0 | 280.1 | 24.39 | 60289 | 24.23 | 24.2012
0 | 0 | 278.9 | 48.44 | 70328.8 | 22.94 | 22.9563 | 0 | 0 | 278.8 | 24.46 | 60477.1 | 24.31 | 24.2707
0 | 0 | 277.7 | 48.98 | 71108.6 | 23.21 | 23.2122 | 0 | 0 | 277.7 | 24.68 | 61005.5 | 24.5 | 24.489
0 | 0 | 276.4 | 49.16 | 71375 | 23.26 | 23.2975 | 0 | 0 | 276.4 | 24.86 | 61469.1 | 24.72 | 24.6676
0 | 0 | 275.3 | 49.64 | 72070.7 | 23.49 | 23.525 | 0 | 0 | 275.2 | 25.02 | 61854 | 24.83 | 24.8264
0 | 0 | 274 | 49.98 | 72566.2 | 23.63 | 23.6861 | 0 | 0 | 274 | 25.22 | 62360 | 25.07 | 25.0248
0 | 0 | 272.8 | 50.31 | 73033.9 | 23.82 | 23.8425 | 0 | 0 | 272.8 | 25.32 | 62610.7 | 25.19 | 25.124
0 | 0 | 271.6 | 50.65 | 73536.6 | 23.96 | 24.0036 | 0 | 0 | 271.6 | 25.54 | 63143.9 | 25.37 | 25.3423
0 | 0 | 270.4 | 50.9 | 73890.7 | 24.13 | 24.1221 | 0 | 0 | 270.4 | 25.66 | 63445.9 | 25.55 | 25.4614
0 | 0 | 269.3 | 51.14 | 74248.6 | 24.23 | 24.2358 | 0 | 0 | 269.2 | 25.84 | 63885.9 | 25.73 | 25.64
0 | 0 | 268 | 51.65 | 74984.1 | 24.49 | 24.4775 | 0 | 0 | 268 | 26.05 | 64400.7 | 25.89 | 25.8484
0 | 0 | 266.8 | 52.03 | 75542 | 24.62 | 24.6576 | 0 | 0 | 266.8 | 26.3 | 65024.5 | 26.08 | 26.0964
0 | 0 | 265.6 | 52.37 | 76033.9 | 24.78 | 24.8187 | 0 | 0 | 265.6 | 26.42 | 65327.7 | 26.21 | 26.2155
0 | 0 | 264.4 | 52.79 | 76645.1 | 25 | 25.0178 | 0 | 0 | 264.4 | 26.53 | 65584.6 | 26.36 | 26.3247
0 | 0 | 263.2 | 53 | 76939.5 | 25.06 | 25.1173 | 0 | 0 | 263.2 | 26.75 | 66146.4 | 26.61 | 26.543
0 | 0 | 262 | 53.52 | 77703.5 | 25.32 | 25.3637 | 0 | 0 | 262 | 26.97 | 66670.8 | 26.79 | 26.7613
0 | 0 | 260.8 | 53.78 | 78075.6 | 25.45 | 25.4869 | 0 | 0 | 260.8 | 27.05 | 66881.8 | 26.88 | 26.8406
0 | 0 | 259.6 | 54.24 | 78747.8 | 25.68 | 25.7049 | 0 | 0 | 259.6 | 27.34 | 67602.7 | 27.15 | 27.1284
0 | 0 | 258.4 | 54.46 | 79062.3 | 25.79 | 25.8092 | 0 | 0 | 258.4 | 27.47 | 67924.8 | 27.26 | 27.2574
0 | 0 | 257.2 | 55.14 | 80046.8 | 26.1 | 26.1315 | 0 | 0 | 257.2 | 27.65 | 68368.3 | 27.45 | 27.436
0 | 0 | 256 | 55.33 | 80327.6 | 26.2 | 26.2215 | 0 | 0 | 256 | 27.85 | 68849.4 | 27.69 | 27.6345
0 | 0 | 254.8 | 55.71 | 80883.4 | 26.39 | 26.4016 | 0 | 0 | 254.8 | 28.04 | 69317.7 | 27.86 | 27.823
0 | 0 | 253.6 | 55.98 | 81278.2 | 26.51 | 26.5295 | 0 | 0 | 253.6 | 28.19 | 69703.5 | 28.04 | 27.9718
0 | 0 | 252.4 | 56.44 | 81936.9 | 26.71 | 26.7475 | 0 | 0 | 252.4 | 28.45 | 70348.8 | 28.25 | 28.2298
0 | 0 | 251.2 | 56.73 | 82362.8 | 26.89 | 26.885 | 0 | 0 | 251.2 | 28.64 | 70815.2 | 28.46 | 28.4183
0 | 0 | 250 | 57.15 | 82969.9 | 27.04 | 27.084 | 0 | 0 | 250 | 28.82 | 71244.8 | 28.62 | 28.5969
0 | 0 | 248.8 | 57.52 | 83512.4 | 27.27 | 27.2594 | 0 | 0 | 248.8 | 29.06 | 71852.1 | 28.86 | 28.8351
0 | 0 | 247.6 | 57.94 | 84120.8 | 27.42 | 27.4584 | 0 | 0 | 247.6 | 29.19 | 72175.6 | 29.03 | 28.9641
0 | 0 | 246.4 | 58.5 | 84923.6 | 27.71 | 27.7238 | 0 | 0 | 246.4 | 29.43 | 72757.2 | 29.3 | 29.2022
0 | 0 | 245.2 | 58.81 | 85385.6 | 27.81 | 27.8707 | 0 | 0 | 245.1 | 29.65 | 73312.1 | 29.45 | 29.4205
0 | 0 | 243.9 | 59.29 | 86073.5 | 28.05 | 28.0982 | 0 | 0 | 244 | 29.9 | 73917.3 | 29.71 | 29.6686
0 | 0 | 242.8 | 59.63 | 86570.4 | 28.17 | 28.2593 | 0 | 0 | 242.7 | 30.03 | 74247.9 | 29.82 | 29.7976
0 | 0 | 241.6 | 60.05 | 87184 | 28.41 | 28.4584 | 0 | 0 | 241.6 | 30.24 | 74771.8 | 30.06 | 30.006
0 | 0 | 240.3 | 60.33 | 87588.9 | 28.55 | 28.5911 | 0 | 0 | 240.3 | 30.32 | 74956.5 | 30.08 | 30.0853
0 | 0 | 239.2 | 60.86 | 88352.7 | 28.78 | 28.8422 | 0 | 0 | 239.1 | 30.57 | 75581.7 | 30.35 | 30.3334
0 | 0 | 237.9 | 61.24 | 88913.9 | 28.99 | 29.0223 | 0 | 0 | 237.9 | 30.83 | 76219.8 | 30.63 | 30.5914
0 | 0 | 236.7 | 61.9 | 89863.1 | 29.28 | 29.3351 | 0 | 0 | 236.7 | 31.01 | 76674 | 30.79 | 30.77
0 | 0 | 235.5 | 62.27 | 90396.4 | 29.45 | 29.5104 | 0 | 0 | 235.5 | 31.17 | 77051.9 | 30.99 | 30.9288
0 | 0 | 234.3 | 62.56 | 90827.4 | 29.63 | 29.6479 | 0 | 0 | 234.3 | 31.43 | 77707.8 | 31.23 | 31.1867
0 | 0 | 233.1 | 63.08 | 91578.1 | 29.84 | 29.8943 | 0 | 0 | 233.1 | 31.52 | 77936.1 | 31.35 | 31.276
0 | 0 | 231.9 | 63.52 | 92217.3 | 30.05 | 30.1028 | 0 | 0 | 231.9 | 31.82 | 78671.4 | 31.59 | 31.5737
0 | 0 | 230.7 | 63.91 | 92783.9 | 30.24 | 30.2877 | 0 | 0 | 230.7 | 32.12 | 79421.2 | 31.9 | 31.8714
0 | 0 | 229.5 | 64.42 | 93528.9 | 30.49 | 30.5294 | 0 | 0 | 229.5 | 32.33 | 79929.1 | 32.1 | 32.0798
0 | 0 | 228.3 | 64.79 | 94056 | 30.68 | 30.7047 | 0 | 0 | 228.3 | 32.5 | 80362.2 | 32.32 | 32.2485
0 | 0 | 227.1 | 65.3 | 94800.1 | 30.89 | 30.9464 | 0 | 0 | 227.1 | 32.75 | 80967.7 | 32.54 | 32.4965
0 | 0 | 225.9 | 65.74 | 95444.9 | 31.1 | 31.1549 | 0 | 0 | 225.9 | 32.97 | 81509.4 | 32.75 | 32.7148
0 | 0 | 224.7 | 66.24 | 96168.1 | 31.36 | 31.3919 | 0 | 0 | 224.7 | 33.21 | 82105.1 | 32.97 | 32.953
0 | 0 | 223.5 | 66.65 | 96757.7 | 31.53 | 31.5862 | 0 | 0 | 223.5 | 33.45 | 82702.7 | 33.24 | 33.1911
0 | 0 | 222.2 | 67.16 | 97496.5 | 31.76 | 31.8279 | 0 | 0 | 222.2 | 33.64 | 83165.5 | 33.38 | 33.3796
0 | 0 | 221.1 | 67.58 | 98105.6 | 31.97 | 32.0269 | 0 | 0 | 221 | 33.91 | 83846.6 | 33.69 | 33.6475
0 | 0 | 219.9 | 68.11 | 98884.4 | 32.28 | 32.2781 | 0 | 0 | 219.9 | 34.1 | 84313.7 | 33.85 | 33.8361
0 | 0 | 218.7 | 68.62 | 99615.2 | 32.52 | 32.5198 | 0 | 0 | 218.6 | 34.31 | 84814.5 | 34.1 | 34.0445
0 | 0 | 217.5 | 69.19 | 100445.6 | 32.67 | 32.7899 | 0 | 0 | 217.5 | 34.65 | 85665 | 34.4 | 34.3818
0 | 0 | 216.3 | 69.66 | 101128.2 | 32.97 | 33.0127 | 0 | 0 | 216.2 | 34.78 | 85995 | 34.58 | 34.5108
0 | 0 | 215.1 | 69.97 | 101585.8 | 33.11 | 33.1596 | 0 | 0 | 215 | 35.02 | 86580.8 | 34.76 | 34.749
0 | 0 | 213.8 | 70.46 | 102297.9 | 33.38 | 33.3918 | 0 | 0 | 213.8 | 35.24 | 87132.5 | 34.98 | 34.9673
0 | 0 | 212.6 | 71.02 | 103101.1 | 33.65 | 33.6572 | 0 | 0 | 212.6 | 35.45 | 87656.8 | 35.26 | 35.1756
0 | 0 | 211.4 | 71.58 | 103912.4 | 33.85 | 33.9226 | 0 | 0 | 211.4 | 35.79 | 88497.1 | 35.59 | 35.513
0 | 0 | 210.2 | 71.97 | 104491.1 | 34.07 | 34.1074 | 0 | 0 | 210.2 | 35.96 | 88911.9 | 35.76 | 35.6817
0 | 0 | 209 | 72.39 | 105093.4 | 34.29 | 34.3064 | 0 | 0 | 209 | 36.22 | 89556.4 | 35.98 | 35.9397
0 | 0 | 207.8 | 72.96 | 105922.5 | 34.53 | 34.5766 | 0 | 0 | 207.8 | 36.44 | 90090.1 | 36.21 | 36.158
0 | 0 | 206.6 | 73.64 | 106905.1 | 34.86 | 34.8988 | 0 | 0 | 206.6 | 36.83 | 91054.7 | 36.49 | 36.5449
0 | 0 | 205.4 | 73.87 | 107249.3 | 34.99 | 35.0078 | 0 | 0 | 205.4 | 37 | 91482 | 36.81 | 36.7136
0 | 0 | 204.2 | 74.43 | 108059.6 | 35.24 | 35.2732 | 0 | 0 | 204.2 | 37.19 | 91958.6 | 36.95 | 36.9022
0 | 0 | 203 | 74.9 | 108736.9 | 35.43 | 35.4959 | 0 | 0 | 203 | 37.47 | 92627.4 | 37.21 | 37.18
0 | 0 | 201.8 | 75.46 | 109550 | 35.71 | 35.7613 | 0 | 0 | 201.8 | 37.77 | 93378.5 | 37.5 | 37.4777
0 | 0 | 200.6 | 76.15 | 110555.2 | 36.04 | 36.0883 | 0 | 0 | 200.6 | 38.03 | 94030.3 | 37.81 | 37.7357
0 | 0 | 199.4 | 76.48 | 111031.3 | 36.25 | 36.2447 | 0 | 0 | 199.4 | 38.27 | 94618.1 | 38.03 | 37.9738
0 | 0 | 198.2 | 77.13 | 111973.8 | 36.47 | 36.5528 | 0 | 0 | 198.2 | 38.55 | 95320.8 | 38.28 | 38.2516
0 | 0 | 197 | 77.76 | 112886.5 | 36.79 | 36.8513 | 0 | 0 | 197 | 38.79 | 95913.9 | 38.59 | 38.4898
0 | 0 | 195.8 | 78.13 | 113422 | 36.99 | 37.0267 | 0 | 0 | 195.8 | 39.1 | 96666.5 | 38.8 | 38.7974
0 | 0 | 194.6 | 78.85 | 114471.4 | 37.29 | 37.3679 | 0 | 0 | 194.6 | 39.25 | 97033 | 38.94 | 38.9462
0 | 0 | 193.4 | 79.19 | 114967 | 37.47 | 37.529 | 0 | 0 | 193.4 | 39.64 | 98008.2 | 39.35 | 39.3332
0 | 0 | 192.2 | 79.91 | 116013.3 | 37.84 | 37.8702 | 0 | 0 | 192.2 | 39.93 | 98723.7 | 39.68 | 39.621
0 | 0 | 191 | 80.4 | 116730.1 | 38.06 | 38.1025 | 0 | 0 | 191 | 40.15 | 99276.6 | 39.92 | 39.8393
0 | 0 | 189.8 | 81.12 | 117762.9 | 38.43 | 38.4437 | 0 | 0 | 189.8 | 40.52 | 100176.2 | 40.22 | 40.2064
0 | 0 | 188.6 | 81.55 | 118392.4 | 38.6 | 38.6475 | 0 | 0 | 188.5 | 40.73 | 100696.1 | 40.48 | 40.4148
0 | 0 | 187.3 | 82.14 | 119245 | 38.87 | 38.9271 | 0 | 0 | 187.4 | 41.03 | 101429 | 40.76 | 40.7124
0 | 0 | 186.2 | 82.65 | 119997 | 39.15 | 39.1688 | 0 | 0 | 186.1 | 41.28 | 102046.6 | 40.98 | 40.9605
0 | 0 | 185 | 83.54 | 121275.6 | 39.54 | 39.5905 | 0 | 0 | 184.9 | 41.5 | 102609.4 | 41.24 | 41.1788
0 | 0 | 183.7 | 83.97 | 121910 | 39.74 | 39.7943 | 0 | 0 | 183.7 | 41.87 | 103520.2 | 41.53 | 41.5459
0 | 0 | 182.6 | 84.58 | 122791.7 | 40 | 40.0834 | 0 | 0 | 182.5 | 42.14 | 104174.2 | 41.87 | 41.8139
0 | 0 | 181.3 | 85.21 | 123700.2 | 40.31 | 40.382 | 0 | 0 | 181.3 | 42.38 | 104788.9 | 42.1 | 42.052
0 | 0 | 180.1 | 85.61 | 124289.8 | 40.55 | 40.5715 | 0 | 0 | 180.1 | 42.66 | 105475.2 | 42.38 | 42.3298
0 | 0 | 178.9 | 86.41 | 125442.4 | 40.87 | 40.9507 | 0 | 0 | 178.9 | 42.95 | 106175.7 | 42.69 | 42.6176
0 | 0 | 177.7 | 86.87 | 126115 | 41.12 | 41.1687 | 0 | 0 | 177.7 | 43.3 | 107043.9 | 43.02 | 42.9649
0 | 0 | 176.5 | 87.5 | 127037 | 41.44 | 41.4672 | 0 | 0 | 176.5 | 43.56 | 107694.7 | 43.3 | 43.2229
0 | 0 | 175.3 | 88.04 | 127817.8 | 41.69 | 41.7231 | 0 | 0 | 175.3 | 43.95 | 108666.8 | 43.64 | 43.6098
0 | 0 | 174.1 | 88.73 | 128810 | 41.97 | 42.0501 | 0 | 0 | 174.1 | 44.18 | 109239.1 | 43.86 | 43.8381
0 | 0 | 172.9 | 89.35 | 129717.3 | 42.26 | 42.344 | 0 | 0 | 172.9 | 44.42 | 109830.9 | 44.13 | 44.0762
0 | 0 | 171.7 | 89.99 | 130641.5 | 42.6 | 42.6473 | 0 | 0 | 171.7 | 44.82 | 110804.1 | 44.5 | 44.4731
0 | 0 | 170.5 | 90.63 | 131577.5 | 42.89 | 42.9506 | 0 | 0 | 170.5 | 45.09 | 111469.7 | 44.85 | 44.741
0 | 0 | 169.3 | 91.19 | 132389.4 | 43.16 | 43.216 | 0 | 0 | 169.3 | 45.38 | 112194.3 | 45.07 | 45.0288
0 | 0 | 168.1 | 92.2 | 133855.7 | 43.57 | 43.6946 | 0 | 0 | 168.1 | 45.79 | 113207.6 | 45.41 | 45.4356
0 | 0 | 166.9 | 92.6 | 134437.3 | 43.73 | 43.8842 | 0 | 0 | 166.9 | 46.02 | 113773.1 | 45.71 | 45.6638
0 | 0 | 165.7 | 93.42 | 135625.2 | 44.24 | 44.2728 | 0 | 0 | 165.7 | 46.26 | 114381.3 | 45.98 | 45.902
0 | 0 | 164.5 | 93.95 | 136388.9 | 44.43 | 44.524 | 0 | 0 | 164.4 | 46.66 | 115368.9 | 46.4 | 46.2989
0 | 0 | 163.3 | 94.77 | 137581.2 | 44.88 | 44.9126 | 0 | 0 | 163.3 | 46.97 | 116120.2 | 46.7 | 46.6065
0 | 0 | 162.1 | 95.21 | 138223.7 | 45.05 | 45.1211 | 0 | 0 | 162 | 47.22 | 116749.8 | 46.99 | 46.8545
0 | 0 | 160.9 | 95.85 | 139159.6 | 45.39 | 45.4244 | 0 | 0 | 160.9 | 47.61 | 117715.4 | 47.29 | 47.2415
0 | 0 | 159.6 | 96.73 | 140425.7 | 45.78 | 45.8414 | 0 | 0 | 159.6 | 47.99 | 118644.5 | 47.77 | 47.6186
0 | 0 | 158.5 | 97.38 | 141371.8 | 46.08 | 46.1495 | 0 | 0 | 158.4 | 48.28 | 119363.4 | 47.94 | 47.9063
0 | 0 | 157.2 | 98 | 142276.3 | 46.36 | 46.4433 | 0 | 0 | 157.2 | 48.55 | 120029.6 | 48.28 | 48.1742
0 | 0 | 156.1 | 98.78 | 143410.4 | 46.7 | 46.8129 | 0 | 0 | 156 | 48.92 | 120939.2 | 48.58 | 48.5414
0 | 0 | 154.9 | 99.37 | 144262.8 | 47.02 | 47.0926 | 0 | 0 | 154.8 | 49.25 | 121772.5 | 48.95 | 48.8688
0 | 0 | 153.6 | 99.96 | 145122.4 | 47.27 | 47.3722 | 0 | 0 | 153.6 | 49.62 | 122682.4 | 49.35 | 49.236
0 | 0 | 152.4 | 100.87 | 146434.9 | 47.66 | 47.8034 | 0 | 0 | 152.4 | 50.03 | 123703.9 | 49.6 | 49.6428
0 | 0 | 151.2 | 101.37 | 147170.3 | 48.03 | 48.0404 | 0 | 0 | 151.2 | 50.22 | 124168.6 | 49.91 | 49.8313
0 | 0 | 150 | 102.22 | 148407.1 | 48.37 | 48.4432 | 0 | 0 | 150 | 50.46 | 124759.1 | 50.18 | 50.0695
0 | 0 | 148.8 | 103.03 | 149578.5 | 48.78 | 48.8271 | 0 | 0 | 148.8 | 50.87 | 125765.6 | 50.56 | 50.4763
0 | 0 | 147.6 | 103.64 | 150456.1 | 49.11 | 49.1162 | 0 | 0 | 147.6 | 51.25 | 126705.6 | 50.86 | 50.8533
0 | 0 | 146.4 | 104.43 | 151605.4 | 49.39 | 49.4905 | 0 | 0 | 146.4 | 51.59 | 127539 | 51.15 | 51.1907
0 | 0 | 145.2 | 104.9 | 152296.6 | 49.66 | 49.7133 | 0 | 0 | 145.2 | 51.97 | 128493.7 | 51.62 | 51.5678
0 | 0 | 144 | 105.89 | 153736.7 | 50.08 | 50.1825 | 0 | 0 | 144 | 52.33 | 129378.5 | 52.01 | 51.925
0 | 0 | 142.8 | 106.73 | 154952.4 | 50.52 | 50.5805 | 0 | 0 | 142.8 | 52.72 | 130336.2 | 52.47 | 52.312
0 | 0 | 141.6 | 107.3 | 155774.4 | 50.86 | 50.8507 | 0 | 0 | 141.6 | 53.01 | 131055.7 | 52.71 | 52.5997
0 | 0 | 140.4 | 108 | 156788.6 | 51.07 | 51.1824 | 0 | 0 | 140.4 | 53.4 | 132017.8 | 52.97 | 52.9867
0 | 0 | 139.2 | 108.96 | 158189.7 | 51.52 | 51.6374 | 0 | 0 | 139.2 | 53.66 | 132656.7 | 53.35 | 53.2447
0 | 0 | 138 | 109.71 | 159277.8 | 51.92 | 51.9928 | 0 | 0 | 138 | 54.05 | 133641 | 53.65 | 53.6317
0 | 0 | 136.8 | 110.21 | 160006.3 | 52.18 | 52.2298 | 0 | 0 | 136.8 | 54.43 | 134578.9 | 54.14 | 54.0087
0 | 0 | 135.6 | 111.1 | 161298.8 | 52.61 | 52.6515 | 0 | 0 | 135.6 | 54.75 | 135369.7 | 54.45 | 54.3263
0 | 0 | 134.4 | 112 | 162592.8 | 53 | 53.0781 | 0 | 0 | 134.4 | 55.14 | 136318.4 | 54.76 | 54.7132
0 | 0 | 133.2 | 112.76 | 163702.6 | 53.4 | 53.4382 | 0 | 0 | 133.2 | 55.49 | 137188.5 | 55.1 | 55.0605
0 | 0 | 131.9 | 113.45 | 164704.9 | 53.64 | 53.7652 | 0 | 0 | 131.9 | 55.85 | 138072 | 55.48 | 55.4177
0 | 0 | 130.8 | 114.19 | 165784.7 | 54.1 | 54.1159 | 0 | 0 | 130.8 | 56.25 | 139075.6 | 55.89 | 55.8146
0 | 0 | 129.6 | 114.97 | 166912.9 | 54.39 | 54.4856 | 0 | 0 | 129.6 | 56.71 | 140213 | 56.4 | 56.2711
0 | 0 | 128.4 | 116.13 | 168596.7 | 54.93 | 55.0353 | 0 | 0 | 128.3 | 57.02 | 140961.3 | 56.57 | 56.5787
0 | 0 | 127.1 | 116.63 | 169317.3 | 55.21 | 55.2723 | 0 | 0 | 127.1 | 57.39 | 141899.8 | 56.96 | 56.9458
0 | 0 | 125.9 | 117.56 | 170670.3 | 55.64 | 55.713 | 0 | 0 | 126 | 57.87 | 143076.8 | 57.42 | 57.4221
0 | 0 | 124.7 | 118.39 | 171874.6 | 56.07 | 56.1063 | 0 | 0 | 124.7 | 58.23 | 143964.5 | 57.84 | 57.7793
0 | 0 | 123.5 | 119.19 | 173037 | 56.42 | 56.4855 | 0 | 0 | 123.5 | 58.57 | 144802.3 | 58.19 | 58.1167
0 | 0 | 122.4 | 120.05 | 174291.1 | 56.78 | 56.893 | 0 | 0 | 122.3 | 59.13 | 146193.7 | 58.73 | 58.6724
0 | 0 | 121.1 | 120.69 | 175212.6 | 57.15 | 57.1963 | 0 | 0 | 121.1 | 59.48 | 147066 | 59.09 | 59.0196
0 | 0 | 119.9 | 121.69 | 176662.6 | 57.61 | 57.6703 | 0 | 0 | 119.9 | 59.85 | 147964.3 | 59.48 | 59.3868
0 | 0 | 118.7 | 122.55 | 177919.5 | 58.02 | 58.0778 | 0 | 0 | 118.7 | 60.24 | 148934.6 | 59.89 | 59.7738
0 | 0 | 117.5 | 123.45 | 179221.3 | 58.43 | 58.5043 | 0 | 0 | 117.5 | 60.6 | 149836.2 | 60.29 | 60.131
0 | 0 | 116.3 | 124.19 | 180302.6 | 58.73 | 58.855 | 0 | 0 | 116.3 | 60.88 | 150522.4 | 60.55 | 60.4088
0 | 0 | 115.1 | 125.06 | 181561.8 | 59.12 | 59.2673 | 0 | 0 | 115.1 | 61.27 | 151475.4 | 60.89 | 60.7958
0 | 0 | 113.9 | 126.04 | 182977.2 | 59.6 | 59.7318 | 0 | 0 | 113.9 | 61.79 | 152758.1 | 61.33 | 61.3118
0 | 0 | 112.7 | 126.83 | 184134.5 | 60.01 | 60.1062 | 0 | 0 | 112.7 | 62.09 | 153506.7 | 61.68 | 61.6094
0 | 0 | 111.5 | 127.63 | 185297.5 | 60.39 | 60.4853 | 0 | 0 | 111.5 | 62.54 | 154614.8 | 62.13 | 62.056
0 | 0 | 110.3 | 128.57 | 186649.5 | 60.84 | 60.9308 | 0 | 0 | 110.3 | 63.01 | 155792.1 | 62.56 | 62.5223
0 | 0 | 109.1 | 129.4 | 187862.6 | 61.2 | 61.3241 | 0 | 0 | 109.1 | 63.51 | 157019.6 | 63 | 63.0185
0 | 0 | 107.9 | 130.6 | 189600 | 61.72 | 61.8928 | 0 | 0 | 107.9 | 64.02 | 158289.7 | 63.56 | 63.5245
0 | 0 | 106.7 | 131.27 | 190574.2 | 62.14 | 62.2103 | 0 | 0 | 106.7 | 64.51 | 159496.6 | 63.84 | 64.0107
0 | 0 | 105.5 | 132.19 | 191910.1 | 62.59 | 62.6463 | 0 | 0 | 105.5 | 64.81 | 160233.8 | 64.36 | 64.3084
0 | 0 | 104.3 | 133.44 | 193728.2 | 63.13 | 63.2387 | 0 | 0 | 104.2 | 65.18 | 161158.2 | 64.88 | 64.6755
0 | 0 | 103.1 | 133.98 | 194517 | 63.45 | 63.4946 | 0 | 0 | 103.1 | 65.66 | 162345.9 | 65.19 | 65.1518
0 | 0 | 101.9 | 135.14 | 196198.8 | 63.94 | 64.0444 | 0 | 0 | 101.9 | 66.12 | 163470.3 | 65.63 | 65.6083
0 | 0 | 100.7 | 135.78 | 197127.1 | 64.26 | 64.3477 | 0 | 0 | 100.7 | 66.56 | 164553.7 | 66.23 | 66.0449
0 | 0 | 99.5 | 136.9 | 198750.5 | 64.77 | 64.8784 | 0 | 0 | 99.5 | 66.8 | 165151.1 | 66.43 | 66.283
0 | 0 | 98.3 | 137.76 | 199994.9 | 65.16 | 65.286 | 0 | 0 | 98.3 | 67.23 | 166216.7 | 66.84 | 66.7097
0 | 0 | 97.1 | 138.48 | 201046.4 | 65.61 | 65.6272 | 0 | 0 | 97.1 | 67.74 | 167469.8 | 67.24 | 67.2157
0 | 0 | 95.9 | 139.87 | 203059.4 | 66.19 | 66.286 | 0 | 0 | 95.8 | 68.06 | 168266.5 | 67.73 | 67.5332
0 | 0 | 94.7 | 140.69 | 204247.7 | 66.6 | 66.6746 | 0 | 0 | 94.7 | 68.39 | 169076.6 | 68.15 | 67.8607
0 | 0 | 93.5 | 141.54 | 205488.1 | 66.96 | 67.0774 | 0 | 0 | 93.5 | 68.97 | 170521.4 | 68.6 | 68.4362
0 | 0 | 92.3 | 142.54 | 206939.6 | 67.5 | 67.5513 | 0 | 0 | 92.2 | 69.61 | 172091.4 | 68.98 | 69.0712
0 | 0 | 91.1 | 143.48 | 208300.9 | 67.93 | 67.9968 | 0 | 0 | 91.1 | 69.94 | 172923.4 | 69.54 | 69.3987
0 | 0 | 89.9 | 144.3 | 209495.2 | 68.28 | 68.3854 | 0 | 0 | 89.8 | 70.5 | 174294.8 | 70.01 | 69.9544
0 | 0 | 88.6 | 145.4 | 211096 | 68.79 | 68.9067 | 0 | 0 | 88.6 | 71.05 | 175650.2 | 70.44 | 70.5001
0 | 0 | 87.4 | 146.51 | 212696.4 | 69.41 | 69.4327 | 0 | 0 | 87.4 | 71.54 | 176880.2 | 71.06 | 70.9863
0 | 0 | 86.3 | 147.71 | 214436.6 | 69.86 | 70.0014 | 0 | 0 | 86.2 | 71.94 | 177867.6 | 71.55 | 71.3832
0 | 0 | 85 | 148.46 | 215534.5 | 70.32 | 70.3569 | 0 | 0 | 85 | 72.45 | 179111.9 | 71.91 | 71.8893
0 | 0 | 83.8 | 149.56 | 217121.9 | 70.85 | 70.8782 | 0 | 0 | 83.8 | 72.91 | 180252.2 | 72.35 | 72.3457
0 | 0 | 82.6 | 150.52 | 218519.1 | 71.24 | 71.3331 | 0 | 0 | 82.6 | 73.37 | 181404 | 72.9 | 72.8021
0 | 0 | 81.4 | 151.42 | 219823.4 | 71.69 | 71.7596 | 0 | 0 | 81.4 | 73.82 | 182515.6 | 73.36 | 73.2487
0 | 0 | 80.2 | 152.63 | 221585.6 | 72.21 | 72.3331 | 0 | 0 | 80.2 | 74.32 | 183742 | 73.79 | 73.7448
0 | 0 | 79 | 153.67 | 223092.7 | 72.71 | 72.8259 | 0 | 0 | 79 | 74.68 | 184629.3 | 74.18 | 74.102
0 | 0 | 77.8 | 154.55 | 224378.8 | 73.17 | 73.243 | 0 | 0 | 77.8 | 75.26 | 186064.7 | 74.7 | 74.6775
0 | 0 | 76.6 | 155.71 | 226062.6 | 73.67 | 73.7927 | 0 | 0 | 76.6 | 75.69 | 187144.5 | 75.09 | 75.1042
0 | 0 | 75.3 | 156.64 | 227413.5 | 74.17 | 74.2334 | 0 | 0 | 75.4 | 75.94 | 187756.6 | 75.41 | 75.3523
0 | 0 | 74.2 | 157.7 | 228940.5 | 74.55 | 74.7358 | 0 | 0 | 74.1 | 76.42 | 188938 | 75.96 | 75.8285
0 | 0 | 72.9 | 158.65 | 230321.4 | 75.09 | 75.186 | 0 | 0 | 73 | 76.9 | 190125.5 | 76.55 | 76.3048
0 | 0 | 71.7 | 160 | 232281.4 | 75.64 | 75.8258 | 0 | 0 | 71.7 | 77.54 | 191710 | 77.05 | 76.9399
0 | 0 | 70.5 | 160.96 | 233682.5 | 76.12 | 76.2807 | 0 | 0 | 70.5 | 78.1 | 193090.9 | 77.55 | 77.4955
0 | 0 | 69.3 | 161.96 | 235138.2 | 76.7 | 76.7547 | 0 | 0 | 69.3 | 78.72 | 194624 | 77.98 | 78.1107
0 | 0 | 68.1 | 163.17 | 236881.3 | 77.22 | 77.3281 | 0 | 0 | 68.1 | 79.04 | 195410 | 78.39 | 78.4283
0 | 0 | 66.9 | 164.03 | 238136.4 | 77.67 | 77.7357 | 0 | 0 | 66.9 | 79.59 | 196786.3 | 79.02 | 78.974
0 | 0 | 65.7 | 165.11 | 239708.1 | 78.13 | 78.2475 | 0 | 0 | 65.7 | 80.07 | 197953.7 | 79.48 | 79.4503
0 | 0 | 64.5 | 166.33 | 241470.9 | 78.73 | 78.8256 | 0 | 0 | 64.5 | 80.58 | 199215.5 | 79.89 | 79.9563
0 | 0 | 63.3 | 167.25 | 242814.6 | 79.13 | 79.2616 | 0 | 0 | 63.3 | 81.05 | 200396.4 | 80.58 | 80.4227
0 | 0 | 62.1 | 168.4 | 244486.1 | 79.68 | 79.8066 | 0 | 0 | 62 | 81.7 | 201999.9 | 81.11 | 81.0677
0 | 0 | 60.9 | 169.44 | 245987.7 | 80.24 | 80.2995 | 0 | 0 | 60.9 | 82.02 | 202779.4 | 81.45 | 81.3852
0 | 0 | 59.7 | 170.46 | 247470.6 | 80.65 | 80.7829 | 0 | 0 | 59.6 | 82.47 | 203895.1 | 82.01 | 81.8317
0 | 0 | 58.5 | 171.63 | 249173.1 | 81.23 | 81.3374 | 0 | 0 | 58.5 | 83.17 | 205630.1 | 82.51 | 82.5263
0 | 0 | 57.3 | 172.44 | 250351 | 81.52 | 81.7212 | 0 | 0 | 57.2 | 83.78 | 207122.8 | 83.14 | 83.1316
0 | 0 | 56.1 | 174.18 | 252874.2 | 82.43 | 82.5459 | 0 | 0 | 56 | 84.05 | 207801.8 | 83.42 | 83.3995
0 | 0 | 54.9 | 174.94 | 253973.2 | 82.76 | 82.906 | 0 | 0 | 54.8 | 84.53 | 208993.7 | 83.93 | 83.8758
0 | 0 | 53.6 | 175.94 | 255420 | 83.25 | 83.3799 | 0 | 0 | 53.6 | 84.95 | 210038.5 | 84.43 | 84.2925
0 | 0 | 52.4 | 177.13 | 257148.5 | 83.8 | 83.9439 | 0 | 0 | 52.4 | 85.63 | 211702.6 | 85.05 | 84.9673
0 | 0 | 51.2 | 178.72 | 259464.8 | 84.56 | 84.6974 | 0 | 0 | 51.2 | 86.21 | 213142.3 | 85.4 | 85.5428
0 | 0 | 50 | 179.63 | 260778.7 | 85.05 | 85.1287 | 0 | 0 | 50 | 86.73 | 214435.1 | 86.13 | 86.0587
0 | 0 | 48.8 | 180.58 | 262169.2 | 85.49 | 85.5789 | 0 | 0 | 48.8 | 87.06 | 215249.2 | 86.55 | 86.3862
0 | 0 | 47.6 | 181.75 | 263856.9 | 86.08 | 86.1334 | 0 | 0 | 47.6 | 87.57 | 216492.8 | 86.98 | 86.8922
0 | 0 | 46.4 | 182.97 | 265633.1 | 86.67 | 86.7115 | 0 | 0 | 46.4 | 88.17 | 217993.2 | 87.38 | 87.4876
0 | 0 | 45.2 | 184.2 | 267422.8 | 87.16 | 87.2944 | 0 | 0 | 45.2 | 88.51 | 218825 | 87.97 | 87.825
0 | 0 | 44 | 185.22 | 268899.1 | 87.65 | 87.7778 | 0 | 0 | 44 | 89.08 | 220237.2 | 88.49 | 88.3906
0 | 0 | 42.8 | 186.31 | 270483.9 | 88.15 | 88.2944 | 0 | 0 | 42.8 | 89.42 | 221077.5 | 88.88 | 88.7279
0 | 0 | 41.6 | 187.34 | 271983.5 | 88.65 | 88.7825 | 0 | 0 | 41.6 | 89.98 | 222465.5 | 89.46 | 89.2836
0 | 0 | 40.3 | 188.54 | 273723.3 | 89.2 | 89.3512 | 0 | 0 | 40.4 | 90.66 | 224139.6 | 89.96 | 89.9583
0 | 0 | 39.2 | 189.56 | 275203.6 | 89.67 | 89.8346 | 0 | 0 | 39.1 | 91.34 | 225820.5 | 90.54 | 90.6331
0 | 0 | 37.9 | 190.69 | 276833.9 | 90.28 | 90.3701 | 0 | 0 | 38 | 91.63 | 226542.2 | 90.99 | 90.9208
0 | 0 | 36.7 | 192.02 | 278771.3 | 90.82 | 91.0004 | 0 | 0 | 36.7 | 92.09 | 227684.5 | 91.53 | 91.3773
0 | 0 | 35.5 | 193.21 | 280501.2 | 91.48 | 91.5644 | 0 | 0 | 35.5 | 92.62 | 228982 | 91.93 | 91.9032
0 | 0 | 34.3 | 194.22 | 281964.1 | 91.99 | 92.043 | 0 | 0 | 34.3 | 93.2 | 230429.1 | 92.41 | 92.4787
0 | 0 | 33.1 | 195.52 | 283853.9 | 92.4 | 92.6591 | 0 | 0 | 33.1 | 93.86 | 232066.6 | 93 | 93.1336
0 | 0 | 31.9 | 196.68 | 285535.1 | 93 | 93.2089 | 0 | 0 | 31.9 | 94.45 | 233507.6 | 93.77 | 93.719
0 | 0 | 30.7 | 197.81 | 287179.5 | 93.69 | 93.7444 | 0 | 0 | 30.7 | 94.91 | 234651 | 94.29 | 94.1754
0 | 0 | 29.5 | 198.9 | 288756 | 94.15 | 94.2609 | 0 | 0 | 29.5 | 95.33 | 235691.5 | 94.68 | 94.5922
0 | 0 | 28.3 | 199.64 | 289841.2 | 94.55 | 94.6116 | 0 | 0 | 28.3 | 95.98 | 237296.9 | 95.29 | 95.2372
0 | 0 | 27.1 | 201.44 | 292445 | 95.41 | 95.4647 | 0 | 0 | 27.1 | 96.25 | 237958.1 | 95.63 | 95.5051
0 | 0 | 25.9 | 202.17 | 293514.1 | 95.75 | 95.8106 | 0 | 0 | 25.9 | 96.81 | 239336 | 96.15 | 96.0607
0 | 0 | 24.7 | 203.41 | 295313.4 | 96.29 | 96.3983 | 0 | 0 | 24.7 | 97.32 | 240618.6 | 96.6 | 96.5668
0 | 0 | 23.5 | 204.45 | 296823.8 | 96.85 | 96.8911 | 0 | 0 | 23.5 | 98.01 | 242304.1 | 97.24 | 97.2514
0 | 0 | 22.3 | 205.74 | 298691.9 | 97.31 | 97.5025 | 0 | 0 | 22.3 | 98.33 | 243106.8 | 97.82 | 97.569
0 | 0 | 21 | 206.79 | 300215.6 | 97.85 | 98.0001 | 0 | 0 | 21.1 | 99 | 244773.2 | 98.39 | 98.2338
0 | 0 | 19.9 | 207.79 | 301664.2 | 98.26 | 98.474 | 0 | 0 | 19.8 | 99.28 | 245457.3 | 98.79 | 98.5116
0 | 0 | 18.6 | 208.8 | 303133.4 | 98.88 | 98.9527 | 0 | 0 | 18.6 | 99.83 | 246816.3 | 99.35 | 99.0574
0 | 0 | 17.4 | 209.89 | 304709.3 | 99.34 | 99.4692 | 0 | 0 | 17.4 | 99.99 | 247202.3 | 99.27 | 99.2161
0 | 0 | 16.2 | 210.51 | 305608.9 | 99.58 | 99.763 | 0 | 0 | 16.2 | 100.74 | 249072.3 | 99.76 | 99.9603
0 | 0 | 15 | 210.9 | 306187.9 | 99.85 | 99.9479 | 0 | 0 | 15 | 100.58 | 248663 | 99.79 | 99.8015
0 | 0 | 13.8 | 211.01 | 306344.7 | 100 | 100 | 0 | 0 | 13.8 | 100.78 | 249172.3 | 99.99 | 100
0 | 0 | 12.6 | 210.92 | 306206.4 | 99.81 | 99.9573 | 0 | 0 | 12.6 | 100.54 | 248559 | 100 | 99.7619
0 | 0 | 11.4 | 210.2 | 305171.9 | 99.46 | 99.6161 | 0 | 0 | 11.4 | 100.04 | 247344.5 | 99.33 | 99.2657
0 | 0 | 10.2 | 208.41 | 302561.8 | 98.69 | 98.7678 | 0 | 0 | 10.2 | 99.02 | 244807.6 | 98.17 | 98.2536
0 | 0 | 9 | 205.74 | 298686.1 | 97.41 | 97.5025 | 0 | 0 | 9 | 97.79 | 241773.5 | 97.12 | 97.0331
0 | 0 | 7.8 | 201.86 | 293052.4 | 95.55 | 95.6637 | 0 | 0 | 7.8 | 95.33 | 235679.2 | 94.65 | 94.5922
0 | 0 | 6.6 | 195.38 | 283645.2 | 92.43 | 92.5928 | 0 | 0 | 6.6 | 91.96 | 227360.3 | 91.27 | 91.2483
0 | 0 | 5.4 | 186.4 | 270615.9 | 88.22 | 88.337 | 0 | 0 | 5.4 | 86.89 | 214830.9 | 86.3 | 86.2175
0 | 0 | 4.2 | 173.39 | 251721.8 | 82.05 | 82.1715 | 0 | 0 | 4.2 | 79.72 | 197103.6 | 79.19 | 79.103
0 | 0 | 3 | 155.3 | 225461.4 | 73.43 | 73.5984 | 0 | 0 | 3 | 69.42 | 171637.6 | 68.95 | 68.8827
0 | 0 | 1.8 | 131.87 | 191442.4 | 62.41 | 62.4947 | 0 | 0 | 1.7 | 59.81 | 147860.3 | 59.41 | 59.3471
0 | 0 | 0.5 | 116.79 | 169559.7 | 55.23 | 55.3481 | 0 | 0 | 0.6 | 54.3 | 134255.9 | 53.92 | 53.8797
0 | 0 | -0.5 | 109.05 | 158318.2 | 51.63 | 51.68 | 0 | 0 | -0.5 | 50.98 | 126046.2 | 50.65 | 50.5854
Measurement time: | 02/08/2022 19:55:15 | [5 empty] | Measurement time: | 12/08/2022 18:08:37
Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Clinac2 Accelerator
Energy: | 6 MV | [5 empty] | Energy: | 23 MV
Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F | [5 empty] | Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie | [5 empty] | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm
SSD: | 1100 mm | [5 empty] | SSD: | 1000 mm
Field size: | 100 x 100 mm | [5 empty] | Field size: | 100 x 100 mm
Gantry angle: | 0 | [5 empty] | Gantry angle: | 0
Measurement medium: | Water | [5 empty] | Measurement medium: | Water
Scan type: | Beam | [5 empty] | Scan type: | Beam
Scan mode: | Continuous | [5 empty] | Scan mode: | Continuous
Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 310 | 30.94 | 44917.4 | 21.89 | 21.6712 | 0 | 0 | 310 | 33.1 | 81833.9 | 33.6 | 32.3495
0 | 0 | 309.1 | 31.25 | 45361.8 | 21.85 | 21.8884 | 0 | 0 | 309.1 | 34.77 | 85965.8 | 33.86 | 33.9816
0 | 0 | 307.6 | 31.67 | 45978.5 | 22.12 | 22.1825 | 0 | 0 | 307.6 | 35.14 | 86867.2 | 34.19 | 34.3432
0 | 0 | 306.6 | 31.73 | 46068.9 | 22.19 | 22.2246 | 0 | 0 | 306.6 | 35.13 | 86861.2 | 34.24 | 34.3335
0 | 0 | 305.3 | 31.79 | 46151.5 | 22.23 | 22.2666 | 0 | 0 | 305.3 | 35.24 | 87125.6 | 34.33 | 34.441
0 | 0 | 304.1 | 32.17 | 46706.6 | 22.5 | 22.5327 | 0 | 0 | 304.2 | 35.38 | 87477.2 | 34.52 | 34.5778
0 | 0 | 302.9 | 32.41 | 47054.8 | 22.67 | 22.7008 | 0 | 0 | 302.9 | 35.77 | 88431.4 | 34.67 | 34.959
0 | 0 | 301.7 | 32.61 | 47344.9 | 22.82 | 22.8409 | 0 | 0 | 301.7 | 35.69 | 88236.2 | 34.83 | 34.8808
0 | 0 | 300.5 | 32.79 | 47601.7 | 22.9 | 22.967 | 0 | 0 | 300.5 | 35.81 | 88538.9 | 35.03 | 34.998
0 | 0 | 299.3 | 33.03 | 47954.8 | 23.12 | 23.1351 | 0 | 0 | 299.3 | 36.29 | 89714.5 | 35.26 | 35.4672
0 | 0 | 298.1 | 33.21 | 48212 | 23.23 | 23.2612 | 0 | 0 | 298.2 | 36.3 | 89741.2 | 35.37 | 35.4769
0 | 0 | 296.9 | 33.41 | 48509.7 | 23.35 | 23.4013 | 0 | 0 | 296.9 | 36.37 | 89909.5 | 35.49 | 35.5453
0 | 0 | 295.7 | 33.57 | 48732.5 | 23.45 | 23.5133 | 0 | 0 | 295.7 | 36.65 | 90610.7 | 35.73 | 35.819
0 | 0 | 294.5 | 33.85 | 49145 | 23.64 | 23.7095 | 0 | 0 | 294.5 | 36.81 | 91002.6 | 35.89 | 35.9754
0 | 0 | 293.3 | 34.03 | 49397.4 | 23.84 | 23.8355 | 0 | 0 | 293.3 | 37.02 | 91532 | 36.11 | 36.1806
0 | 0 | 292.1 | 34.23 | 49695.7 | 23.89 | 23.9756 | 0 | 0 | 292.1 | 37.24 | 92073.8 | 36.22 | 36.3956
0 | 0 | 290.9 | 34.49 | 50067 | 24.09 | 24.1577 | 0 | 0 | 290.9 | 37.22 | 92019.1 | 36.38 | 36.3761
0 | 0 | 289.7 | 34.71 | 50391.7 | 24.27 | 24.3118 | 0 | 0 | 289.7 | 37.58 | 92909.5 | 36.55 | 36.7279
0 | 0 | 288.5 | 34.95 | 50736.5 | 24.41 | 24.4799 | 0 | 0 | 288.5 | 37.62 | 93011.1 | 36.73 | 36.767
0 | 0 | 287.3 | 35.15 | 51032.5 | 24.54 | 24.62 | 0 | 0 | 287.3 | 37.97 | 93882.8 | 36.92 | 37.1091
0 | 0 | 286.1 | 35.33 | 51295.8 | 24.71 | 24.7461 | 0 | 0 | 286.1 | 37.97 | 93884.1 | 37.04 | 37.1091
0 | 0 | 284.9 | 35.57 | 51646.8 | 24.9 | 24.9142 | 0 | 0 | 284.9 | 38.3 | 94692.1 | 37.31 | 37.4316
0 | 0 | 283.7 | 35.9 | 52125.1 | 25.06 | 25.1453 | 0 | 0 | 283.7 | 38.39 | 94907.2 | 37.48 | 37.5195
0 | 0 | 282.5 | 36.08 | 52375.2 | 25.25 | 25.2714 | 0 | 0 | 282.5 | 38.67 | 95602.5 | 37.49 | 37.7932
0 | 0 | 281.3 | 36.33 | 52749.2 | 25.38 | 25.4465 | 0 | 0 | 281.3 | 38.79 | 95903.7 | 37.79 | 37.9105
0 | 0 | 280.1 | 36.5 | 52988 | 25.53 | 25.5656 | 0 | 0 | 280.1 | 39.1 | 96663.6 | 38.05 | 38.2134
0 | 0 | 278.9 | 36.79 | 53412 | 25.71 | 25.7687 | 0 | 0 | 278.9 | 39.2 | 96917.8 | 38.26 | 38.3112
0 | 0 | 277.7 | 36.97 | 53678 | 25.88 | 25.8948 | 0 | 0 | 277.7 | 39.33 | 97229.6 | 38.41 | 38.4382
0 | 0 | 276.4 | 37.34 | 54206.7 | 26.12 | 26.154 | 0 | 0 | 276.5 | 39.57 | 97820.6 | 38.54 | 38.6728
0 | 0 | 275.3 | 37.52 | 54469.6 | 26.26 | 26.28 | 0 | 0 | 275.2 | 39.79 | 98383.3 | 38.67 | 38.8878
0 | 0 | 274 | 37.82 | 54903.1 | 26.43 | 26.4902 | 0 | 0 | 274.1 | 39.92 | 98698.7 | 38.95 | 39.0149
0 | 0 | 272.9 | 38 | 55162.4 | 26.54 | 26.6162 | 0 | 0 | 272.8 | 40.23 | 99458.6 | 39.24 | 39.3178
0 | 0 | 271.7 | 38.25 | 55523.8 | 26.77 | 26.7913 | 0 | 0 | 271.6 | 40.47 | 100065.5 | 39.41 | 39.5524
0 | 0 | 270.4 | 38.53 | 55937.8 | 26.88 | 26.9875 | 0 | 0 | 270.4 | 40.53 | 100198 | 39.53 | 39.611
0 | 0 | 269.3 | 38.62 | 56061.1 | 27 | 27.0505 | 0 | 0 | 269.2 | 40.75 | 100740.1 | 39.78 | 39.826
0 | 0 | 268 | 39.11 | 56786 | 27.36 | 27.3937 | 0 | 0 | 268 | 40.83 | 100949.8 | 39.95 | 39.9042
0 | 0 | 266.8 | 39.29 | 57037 | 27.45 | 27.5198 | 0 | 0 | 266.8 | 41.2 | 101872.6 | 40.18 | 40.2658
0 | 0 | 265.6 | 39.45 | 57275.2 | 27.56 | 27.6319 | 0 | 0 | 265.6 | 41.44 | 102466.2 | 40.27 | 40.5004
0 | 0 | 264.4 | 39.82 | 57810.3 | 27.83 | 27.891 | 0 | 0 | 264.4 | 41.48 | 102555.3 | 40.55 | 40.5395
0 | 0 | 263.2 | 40.2 | 58367.2 | 28.12 | 28.1572 | 0 | 0 | 263.2 | 41.73 | 103173.7 | 40.79 | 40.7838
0 | 0 | 262 | 40.3 | 58504.9 | 28.13 | 28.2272 | 0 | 0 | 262 | 41.8 | 103346.9 | 40.88 | 40.8522
0 | 0 | 260.8 | 40.49 | 58784.9 | 28.35 | 28.3603 | 0 | 0 | 260.8 | 42.34 | 104690.9 | 41.15 | 41.38
0 | 0 | 259.6 | 40.82 | 59260.2 | 28.52 | 28.5914 | 0 | 0 | 259.6 | 42.28 | 104524.2 | 41.23 | 41.3213
0 | 0 | 258.4 | 41.1 | 59667.3 | 28.77 | 28.7876 | 0 | 0 | 258.4 | 42.52 | 105128.6 | 41.43 | 41.5559
0 | 0 | 257.2 | 41.39 | 60085.3 | 28.93 | 28.9907 | 0 | 0 | 257.2 | 42.7 | 105580.4 | 41.63 | 41.7318
0 | 0 | 256 | 41.61 | 60415 | 29.12 | 29.1448 | 0 | 0 | 256 | 42.91 | 106100.2 | 41.91 | 41.9371
0 | 0 | 254.8 | 41.94 | 60883 | 29.31 | 29.3759 | 0 | 0 | 254.8 | 43.13 | 106644.1 | 42.11 | 42.1521
0 | 0 | 253.6 | 42.11 | 61139.7 | 29.48 | 29.495 | 0 | 0 | 253.6 | 43.35 | 107175.1 | 42.24 | 42.3671
0 | 0 | 252.3 | 42.51 | 61711.9 | 29.7 | 29.7752 | 0 | 0 | 252.4 | 43.66 | 107930.9 | 42.49 | 42.6701
0 | 0 | 251.2 | 42.84 | 62201.3 | 29.97 | 30.0063 | 0 | 0 | 251.2 | 43.79 | 108254.8 | 42.69 | 42.7971
0 | 0 | 249.9 | 43.13 | 62621.9 | 30.11 | 30.2094 | 0 | 0 | 250 | 44.01 | 108796.4 | 42.94 | 43.0121
0 | 0 | 248.7 | 43.47 | 63115.6 | 30.44 | 30.4476 | 0 | 0 | 248.8 | 44.22 | 109320.4 | 43.14 | 43.2174
0 | 0 | 247.5 | 43.69 | 63432.1 | 30.53 | 30.6017 | 0 | 0 | 247.6 | 44.34 | 109623.4 | 43.41 | 43.3346
0 | 0 | 246.3 | 43.9 | 63729.5 | 30.72 | 30.7488 | 0 | 0 | 246.4 | 44.62 | 110311.2 | 43.65 | 43.6083
0 | 0 | 245.1 | 44.24 | 64222.7 | 30.96 | 30.9869 | 0 | 0 | 245.2 | 45.02 | 111302.9 | 43.86 | 43.9992
0 | 0 | 243.9 | 44.38 | 64435 | 31.08 | 31.085 | 0 | 0 | 244 | 45.16 | 111655.1 | 44.01 | 44.136
0 | 0 | 242.7 | 44.78 | 65017.8 | 31.31 | 31.3651 | 0 | 0 | 242.7 | 45.19 | 111717 | 44.19 | 44.1654
0 | 0 | 241.5 | 44.96 | 65272.5 | 31.46 | 31.4912 | 0 | 0 | 241.6 | 45.55 | 112607.3 | 44.5 | 44.5172
0 | 0 | 240.3 | 45.31 | 65773.6 | 31.67 | 31.7364 | 0 | 0 | 240.4 | 45.78 | 113175.7 | 44.69 | 44.742
0 | 0 | 239.1 | 45.67 | 66306.2 | 31.98 | 31.9885 | 0 | 0 | 239.1 | 46.12 | 114034.2 | 44.98 | 45.0743
0 | 0 | 237.9 | 45.87 | 66591.6 | 32.06 | 32.1286 | 0 | 0 | 238 | 46.3 | 114477.2 | 45.1 | 45.2502
0 | 0 | 236.7 | 46.04 | 66841 | 32.2 | 32.2477 | 0 | 0 | 236.7 | 46.45 | 114836.5 | 45.43 | 45.3968
0 | 0 | 235.5 | 46.65 | 67728 | 32.63 | 32.6749 | 0 | 0 | 235.5 | 46.76 | 115595.3 | 45.64 | 45.6998
0 | 0 | 234.3 | 46.97 | 68185.5 | 32.87 | 32.8991 | 0 | 0 | 234.3 | 46.96 | 116109.8 | 45.78 | 45.8952
0 | 0 | 233.1 | 47.12 | 68411.2 | 32.97 | 33.0041 | 0 | 0 | 233.1 | 47.15 | 116568.6 | 45.89 | 46.0809
0 | 0 | 231.8 | 47.49 | 68939.4 | 33.18 | 33.2633 | 0 | 0 | 231.9 | 47.3 | 116937.3 | 46.25 | 46.2275
0 | 0 | 230.7 | 47.86 | 69477.8 | 33.53 | 33.5224 | 0 | 0 | 230.7 | 47.51 | 117469.4 | 46.53 | 46.4328
0 | 0 | 229.5 | 48.17 | 69939.6 | 33.68 | 33.7396 | 0 | 0 | 229.5 | 47.89 | 118410.4 | 46.69 | 46.8041
0 | 0 | 228.3 | 48.46 | 70355 | 33.87 | 33.9427 | 0 | 0 | 228.3 | 48.21 | 119182.8 | 47.01 | 47.1169
0 | 0 | 227.1 | 48.79 | 70836.1 | 34.12 | 34.1738 | 0 | 0 | 227.1 | 48.32 | 119459.3 | 47.15 | 47.2244
0 | 0 | 225.8 | 49.08 | 71247.6 | 34.31 | 34.377 | 0 | 0 | 225.9 | 48.46 | 119810.1 | 47.28 | 47.3612
0 | 0 | 224.7 | 49.4 | 71711.5 | 34.55 | 34.6011 | 0 | 0 | 224.7 | 48.7 | 120409.2 | 47.73 | 47.5958
0 | 0 | 223.4 | 49.82 | 72321.9 | 34.86 | 34.8953 | 0 | 0 | 223.5 | 49.19 | 121612.5 | 48.03 | 48.0747
0 | 0 | 222.2 | 50 | 72593.7 | 35.02 | 35.0214 | 0 | 0 | 222.3 | 49.26 | 121778.5 | 48.13 | 48.1431
0 | 0 | 221.1 | 50.33 | 73062.2 | 35.2 | 35.2525 | 0 | 0 | 221.1 | 49.43 | 122205.7 | 48.38 | 48.3092
0 | 0 | 219.8 | 50.7 | 73604.7 | 35.42 | 35.5117 | 0 | 0 | 219.9 | 49.69 | 122855.5 | 48.69 | 48.5633
0 | 0 | 218.7 | 51.02 | 74068.7 | 35.72 | 35.7358 | 0 | 0 | 218.7 | 49.99 | 123599.2 | 48.83 | 48.8565
0 | 0 | 217.4 | 51.47 | 74724.6 | 35.97 | 36.051 | 0 | 0 | 217.5 | 50.5 | 124851.1 | 49.18 | 49.355
0 | 0 | 216.2 | 51.87 | 75305.9 | 36.22 | 36.3312 | 0 | 0 | 216.3 | 50.43 | 124683.7 | 49.35 | 49.2866
0 | 0 | 215 | 52.11 | 75659 | 36.45 | 36.4993 | 0 | 0 | 215.1 | 50.86 | 125741.9 | 49.67 | 49.7068
0 | 0 | 213.8 | 52.4 | 76076.5 | 36.64 | 36.7024 | 0 | 0 | 213.9 | 51.04 | 126201.1 | 49.76 | 49.8827
0 | 0 | 212.6 | 52.73 | 76554.4 | 36.85 | 36.9335 | 0 | 0 | 212.6 | 51.42 | 127121 | 50.07 | 50.2541
0 | 0 | 211.4 | 53.16 | 77174 | 37.13 | 37.2347 | 0 | 0 | 211.5 | 51.59 | 127555.2 | 50.47 | 50.4203
0 | 0 | 210.2 | 53.51 | 77688 | 37.43 | 37.4799 | 0 | 0 | 210.2 | 51.87 | 128241.4 | 50.55 | 50.6939
0 | 0 | 209 | 53.89 | 78231.7 | 37.65 | 37.746 | 0 | 0 | 209 | 51.9 | 128327.1 | 50.86 | 50.7232
0 | 0 | 207.8 | 54.17 | 78640.7 | 37.89 | 37.9421 | 0 | 0 | 207.9 | 52.28 | 129256.6 | 50.93 | 51.0946
0 | 0 | 206.6 | 54.43 | 79021.9 | 38.07 | 38.1243 | 0 | 0 | 206.6 | 52.55 | 129922.6 | 51.35 | 51.3585
0 | 0 | 205.4 | 54.96 | 79783 | 38.43 | 38.4955 | 0 | 0 | 205.4 | 52.72 | 130336.3 | 51.58 | 51.5246
0 | 0 | 204.2 | 55.36 | 80377.2 | 38.71 | 38.7757 | 0 | 0 | 204.2 | 52.84 | 130634.9 | 51.9 | 51.6419
0 | 0 | 203 | 55.56 | 80654.1 | 38.88 | 38.9157 | 0 | 0 | 203 | 53.25 | 131647.5 | 52.15 | 52.0426
0 | 0 | 201.8 | 56.02 | 81330.8 | 39.09 | 39.2379 | 0 | 0 | 201.8 | 53.68 | 132719 | 52.34 | 52.4629
0 | 0 | 200.6 | 56.37 | 81838.6 | 39.44 | 39.4831 | 0 | 0 | 200.6 | 53.82 | 133070.9 | 52.63 | 52.5997
0 | 0 | 199.4 | 56.56 | 82117.4 | 39.56 | 39.6162 | 0 | 0 | 199.4 | 54.28 | 134208.1 | 52.99 | 53.0493
0 | 0 | 198.2 | 57.01 | 82772.7 | 39.89 | 39.9314 | 0 | 0 | 198.2 | 54.39 | 134472.9 | 53.13 | 53.1568
0 | 0 | 197 | 57.56 | 83560.4 | 40.22 | 40.3166 | 0 | 0 | 197 | 54.53 | 134812.3 | 53.31 | 53.2936
0 | 0 | 195.7 | 57.86 | 84005.1 | 40.47 | 40.5267 | 0 | 0 | 195.8 | 54.82 | 135523.3 | 53.4 | 53.577
0 | 0 | 194.6 | 58.22 | 84520.3 | 40.75 | 40.7789 | 0 | 0 | 194.6 | 55.28 | 136675.6 | 53.86 | 54.0266
0 | 0 | 193.4 | 58.7 | 85223.9 | 41.06 | 41.1151 | 0 | 0 | 193.4 | 55.44 | 137059.9 | 54.27 | 54.183
0 | 0 | 192.1 | 59.14 | 85858.4 | 41.4 | 41.4233 | 0 | 0 | 192.2 | 55.94 | 138313.7 | 54.42 | 54.6716
0 | 0 | 191 | 59.29 | 86078 | 41.51 | 41.5283 | 0 | 0 | 191 | 55.89 | 138173 | 54.63 | 54.6228
0 | 0 | 189.7 | 59.81 | 86835.1 | 41.77 | 41.8926 | 0 | 0 | 189.8 | 56.37 | 139370.8 | 54.85 | 55.0919
0 | 0 | 188.5 | 60.05 | 87183 | 42 | 42.0607 | 0 | 0 | 188.5 | 56.54 | 139782.9 | 55.28 | 55.258
0 | 0 | 187.3 | 60.52 | 87856.2 | 42.33 | 42.3899 | 0 | 0 | 187.4 | 56.76 | 140334.3 | 55.51 | 55.473
0 | 0 | 186.1 | 60.98 | 88522.9 | 42.69 | 42.7121 | 0 | 0 | 186.2 | 57.11 | 141202.1 | 55.66 | 55.8151
0 | 0 | 184.9 | 61.29 | 88984.4 | 42.89 | 42.9292 | 0 | 0 | 185 | 57.36 | 141825.2 | 55.97 | 56.0594
0 | 0 | 183.7 | 61.65 | 89495.5 | 43.12 | 43.1813 | 0 | 0 | 183.8 | 57.51 | 142187.8 | 56.27 | 56.206
0 | 0 | 182.5 | 62.18 | 90271.8 | 43.46 | 43.5526 | 0 | 0 | 182.6 | 57.95 | 143262.1 | 56.66 | 56.636
0 | 0 | 181.3 | 62.58 | 90849.3 | 43.75 | 43.8327 | 0 | 0 | 181.4 | 58.36 | 144276.7 | 57 | 57.0367
0 | 0 | 180.1 | 62.82 | 91196.6 | 43.91 | 44.0008 | 0 | 0 | 180.1 | 58.67 | 145054.6 | 57.29 | 57.3397
0 | 0 | 178.9 | 63.45 | 92117.6 | 44.32 | 44.4421 | 0 | 0 | 178.9 | 58.77 | 145298.2 | 57.48 | 57.4375
0 | 0 | 177.7 | 63.93 | 92817.5 | 44.69 | 44.7783 | 0 | 0 | 177.8 | 59.17 | 146297.5 | 57.81 | 57.8284
0 | 0 | 176.5 | 64.09 | 93043.7 | 44.74 | 44.8904 | 0 | 0 | 176.5 | 59.38 | 146817.3 | 58.12 | 58.0336
0 | 0 | 175.3 | 64.58 | 93753.6 | 45.18 | 45.2336 | 0 | 0 | 175.3 | 59.84 | 147933.2 | 58.26 | 58.4832
0 | 0 | 174.1 | 65.1 | 94518.1 | 45.43 | 45.5978 | 0 | 0 | 174.1 | 60.06 | 148496.6 | 58.65 | 58.6982
0 | 0 | 172.9 | 65.42 | 94981.8 | 45.72 | 45.822 | 0 | 0 | 172.9 | 60.06 | 148491.9 | 58.88 | 58.6982
0 | 0 | 171.6 | 65.97 | 95772.7 | 46.11 | 46.2072 | 0 | 0 | 171.7 | 60.57 | 149746.6 | 59.32 | 59.1966
0 | 0 | 170.5 | 66.49 | 96523.8 | 46.46 | 46.5714 | 0 | 0 | 170.5 | 60.97 | 150738.4 | 59.52 | 59.5876
0 | 0 | 169.3 | 66.77 | 96939.1 | 46.65 | 46.7675 | 0 | 0 | 169.3 | 61.38 | 151744.6 | 59.9 | 59.9883
0 | 0 | 168 | 67.09 | 97401.8 | 46.93 | 46.9917 | 0 | 0 | 168.1 | 61.46 | 151941.3 | 60.12 | 60.0665
0 | 0 | 166.8 | 67.68 | 98253.1 | 47.29 | 47.4049 | 0 | 0 | 166.9 | 61.85 | 152921.9 | 60.38 | 60.4476
0 | 0 | 165.6 | 68.08 | 98843.3 | 47.58 | 47.6851 | 0 | 0 | 165.7 | 61.84 | 152889.6 | 60.61 | 60.4378
0 | 0 | 164.5 | 68.66 | 99681.1 | 48.05 | 48.0913 | 0 | 0 | 164.5 | 62.51 | 154552.7 | 60.86 | 61.0927
0 | 0 | 163.2 | 69.06 | 100257.9 | 48.32 | 48.3715 | 0 | 0 | 163.3 | 62.51 | 154555.2 | 61.28 | 61.0927
0 | 0 | 162 | 69.45 | 100823.8 | 48.53 | 48.6447 | 0 | 0 | 162.1 | 62.99 | 155721.9 | 61.58 | 61.5618
0 | 0 | 160.8 | 69.91 | 101496.9 | 48.86 | 48.9669 | 0 | 0 | 160.9 | 63.37 | 156677.2 | 61.78 | 61.9332
0 | 0 | 159.6 | 70.25 | 101986.5 | 49.11 | 49.205 | 0 | 0 | 159.7 | 63.84 | 157839.8 | 62.2 | 62.3925
0 | 0 | 158.4 | 70.99 | 103061.5 | 49.71 | 49.7233 | 0 | 0 | 158.5 | 63.96 | 158125.9 | 62.42 | 62.5098
0 | 0 | 157.2 | 71.32 | 103534.8 | 49.94 | 49.9545 | 0 | 0 | 157.3 | 64.15 | 158597.7 | 62.75 | 62.6955
0 | 0 | 156 | 71.86 | 104330.6 | 50.23 | 50.3327 | 0 | 0 | 156.1 | 64.65 | 159834.5 | 63.14 | 63.1841
0 | 0 | 154.8 | 72.22 | 104842.6 | 50.53 | 50.5849 | 0 | 0 | 154.9 | 64.74 | 160059.3 | 63.38 | 63.2721
0 | 0 | 153.6 | 72.61 | 105414.4 | 50.73 | 50.858 | 0 | 0 | 153.6 | 65.28 | 161406 | 63.55 | 63.7998
0 | 0 | 152.4 | 73.01 | 106002 | 51.04 | 51.1382 | 0 | 0 | 152.4 | 65.49 | 161908.2 | 63.89 | 64.0051
0 | 0 | 151.2 | 73.82 | 107168.7 | 51.58 | 51.7055 | 0 | 0 | 151.3 | 65.75 | 162559.3 | 64.34 | 64.2592
0 | 0 | 150 | 74.12 | 107605.2 | 51.81 | 51.9157 | 0 | 0 | 150 | 66.03 | 163240.5 | 64.61 | 64.5328
0 | 0 | 148.8 | 74.67 | 108405 | 52.23 | 52.3009 | 0 | 0 | 148.8 | 66.62 | 164701.8 | 65.13 | 65.1095
0 | 0 | 147.6 | 74.99 | 108869.6 | 52.4 | 52.525 | 0 | 0 | 147.6 | 66.91 | 165421.5 | 65.34 | 65.3929
0 | 0 | 146.4 | 75.53 | 109657 | 52.73 | 52.9033 | 0 | 0 | 146.4 | 67.3 | 166391.9 | 65.69 | 65.774
0 | 0 | 145.2 | 76.11 | 110500.1 | 53.28 | 53.3095 | 0 | 0 | 145.2 | 67.48 | 166823.8 | 66.01 | 65.95
0 | 0 | 144 | 76.64 | 111265.2 | 53.61 | 53.6807 | 0 | 0 | 144 | 67.74 | 167478.5 | 66.09 | 66.2041
0 | 0 | 142.8 | 77.08 | 111897.5 | 53.89 | 53.9889 | 0 | 0 | 142.8 | 68.22 | 168672.6 | 66.45 | 66.6732
0 | 0 | 141.6 | 77.57 | 112612.3 | 54.22 | 54.3321 | 0 | 0 | 141.6 | 68.57 | 169526.7 | 66.87 | 67.0152
0 | 0 | 140.4 | 78.09 | 113363.1 | 54.68 | 54.6964 | 0 | 0 | 140.4 | 68.72 | 169893.8 | 67.4 | 67.1618
0 | 0 | 139.2 | 78.55 | 114032.3 | 54.91 | 55.0186 | 0 | 0 | 139.2 | 69.14 | 170932.3 | 67.53 | 67.5723
0 | 0 | 138 | 78.99 | 114670.5 | 55.19 | 55.3267 | 0 | 0 | 138 | 69.66 | 172223.8 | 67.83 | 68.0805
0 | 0 | 136.8 | 79.62 | 115584.1 | 55.65 | 55.768 | 0 | 0 | 136.8 | 69.83 | 172652.2 | 68.15 | 68.2467
0 | 0 | 135.6 | 80.21 | 116452.4 | 56.12 | 56.1813 | 0 | 0 | 135.6 | 70.25 | 173688.4 | 68.49 | 68.6572
0 | 0 | 134.4 | 80.58 | 116986 | 56.26 | 56.4404 | 0 | 0 | 134.4 | 70.56 | 174440.5 | 69.08 | 68.9601
0 | 0 | 133.2 | 81.13 | 117779.3 | 56.72 | 56.8257 | 0 | 0 | 133.2 | 71.13 | 175854.2 | 69.38 | 69.5172
0 | 0 | 131.9 | 81.77 | 118714.2 | 57.21 | 57.2739 | 0 | 0 | 132 | 71.28 | 176236.4 | 69.61 | 69.6638
0 | 0 | 130.7 | 82.13 | 119234.6 | 57.36 | 57.5261 | 0 | 0 | 130.8 | 71.66 | 177173.7 | 70.09 | 70.0352
0 | 0 | 129.5 | 82.77 | 120162 | 57.88 | 57.9744 | 0 | 0 | 129.6 | 71.98 | 177967.1 | 70.42 | 70.3479
0 | 0 | 128.3 | 83.39 | 121063.8 | 58.31 | 58.4086 | 0 | 0 | 128.4 | 72.48 | 179197.9 | 70.78 | 70.8366
0 | 0 | 127.1 | 83.85 | 121738.2 | 58.62 | 58.7308 | 0 | 0 | 127.2 | 73.11 | 180757.5 | 71.13 | 71.4523
0 | 0 | 125.9 | 84.38 | 122500.7 | 58.85 | 59.1021 | 0 | 0 | 126 | 72.71 | 179764.5 | 71.32 | 71.0614
0 | 0 | 124.7 | 85.05 | 123472.9 | 59.42 | 59.5713 | 0 | 0 | 124.8 | 73.56 | 181862.9 | 71.68 | 71.8921
0 | 0 | 123.5 | 85.4 | 123989.4 | 59.82 | 59.8165 | 0 | 0 | 123.6 | 73.67 | 182138.4 | 72.07 | 71.9996
0 | 0 | 122.3 | 85.95 | 124785.3 | 60.09 | 60.2017 | 0 | 0 | 122.3 | 74.12 | 183260.5 | 72.59 | 72.4394
0 | 0 | 121.1 | 86.52 | 125603.7 | 60.5 | 60.601 | 0 | 0 | 121.1 | 74.59 | 184417.5 | 72.88 | 72.8987
0 | 0 | 119.9 | 87.03 | 126344.6 | 60.98 | 60.9582 | 0 | 0 | 119.9 | 74.86 | 185087.5 | 73.09 | 73.1626
0 | 0 | 118.7 | 87.75 | 127388 | 61.25 | 61.4625 | 0 | 0 | 118.7 | 75.32 | 186226.4 | 73.51 | 73.6122
0 | 0 | 117.5 | 88.11 | 127920.5 | 61.63 | 61.7146 | 0 | 0 | 117.6 | 75.32 | 186214.2 | 73.74 | 73.6122
0 | 0 | 116.3 | 88.84 | 128972.3 | 62.07 | 62.226 | 0 | 0 | 116.3 | 76.08 | 188090.1 | 74.31 | 74.355
0 | 0 | 115.1 | 89.3 | 129641.7 | 62.4 | 62.5482 | 0 | 0 | 115.1 | 76.29 | 188617.7 | 74.61 | 74.5602
0 | 0 | 113.9 | 89.64 | 130143.3 | 62.72 | 62.7863 | 0 | 0 | 113.9 | 76.78 | 189822.8 | 74.99 | 75.0391
0 | 0 | 112.7 | 90.48 | 131361.9 | 63.2 | 63.3747 | 0 | 0 | 112.7 | 77.11 | 190648.5 | 75.32 | 75.3616
0 | 0 | 111.5 | 91.04 | 132164.1 | 63.61 | 63.7669 | 0 | 0 | 111.5 | 77.32 | 191166.2 | 75.6 | 75.5668
0 | 0 | 110.3 | 91.52 | 132869.8 | 63.95 | 64.1031 | 0 | 0 | 110.3 | 77.99 | 192828.2 | 75.93 | 76.2217
0 | 0 | 109.1 | 92.19 | 133838.7 | 64.49 | 64.5724 | 0 | 0 | 109.1 | 78.05 | 192974.6 | 76.55 | 76.2803
0 | 0 | 107.9 | 92.7 | 134587.6 | 64.76 | 64.9296 | 0 | 0 | 107.9 | 78.49 | 194044.7 | 76.75 | 76.7103
0 | 0 | 106.7 | 93.25 | 135378.1 | 65.31 | 65.3148 | 0 | 0 | 106.7 | 79.09 | 195531.7 | 77.14 | 77.2967
0 | 0 | 105.4 | 93.91 | 136333 | 65.72 | 65.7771 | 0 | 0 | 105.5 | 79.45 | 196422.9 | 77.31 | 77.6486
0 | 0 | 104.3 | 94.44 | 137105.8 | 66.01 | 66.1484 | 0 | 0 | 104.3 | 79.64 | 196899.2 | 77.78 | 77.8342
0 | 0 | 103.1 | 95.18 | 138180 | 66.51 | 66.6667 | 0 | 0 | 103.1 | 79.85 | 197410 | 78.43 | 78.0395
0 | 0 | 101.9 | 95.61 | 138811.5 | 66.78 | 66.9679 | 0 | 0 | 101.9 | 80.69 | 199491.6 | 78.52 | 78.8604
0 | 0 | 100.7 | 96.43 | 139993.6 | 67.44 | 67.5422 | 0 | 0 | 100.7 | 80.88 | 199964.6 | 79.08 | 79.0461
0 | 0 | 99.5 | 96.92 | 140705.4 | 67.73 | 67.8854 | 0 | 0 | 99.5 | 81.22 | 200802.8 | 79.39 | 79.3784
0 | 0 | 98.3 | 97.49 | 141540.1 | 68.19 | 68.2847 | 0 | 0 | 98.3 | 81.66 | 201899.7 | 79.8 | 79.8084
0 | 0 | 97.1 | 98.19 | 142544.2 | 68.49 | 68.775 | 0 | 0 | 97.1 | 82.06 | 202874.7 | 80.12 | 80.1994
0 | 0 | 95.9 | 98.62 | 143173.2 | 68.9 | 69.0761 | 0 | 0 | 95.9 | 82.34 | 203578.6 | 80.47 | 80.473
0 | 0 | 94.7 | 99.35 | 144235.3 | 69.43 | 69.5874 | 0 | 0 | 94.7 | 82.86 | 204863.3 | 81.08 | 80.9812
0 | 0 | 93.4 | 100.1 | 145317.4 | 69.88 | 70.1128 | 0 | 0 | 93.5 | 83.28 | 205906.9 | 81.37 | 81.3917
0 | 0 | 92.2 | 100.61 | 146060.3 | 70.28 | 70.47 | 0 | 0 | 92.2 | 83.68 | 206889.1 | 82.04 | 81.7826
0 | 0 | 91.1 | 101.28 | 147034.5 | 70.77 | 70.9393 | 0 | 0 | 91.1 | 84.36 | 208557.7 | 82.21 | 82.4472
0 | 0 | 89.9 | 101.85 | 147858.6 | 71.33 | 71.3385 | 0 | 0 | 89.9 | 84.53 | 208986.4 | 82.73 | 82.6134
0 | 0 | 88.6 | 102.55 | 148874.2 | 71.68 | 71.8288 | 0 | 0 | 88.7 | 84.97 | 210064.3 | 82.89 | 83.0434
0 | 0 | 87.4 | 103.16 | 149762.3 | 71.91 | 72.2561 | 0 | 0 | 87.5 | 85.23 | 210707.1 | 83.32 | 83.2975
0 | 0 | 86.2 | 103.85 | 150765.7 | 72.62 | 72.7394 | 0 | 0 | 86.2 | 85.86 | 212269.2 | 83.93 | 83.9132
0 | 0 | 85 | 104.34 | 151472.7 | 72.85 | 73.0826 | 0 | 0 | 85.1 | 86.28 | 213309.7 | 84.21 | 84.3237
0 | 0 | 83.8 | 104.98 | 152402.7 | 73.23 | 73.5309 | 0 | 0 | 83.8 | 86.51 | 213875.1 | 84.55 | 84.5485
0 | 0 | 82.6 | 105.81 | 153618.2 | 73.93 | 74.1122 | 0 | 0 | 82.6 | 87.07 | 215264.3 | 85.1 | 85.0958
0 | 0 | 81.4 | 106.37 | 154431.2 | 74.34 | 74.5044 | 0 | 0 | 81.4 | 87.27 | 215761.8 | 85.35 | 85.2912
0 | 0 | 80.2 | 107.16 | 155568.5 | 74.9 | 75.0578 | 0 | 0 | 80.2 | 87.72 | 216866.4 | 85.79 | 85.731
0 | 0 | 79 | 107.61 | 156229.8 | 75.13 | 75.373 | 0 | 0 | 79 | 88.27 | 218232.7 | 86.04 | 86.2686
0 | 0 | 77.8 | 108.32 | 157261.9 | 75.67 | 75.8703 | 0 | 0 | 77.8 | 88.51 | 218832.3 | 86.46 | 86.5031
0 | 0 | 76.6 | 108.91 | 158118.2 | 75.98 | 76.2835 | 0 | 0 | 76.6 | 88.76 | 219437.7 | 86.88 | 86.7475
0 | 0 | 75.4 | 109.51 | 158986.9 | 76.71 | 76.7038 | 0 | 0 | 75.4 | 89.51 | 221307.9 | 87.31 | 87.4805
0 | 0 | 74.2 | 110.28 | 160100.6 | 77.04 | 77.2431 | 0 | 0 | 74.2 | 89.81 | 222037 | 87.8 | 87.7737
0 | 0 | 72.9 | 110.99 | 161137.9 | 77.59 | 77.7404 | 0 | 0 | 73 | 90.55 | 223876.1 | 88.32 | 88.4969
0 | 0 | 71.7 | 111.39 | 161721.3 | 77.81 | 78.0206 | 0 | 0 | 71.8 | 90.45 | 223614.5 | 88.58 | 88.3991
0 | 0 | 70.5 | 112.38 | 163151.7 | 78.54 | 78.714 | 0 | 0 | 70.5 | 91.08 | 225173 | 89.17 | 89.0149
0 | 0 | 69.3 | 112.77 | 163724.3 | 78.72 | 78.9872 | 0 | 0 | 69.3 | 91.66 | 226607.8 | 89.55 | 89.5817
0 | 0 | 68.1 | 113.57 | 164875.9 | 79.4 | 79.5475 | 0 | 0 | 68.1 | 91.97 | 227391.5 | 89.73 | 89.8847
0 | 0 | 66.9 | 114.51 | 166239.7 | 80 | 80.2059 | 0 | 0 | 66.9 | 92.47 | 228615.1 | 90.37 | 90.3733
0 | 0 | 65.7 | 114.76 | 166610.4 | 80.25 | 80.381 | 0 | 0 | 65.7 | 92.99 | 229906.7 | 91.01 | 90.8815
0 | 0 | 64.5 | 115.68 | 167948.8 | 80.86 | 81.0254 | 0 | 0 | 64.5 | 93.38 | 230873.5 | 91.29 | 91.2627
0 | 0 | 63.3 | 116.27 | 168801.8 | 81.17 | 81.4387 | 0 | 0 | 63.3 | 93.72 | 231716.6 | 91.52 | 91.595
0 | 0 | 62.1 | 116.99 | 169848.5 | 81.8 | 81.943 | 0 | 0 | 62.1 | 94.17 | 232830 | 91.95 | 92.0348
0 | 0 | 60.9 | 117.64 | 170780.8 | 82.21 | 82.3983 | 0 | 0 | 60.9 | 94.7 | 234124.4 | 92.46 | 92.5528
0 | 0 | 59.7 | 118.3 | 171747.9 | 82.67 | 82.8605 | 0 | 0 | 59.7 | 94.88 | 234581.7 | 92.85 | 92.7287
0 | 0 | 58.5 | 119.23 | 173098.9 | 83.27 | 83.5119 | 0 | 0 | 58.5 | 95.4 | 235855.4 | 93.28 | 93.2369
0 | 0 | 57.3 | 119.85 | 173998.7 | 83.69 | 83.9462 | 0 | 0 | 57.3 | 96.07 | 237512.2 | 93.54 | 93.8917
0 | 0 | 56.1 | 120.4 | 174801.8 | 84.07 | 84.3314 | 0 | 0 | 56.1 | 96.47 | 238510.3 | 94.11 | 94.2826
0 | 0 | 54.8 | 121.15 | 175877.5 | 84.68 | 84.8568 | 0 | 0 | 54.9 | 96.46 | 238493.9 | 94.7 | 94.2729
0 | 0 | 53.6 | 121.74 | 176738.1 | 85.08 | 85.27 | 0 | 0 | 53.6 | 97.28 | 240511.1 | 94.85 | 95.0743
0 | 0 | 52.4 | 122.38 | 177668.2 | 85.36 | 85.7183 | 0 | 0 | 52.4 | 97.52 | 241096.7 | 95.35 | 95.3088
0 | 0 | 51.2 | 123.07 | 178669.9 | 86.06 | 86.2016 | 0 | 0 | 51.2 | 97.99 | 242256.7 | 95.91 | 95.7682
0 | 0 | 50 | 123.83 | 179772.5 | 86.59 | 86.7339 | 0 | 0 | 50 | 98.06 | 242449.4 | 96.14 | 95.8366
0 | 0 | 48.8 | 124.77 | 181142.5 | 87.07 | 87.3923 | 0 | 0 | 48.8 | 98.69 | 243986.8 | 96.36 | 96.4523
0 | 0 | 47.6 | 125.11 | 181625.9 | 87.54 | 87.6305 | 0 | 0 | 47.6 | 99.07 | 244945.8 | 96.78 | 96.8237
0 | 0 | 46.4 | 125.85 | 182710.8 | 87.98 | 88.1488 | 0 | 0 | 46.4 | 99.32 | 245552.9 | 96.96 | 97.068
0 | 0 | 45.2 | 126.6 | 183794.3 | 88.45 | 88.6741 | 0 | 0 | 45.2 | 99.71 | 246528.1 | 97.43 | 97.4492
0 | 0 | 44 | 127.27 | 184774.3 | 88.83 | 89.1434 | 0 | 0 | 44 | 100.15 | 247596.8 | 97.82 | 97.8792
0 | 0 | 42.7 | 127.98 | 185798.8 | 89.39 | 89.6407 | 0 | 0 | 42.8 | 100.36 | 248131.4 | 98.12 | 98.0844
0 | 0 | 41.5 | 128.97 | 187232 | 90.09 | 90.3341 | 0 | 0 | 41.6 | 100.88 | 249405.3 | 98.3 | 98.5927
0 | 0 | 40.3 | 129.61 | 188169.9 | 90.46 | 90.7824 | 0 | 0 | 40.4 | 101.24 | 250306.6 | 98.7 | 98.9445
0 | 0 | 39.1 | 130.18 | 188996.4 | 90.91 | 91.1816 | 0 | 0 | 39.2 | 100.95 | 249574.9 | 99.27 | 98.6611
0 | 0 | 37.9 | 130.93 | 190085.1 | 91.53 | 91.7069 | 0 | 0 | 38 | 101.5 | 250941.2 | 99.27 | 99.1986
0 | 0 | 36.7 | 131.61 | 191070.2 | 92.08 | 92.1832 | 0 | 0 | 36.8 | 102.1 | 252438.1 | 99.52 | 99.785
0 | 0 | 35.5 | 132.16 | 191861.3 | 92.31 | 92.5685 | 0 | 0 | 35.5 | 101.93 | 252018.7 | 99.72 | 99.6188
0 | 0 | 34.3 | 132.93 | 192989.1 | 92.88 | 93.1078 | 0 | 0 | 34.3 | 102.32 | 252970.9 | 99.96 | 100
0 | 0 | 33.1 | 133.9 | 194400.5 | 93.62 | 93.7872 | 0 | 0 | 33.1 | 102.06 | 252335.1 | 100 | 99.7459
0 | 0 | 31.9 | 134.64 | 195470.3 | 93.96 | 94.3055 | 0 | 0 | 31.9 | 102.23 | 252756.5 | 99.91 | 99.912
0 | 0 | 30.7 | 135.11 | 196153.1 | 94.32 | 94.6347 | 0 | 0 | 30.7 | 102.3 | 252930.6 | 99.7 | 99.9805
0 | 0 | 29.5 | 135.95 | 197367.6 | 95.02 | 95.2231 | 0 | 0 | 29.5 | 102.27 | 252848.2 | 99.79 | 99.9511
0 | 0 | 28.3 | 136.38 | 197990.6 | 95.26 | 95.5243 | 0 | 0 | 28.3 | 101.94 | 252031.8 | 99.74 | 99.6286
0 | 0 | 27.1 | 136.93 | 198794.2 | 95.69 | 95.9095 | 0 | 0 | 27.1 | 101.62 | 251250.9 | 99.46 | 99.3159
0 | 0 | 25.9 | 137.93 | 200239.7 | 96.39 | 96.6099 | 0 | 0 | 25.9 | 101.66 | 251350.4 | 99.28 | 99.355
0 | 0 | 24.6 | 138.64 | 201278.1 | 96.89 | 97.1072 | 0 | 0 | 24.7 | 100.87 | 249397.5 | 98.66 | 98.5829
0 | 0 | 23.4 | 139.19 | 202068.9 | 97.3 | 97.4925 | 0 | 0 | 23.5 | 100.65 | 248854.2 | 98.45 | 98.3679
0 | 0 | 22.2 | 139.89 | 203095.9 | 97.57 | 97.9828 | 0 | 0 | 22.3 | 99.78 | 246694.8 | 97.51 | 97.5176
0 | 0 | 21 | 140.6 | 204115.4 | 98.15 | 98.4801 | 0 | 0 | 21 | 98.69 | 243992 | 96.83 | 96.4523
0 | 0 | 19.8 | 141.25 | 205059.8 | 98.77 | 98.9354 | 0 | 0 | 19.8 | 98.07 | 242464.2 | 95.77 | 95.8464
0 | 0 | 18.6 | 141.54 | 205487.5 | 98.92 | 99.1385 | 0 | 0 | 18.7 | 97 | 239813.6 | 94.61 | 94.8006
0 | 0 | 17.4 | 142.05 | 206220.1 | 99.24 | 99.4957 | 0 | 0 | 17.4 | 95.7 | 236611 | 93.43 | 93.5301
0 | 0 | 16.2 | 142.52 | 206907.2 | 99.62 | 99.8249 | 0 | 0 | 16.3 | 94 | 232402.7 | 91.95 | 91.8686
0 | 0 | 15 | 142.76 | 207256.8 | 99.86 | 99.993 | 0 | 0 | 15 | 92.08 | 227646.9 | 90 | 89.9922
0 | 0 | 13.8 | 142.74 | 207235.1 | 99.75 | 99.979 | 0 | 0 | 13.8 | 90.06 | 222648.8 | 87.87 | 88.018
0 | 0 | 12.6 | 142.77 | 207276.4 | 100 | 100 | 0 | 0 | 12.6 | 87.56 | 216481.9 | 85.42 | 85.5747
0 | 0 | 11.4 | 141.89 | 206001 | 99.13 | 99.3836 | 0 | 0 | 11.4 | 84.36 | 208571.4 | 82.84 | 82.4472
0 | 0 | 10.2 | 140.74 | 204326.8 | 98.35 | 98.5781 | 0 | 0 | 10.2 | 81.48 | 201459.3 | 79.5 | 79.6325
0 | 0 | 8.9 | 138.45 | 200998.3 | 96.91 | 96.9742 | 0 | 0 | 9 | 77.7 | 192105.4 | 75.86 | 75.9382
0 | 0 | 7.8 | 135.66 | 196953.6 | 94.73 | 95.02 | 0 | 0 | 7.8 | 72.85 | 180115.4 | 71.02 | 71.1982
0 | 0 | 6.5 | 131.62 | 191089.5 | 92.03 | 92.1902 | 0 | 0 | 6.6 | 67.39 | 166609.8 | 65.82 | 65.862
0 | 0 | 5.4 | 125.44 | 182116.9 | 87.62 | 87.8616 | 0 | 0 | 5.4 | 60.8 | 150329.2 | 59.23 | 59.4214
0 | 0 | 4.1 | 116.45 | 169055.6 | 81.43 | 81.5648 | 0 | 0 | 4.2 | 53.13 | 131355.7 | 51.76 | 51.9253
0 | 0 | 2.9 | 103.87 | 150801 | 72.71 | 72.7534 | 0 | 0 | 3 | 44.31 | 109548.9 | 43.46 | 43.3053
0 | 0 | 1.7 | 88.71 | 128789.8 | 62.02 | 62.1349 | 0 | 0 | 1.8 | 37.36 | 92377 | 36.44 | 36.5129
0 | 0 | 0.5 | 78.47 | 113915.8 | 54.8 | 54.9625 | 0 | 0 | 0.6 | 33.46 | 82728.8 | 32.77 | 32.7013
0 | 0 | -0.5 | 72.97 | 105942.3 | 51.03 | 51.1102 | 0 | 0 | -0.5 | 31.45 | 77754.3 | 30.64 | 30.7369
Measurement time: | 02/08/2022 19:59:54
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 306.6 | 35.9 | 52112.8 | 20.92 | 20.8903
0 | 0 | 305.3 | 36.3 | 52694 | 21.17 | 21.1231
0 | 0 | 304.1 | 36.46 | 52937.8 | 21.27 | 21.2162
0 | 0 | 302.9 | 36.69 | 53262.2 | 21.34 | 21.35
0 | 0 | 301.7 | 37.04 | 53780.3 | 21.59 | 21.5537
0 | 0 | 300.5 | 37.17 | 53956.3 | 21.64 | 21.6293
0 | 0 | 299.3 | 37.49 | 54424.6 | 21.86 | 21.8155
0 | 0 | 298.1 | 37.75 | 54811.3 | 22 | 21.9668
0 | 0 | 296.9 | 38.05 | 55242 | 22.13 | 22.1414
0 | 0 | 295.7 | 38.27 | 55552.8 | 22.32 | 22.2694
0 | 0 | 294.5 | 38.52 | 55922.7 | 22.47 | 22.4149
0 | 0 | 293.3 | 38.72 | 56211.9 | 22.52 | 22.5313
0 | 0 | 292.1 | 39.09 | 56745.1 | 22.75 | 22.7466
0 | 0 | 290.9 | 39.2 | 56910.8 | 22.81 | 22.8106
0 | 0 | 289.7 | 39.64 | 57551.2 | 23.1 | 23.0666
0 | 0 | 288.5 | 39.77 | 57732.3 | 23.12 | 23.1423
0 | 0 | 287.3 | 40.01 | 58090.4 | 23.26 | 23.2819
0 | 0 | 286.1 | 40.42 | 58685 | 23.59 | 23.5205
0 | 0 | 284.9 | 40.6 | 58936 | 23.62 | 23.6253
0 | 0 | 283.7 | 40.89 | 59368.1 | 23.77 | 23.794
0 | 0 | 282.4 | 41.18 | 59790.2 | 23.98 | 23.9628
0 | 0 | 281.3 | 41.42 | 60126.6 | 24.11 | 24.1024
0 | 0 | 280 | 41.64 | 60447.9 | 24.21 | 24.2304
0 | 0 | 278.9 | 41.94 | 60885 | 24.41 | 24.405
0 | 0 | 277.6 | 42.26 | 61345.6 | 24.54 | 24.5912
0 | 0 | 276.4 | 42.66 | 61939 | 24.8 | 24.824
0 | 0 | 275.3 | 42.83 | 62185.5 | 24.89 | 24.9229
0 | 0 | 274 | 43.17 | 62678.4 | 25.11 | 25.1207
0 | 0 | 272.9 | 43.41 | 63018.3 | 25.28 | 25.2604
0 | 0 | 271.7 | 43.61 | 63313.8 | 25.4 | 25.3768
0 | 0 | 270.4 | 44 | 63873.3 | 25.61 | 25.6037
0 | 0 | 269.3 | 44.3 | 64311.4 | 25.72 | 25.7783
0 | 0 | 268 | 44.51 | 64623.9 | 25.91 | 25.9005
0 | 0 | 266.8 | 44.96 | 65265.1 | 26.19 | 26.1624
0 | 0 | 265.6 | 45.07 | 65431.4 | 26.21 | 26.2264
0 | 0 | 264.4 | 45.41 | 65922.1 | 26.36 | 26.4242
0 | 0 | 263.2 | 45.75 | 66422.7 | 26.67 | 26.6221
0 | 0 | 262 | 46.12 | 66960.4 | 26.87 | 26.8374
0 | 0 | 260.9 | 46.47 | 67463.1 | 27.04 | 27.041
0 | 0 | 259.6 | 46.62 | 67676 | 27.15 | 27.1283
0 | 0 | 258.4 | 47.03 | 68279.6 | 27.39 | 27.3669
0 | 0 | 257.2 | 47.39 | 68806.7 | 27.61 | 27.5764
0 | 0 | 256 | 47.56 | 69051.2 | 27.7 | 27.6753
0 | 0 | 254.8 | 47.98 | 69662.7 | 27.87 | 27.9197
0 | 0 | 253.6 | 48.3 | 70117.9 | 28.15 | 28.1059
0 | 0 | 252.4 | 48.66 | 70643.6 | 28.37 | 28.3154
0 | 0 | 251.2 | 48.92 | 71021.2 | 28.47 | 28.4667
0 | 0 | 250 | 49.33 | 71619.5 | 28.71 | 28.7053
0 | 0 | 248.8 | 49.62 | 72034.4 | 28.9 | 28.874
0 | 0 | 247.6 | 50.01 | 72598.3 | 29.08 | 29.101
0 | 0 | 246.4 | 50.14 | 72797 | 29.15 | 29.1766
0 | 0 | 245.2 | 50.74 | 73659.4 | 29.52 | 29.5257
0 | 0 | 244 | 51.06 | 74135.1 | 29.81 | 29.712
0 | 0 | 242.8 | 51.41 | 74635.4 | 29.95 | 29.9156
0 | 0 | 241.6 | 51.6 | 74917.7 | 30 | 30.0262
0 | 0 | 240.3 | 52.03 | 75530.9 | 30.3 | 30.2764
0 | 0 | 239.2 | 52.4 | 76073.4 | 30.54 | 30.4917
0 | 0 | 237.9 | 52.73 | 76556.9 | 30.72 | 30.6837
0 | 0 | 236.8 | 53.02 | 76980.4 | 30.87 | 30.8525
0 | 0 | 235.5 | 53.31 | 77401.2 | 31.06 | 31.0212
0 | 0 | 234.3 | 53.74 | 78023.6 | 31.31 | 31.2715
0 | 0 | 233.1 | 54.03 | 78443.9 | 31.47 | 31.4402
0 | 0 | 231.9 | 54.46 | 79070.5 | 31.73 | 31.6904
0 | 0 | 230.7 | 54.79 | 79543.7 | 31.88 | 31.8825
0 | 0 | 229.5 | 55.3 | 80280.5 | 32.17 | 32.1792
0 | 0 | 228.3 | 55.54 | 80627.6 | 32.35 | 32.3189
0 | 0 | 227.1 | 56.12 | 81467.7 | 32.64 | 32.6564
0 | 0 | 225.9 | 56.41 | 81896.4 | 32.86 | 32.8251
0 | 0 | 224.7 | 56.71 | 82324.2 | 33.06 | 32.9997
0 | 0 | 223.5 | 57.14 | 82953.9 | 33.29 | 33.2499
0 | 0 | 222.3 | 57.44 | 83397.7 | 33.42 | 33.4245
0 | 0 | 221.1 | 58.05 | 84273.8 | 33.84 | 33.7795
0 | 0 | 219.9 | 58.41 | 84797.4 | 34.01 | 33.9889
0 | 0 | 218.7 | 58.67 | 85182.3 | 34.15 | 34.1402
0 | 0 | 217.5 | 59.08 | 85766.2 | 34.42 | 34.3788
0 | 0 | 216.3 | 59.62 | 86552.6 | 34.65 | 34.693
0 | 0 | 215.1 | 59.95 | 87032.8 | 34.93 | 34.8851
0 | 0 | 213.9 | 60.16 | 87335.7 | 35.03 | 35.0073
0 | 0 | 212.7 | 60.72 | 88155.7 | 35.27 | 35.3331
0 | 0 | 211.5 | 61.06 | 88647.7 | 35.55 | 35.531
0 | 0 | 210.3 | 61.44 | 89204.3 | 35.79 | 35.7521
0 | 0 | 209.1 | 61.89 | 89853.6 | 36.02 | 36.014
0 | 0 | 207.9 | 62.33 | 90487.9 | 36.3 | 36.27
0 | 0 | 206.7 | 62.74 | 91085.1 | 36.54 | 36.5086
0 | 0 | 205.5 | 63.15 | 91680.9 | 36.74 | 36.7472
0 | 0 | 204.2 | 63.64 | 92393.1 | 37.05 | 37.0323
0 | 0 | 203.1 | 64.01 | 92923.4 | 37.26 | 37.2476
0 | 0 | 201.8 | 64.58 | 93761.3 | 37.65 | 37.5793
0 | 0 | 200.6 | 64.9 | 94213.8 | 37.77 | 37.7655
0 | 0 | 199.5 | 65.38 | 94913.9 | 38.05 | 38.0448
0 | 0 | 198.2 | 65.8 | 95534.1 | 38.28 | 38.2892
0 | 0 | 197 | 66.13 | 96010.8 | 38.54 | 38.4812
0 | 0 | 195.8 | 66.73 | 96880.1 | 38.84 | 38.8304
0 | 0 | 194.6 | 67.01 | 97289.3 | 38.9 | 38.9933
0 | 0 | 193.4 | 67.61 | 98160.7 | 39.33 | 39.3424
0 | 0 | 192.2 | 68.21 | 99024.1 | 39.7 | 39.6916
0 | 0 | 191 | 68.44 | 99360.7 | 39.87 | 39.8254
0 | 0 | 189.8 | 69.04 | 100227.4 | 40.14 | 40.1746
0 | 0 | 188.6 | 69.46 | 100842.7 | 40.47 | 40.419
0 | 0 | 187.4 | 69.96 | 101571.8 | 40.7 | 40.7099
0 | 0 | 186.2 | 70.51 | 102360.5 | 41.02 | 41.03
0 | 0 | 185 | 70.9 | 102934.4 | 41.21 | 41.2569
0 | 0 | 183.8 | 71.54 | 103854.3 | 41.63 | 41.6293
0 | 0 | 182.6 | 72.14 | 104728.5 | 42.03 | 41.9785
0 | 0 | 181.4 | 72.36 | 105056.8 | 42.15 | 42.1065
0 | 0 | 180.2 | 72.86 | 105770.1 | 42.36 | 42.3974
0 | 0 | 178.9 | 73.18 | 106237 | 42.62 | 42.5836
0 | 0 | 177.7 | 74.03 | 107468.7 | 43.12 | 43.0783
0 | 0 | 176.6 | 74.56 | 108246.6 | 43.39 | 43.3867
0 | 0 | 175.3 | 74.82 | 108620.6 | 43.53 | 43.538
0 | 0 | 174.1 | 75.56 | 109692.9 | 43.96 | 43.9686
0 | 0 | 173 | 75.96 | 110277.5 | 44.23 | 44.2013
0 | 0 | 171.7 | 76.42 | 110939.8 | 44.45 | 44.469
0 | 0 | 170.5 | 77.01 | 111795.4 | 44.79 | 44.8123
0 | 0 | 169.3 | 77.33 | 112271.3 | 45.03 | 44.9985
0 | 0 | 168.1 | 77.88 | 113064.8 | 45.32 | 45.3186
0 | 0 | 166.9 | 78.62 | 114144.4 | 45.71 | 45.7492
0 | 0 | 165.7 | 79.13 | 114882.9 | 46.06 | 46.046
0 | 0 | 164.5 | 79.66 | 115654 | 46.38 | 46.3544
0 | 0 | 163.3 | 80.27 | 116533.6 | 46.75 | 46.7093
0 | 0 | 162.1 | 80.83 | 117345.8 | 47.04 | 47.0352
0 | 0 | 160.9 | 81.39 | 118162.6 | 47.35 | 47.3611
0 | 0 | 159.7 | 81.83 | 118796 | 47.69 | 47.6171
0 | 0 | 158.5 | 82.46 | 119712.1 | 47.92 | 47.9837
0 | 0 | 157.3 | 82.94 | 120410.9 | 48.31 | 48.263
0 | 0 | 156.1 | 83.28 | 120905.1 | 48.5 | 48.4609
0 | 0 | 154.9 | 84.03 | 121986.5 | 48.99 | 48.8973
0 | 0 | 153.7 | 84.52 | 122702.7 | 49.2 | 49.1824
0 | 0 | 152.5 | 85.18 | 123662.3 | 49.49 | 49.5665
0 | 0 | 151.3 | 85.87 | 124670.2 | 49.95 | 49.968
0 | 0 | 150.1 | 86.28 | 125254.4 | 50.19 | 50.2066
0 | 0 | 148.9 | 86.78 | 125988.2 | 50.6 | 50.4975
0 | 0 | 147.6 | 87.62 | 127203.7 | 50.94 | 50.9863
0 | 0 | 146.5 | 88.01 | 127768.7 | 51.25 | 51.2133
0 | 0 | 145.2 | 88.79 | 128898.3 | 51.65 | 51.6672
0 | 0 | 144 | 89.25 | 129574.6 | 52.02 | 51.9348
0 | 0 | 142.8 | 89.81 | 130382.3 | 52.23 | 52.2607
0 | 0 | 141.6 | 90.56 | 131469 | 52.65 | 52.6971
0 | 0 | 140.4 | 91.08 | 132222.4 | 53.02 | 52.9997
0 | 0 | 139.2 | 91.9 | 133416.2 | 53.5 | 53.4769
0 | 0 | 138 | 92.3 | 133993.4 | 53.78 | 53.7096
0 | 0 | 136.8 | 92.91 | 134888.3 | 54.06 | 54.0646
0 | 0 | 135.6 | 93.82 | 136212.8 | 54.59 | 54.5941
0 | 0 | 134.4 | 94.23 | 136800.4 | 54.8 | 54.8327
0 | 0 | 133.2 | 94.93 | 137817.8 | 55.23 | 55.24
0 | 0 | 132 | 95.59 | 138779.5 | 55.67 | 55.6241
0 | 0 | 130.8 | 96.02 | 139402.8 | 55.92 | 55.8743
0 | 0 | 129.6 | 96.9 | 140674.7 | 56.46 | 56.3864
0 | 0 | 128.4 | 97.27 | 141221.4 | 56.54 | 56.6017
0 | 0 | 127.2 | 97.96 | 142219.3 | 57.08 | 57.0032
0 | 0 | 126 | 98.72 | 143318.6 | 57.5 | 57.4454
0 | 0 | 124.8 | 99.41 | 144320.1 | 57.75 | 57.847
0 | 0 | 123.5 | 99.99 | 145163.7 | 58.15 | 58.1845
0 | 0 | 122.4 | 100.76 | 146276.7 | 58.6 | 58.6325
0 | 0 | 121.1 | 101.21 | 146937.7 | 58.93 | 58.8944
0 | 0 | 120 | 101.95 | 148003.4 | 59.35 | 59.325
0 | 0 | 118.8 | 102.74 | 149159.3 | 59.72 | 59.7847
0 | 0 | 117.5 | 103.56 | 150349.3 | 60.29 | 60.2619
0 | 0 | 116.4 | 104.07 | 151091.1 | 60.66 | 60.5586
0 | 0 | 115.1 | 104.9 | 152295.1 | 61.08 | 61.0416
0 | 0 | 113.9 | 105.59 | 153300.8 | 61.41 | 61.4431
0 | 0 | 112.7 | 106.29 | 154317.1 | 61.79 | 61.8505
0 | 0 | 111.5 | 106.87 | 155146 | 62.24 | 62.188
0 | 0 | 110.3 | 107.26 | 155722 | 62.45 | 62.4149
0 | 0 | 109.1 | 108.49 | 157501.4 | 63.02 | 63.1306
0 | 0 | 107.9 | 109.12 | 158422.2 | 63.44 | 63.4972
0 | 0 | 106.7 | 109.75 | 159332.4 | 63.87 | 63.8638
0 | 0 | 105.5 | 110.64 | 160619.9 | 64.33 | 64.3817
0 | 0 | 104.3 | 111.24 | 161492.2 | 64.76 | 64.7309
0 | 0 | 103.1 | 111.99 | 162586.2 | 65.18 | 65.1673
0 | 0 | 101.9 | 112.64 | 163534.1 | 65.52 | 65.5455
0 | 0 | 100.7 | 113.48 | 164753.7 | 65.92 | 66.0343
0 | 0 | 99.5 | 114.3 | 165939.6 | 66.58 | 66.5115
0 | 0 | 98.3 | 114.97 | 166918.8 | 67.02 | 66.9014
0 | 0 | 97.1 | 115.51 | 167701.3 | 67.26 | 67.2156
0 | 0 | 95.9 | 116.57 | 169230 | 67.78 | 67.8324
0 | 0 | 94.7 | 117.16 | 170094.4 | 68.06 | 68.1757
0 | 0 | 93.5 | 118.05 | 171390.5 | 68.75 | 68.6936
0 | 0 | 92.3 | 118.83 | 172510.5 | 69.24 | 69.1475
0 | 0 | 91.1 | 119.55 | 173565.2 | 69.46 | 69.5665
0 | 0 | 89.9 | 120.29 | 174633.6 | 70.11 | 69.9971
0 | 0 | 88.7 | 120.95 | 175594.2 | 70.35 | 70.3811
0 | 0 | 87.5 | 121.82 | 176860.3 | 70.94 | 70.8874
0 | 0 | 86.3 | 122.81 | 178287.6 | 71.55 | 71.4635
0 | 0 | 85.1 | 123.39 | 179131.8 | 71.81 | 71.801
0 | 0 | 83.8 | 124.32 | 180488.9 | 72.39 | 72.3422
0 | 0 | 82.7 | 125.14 | 181673.7 | 72.83 | 72.8193
0 | 0 | 81.4 | 125.95 | 182851.2 | 73.27 | 73.2907
0 | 0 | 80.3 | 126.78 | 184061.4 | 73.8 | 73.7736
0 | 0 | 79 | 127.52 | 185131 | 74.19 | 74.2042
0 | 0 | 77.8 | 128.19 | 186103.7 | 74.68 | 74.5941
0 | 0 | 76.6 | 128.91 | 187143.1 | 74.98 | 75.0131
0 | 0 | 75.4 | 129.69 | 188278.1 | 75.46 | 75.467
0 | 0 | 74.2 | 130.73 | 189789.1 | 76.27 | 76.0722
0 | 0 | 73 | 131.6 | 191058.6 | 76.53 | 76.5784
0 | 0 | 71.8 | 132.07 | 191744.3 | 76.93 | 76.8519
0 | 0 | 70.6 | 133.11 | 193250.1 | 77.56 | 77.4571
0 | 0 | 69.3 | 134.08 | 194651.6 | 78.14 | 78.0215
0 | 0 | 68.2 | 134.76 | 195644.1 | 78.41 | 78.4172
0 | 0 | 66.9 | 135.62 | 196889.2 | 78.78 | 78.9177
0 | 0 | 65.7 | 136.37 | 197985.3 | 79.36 | 79.3541
0 | 0 | 64.5 | 137.26 | 199274.4 | 79.84 | 79.872
0 | 0 | 63.3 | 138.01 | 200355.7 | 80.41 | 80.3084
0 | 0 | 62.1 | 139.17 | 202040 | 81.05 | 80.9834
0 | 0 | 60.9 | 139.89 | 203095 | 81.56 | 81.4024
0 | 0 | 59.7 | 140.75 | 204331.7 | 81.97 | 81.9028
0 | 0 | 58.5 | 141.57 | 205527.5 | 82.23 | 82.38
0 | 0 | 57.3 | 142.46 | 206826.9 | 82.84 | 82.8979
0 | 0 | 56.1 | 143.25 | 207967.5 | 83.21 | 83.3576
0 | 0 | 54.9 | 144.28 | 209465.1 | 84.11 | 83.9569
0 | 0 | 53.7 | 145.27 | 210902.9 | 84.54 | 84.533
0 | 0 | 52.4 | 145.99 | 211945.5 | 84.95 | 84.952
0 | 0 | 51.2 | 147.02 | 213442.1 | 85.43 | 85.5514
0 | 0 | 50 | 147.54 | 214191.4 | 85.93 | 85.8539
0 | 0 | 48.8 | 148.65 | 215813.2 | 86.38 | 86.4999
0 | 0 | 47.6 | 149.43 | 216936.6 | 86.9 | 86.9537
0 | 0 | 46.4 | 150.36 | 218287.7 | 87.5 | 87.4949
0 | 0 | 45.2 | 151.51 | 219966.9 | 88.3 | 88.1641
0 | 0 | 44 | 152.18 | 220929.3 | 88.57 | 88.554
0 | 0 | 42.8 | 152.99 | 222110.7 | 88.95 | 89.0253
0 | 0 | 41.6 | 153.8 | 223290 | 89.42 | 89.4967
0 | 0 | 40.4 | 154.85 | 224802.1 | 90.09 | 90.1077
0 | 0 | 39.2 | 155.37 | 225569.1 | 90.44 | 90.4102
0 | 0 | 38 | 156.65 | 227420.7 | 90.93 | 91.1551
0 | 0 | 36.8 | 157.58 | 228765.7 | 91.6 | 91.6962
0 | 0 | 35.6 | 158.22 | 229697.3 | 92.08 | 92.0687
0 | 0 | 34.4 | 159.13 | 231016.3 | 92.56 | 92.5982
0 | 0 | 33.2 | 159.99 | 232266.4 | 93.12 | 93.0986
0 | 0 | 31.9 | 161.13 | 233922.3 | 93.75 | 93.762
0 | 0 | 30.7 | 161.59 | 234599.8 | 94.06 | 94.0297
0 | 0 | 29.5 | 162.66 | 236144.5 | 94.61 | 94.6523
0 | 0 | 28.3 | 163.73 | 237700.9 | 95.11 | 95.2749
0 | 0 | 27.1 | 164.46 | 238763.4 | 95.86 | 95.6997
0 | 0 | 25.9 | 165.55 | 240343.3 | 96.37 | 96.334
0 | 0 | 24.7 | 166.05 | 241071.7 | 96.55 | 96.625
0 | 0 | 23.5 | 166.94 | 242355 | 96.92 | 97.1429
0 | 0 | 22.3 | 168.07 | 243994.7 | 97.84 | 97.8004
0 | 0 | 21.1 | 169 | 245357.1 | 98.53 | 98.3416
0 | 0 | 19.9 | 169.57 | 246183.8 | 98.76 | 98.6733
0 | 0 | 18.6 | 170.16 | 247039.7 | 98.94 | 99.0166
0 | 0 | 17.5 | 170.66 | 247760.6 | 99.46 | 99.3075
0 | 0 | 16.3 | 171.42 | 248859.7 | 99.83 | 99.7498
0 | 0 | 15 | 171.85 | 249495.2 | 100 | 100
0 | 0 | 13.8 | 171.71 | 249279.2 | 99.89 | 99.9185
0 | 0 | 12.6 | 171.41 | 248847.4 | 99.83 | 99.744
0 | 0 | 11.4 | 170.57 | 247634.2 | 99.35 | 99.2552
0 | 0 | 10.2 | 169.23 | 245687.1 | 98.44 | 98.4754
0 | 0 | 9 | 166.83 | 242200.6 | 97.14 | 97.0788
0 | 0 | 7.8 | 163.49 | 237352.5 | 95.12 | 95.1353
0 | 0 | 6.6 | 158.18 | 229640.6 | 92.08 | 92.0454
0 | 0 | 5.4 | 151.17 | 219472.1 | 87.82 | 87.9662
0 | 0 | 4.2 | 141.14 | 204899.7 | 82.02 | 82.1298
0 | 0 | 2.9 | 123.89 | 179867.1 | 72.25 | 72.0919
0 | 0 | 1.8 | 106.67 | 154861.2 | 62.09 | 62.0716
0 | 0 | 0.5 | 94.37 | 137009 | 54.83 | 54.9142
0 | 0 | -0.5 | 87.75 | 127391.6 | 51.08 | 51.062
Measurement time: | 02/08/2022 20:01:19
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF)
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 306.6 | 111.74 | 162225.7 | 18.59 | 18.535
0 | 0 | 305.3 | 112.81 | 163771.1 | 18.72 | 18.7125
0 | 0 | 304.1 | 113.2 | 164342.4 | 18.82 | 18.7772
0 | 0 | 302.9 | 113.96 | 165444.8 | 18.96 | 18.9032
0 | 0 | 301.7 | 114.83 | 166714.3 | 19.1 | 19.0475
0 | 0 | 300.5 | 115.65 | 167901.9 | 19.23 | 19.1836
0 | 0 | 299.3 | 116.57 | 169227.8 | 19.38 | 19.3362
0 | 0 | 298.1 | 117.21 | 170159.3 | 19.48 | 19.4423
0 | 0 | 296.9 | 117.93 | 171215.8 | 19.6 | 19.5618
0 | 0 | 295.7 | 118.78 | 172437.6 | 19.73 | 19.7028
0 | 0 | 294.5 | 119.35 | 173272.6 | 19.86 | 19.7973
0 | 0 | 293.3 | 120.56 | 175029.7 | 20 | 19.998
0 | 0 | 292.1 | 121.35 | 176170.3 | 20.13 | 20.1291
0 | 0 | 290.9 | 122.01 | 177133.4 | 20.25 | 20.2385
0 | 0 | 289.7 | 122.91 | 178443.7 | 20.43 | 20.3878
0 | 0 | 288.5 | 123.6 | 179440.1 | 20.58 | 20.5023
0 | 0 | 287.3 | 124.86 | 181266.2 | 20.7 | 20.7113
0 | 0 | 286.1 | 125.49 | 182183.3 | 20.9 | 20.8158
0 | 0 | 284.9 | 126.62 | 183820.8 | 21.02 | 21.0032
0 | 0 | 283.7 | 127.09 | 184510.2 | 21.15 | 21.0812
0 | 0 | 282.4 | 128.22 | 186151.4 | 21.31 | 21.2686
0 | 0 | 281.2 | 129.09 | 187406.6 | 21.43 | 21.4129
0 | 0 | 280 | 129.92 | 188617.8 | 21.6 | 21.5506
0 | 0 | 278.8 | 130.73 | 189794.5 | 21.72 | 21.685
0 | 0 | 277.6 | 131.71 | 191216 | 21.89 | 21.8475
0 | 0 | 276.4 | 132.78 | 192775.1 | 22.07 | 22.025
0 | 0 | 275.2 | 133.83 | 194292.6 | 22.21 | 22.1992
0 | 0 | 274 | 134.77 | 195654.4 | 22.36 | 22.3551
0 | 0 | 272.8 | 135.38 | 196545.3 | 22.53 | 22.4563
0 | 0 | 271.6 | 136.33 | 197921.6 | 22.66 | 22.6139
0 | 0 | 270.4 | 137.47 | 199570.4 | 22.81 | 22.803
0 | 0 | 269.2 | 138.41 | 200937.5 | 23 | 22.9589
0 | 0 | 268 | 139.34 | 202291.2 | 23.13 | 23.1132
0 | 0 | 266.8 | 140.54 | 204039.6 | 23.33 | 23.3122
0 | 0 | 265.6 | 141.23 | 205030.8 | 23.47 | 23.4267
0 | 0 | 264.4 | 142.48 | 206843.9 | 23.65 | 23.634
0 | 0 | 263.2 | 143.55 | 208406 | 23.85 | 23.8115
0 | 0 | 262 | 144.26 | 209438.6 | 23.99 | 23.9293
0 | 0 | 260.8 | 145.4 | 211087.9 | 24.14 | 24.1184
0 | 0 | 259.6 | 146.29 | 212379.4 | 24.31 | 24.266
0 | 0 | 258.4 | 147.3 | 213853.5 | 24.49 | 24.4335
0 | 0 | 257.2 | 148.48 | 215554.3 | 24.64 | 24.6293
0 | 0 | 256 | 149.66 | 217276.2 | 24.86 | 24.825
0 | 0 | 254.8 | 150.54 | 218553.8 | 24.99 | 24.971
0 | 0 | 253.5 | 151.77 | 220343.2 | 25.22 | 25.175
0 | 0 | 252.3 | 153.04 | 222184.4 | 25.38 | 25.3857
0 | 0 | 251.2 | 153.9 | 223426.4 | 25.53 | 25.5283
0 | 0 | 249.9 | 154.77 | 224688.3 | 25.74 | 25.6726
0 | 0 | 248.7 | 155.78 | 226158.8 | 25.9 | 25.8402
0 | 0 | 247.5 | 157.1 | 228081.8 | 26.07 | 26.0591
0 | 0 | 246.3 | 158.43 | 230003.4 | 26.25 | 26.2797
0 | 0 | 245.1 | 159.22 | 231152.8 | 26.47 | 26.4108
0 | 0 | 243.9 | 160.09 | 232420.3 | 26.65 | 26.5551
0 | 0 | 242.7 | 161.48 | 234437.2 | 26.78 | 26.7857
0 | 0 | 241.5 | 163.37 | 237180 | 27.08 | 27.0992
0 | 0 | 240.3 | 163.54 | 237423 | 27.23 | 27.1274
0 | 0 | 239.1 | 164.98 | 239511.4 | 27.41 | 27.3662
0 | 0 | 237.9 | 166.77 | 242109.3 | 27.6 | 27.6631
0 | 0 | 236.7 | 166.94 | 242359.6 | 27.79 | 27.6913
0 | 0 | 235.5 | 168.9 | 245208 | 28.03 | 28.0165
0 | 0 | 234.3 | 169.91 | 246672 | 28.24 | 28.184
0 | 0 | 233.1 | 170.86 | 248055.7 | 28.41 | 28.3416
0 | 0 | 231.9 | 172.35 | 250210.4 | 28.56 | 28.5887
0 | 0 | 230.7 | 173.14 | 251360.6 | 28.81 | 28.7198
0 | 0 | 229.5 | 174.85 | 253841.7 | 29.03 | 29.0034
0 | 0 | 228.3 | 176.13 | 255705.6 | 29.26 | 29.2157
0 | 0 | 227.1 | 177.04 | 257031 | 29.43 | 29.3667
0 | 0 | 225.9 | 178.76 | 259515.2 | 29.68 | 29.652
0 | 0 | 224.7 | 180.3 | 261760 | 29.91 | 29.9074
0 | 0 | 223.5 | 180.94 | 262682.2 | 30.06 | 30.0136
0 | 0 | 222.2 | 182.82 | 265419.4 | 30.3 | 30.3254
0 | 0 | 221.1 | 183.47 | 266363 | 30.5 | 30.4333
0 | 0 | 219.8 | 185.15 | 268792.2 | 30.74 | 30.7119
0 | 0 | 218.7 | 186.68 | 271015.7 | 30.98 | 30.9657
0 | 0 | 217.4 | 187.46 | 272154.6 | 31.16 | 31.0951
0 | 0 | 216.2 | 189.48 | 275087.5 | 31.4 | 31.4302
0 | 0 | 215 | 189.97 | 275801.7 | 31.58 | 31.5115
0 | 0 | 213.8 | 191.8 | 278453.1 | 31.88 | 31.815
0 | 0 | 212.6 | 193.91 | 281513.2 | 32.16 | 32.165
0 | 0 | 211.4 | 194.68 | 282629.8 | 32.34 | 32.2927
0 | 0 | 210.2 | 195.62 | 283996.8 | 32.55 | 32.4487
0 | 0 | 209 | 197.86 | 287249.9 | 32.77 | 32.8202
0 | 0 | 207.8 | 198.64 | 288379.7 | 33.04 | 32.9496
0 | 0 | 206.6 | 200.17 | 290606.7 | 33.3 | 33.2034
0 | 0 | 205.4 | 201.46 | 292480.6 | 33.49 | 33.4174
0 | 0 | 204.2 | 203.62 | 295605.8 | 33.78 | 33.7757
0 | 0 | 203 | 204.71 | 297201.5 | 33.94 | 33.9565
0 | 0 | 201.8 | 205.98 | 299038.8 | 34.24 | 34.1671
0 | 0 | 200.6 | 207.24 | 300866 | 34.51 | 34.3761
0 | 0 | 199.4 | 209.61 | 304311.9 | 34.79 | 34.7693
0 | 0 | 198.2 | 210.61 | 305756.9 | 34.94 | 34.9351
0 | 0 | 197 | 212.14 | 307988.8 | 35.24 | 35.1889
0 | 0 | 195.7 | 214.39 | 311241.5 | 35.57 | 35.5622
0 | 0 | 194.6 | 215.17 | 312384.6 | 35.8 | 35.6915
0 | 0 | 193.3 | 217.03 | 315084 | 36.04 | 36.0001
0 | 0 | 192.1 | 218.33 | 316972 | 36.29 | 36.2157
0 | 0 | 191 | 220.54 | 320183.6 | 36.55 | 36.5823
0 | 0 | 189.7 | 221.67 | 321823.2 | 36.83 | 36.7697
0 | 0 | 188.6 | 224.02 | 325224.6 | 37.12 | 37.1595
0 | 0 | 187.3 | 224.89 | 326488 | 37.35 | 37.3039
0 | 0 | 186.1 | 226.67 | 329070.2 | 37.62 | 37.5991
0 | 0 | 184.9 | 227.45 | 330212.9 | 37.82 | 37.7285
0 | 0 | 183.7 | 230.62 | 334806.5 | 38.23 | 38.2543
0 | 0 | 182.5 | 231.32 | 335823.2 | 38.5 | 38.3704
0 | 0 | 181.3 | 233.59 | 339121.9 | 38.75 | 38.747
0 | 0 | 180.1 | 235.31 | 341622.5 | 39.02 | 39.0323
0 | 0 | 178.9 | 236.5 | 343350.5 | 39.22 | 39.2297
0 | 0 | 177.7 | 238.55 | 346320.9 | 39.65 | 39.5697
0 | 0 | 176.5 | 240.69 | 349436.1 | 39.93 | 39.9247
0 | 0 | 175.3 | 242.25 | 351694.6 | 40.23 | 40.1835
0 | 0 | 174.1 | 243.97 | 354186.9 | 40.47 | 40.4688
0 | 0 | 172.9 | 245.68 | 356676.8 | 40.74 | 40.7524
0 | 0 | 171.7 | 247.15 | 358805.4 | 41.04 | 40.9963
0 | 0 | 170.5 | 249.34 | 361984.9 | 41.44 | 41.3595
0 | 0 | 169.3 | 251.05 | 364468.5 | 41.68 | 41.6432
0 | 0 | 168.1 | 253.1 | 367445.2 | 41.97 | 41.9832
0 | 0 | 166.8 | 254.38 | 369300.7 | 42.27 | 42.1955
0 | 0 | 165.6 | 256.39 | 372222.3 | 42.61 | 42.5289
0 | 0 | 164.4 | 259.19 | 376287.6 | 42.95 | 42.9934
0 | 0 | 163.2 | 260.31 | 377908 | 43.18 | 43.1792
0 | 0 | 162 | 262.37 | 380908.2 | 43.56 | 43.5209
0 | 0 | 160.8 | 264.37 | 383812.7 | 43.83 | 43.8526
0 | 0 | 159.6 | 265.68 | 385710.6 | 44.23 | 44.0699
0 | 0 | 158.4 | 268.6 | 389956.2 | 44.55 | 44.5543
0 | 0 | 157.2 | 269.92 | 391865.2 | 44.84 | 44.7732
0 | 0 | 156 | 272.14 | 395091.5 | 45.17 | 45.1415
0 | 0 | 154.8 | 274.02 | 397821.9 | 45.48 | 45.4533
0 | 0 | 153.6 | 275.81 | 400412.8 | 45.85 | 45.7503
0 | 0 | 152.4 | 278.22 | 403909 | 46.17 | 46.15
0 | 0 | 151.2 | 280.12 | 406679.9 | 46.51 | 46.4652
0 | 0 | 150 | 282.06 | 409492.2 | 46.85 | 46.787
0 | 0 | 148.8 | 284.27 | 412703.5 | 47.16 | 47.1536
0 | 0 | 147.6 | 285.91 | 415076.9 | 47.48 | 47.4256
0 | 0 | 146.4 | 288.74 | 419183.7 | 47.92 | 47.895
0 | 0 | 145.2 | 290.27 | 421406.7 | 48.26 | 48.1488
0 | 0 | 144 | 292.6 | 424799.3 | 48.57 | 48.5353
0 | 0 | 142.8 | 294.57 | 427656 | 48.94 | 48.8621
0 | 0 | 141.5 | 296.97 | 431131.3 | 49.33 | 49.2602
0 | 0 | 140.3 | 299.49 | 434798.1 | 49.69 | 49.6782
0 | 0 | 139.1 | 301.22 | 437311.3 | 49.99 | 49.9652
0 | 0 | 137.9 | 303.48 | 440594.1 | 50.35 | 50.34
0 | 0 | 136.8 | 306.19 | 444520.5 | 50.83 | 50.7896
0 | 0 | 135.5 | 307.74 | 446770.8 | 51.15 | 51.0467
0 | 0 | 134.4 | 310.34 | 450541.6 | 51.51 | 51.478
0 | 0 | 133.1 | 312.63 | 453878.3 | 51.88 | 51.8578
0 | 0 | 131.9 | 314.62 | 456765.7 | 52.24 | 52.1879
0 | 0 | 130.7 | 317.83 | 461425.7 | 52.74 | 52.7204
0 | 0 | 129.5 | 320.1 | 464720.7 | 53.04 | 53.0969
0 | 0 | 128.3 | 321.53 | 466798.5 | 53.43 | 53.3341
0 | 0 | 127.1 | 324.39 | 470948.6 | 53.8 | 53.8085
0 | 0 | 125.9 | 326.18 | 473549 | 54.21 | 54.1054
0 | 0 | 124.7 | 329.03 | 477679.4 | 54.7 | 54.5782
0 | 0 | 123.5 | 331.08 | 480662 | 54.92 | 54.9182
0 | 0 | 122.3 | 333.97 | 484856.5 | 55.41 | 55.3976
0 | 0 | 121.1 | 336.5 | 488532.2 | 55.83 | 55.8173
0 | 0 | 119.9 | 338.76 | 491813 | 56.23 | 56.1922
0 | 0 | 118.7 | 340.94 | 494977.4 | 56.6 | 56.5538
0 | 0 | 117.5 | 343.25 | 498330.2 | 56.95 | 56.9369
0 | 0 | 116.3 | 345.91 | 502186 | 57.4 | 57.3782
0 | 0 | 115.1 | 348.3 | 505655.9 | 57.71 | 57.7746
0 | 0 | 113.9 | 351.21 | 509880.9 | 58.32 | 58.2573
0 | 0 | 112.7 | 353.49 | 513190.6 | 58.7 | 58.6355
0 | 0 | 111.5 | 355.94 | 516741.9 | 59.13 | 59.0419
0 | 0 | 110.3 | 358.66 | 520700.8 | 59.56 | 59.4931
0 | 0 | 109.1 | 361.09 | 524229.2 | 59.98 | 59.8962
0 | 0 | 107.9 | 363.79 | 528151.8 | 60.38 | 60.344
0 | 0 | 106.7 | 366.79 | 532496.2 | 60.88 | 60.8417
0 | 0 | 105.4 | 369.47 | 536387.1 | 61.34 | 61.2862
0 | 0 | 104.3 | 371.8 | 539772.6 | 61.69 | 61.6727
0 | 0 | 103.1 | 374.21 | 543277.1 | 62.19 | 62.0725
0 | 0 | 101.9 | 377.77 | 548446 | 62.68 | 62.663
0 | 0 | 100.7 | 380.02 | 551702.5 | 63.11 | 63.0362
0 | 0 | 99.5 | 383.42 | 556642.5 | 63.55 | 63.6002
0 | 0 | 98.3 | 385.7 | 559948.1 | 64.07 | 63.9784
0 | 0 | 97 | 387.79 | 562982.5 | 64.49 | 64.3251
0 | 0 | 95.9 | 390.95 | 567578.1 | 64.9 | 64.8492
0 | 0 | 94.7 | 394.27 | 572392 | 65.4 | 65.3999
0 | 0 | 93.4 | 396.82 | 576096.7 | 65.91 | 65.8229
0 | 0 | 92.3 | 399.79 | 580415.5 | 66.37 | 66.3156
0 | 0 | 91 | 403.8 | 586229 | 66.89 | 66.9807
0 | 0 | 89.9 | 405.74 | 589048 | 67.42 | 67.3025
0 | 0 | 88.6 | 408.53 | 593103.1 | 67.83 | 67.7653
0 | 0 | 87.4 | 411 | 596679.8 | 68.26 | 68.175
0 | 0 | 86.2 | 414.05 | 601116.4 | 68.73 | 68.681
0 | 0 | 85 | 416.58 | 604790.8 | 69.14 | 69.1006
0 | 0 | 83.8 | 419.17 | 608547 | 69.65 | 69.5302
0 | 0 | 82.6 | 423.5 | 614829.6 | 70.23 | 70.2485
0 | 0 | 81.4 | 425.73 | 618065.8 | 70.69 | 70.6184
0 | 0 | 80.2 | 429.31 | 623266.3 | 71.3 | 71.2122
0 | 0 | 79 | 432.84 | 628396.2 | 71.79 | 71.7978
0 | 0 | 77.8 | 435.05 | 631601 | 72.24 | 72.1643
0 | 0 | 76.6 | 438.44 | 636522 | 72.84 | 72.7267
0 | 0 | 75.4 | 440.76 | 639889.9 | 73.28 | 73.1115
0 | 0 | 74.2 | 444.51 | 645328.9 | 73.74 | 73.7335
0 | 0 | 73 | 446.51 | 648231.1 | 74.19 | 74.0653
0 | 0 | 71.8 | 450.54 | 654087.1 | 74.77 | 74.7338
0 | 0 | 70.5 | 453.7 | 658676.6 | 75.32 | 75.2579
0 | 0 | 69.3 | 457.11 | 663630.7 | 75.74 | 75.8236
0 | 0 | 68.1 | 459.17 | 666611.1 | 76.31 | 76.1653
0 | 0 | 66.9 | 463.81 | 673356.4 | 76.87 | 76.9349
0 | 0 | 65.7 | 465.6 | 675945.5 | 77.44 | 77.2319
0 | 0 | 64.5 | 469.76 | 681988.5 | 77.98 | 77.9219
0 | 0 | 63.3 | 472.64 | 686168.5 | 78.51 | 78.3996
0 | 0 | 62.1 | 476.05 | 691125.1 | 79.01 | 78.9653
0 | 0 | 60.9 | 478.65 | 694898.1 | 79.53 | 79.3965
0 | 0 | 59.7 | 483.97 | 702625.7 | 80.06 | 80.279
0 | 0 | 58.4 | 484.93 | 704015 | 80.58 | 80.4382
0 | 0 | 57.2 | 488.16 | 708705.9 | 81.17 | 80.974
0 | 0 | 56.1 | 492.75 | 715360.9 | 81.68 | 81.7354
0 | 0 | 54.8 | 494.52 | 717941 | 82.22 | 82.029
0 | 0 | 53.6 | 498.67 | 723956.5 | 82.74 | 82.7174
0 | 0 | 52.4 | 501.95 | 728730.9 | 83.37 | 83.2615
0 | 0 | 51.2 | 504.82 | 732892.5 | 83.84 | 83.7375
0 | 0 | 50 | 508.67 | 738486.5 | 84.42 | 84.3761
0 | 0 | 48.8 | 511.99 | 743296.8 | 85.01 | 84.9268
0 | 0 | 47.6 | 516.07 | 749220.7 | 85.53 | 85.6036
0 | 0 | 46.4 | 517.84 | 751799.1 | 86.11 | 85.8972
0 | 0 | 45.2 | 522.29 | 758253.7 | 86.58 | 86.6354
0 | 0 | 43.9 | 524.7 | 761753.9 | 87.14 | 87.0351
0 | 0 | 42.7 | 528.39 | 767106.6 | 87.77 | 87.6472
0 | 0 | 41.6 | 531.7 | 771919.7 | 88.27 | 88.1963
0 | 0 | 40.3 | 534.95 | 776635 | 88.87 | 88.7354
0 | 0 | 39.1 | 539.29 | 782932.7 | 89.51 | 89.4553
0 | 0 | 37.9 | 541.25 | 785778.4 | 89.95 | 89.7804
0 | 0 | 36.7 | 545.04 | 791282.1 | 90.42 | 90.4091
0 | 0 | 35.5 | 548.76 | 796676.8 | 91.03 | 91.0261
0 | 0 | 34.3 | 552.58 | 802220.7 | 91.64 | 91.6598
0 | 0 | 33.1 | 555.23 | 806077.5 | 92.31 | 92.0993
0 | 0 | 31.9 | 558.68 | 811082.2 | 92.73 | 92.6716
0 | 0 | 30.7 | 561.76 | 815555.4 | 93.32 | 93.1825
0 | 0 | 29.5 | 565.22 | 820571.2 | 93.88 | 93.7564
0 | 0 | 28.3 | 568.99 | 826047.5 | 94.51 | 94.3818
0 | 0 | 27 | 573.36 | 832395 | 95.14 | 95.1067
0 | 0 | 25.8 | 575.93 | 836120.4 | 95.55 | 95.533
0 | 0 | 24.6 | 578.07 | 839235.9 | 96.12 | 95.8879
0 | 0 | 23.4 | 581.22 | 843809.6 | 96.63 | 96.4104
0 | 0 | 22.2 | 585.29 | 849716.7 | 97.16 | 97.0856
0 | 0 | 21 | 589.94 | 856462.4 | 97.7 | 97.8569
0 | 0 | 19.8 | 590.36 | 857074.8 | 98.2 | 97.9266
0 | 0 | 18.6 | 594.68 | 863355 | 98.66 | 98.6431
0 | 0 | 17.4 | 596.07 | 865368.5 | 99.12 | 98.8737
0 | 0 | 16.2 | 598.77 | 869288.2 | 99.52 | 99.3216
0 | 0 | 15 | 601.69 | 873517.7 | 99.8 | 99.8059
0 | 0 | 13.8 | 602.3 | 874405.4 | 100 | 99.9071
0 | 0 | 12.6 | 602.86 | 875227.2 | 99.92 | 100
0 | 0 | 11.4 | 600.07 | 871167.4 | 99.76 | 99.5372
0 | 0 | 10.2 | 597.5 | 867437.9 | 99.3 | 99.1109
0 | 0 | 8.9 | 591.63 | 858914 | 98.35 | 98.1372
0 | 0 | 7.7 | 582.52 | 845697.6 | 96.73 | 96.6261
0 | 0 | 6.5 | 569.69 | 827073.8 | 94.58 | 94.4979
0 | 0 | 5.3 | 551.11 | 800087.8 | 91.3 | 91.4159
0 | 0 | 4.1 | 518.67 | 752991 | 86.28 | 86.0349
0 | 0 | 2.9 | 474.03 | 688188.8 | 78.79 | 78.6302
0 | 0 | 1.7 | 418.33 | 607320.3 | 69.51 | 69.3909
0 | 0 | 0.5 | 379.58 | 551070.8 | 63.02 | 62.9632
0 | 0 | -0.5 | 359.95 | 522565.5 | 59.74 | 59.7071
Measurement time: | 02/08/2022 20:03:25
Radiation device: | Novalis TB Accelerator
Energy: | 10 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 49.29 | 71552.7 | 27.58 | 28.4272
0 | 0 | 309.1 | 48.4 | 70263.2 | 27.91 | 27.914
0 | 0 | 307.6 | 48.74 | 70762.8 | 28.13 | 28.11
0 | 0 | 306.6 | 49.05 | 71204.4 | 28.35 | 28.2888
0 | 0 | 305.3 | 49.3 | 71569 | 28.42 | 28.433
0 | 0 | 304.1 | 49.59 | 71996.4 | 28.56 | 28.6003
0 | 0 | 302.9 | 49.95 | 72513.1 | 28.8 | 28.8079
0 | 0 | 301.7 | 50.24 | 72933.7 | 28.92 | 28.9751
0 | 0 | 300.5 | 50.38 | 73137.4 | 29.03 | 29.0559
0 | 0 | 299.3 | 50.68 | 73576.9 | 29.25 | 29.2289
0 | 0 | 298.1 | 50.95 | 73974.9 | 29.34 | 29.3846
0 | 0 | 296.9 | 51.27 | 74427.4 | 29.57 | 29.5692
0 | 0 | 295.7 | 51.51 | 74783 | 29.72 | 29.7076
0 | 0 | 294.5 | 51.81 | 75220.4 | 29.86 | 29.8806
0 | 0 | 293.3 | 52.08 | 75612.5 | 30.02 | 30.0363
0 | 0 | 292.1 | 52.4 | 76079.7 | 30.14 | 30.2209
0 | 0 | 290.9 | 52.55 | 76295.2 | 30.31 | 30.3074
0 | 0 | 289.7 | 52.99 | 76922.9 | 30.55 | 30.5612
0 | 0 | 288.5 | 53.24 | 77295.2 | 30.66 | 30.7053
0 | 0 | 287.3 | 53.56 | 77754.1 | 30.86 | 30.8899
0 | 0 | 286.1 | 53.82 | 78140 | 31.01 | 31.0399
0 | 0 | 284.9 | 54.13 | 78587.9 | 31.22 | 31.2186
0 | 0 | 283.7 | 54.53 | 79172.7 | 31.41 | 31.4493
0 | 0 | 282.5 | 54.65 | 79343.2 | 31.52 | 31.5185
0 | 0 | 281.3 | 54.9 | 79698.8 | 31.67 | 31.6627
0 | 0 | 280.1 | 55.34 | 80347.3 | 31.98 | 31.9165
0 | 0 | 278.9 | 55.64 | 80783.8 | 32.07 | 32.0895
0 | 0 | 277.6 | 55.96 | 81238.7 | 32.27 | 32.2741
0 | 0 | 276.4 | 56.3 | 81729.3 | 32.43 | 32.4702
0 | 0 | 275.3 | 56.6 | 82167.2 | 32.63 | 32.6432
0 | 0 | 274.1 | 56.86 | 82552 | 32.75 | 32.7931
0 | 0 | 272.8 | 57.29 | 83166.9 | 33 | 33.0411
0 | 0 | 271.6 | 57.44 | 83393.1 | 33.11 | 33.1276
0 | 0 | 270.4 | 57.98 | 84170.3 | 33.42 | 33.4391
0 | 0 | 269.2 | 58.22 | 84519 | 33.51 | 33.5775
0 | 0 | 268 | 58.49 | 84918.5 | 33.78 | 33.7332
0 | 0 | 266.8 | 58.76 | 85303.4 | 33.85 | 33.8889
0 | 0 | 265.6 | 59.19 | 85930.4 | 34.13 | 34.1369
0 | 0 | 264.4 | 59.58 | 86490.9 | 34.3 | 34.3618
0 | 0 | 263.2 | 59.83 | 86864.9 | 34.46 | 34.506
0 | 0 | 262 | 60.26 | 87479.4 | 34.72 | 34.754
0 | 0 | 260.8 | 60.57 | 87932.4 | 34.89 | 34.9328
0 | 0 | 259.6 | 60.86 | 88357.4 | 35.04 | 35.1001
0 | 0 | 258.4 | 61.28 | 88966.2 | 35.34 | 35.3423
0 | 0 | 257.2 | 61.48 | 89249 | 35.49 | 35.4576
0 | 0 | 256 | 61.85 | 89798.3 | 35.69 | 35.671
0 | 0 | 254.8 | 62.21 | 90320.2 | 35.9 | 35.8787
0 | 0 | 253.6 | 62.6 | 90888 | 36.06 | 36.1036
0 | 0 | 252.4 | 62.86 | 91255.7 | 36.26 | 36.2535
0 | 0 | 251.2 | 63.3 | 91902 | 36.46 | 36.5073
0 | 0 | 250 | 63.74 | 92542.2 | 36.73 | 36.7611
0 | 0 | 248.7 | 63.9 | 92775.8 | 36.89 | 36.8533
0 | 0 | 247.5 | 64.32 | 93379.8 | 37.05 | 37.0956
0 | 0 | 246.3 | 64.67 | 93888.3 | 37.27 | 37.2974
0 | 0 | 245.1 | 64.98 | 94335.6 | 37.42 | 37.4762
0 | 0 | 243.9 | 65.36 | 94890.9 | 37.65 | 37.6954
0 | 0 | 242.7 | 65.83 | 95574.8 | 37.92 | 37.9664
0 | 0 | 241.5 | 66.09 | 95955 | 38.23 | 38.1164
0 | 0 | 240.3 | 66.32 | 96275.3 | 38.25 | 38.249
0 | 0 | 239.1 | 66.91 | 97143.9 | 38.6 | 38.5893
0 | 0 | 237.9 | 67.19 | 97538.9 | 38.78 | 38.7508
0 | 0 | 236.7 | 67.65 | 98217.7 | 38.94 | 39.0161
0 | 0 | 235.5 | 68.04 | 98781.3 | 39.22 | 39.241
0 | 0 | 234.3 | 68.41 | 99319.7 | 39.41 | 39.4544
0 | 0 | 233.1 | 68.88 | 99995 | 39.69 | 39.7255
0 | 0 | 231.9 | 69.09 | 100310.9 | 39.79 | 39.8466
0 | 0 | 230.7 | 69.52 | 100924.2 | 40 | 40.0946
0 | 0 | 229.5 | 69.91 | 101501 | 40.35 | 40.3195
0 | 0 | 228.3 | 70.32 | 102094.3 | 40.46 | 40.556
0 | 0 | 227.1 | 70.63 | 102538.7 | 40.69 | 40.7348
0 | 0 | 225.9 | 71.23 | 103410.2 | 41.05 | 41.0808
0 | 0 | 224.6 | 71.69 | 104083.2 | 41.29 | 41.3461
0 | 0 | 223.4 | 71.84 | 104300.4 | 41.38 | 41.4326
0 | 0 | 222.2 | 72.3 | 104966.5 | 41.64 | 41.6979
0 | 0 | 221 | 72.61 | 105416.1 | 41.85 | 41.8767
0 | 0 | 219.8 | 73.19 | 106260.6 | 42.14 | 42.2112
0 | 0 | 218.6 | 73.45 | 106640.2 | 42.38 | 42.3612
0 | 0 | 217.4 | 74.04 | 107495.2 | 42.65 | 42.7014
0 | 0 | 216.2 | 74.34 | 107932 | 42.83 | 42.8744
0 | 0 | 215 | 74.84 | 108649.1 | 43.09 | 43.1628
0 | 0 | 213.8 | 75.24 | 109227.3 | 43.44 | 43.3935
0 | 0 | 212.6 | 75.59 | 109734.3 | 43.6 | 43.5954
0 | 0 | 211.4 | 76.13 | 110517.5 | 43.93 | 43.9068
0 | 0 | 210.2 | 76.45 | 110984 | 44.03 | 44.0914
0 | 0 | 209 | 76.98 | 111753.8 | 44.34 | 44.397
0 | 0 | 207.8 | 77.34 | 112280.7 | 44.54 | 44.6046
0 | 0 | 206.6 | 77.81 | 112962.9 | 44.85 | 44.8757
0 | 0 | 205.4 | 78.16 | 113468.3 | 45.1 | 45.0776
0 | 0 | 204.2 | 78.61 | 114130.8 | 45.3 | 45.3371
0 | 0 | 203 | 79.11 | 114848.4 | 45.67 | 45.6255
0 | 0 | 201.8 | 79.45 | 115344.1 | 45.87 | 45.8216
0 | 0 | 200.6 | 80.06 | 116234.6 | 46.15 | 46.1734
0 | 0 | 199.4 | 80.45 | 116791.4 | 46.34 | 46.3983
0 | 0 | 198.1 | 80.91 | 117459.3 | 46.61 | 46.6636
0 | 0 | 197 | 81.38 | 118152.5 | 46.92 | 46.9347
0 | 0 | 195.7 | 81.86 | 118837.7 | 47.09 | 47.2115
0 | 0 | 194.6 | 82.38 | 119597.5 | 47.44 | 47.5114
0 | 0 | 193.4 | 82.65 | 119986.9 | 47.59 | 47.6671
0 | 0 | 192.1 | 83.29 | 120923.2 | 48 | 48.0362
0 | 0 | 191 | 83.58 | 121340.6 | 48.18 | 48.2035
0 | 0 | 189.7 | 83.97 | 121909.8 | 48.42 | 48.4284
0 | 0 | 188.5 | 84.53 | 122719 | 48.76 | 48.7514
0 | 0 | 187.3 | 85.16 | 123639.6 | 49.04 | 49.1147
0 | 0 | 186.1 | 85.45 | 124060.2 | 49.29 | 49.282
0 | 0 | 184.9 | 86.13 | 125039.1 | 49.7 | 49.6741
0 | 0 | 183.7 | 86.59 | 125709.2 | 49.8 | 49.9394
0 | 0 | 182.5 | 87.07 | 126404.3 | 50.21 | 50.2163
0 | 0 | 181.3 | 87.43 | 126936.7 | 50.4 | 50.4239
0 | 0 | 180.1 | 88.01 | 127772.5 | 50.65 | 50.7584
0 | 0 | 178.9 | 88.54 | 128540.3 | 51.09 | 51.0641
0 | 0 | 177.7 | 89.01 | 129221.2 | 51.32 | 51.3351
0 | 0 | 176.5 | 89.37 | 129740.7 | 51.58 | 51.5428
0 | 0 | 175.3 | 89.94 | 130568.3 | 51.85 | 51.8715
0 | 0 | 174.1 | 90.38 | 131212.8 | 52.1 | 52.1253
0 | 0 | 172.9 | 90.98 | 132085.8 | 52.44 | 52.4713
0 | 0 | 171.7 | 91.56 | 132926 | 52.71 | 52.8058
0 | 0 | 170.5 | 92.11 | 133716.8 | 53.17 | 53.123
0 | 0 | 169.3 | 92.46 | 134227.2 | 53.29 | 53.3249
0 | 0 | 168.1 | 93.12 | 135196.4 | 53.65 | 53.7055
0 | 0 | 166.9 | 93.52 | 135772.3 | 53.91 | 53.9362
0 | 0 | 165.6 | 94.18 | 136736.4 | 54.29 | 54.3169
0 | 0 | 164.5 | 94.74 | 137539.1 | 54.63 | 54.6398
0 | 0 | 163.2 | 95.1 | 138071.5 | 54.72 | 54.8475
0 | 0 | 162 | 95.69 | 138918.2 | 55.14 | 55.1877
0 | 0 | 160.8 | 96.26 | 139754.2 | 55.52 | 55.5165
0 | 0 | 159.6 | 96.77 | 140495.6 | 55.73 | 55.8106
0 | 0 | 158.4 | 97.32 | 141292.7 | 56.12 | 56.1278
0 | 0 | 157.2 | 97.92 | 142157.9 | 56.44 | 56.4738
0 | 0 | 156 | 98.32 | 142741.7 | 56.62 | 56.7045
0 | 0 | 154.8 | 98.9 | 143580.5 | 57.04 | 57.039
0 | 0 | 153.6 | 99.6 | 144600.4 | 57.4 | 57.4428
0 | 0 | 152.4 | 99.93 | 145082.4 | 57.55 | 57.6331
0 | 0 | 151.2 | 100.59 | 146030.6 | 57.85 | 58.0137
0 | 0 | 150 | 101.17 | 146880.3 | 58.32 | 58.3482
0 | 0 | 148.8 | 101.77 | 147751.5 | 58.63 | 58.6943
0 | 0 | 147.6 | 102.3 | 148513.3 | 59.02 | 58.9999
0 | 0 | 146.4 | 102.77 | 149206.9 | 59.23 | 59.271
0 | 0 | 145.2 | 103.42 | 150147.6 | 59.53 | 59.6459
0 | 0 | 144 | 104.11 | 151142.6 | 59.95 | 60.0438
0 | 0 | 142.8 | 104.57 | 151817.7 | 60.17 | 60.3091
0 | 0 | 141.6 | 105.27 | 152834.2 | 60.65 | 60.7128
0 | 0 | 140.4 | 105.8 | 153598.7 | 60.96 | 61.0185
0 | 0 | 139.2 | 106.52 | 154638.3 | 61.35 | 61.4338
0 | 0 | 138 | 107.09 | 155465 | 61.63 | 61.7625
0 | 0 | 136.8 | 107.49 | 156049 | 61.99 | 61.9932
0 | 0 | 135.5 | 108.3 | 157222.9 | 62.43 | 62.4603
0 | 0 | 134.4 | 108.73 | 157851.7 | 62.78 | 62.7083
0 | 0 | 133.1 | 109.41 | 158841.6 | 63.17 | 63.1005
0 | 0 | 131.9 | 110.06 | 159787.1 | 63.48 | 63.4754
0 | 0 | 130.7 | 110.61 | 160589.2 | 63.68 | 63.7926
0 | 0 | 129.5 | 111.36 | 161670.9 | 64.17 | 64.2252
0 | 0 | 128.3 | 111.87 | 162418.1 | 64.59 | 64.5193
0 | 0 | 127.1 | 112.61 | 163486.9 | 64.86 | 64.9461
0 | 0 | 125.9 | 113.11 | 164216.1 | 65.18 | 65.2344
0 | 0 | 124.7 | 113.71 | 165089.5 | 65.65 | 65.5805
0 | 0 | 123.5 | 114.36 | 166027.1 | 65.89 | 65.9554
0 | 0 | 122.3 | 115.12 | 167133.2 | 66.27 | 66.3937
0 | 0 | 121.1 | 115.68 | 167947.5 | 66.64 | 66.7167
0 | 0 | 119.9 | 116.26 | 168790.1 | 66.97 | 67.0512
0 | 0 | 118.7 | 117.01 | 169876.7 | 67.41 | 67.4837
0 | 0 | 117.5 | 117.63 | 170775.5 | 67.7 | 67.8413
0 | 0 | 116.3 | 118.27 | 171704.4 | 68.01 | 68.2104
0 | 0 | 115.1 | 119.01 | 172779.6 | 68.47 | 68.6372
0 | 0 | 113.9 | 119.59 | 173625.9 | 68.89 | 68.9717
0 | 0 | 112.7 | 120.37 | 174755.2 | 69.22 | 69.4215
0 | 0 | 111.4 | 120.87 | 175480.5 | 69.62 | 69.7099
0 | 0 | 110.3 | 121.46 | 176332.6 | 69.96 | 70.0502
0 | 0 | 109 | 122.1 | 177269 | 70.33 | 70.4193
0 | 0 | 107.8 | 122.8 | 178273.4 | 70.88 | 70.823
0 | 0 | 106.7 | 123.54 | 179353.7 | 71.07 | 71.2498
0 | 0 | 105.4 | 124.25 | 180385.6 | 71.6 | 71.6593
0 | 0 | 104.3 | 125.04 | 181533.4 | 71.95 | 72.1149
0 | 0 | 103 | 125.76 | 182576.7 | 72.47 | 72.5301
0 | 0 | 101.8 | 126.23 | 183258.2 | 72.83 | 72.8012
0 | 0 | 100.6 | 126.93 | 184278.3 | 73.3 | 73.2049
0 | 0 | 99.4 | 127.51 | 185110.9 | 73.44 | 73.5394
0 | 0 | 98.2 | 128.19 | 186100.5 | 73.81 | 73.9316
0 | 0 | 97 | 129.17 | 187532 | 74.43 | 74.4968
0 | 0 | 95.8 | 129.74 | 188351.4 | 74.83 | 74.8255
0 | 0 | 94.6 | 130.35 | 189247.2 | 75.04 | 75.1773
0 | 0 | 93.4 | 131.2 | 190474.1 | 75.56 | 75.6676
0 | 0 | 92.2 | 131.97 | 191587.2 | 76.03 | 76.1117
0 | 0 | 91 | 132.71 | 192661.1 | 76.38 | 76.5384
0 | 0 | 89.8 | 133.32 | 193546 | 76.91 | 76.8902
0 | 0 | 88.6 | 134.1 | 194685.9 | 77.18 | 77.3401
0 | 0 | 87.4 | 134.78 | 195676.2 | 77.54 | 77.7323
0 | 0 | 86.2 | 135.39 | 196552.7 | 78.19 | 78.0841
0 | 0 | 85 | 136.36 | 197966.7 | 78.55 | 78.6435
0 | 0 | 83.8 | 137.04 | 198945.8 | 78.81 | 79.0357
0 | 0 | 82.6 | 137.96 | 200286.9 | 79.54 | 79.5663
0 | 0 | 81.4 | 138.52 | 201102 | 79.86 | 79.8893
0 | 0 | 80.2 | 139.13 | 201983.2 | 80.19 | 80.2411
0 | 0 | 78.9 | 140.09 | 203373.9 | 80.62 | 80.7947
0 | 0 | 77.7 | 140.78 | 204380 | 81.04 | 81.1927
0 | 0 | 76.5 | 141.38 | 205255.1 | 81.44 | 81.5387
0 | 0 | 75.3 | 142.2 | 206450.6 | 82.06 | 82.0117
0 | 0 | 74.1 | 143.01 | 207623.9 | 82.44 | 82.4788
0 | 0 | 72.9 | 143.67 | 208575.7 | 82.75 | 82.8594
0 | 0 | 71.7 | 144.46 | 209731.5 | 83.24 | 83.3151
0 | 0 | 70.5 | 145.18 | 210763.6 | 83.59 | 83.7303
0 | 0 | 69.3 | 145.85 | 211741.5 | 84.04 | 84.1167
0 | 0 | 68.1 | 146.54 | 212751.6 | 84.45 | 84.5147
0 | 0 | 66.9 | 147.48 | 214113.3 | 84.99 | 85.0568
0 | 0 | 65.7 | 148.2 | 215156.2 | 85.34 | 85.4721
0 | 0 | 64.5 | 149.16 | 216547.8 | 86.13 | 86.0257
0 | 0 | 63.3 | 149.88 | 217593.4 | 86.35 | 86.441
0 | 0 | 62.1 | 150.44 | 218405.1 | 86.84 | 86.7639
0 | 0 | 60.8 | 151.5 | 219941.2 | 87.29 | 87.3753
0 | 0 | 59.7 | 152.06 | 220762 | 87.65 | 87.6983
0 | 0 | 58.4 | 153.03 | 222167.9 | 88.16 | 88.2577
0 | 0 | 57.2 | 153.61 | 223004.1 | 88.52 | 88.5922
0 | 0 | 56 | 154.47 | 224262.7 | 88.91 | 89.0882
0 | 0 | 54.8 | 155.18 | 225289.3 | 89.32 | 89.4977
0 | 0 | 53.6 | 156.17 | 226728.4 | 89.97 | 90.0686
0 | 0 | 52.4 | 157.01 | 227947.7 | 90.52 | 90.5531
0 | 0 | 51.2 | 157.64 | 228865.1 | 90.79 | 90.9164
0 | 0 | 50 | 158.57 | 230212.7 | 91.41 | 91.4528
0 | 0 | 48.8 | 159.59 | 231688 | 91.95 | 92.0411
0 | 0 | 47.6 | 160.03 | 232335.3 | 92.22 | 92.2948
0 | 0 | 46.4 | 160.93 | 233640.4 | 92.87 | 92.8139
0 | 0 | 45.1 | 161.79 | 234891.6 | 93.22 | 93.3099
0 | 0 | 43.9 | 162.64 | 236111.5 | 93.66 | 93.8001
0 | 0 | 42.7 | 163.55 | 237446.4 | 94.29 | 94.3249
0 | 0 | 41.5 | 164.19 | 238370.9 | 94.69 | 94.694
0 | 0 | 40.3 | 165.2 | 239832.1 | 95.18 | 95.2765
0 | 0 | 39.1 | 166.15 | 241216 | 95.7 | 95.8244
0 | 0 | 37.9 | 166.51 | 241734.3 | 95.84 | 96.0321
0 | 0 | 36.7 | 167.56 | 243258.7 | 96.4 | 96.6376
0 | 0 | 35.5 | 168.14 | 244104.4 | 96.87 | 96.9721
0 | 0 | 34.2 | 169.2 | 245641.2 | 97.45 | 97.5835
0 | 0 | 33.1 | 169.77 | 246469 | 97.79 | 97.9122
0 | 0 | 31.9 | 170.13 | 246988.9 | 98.11 | 98.1198
0 | 0 | 30.6 | 171.08 | 248364.5 | 98.52 | 98.6677
0 | 0 | 29.4 | 171.54 | 249035.2 | 98.79 | 98.933
0 | 0 | 28.2 | 172.22 | 250028.7 | 99.12 | 99.3252
0 | 0 | 27 | 172.69 | 250715 | 99.48 | 99.5963
0 | 0 | 25.8 | 172.92 | 251039.5 | 99.45 | 99.7289
0 | 0 | 24.6 | 173.27 | 251545.2 | 99.75 | 99.9308
0 | 0 | 23.4 | 173.25 | 251529.1 | 100 | 99.9193
0 | 0 | 22.2 | 173.39 | 251730.5 | 99.86 | 100
0 | 0 | 21 | 173.24 | 251511.7 | 99.68 | 99.9135
0 | 0 | 19.8 | 172.87 | 250965.2 | 99.58 | 99.7001
0 | 0 | 18.6 | 172.03 | 249754.7 | 99.14 | 99.2156
0 | 0 | 17.4 | 171.16 | 248482.3 | 98.58 | 98.7139
0 | 0 | 16.2 | 169.65 | 246290.4 | 97.75 | 97.843
0 | 0 | 15 | 167.95 | 243830.2 | 96.62 | 96.8626
0 | 0 | 13.8 | 165.5 | 240277.7 | 95.35 | 95.4496
0 | 0 | 12.6 | 162.65 | 236132.2 | 93.69 | 93.8059
0 | 0 | 11.4 | 159.1 | 230978.3 | 91.57 | 91.7585
0 | 0 | 10.1 | 154.45 | 224235.5 | 88.89 | 89.0766
0 | 0 | 8.9 | 148.87 | 216132.4 | 85.66 | 85.8585
0 | 0 | 7.7 | 142.11 | 206307.2 | 82 | 81.9597
0 | 0 | 6.5 | 133.31 | 193540.5 | 76.92 | 76.8845
0 | 0 | 5.3 | 123.03 | 178607.1 | 70.92 | 70.9556
0 | 0 | 4.1 | 109.63 | 159165.1 | 63.09 | 63.2274
0 | 0 | 2.9 | 93.2 | 135307.1 | 53.76 | 53.7517
0 | 0 | 1.7 | 76.35 | 110850.9 | 43.95 | 44.0337
0 | 0 | 0.5 | 65.45 | 95023.8 | 37.6 | 37.7473
0 | 0 | -0.5 | 60.46 | 87773.4 | 34.81 | 34.8694
Measurement time: | 02/08/2022 20:05:50
Radiation device: | Novalis TB Accelerator
Energy: | 10 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 73.05 | 106051.8 | 28.13 | 28.039
0 | 0 | 309.1 | 73 | 105985.3 | 28.02 | 28.0198
0 | 0 | 307.6 | 73.64 | 106905.8 | 28.26 | 28.2655
0 | 0 | 306.5 | 73.91 | 107305.9 | 28.37 | 28.3691
0 | 0 | 305.3 | 74.38 | 107989.8 | 28.58 | 28.5495
0 | 0 | 304.1 | 74.73 | 108496.7 | 28.7 | 28.6838
0 | 0 | 302.9 | 75.22 | 109198.1 | 28.9 | 28.8719
0 | 0 | 301.7 | 75.66 | 109835.7 | 29.04 | 29.0408
0 | 0 | 300.5 | 75.91 | 110210.7 | 29.16 | 29.1368
0 | 0 | 299.3 | 76.44 | 110970.4 | 29.34 | 29.3402
0 | 0 | 298.1 | 76.97 | 111748.6 | 29.56 | 29.5436
0 | 0 | 296.9 | 77.34 | 112277.9 | 29.69 | 29.6856
0 | 0 | 295.7 | 77.68 | 112773.6 | 29.81 | 29.8161
0 | 0 | 294.5 | 78.01 | 113257.1 | 29.96 | 29.9428
0 | 0 | 293.3 | 78.39 | 113810.6 | 30.1 | 30.0887
0 | 0 | 292.1 | 78.87 | 114505.6 | 30.26 | 30.2729
0 | 0 | 290.9 | 79.39 | 115260 | 30.5 | 30.4725
0 | 0 | 289.7 | 79.81 | 115863.8 | 30.66 | 30.6337
0 | 0 | 288.5 | 80.19 | 116421.4 | 30.8 | 30.7796
0 | 0 | 287.3 | 80.79 | 117296.8 | 31.04 | 31.0099
0 | 0 | 286.1 | 81.15 | 117817.9 | 31.21 | 31.148
0 | 0 | 284.9 | 81.64 | 118520.3 | 31.34 | 31.3361
0 | 0 | 283.7 | 82.11 | 119212.1 | 31.53 | 31.5165
0 | 0 | 282.4 | 82.61 | 119931.3 | 31.74 | 31.7084
0 | 0 | 281.3 | 83.08 | 120621 | 31.9 | 31.8888
0 | 0 | 280.1 | 83.45 | 121155.5 | 32.01 | 32.0309
0 | 0 | 278.9 | 83.89 | 121797.1 | 32.22 | 32.1997
0 | 0 | 277.7 | 84.43 | 122569.1 | 32.39 | 32.407
0 | 0 | 276.5 | 84.78 | 123088.2 | 32.54 | 32.5414
0 | 0 | 275.3 | 85.22 | 123716.8 | 32.68 | 32.7102
0 | 0 | 274 | 85.66 | 124360.8 | 32.93 | 32.8791
0 | 0 | 272.9 | 86.26 | 125225.2 | 33.11 | 33.1094
0 | 0 | 271.7 | 86.76 | 125952.1 | 33.32 | 33.3013
0 | 0 | 270.5 | 87.25 | 126675 | 33.48 | 33.4894
0 | 0 | 269.3 | 87.66 | 127268.8 | 33.68 | 33.6468
0 | 0 | 268 | 88.07 | 127852.5 | 33.78 | 33.8042
0 | 0 | 266.8 | 88.62 | 128658.6 | 34 | 34.0153
0 | 0 | 265.7 | 89.12 | 129377.9 | 34.23 | 34.2072
0 | 0 | 264.4 | 89.58 | 130049.6 | 34.42 | 34.3838
0 | 0 | 263.2 | 90.17 | 130913.9 | 34.61 | 34.6102
0 | 0 | 262 | 90.64 | 131587.2 | 34.84 | 34.7906
0 | 0 | 260.9 | 91.16 | 132338.3 | 34.99 | 34.9902
0 | 0 | 259.6 | 91.63 | 133028.3 | 35.18 | 35.1706
0 | 0 | 258.4 | 92.08 | 133676.3 | 35.32 | 35.3433
0 | 0 | 257.2 | 92.58 | 134399.5 | 35.5 | 35.5353
0 | 0 | 256 | 93.01 | 135036.5 | 35.71 | 35.7003
0 | 0 | 254.8 | 93.74 | 136091.3 | 36 | 35.9805
0 | 0 | 253.6 | 94.13 | 136655.6 | 36.12 | 36.1302
0 | 0 | 252.4 | 94.62 | 137362.7 | 36.31 | 36.3183
0 | 0 | 251.2 | 95.18 | 138185 | 36.52 | 36.5332
0 | 0 | 250 | 95.69 | 138926 | 36.74 | 36.729
0 | 0 | 248.8 | 96.13 | 139559.2 | 36.88 | 36.8979
0 | 0 | 247.6 | 96.82 | 140568.3 | 37.16 | 37.1627
0 | 0 | 246.4 | 97.31 | 141280.1 | 37.36 | 37.3508
0 | 0 | 245.2 | 98 | 142272.3 | 37.68 | 37.6156
0 | 0 | 244 | 98.45 | 142934.5 | 37.81 | 37.7884
0 | 0 | 242.8 | 99 | 143722.6 | 37.98 | 37.9995
0 | 0 | 241.6 | 99.54 | 144509.7 | 38.23 | 38.2067
0 | 0 | 240.4 | 99.93 | 145075.9 | 38.36 | 38.3564
0 | 0 | 239.2 | 100.71 | 146206.3 | 38.64 | 38.6558
0 | 0 | 238 | 101.22 | 146954.4 | 38.86 | 38.8516
0 | 0 | 236.8 | 101.76 | 147726.8 | 39.05 | 39.0588
0 | 0 | 235.5 | 102.46 | 148746.1 | 39.33 | 39.3275
0 | 0 | 234.4 | 102.99 | 149513.6 | 39.54 | 39.531
0 | 0 | 233.1 | 103.42 | 150148 | 39.71 | 39.696
0 | 0 | 231.9 | 104.13 | 151175.6 | 39.94 | 39.9685
0 | 0 | 230.8 | 104.71 | 152016.6 | 40.22 | 40.1911
0 | 0 | 229.5 | 105.22 | 152754.7 | 40.44 | 40.3869
0 | 0 | 228.3 | 105.87 | 153698.2 | 40.69 | 40.6364
0 | 0 | 227.1 | 106.48 | 154587.8 | 40.84 | 40.8705
0 | 0 | 225.9 | 107.11 | 155493.7 | 41.06 | 41.1123
0 | 0 | 224.7 | 107.64 | 156267.2 | 41.32 | 41.3158
0 | 0 | 223.5 | 108.21 | 157103.6 | 41.54 | 41.5346
0 | 0 | 222.3 | 108.91 | 158117.2 | 41.81 | 41.8032
0 | 0 | 221.1 | 109.37 | 158786.7 | 41.92 | 41.9798
0 | 0 | 219.9 | 110.06 | 159789.7 | 42.22 | 42.2447
0 | 0 | 218.7 | 110.7 | 160716.3 | 42.49 | 42.4903
0 | 0 | 217.5 | 111.18 | 161415.2 | 42.62 | 42.6745
0 | 0 | 216.3 | 111.76 | 162246.7 | 42.86 | 42.8972
0 | 0 | 215.1 | 112.54 | 163380.8 | 43.27 | 43.1966
0 | 0 | 213.9 | 113.13 | 164244.2 | 43.38 | 43.423
0 | 0 | 212.7 | 113.87 | 165321.9 | 43.65 | 43.7071
0 | 0 | 211.5 | 114.44 | 166137.8 | 43.9 | 43.9258
0 | 0 | 210.3 | 114.99 | 166944.7 | 44.18 | 44.137
0 | 0 | 209.1 | 115.77 | 168067.1 | 44.49 | 44.4363
0 | 0 | 207.9 | 116.5 | 169139.4 | 44.72 | 44.7165
0 | 0 | 206.7 | 117.01 | 169872.5 | 44.93 | 44.9123
0 | 0 | 205.4 | 117.68 | 170844.7 | 45.19 | 45.1695
0 | 0 | 204.3 | 118.32 | 171772.9 | 45.49 | 45.4151
0 | 0 | 203.1 | 118.83 | 172510.9 | 45.65 | 45.6109
0 | 0 | 201.8 | 119.59 | 173618.1 | 45.96 | 45.9026
0 | 0 | 200.6 | 120.25 | 174574.2 | 46.23 | 46.1559
0 | 0 | 199.5 | 121.08 | 175776.9 | 46.49 | 46.4745
0 | 0 | 198.3 | 121.64 | 176593.1 | 46.73 | 46.6894
0 | 0 | 197.1 | 122.26 | 177497.9 | 46.97 | 46.9274
0 | 0 | 195.8 | 123.11 | 178729.1 | 47.32 | 47.2537
0 | 0 | 194.7 | 123.7 | 179587.7 | 47.54 | 47.4801
0 | 0 | 193.4 | 124.5 | 180749.5 | 47.84 | 47.7872
0 | 0 | 192.2 | 124.95 | 181402.7 | 47.99 | 47.9599
0 | 0 | 191 | 125.81 | 182646.9 | 48.35 | 48.29
0 | 0 | 189.8 | 126.67 | 183898.1 | 48.66 | 48.6201
0 | 0 | 188.6 | 127.31 | 184832.9 | 48.91 | 48.8658
0 | 0 | 187.4 | 127.92 | 185715.6 | 49.12 | 49.0999
0 | 0 | 186.2 | 128.65 | 186769.6 | 49.36 | 49.3801
0 | 0 | 185.1 | 129.44 | 187922.7 | 49.69 | 49.6833
0 | 0 | 183.8 | 130.07 | 188834.6 | 49.98 | 49.9252
0 | 0 | 182.6 | 130.88 | 190005.1 | 50.29 | 50.2361
0 | 0 | 181.4 | 131.36 | 190707.1 | 50.46 | 50.4203
0 | 0 | 180.2 | 132.55 | 192429.9 | 50.85 | 50.8771
0 | 0 | 179 | 132.87 | 192891.5 | 51.04 | 50.9999
0 | 0 | 177.8 | 133.8 | 194251.8 | 51.39 | 51.3568
0 | 0 | 176.6 | 134.34 | 195035.2 | 51.59 | 51.5641
0 | 0 | 175.4 | 135.37 | 196533.1 | 51.99 | 51.9595
0 | 0 | 174.2 | 135.96 | 197387.6 | 52.2 | 52.1859
0 | 0 | 173 | 136.83 | 198651.2 | 52.52 | 52.5199
0 | 0 | 171.8 | 137.66 | 199846.2 | 52.91 | 52.8384
0 | 0 | 170.6 | 138.36 | 200870.2 | 53.13 | 53.1071
0 | 0 | 169.4 | 139.02 | 201827.8 | 53.38 | 53.3605
0 | 0 | 168.2 | 139.73 | 202853 | 53.68 | 53.633
0 | 0 | 167 | 140.71 | 204280.8 | 54.02 | 54.0091
0 | 0 | 165.8 | 141.61 | 205590.2 | 54.36 | 54.3546
0 | 0 | 164.6 | 142.44 | 206786.7 | 54.62 | 54.6732
0 | 0 | 163.4 | 143.02 | 207637.3 | 54.91 | 54.8958
0 | 0 | 162.2 | 143.79 | 208751.5 | 55.22 | 55.1913
0 | 0 | 161 | 144.72 | 210108.1 | 55.6 | 55.5483
0 | 0 | 159.7 | 145.38 | 211058.7 | 55.79 | 55.8016
0 | 0 | 158.6 | 146.31 | 212405.8 | 56.14 | 56.1586
0 | 0 | 157.3 | 147.14 | 213611.1 | 56.45 | 56.4772
0 | 0 | 156.1 | 147.8 | 214580.6 | 56.81 | 56.7305
0 | 0 | 154.9 | 148.72 | 215914.9 | 57.1 | 57.0836
0 | 0 | 153.7 | 149.64 | 217242 | 57.51 | 57.4368
0 | 0 | 152.5 | 150.24 | 218121.1 | 57.73 | 57.6671
0 | 0 | 151.3 | 151.29 | 219633.5 | 58.06 | 58.0701
0 | 0 | 150.1 | 152.11 | 220828.1 | 58.41 | 58.3848
0 | 0 | 148.9 | 152.96 | 222060.2 | 58.66 | 58.7111
0 | 0 | 147.7 | 153.88 | 223402 | 59.06 | 59.0642
0 | 0 | 146.5 | 154.53 | 224350.4 | 59.27 | 59.3137
0 | 0 | 145.3 | 155.52 | 225780.6 | 59.67 | 59.6937
0 | 0 | 144.1 | 156.2 | 226775.6 | 59.95 | 59.9547
0 | 0 | 142.9 | 157.33 | 228408.5 | 60.35 | 60.3884
0 | 0 | 141.7 | 157.96 | 229320.7 | 60.6 | 60.6303
0 | 0 | 140.5 | 158.89 | 230671.7 | 60.99 | 60.9872
0 | 0 | 139.3 | 159.83 | 232039 | 61.35 | 61.348
0 | 0 | 138.1 | 160.69 | 233284.1 | 61.68 | 61.6781
0 | 0 | 136.9 | 161.61 | 234628.8 | 62.11 | 62.0312
0 | 0 | 135.6 | 162.59 | 236041.6 | 62.45 | 62.4074
0 | 0 | 134.5 | 163.39 | 237208 | 62.78 | 62.7145
0 | 0 | 133.3 | 164.34 | 238593.4 | 63.19 | 63.0791
0 | 0 | 132 | 165.3 | 239975.3 | 63.48 | 63.4476
0 | 0 | 130.8 | 166.34 | 241484.9 | 63.84 | 63.8468
0 | 0 | 129.6 | 167.24 | 242803.4 | 64.17 | 64.1922
0 | 0 | 128.4 | 168.09 | 244024.2 | 64.42 | 64.5185
0 | 0 | 127.2 | 169.11 | 245517.4 | 65.02 | 64.91
0 | 0 | 126 | 170.12 | 246983.4 | 65.29 | 65.2977
0 | 0 | 124.8 | 170.84 | 248023.3 | 65.55 | 65.574
0 | 0 | 123.6 | 171.84 | 249470.3 | 65.86 | 65.9579
0 | 0 | 122.4 | 173.07 | 251257.4 | 66.49 | 66.43
0 | 0 | 121.2 | 173.91 | 252478.6 | 66.76 | 66.7524
0 | 0 | 120 | 174.97 | 254016.8 | 67.24 | 67.1593
0 | 0 | 118.8 | 175.81 | 255238.4 | 67.52 | 67.4817
0 | 0 | 117.6 | 176.81 | 256693.7 | 67.88 | 67.8655
0 | 0 | 116.4 | 177.66 | 257925.9 | 68.28 | 68.1918
0 | 0 | 115.2 | 178.72 | 259468.6 | 68.64 | 68.5986
0 | 0 | 114 | 179.6 | 260744.3 | 68.99 | 68.9364
0 | 0 | 112.8 | 180.62 | 262226.7 | 69.41 | 69.3279
0 | 0 | 111.6 | 181.67 | 263752 | 69.76 | 69.7309
0 | 0 | 110.4 | 182.71 | 265253.6 | 70.09 | 70.1301
0 | 0 | 109.2 | 183.64 | 266604.9 | 70.56 | 70.4871
0 | 0 | 108 | 184.53 | 267896.8 | 70.88 | 70.8287
0 | 0 | 106.8 | 185.55 | 269381.5 | 71.22 | 71.2202
0 | 0 | 105.6 | 186.62 | 270926.1 | 71.61 | 71.6309
0 | 0 | 104.4 | 187.7 | 272505.5 | 72.12 | 72.0454
0 | 0 | 103.2 | 188.6 | 273812.2 | 72.4 | 72.3909
0 | 0 | 102 | 189.58 | 275236 | 72.78 | 72.7671
0 | 0 | 100.8 | 190.89 | 277127.5 | 73.21 | 73.2699
0 | 0 | 99.6 | 191.94 | 278653.6 | 73.76 | 73.6729
0 | 0 | 98.4 | 192.95 | 280116.1 | 74.15 | 74.0606
0 | 0 | 97.2 | 193.8 | 281356.3 | 74.46 | 74.3868
0 | 0 | 96 | 195.01 | 283118.7 | 74.81 | 74.8513
0 | 0 | 94.8 | 195.97 | 284505 | 75.29 | 75.2197
0 | 0 | 93.6 | 197.17 | 286245.7 | 75.65 | 75.6803
0 | 0 | 92.4 | 198.31 | 287899.9 | 76.11 | 76.1179
0 | 0 | 91.2 | 199.12 | 289081 | 76.47 | 76.4288
0 | 0 | 90 | 200.45 | 291013.8 | 76.84 | 76.9393
0 | 0 | 88.8 | 201.31 | 292255.4 | 77.25 | 77.2694
0 | 0 | 87.5 | 202.61 | 294149 | 77.82 | 77.7684
0 | 0 | 86.4 | 203.67 | 295682.8 | 78.11 | 78.1753
0 | 0 | 85.1 | 204.8 | 297330.4 | 78.66 | 78.609
0 | 0 | 83.9 | 205.77 | 298734.9 | 79.13 | 78.9813
0 | 0 | 82.7 | 206.86 | 300315 | 79.41 | 79.3997
0 | 0 | 81.5 | 208.16 | 302200.5 | 79.84 | 79.8987
0 | 0 | 80.3 | 209.05 | 303494.1 | 80.15 | 80.2403
0 | 0 | 79.1 | 210.13 | 305062.6 | 80.66 | 80.6548
0 | 0 | 77.9 | 211.54 | 307106.4 | 81.13 | 81.196
0 | 0 | 76.7 | 212.22 | 308101.7 | 81.51 | 81.457
0 | 0 | 75.5 | 213.59 | 310083.7 | 82.01 | 81.9829
0 | 0 | 74.3 | 214.72 | 311721.8 | 82.42 | 82.4166
0 | 0 | 73.1 | 215.89 | 313424.3 | 82.85 | 82.8657
0 | 0 | 71.8 | 216.98 | 315009.4 | 83.27 | 83.2841
0 | 0 | 70.7 | 218.17 | 316736.9 | 83.75 | 83.7408
0 | 0 | 69.4 | 219.32 | 318404.5 | 84.12 | 84.1822
0 | 0 | 68.2 | 220.49 | 320101.7 | 84.77 | 84.6313
0 | 0 | 67 | 221.52 | 321605.1 | 85.05 | 85.0267
0 | 0 | 65.8 | 222.93 | 323644.9 | 85.5 | 85.5679
0 | 0 | 64.6 | 223.84 | 324961.4 | 85.89 | 85.9172
0 | 0 | 63.4 | 225.06 | 326738 | 86.37 | 86.3854
0 | 0 | 62.2 | 226.09 | 328231.9 | 86.92 | 86.7808
0 | 0 | 61 | 227.54 | 330342.4 | 87.27 | 87.3374
0 | 0 | 59.8 | 228.45 | 331661 | 87.74 | 87.6866
0 | 0 | 58.6 | 229.58 | 333298.6 | 88.14 | 88.1204
0 | 0 | 57.4 | 231.01 | 335380.6 | 88.69 | 88.6693
0 | 0 | 56.2 | 232.24 | 337168.7 | 89.28 | 89.1414
0 | 0 | 55 | 233.37 | 338804 | 89.66 | 89.5751
0 | 0 | 53.7 | 234.54 | 340504.1 | 90.13 | 90.0242
0 | 0 | 52.6 | 235.76 | 342280.1 | 90.46 | 90.4925
0 | 0 | 51.3 | 236.96 | 344022.1 | 91.03 | 90.9531
0 | 0 | 50.1 | 238.23 | 345851.9 | 91.41 | 91.4405
0 | 0 | 48.9 | 239.34 | 347474.9 | 91.92 | 91.8666
0 | 0 | 47.7 | 240.43 | 349059.2 | 92.24 | 92.285
0 | 0 | 46.5 | 241.51 | 350626.3 | 92.74 | 92.6995
0 | 0 | 45.3 | 243.07 | 352888.3 | 93.27 | 93.2983
0 | 0 | 44.1 | 244.13 | 354429.8 | 93.66 | 93.7051
0 | 0 | 42.9 | 245.26 | 356069.2 | 94.13 | 94.1389
0 | 0 | 41.7 | 246.69 | 358145.6 | 94.78 | 94.6878
0 | 0 | 40.5 | 247.7 | 359613.5 | 95.16 | 95.0754
0 | 0 | 39.2 | 248.84 | 361259.5 | 95.4 | 95.513
0 | 0 | 38.1 | 250.35 | 363451.9 | 96.09 | 96.0926
0 | 0 | 36.8 | 251.31 | 364848.6 | 96.61 | 96.4611
0 | 0 | 35.6 | 252.52 | 366605.6 | 97.03 | 96.9255
0 | 0 | 34.4 | 253.45 | 367953.6 | 97.36 | 97.2825
0 | 0 | 33.2 | 254.75 | 369843.4 | 97.8 | 97.7814
0 | 0 | 32 | 255.83 | 371415.7 | 98.24 | 98.196
0 | 0 | 30.8 | 256.74 | 372738.5 | 98.49 | 98.5453
0 | 0 | 29.6 | 257.64 | 374041.4 | 98.83 | 98.8907
0 | 0 | 28.4 | 258.31 | 375010.7 | 99.38 | 99.1479
0 | 0 | 27.2 | 259.32 | 376479.7 | 99.65 | 99.5356
0 | 0 | 26 | 259.86 | 377260.9 | 99.8 | 99.7428
0 | 0 | 24.8 | 260.24 | 377806.8 | 99.95 | 99.8887
0 | 0 | 23.6 | 260.53 | 378229.1 | 99.98 | 100
0 | 0 | 22.4 | 260.46 | 378126.5 | 99.9 | 99.9731
0 | 0 | 21.2 | 260.43 | 378086.5 | 100 | 99.9616
0 | 0 | 20 | 259.85 | 377248.8 | 99.71 | 99.739
0 | 0 | 18.7 | 258.54 | 375347.1 | 99.22 | 99.2362
0 | 0 | 17.6 | 257.32 | 373575.6 | 98.79 | 98.7679
0 | 0 | 16.4 | 255.21 | 370504.6 | 97.92 | 97.958
0 | 0 | 15.1 | 252.81 | 367020.5 | 97.01 | 97.0368
0 | 0 | 13.9 | 249.37 | 362033.1 | 95.7 | 95.7164
0 | 0 | 12.7 | 245.19 | 355965.1 | 94.13 | 94.112
0 | 0 | 11.5 | 239.88 | 348255.3 | 92.03 | 92.0738
0 | 0 | 10.3 | 233.07 | 338360.5 | 89.47 | 89.4599
0 | 0 | 9.1 | 224.98 | 326627.7 | 86.28 | 86.3547
0 | 0 | 7.9 | 215.05 | 312205.8 | 82.46 | 82.5433
0 | 0 | 6.7 | 202.52 | 294015.2 | 77.72 | 77.7339
0 | 0 | 5.5 | 187.32 | 271954 | 71.98 | 71.8996
0 | 0 | 4.3 | 168.12 | 244071.1 | 64.46 | 64.53
0 | 0 | 3.1 | 143.56 | 208423.3 | 55.08 | 55.1031
0 | 0 | 1.9 | 117.4 | 170436 | 45.05 | 45.062
0 | 0 | 0.6 | 100.42 | 145794.3 | 38.55 | 38.5445
0 | 0 | -0.5 | 91.44 | 132752 | 35.1 | 35.0977
Measurement time: | 02/08/2022 20:09:17
Radiation device: | Novalis TB Accelerator
Energy: | 6 MV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 50 x 50 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 310 | 29.78 | 43230.2 | 18.71 | 18.3442
0 | 0 | 309.1 | 28.85 | 41879.5 | 17.78 | 17.7713
0 | 0 | 307.6 | 29.29 | 42521.9 | 18.05 | 18.0424
0 | 0 | 306.5 | 29.39 | 42664.1 | 18.1 | 18.104
0 | 0 | 305.3 | 29.63 | 43023.3 | 18.31 | 18.2518
0 | 0 | 304.1 | 29.74 | 43178.5 | 18.34 | 18.3196
0 | 0 | 302.9 | 29.98 | 43518.4 | 18.51 | 18.4674
0 | 0 | 301.7 | 30.1 | 43701.6 | 18.55 | 18.5413
0 | 0 | 300.5 | 30.45 | 44213.1 | 18.76 | 18.7569
0 | 0 | 299.3 | 30.6 | 44418.8 | 18.81 | 18.8493
0 | 0 | 298.1 | 30.71 | 44591.4 | 18.94 | 18.9171
0 | 0 | 296.9 | 31.01 | 45021.9 | 19.11 | 19.1019
0 | 0 | 295.7 | 31.23 | 45344 | 19.27 | 19.2374
0 | 0 | 294.5 | 31.39 | 45568.4 | 19.32 | 19.336
0 | 0 | 293.3 | 31.57 | 45839.6 | 19.45 | 19.4468
0 | 0 | 292.1 | 31.84 | 46222.1 | 19.64 | 19.6132
0 | 0 | 290.9 | 32.28 | 46860.5 | 19.82 | 19.8842
0 | 0 | 289.7 | 32.38 | 47010 | 20 | 19.9458
0 | 0 | 288.5 | 32.67 | 47427.4 | 20.14 | 20.1244
0 | 0 | 287.3 | 32.85 | 47689.8 | 20.25 | 20.2353
0 | 0 | 286.1 | 33.03 | 47950.9 | 20.41 | 20.3462
0 | 0 | 284.8 | 33.27 | 48294 | 20.5 | 20.494
0 | 0 | 283.6 | 33.59 | 48761.7 | 20.69 | 20.6911
0 | 0 | 282.4 | 33.86 | 49158.3 | 20.89 | 20.8575
0 | 0 | 281.3 | 34.11 | 49523.1 | 21.02 | 21.0115
0 | 0 | 280 | 34.25 | 49725.8 | 21.08 | 21.0977
0 | 0 | 278.8 | 34.52 | 50116.8 | 21.23 | 21.264
0 | 0 | 277.6 | 34.72 | 50401.9 | 21.36 | 21.3872
0 | 0 | 276.4 | 34.94 | 50724.9 | 21.53 | 21.5227
0 | 0 | 275.2 | 35.18 | 51067.1 | 21.66 | 21.6706
0 | 0 | 274 | 35.57 | 51641.7 | 21.92 | 21.9108
0 | 0 | 272.8 | 35.68 | 51801.8 | 21.98 | 21.9786
0 | 0 | 271.6 | 35.97 | 52219.2 | 22.19 | 22.1572
0 | 0 | 270.4 | 36.27 | 52654.4 | 22.32 | 22.342
0 | 0 | 269.2 | 36.52 | 53022 | 22.53 | 22.496
0 | 0 | 268 | 36.69 | 53260.9 | 22.65 | 22.6007
0 | 0 | 266.8 | 36.99 | 53705.5 | 22.8 | 22.7855
0 | 0 | 265.6 | 37.28 | 54118.8 | 23.01 | 22.9641
0 | 0 | 264.4 | 37.52 | 54469.2 | 23.12 | 23.112
0 | 0 | 263.2 | 37.81 | 54897.7 | 23.25 | 23.2906
0 | 0 | 262 | 37.91 | 55041.9 | 23.4 | 23.3522
0 | 0 | 260.8 | 38.45 | 55815 | 23.72 | 23.6849
0 | 0 | 259.6 | 38.62 | 56065.8 | 23.81 | 23.7896
0 | 0 | 258.4 | 38.91 | 56483.9 | 24.01 | 23.9682
0 | 0 | 257.2 | 39.18 | 56882.4 | 24.17 | 24.1345
0 | 0 | 255.9 | 39.38 | 57164.4 | 24.25 | 24.2577
0 | 0 | 254.8 | 39.72 | 57671 | 24.5 | 24.4672
0 | 0 | 253.6 | 39.98 | 58044.8 | 24.65 | 24.6273
0 | 0 | 252.4 | 40.27 | 58465.7 | 24.84 | 24.806
0 | 0 | 251.2 | 40.55 | 58870.4 | 24.96 | 24.9784
0 | 0 | 249.9 | 40.91 | 59385.7 | 25.19 | 25.2002
0 | 0 | 248.7 | 41.22 | 59842.3 | 25.46 | 25.3912
0 | 0 | 247.5 | 41.49 | 60234.7 | 25.57 | 25.5575
0 | 0 | 246.3 | 41.57 | 60346.9 | 25.63 | 25.6068
0 | 0 | 245.1 | 42.01 | 60993.2 | 25.87 | 25.8778
0 | 0 | 243.9 | 42.42 | 61588.1 | 26.22 | 26.1303
0 | 0 | 242.7 | 42.62 | 61872.3 | 26.31 | 26.2535
0 | 0 | 241.5 | 42.95 | 62347.5 | 26.43 | 26.4568
0 | 0 | 240.3 | 43.3 | 62857.7 | 26.62 | 26.6724
0 | 0 | 239.1 | 43.63 | 63334.4 | 26.89 | 26.8757
0 | 0 | 237.9 | 43.81 | 63603.3 | 26.99 | 26.9866
0 | 0 | 236.7 | 44.14 | 64081.6 | 27.16 | 27.1898
0 | 0 | 235.5 | 44.65 | 64815.7 | 27.52 | 27.504
0 | 0 | 234.3 | 44.81 | 65049 | 27.61 | 27.6026
0 | 0 | 233.1 | 45.22 | 65655.4 | 27.88 | 27.8551
0 | 0 | 231.9 | 45.61 | 66220 | 28.07 | 28.0954
0 | 0 | 230.7 | 45.83 | 66535.9 | 28.25 | 28.2309
0 | 0 | 229.5 | 46.1 | 66928.5 | 28.46 | 28.3972
0 | 0 | 228.3 | 46.51 | 67517.7 | 28.59 | 28.6497
0 | 0 | 227.1 | 46.81 | 67960 | 28.79 | 28.8345
0 | 0 | 225.9 | 47.13 | 68429.3 | 28.97 | 29.0317
0 | 0 | 224.7 | 47.48 | 68933 | 29.31 | 29.2473
0 | 0 | 223.5 | 47.91 | 69547.8 | 29.52 | 29.5121
0 | 0 | 222.2 | 48.21 | 69987.8 | 29.65 | 29.6969
0 | 0 | 221.1 | 48.58 | 70521.9 | 29.92 | 29.9248
0 | 0 | 219.8 | 48.91 | 71000.9 | 30.16 | 30.1281
0 | 0 | 218.7 | 49.22 | 71454.8 | 30.35 | 30.3191
0 | 0 | 217.4 | 49.54 | 71926.6 | 30.5 | 30.5162
0 | 0 | 216.2 | 49.9 | 72444.4 | 30.72 | 30.738
0 | 0 | 215.1 | 50.27 | 72984 | 31.03 | 30.9659
0 | 0 | 213.8 | 50.77 | 73712.2 | 31.29 | 31.2739
0 | 0 | 212.6 | 51.03 | 74081.4 | 31.42 | 31.434
0 | 0 | 211.4 | 51.43 | 74672.4 | 31.69 | 31.6804
0 | 0 | 210.2 | 51.84 | 75257.9 | 31.96 | 31.933
0 | 0 | 209 | 52.12 | 75669 | 32.09 | 32.1055
0 | 0 | 207.8 | 52.55 | 76287 | 32.41 | 32.3703
0 | 0 | 206.6 | 52.83 | 76691.9 | 32.57 | 32.5428
0 | 0 | 205.4 | 53.44 | 77588.1 | 32.98 | 32.9186
0 | 0 | 204.2 | 53.75 | 78027.6 | 33.17 | 33.1095
0 | 0 | 203 | 54.06 | 78488.6 | 33.31 | 33.3005
0 | 0 | 201.8 | 54.5 | 79115.6 | 33.57 | 33.5715
0 | 0 | 200.6 | 55.02 | 79875.7 | 33.95 | 33.8918
0 | 0 | 199.4 | 55.26 | 80228.8 | 34.07 | 34.0397
0 | 0 | 198.2 | 55.68 | 80830.1 | 34.26 | 34.2984
0 | 0 | 197 | 56.09 | 81430.1 | 34.64 | 34.5509
0 | 0 | 195.7 | 56.61 | 82188.8 | 34.88 | 34.8713
0 | 0 | 194.6 | 56.94 | 82660.4 | 35.12 | 35.0745
0 | 0 | 193.4 | 57.31 | 83207.7 | 35.28 | 35.3025
0 | 0 | 192.2 | 57.78 | 83879.6 | 35.58 | 35.592
0 | 0 | 191 | 58.11 | 84370.5 | 35.86 | 35.7952
0 | 0 | 189.8 | 58.55 | 85005.6 | 36.07 | 36.0663
0 | 0 | 188.6 | 59.1 | 85797.5 | 36.36 | 36.4051
0 | 0 | 187.3 | 59.55 | 86453.6 | 36.64 | 36.6823
0 | 0 | 186.1 | 59.99 | 87098.7 | 36.96 | 36.9533
0 | 0 | 185 | 60.39 | 87670.2 | 37.16 | 37.1997
0 | 0 | 183.7 | 60.75 | 88199.9 | 37.47 | 37.4215
0 | 0 | 182.6 | 61.21 | 88856.8 | 37.69 | 37.7048
0 | 0 | 181.3 | 61.69 | 89559.2 | 37.91 | 38.0005
0 | 0 | 180.1 | 62.26 | 90392.5 | 38.32 | 38.3516
0 | 0 | 178.9 | 62.5 | 90739.3 | 38.42 | 38.4994
0 | 0 | 177.7 | 62.98 | 91431.6 | 38.81 | 38.7951
0 | 0 | 176.5 | 63.62 | 92366.2 | 39.14 | 39.1894
0 | 0 | 175.3 | 63.94 | 92830.8 | 39.31 | 39.3865
0 | 0 | 174.1 | 64.57 | 93748.7 | 39.71 | 39.7745
0 | 0 | 172.9 | 65 | 94362.5 | 40.05 | 40.0394
0 | 0 | 171.7 | 65.46 | 95031.5 | 40.36 | 40.3228
0 | 0 | 170.5 | 65.9 | 95675.4 | 40.63 | 40.5938
0 | 0 | 169.3 | 66.45 | 96467.3 | 40.9 | 40.9326
0 | 0 | 168.1 | 66.87 | 97079.9 | 41.18 | 41.1913
0 | 0 | 166.9 | 67.36 | 97794.1 | 41.47 | 41.4932
0 | 0 | 165.7 | 67.76 | 98368.9 | 41.75 | 41.7396
0 | 0 | 164.5 | 68.39 | 99284.7 | 42.1 | 42.1276
0 | 0 | 163.3 | 68.87 | 99989.7 | 42.42 | 42.4233
0 | 0 | 162.1 | 69.33 | 100652.9 | 42.72 | 42.7067
0 | 0 | 160.9 | 69.75 | 101256.3 | 42.98 | 42.9654
0 | 0 | 159.6 | 70.35 | 102135.9 | 43.36 | 43.335
0 | 0 | 158.5 | 70.97 | 103034.1 | 43.69 | 43.7169
0 | 0 | 157.3 | 71.56 | 103882.8 | 44.13 | 44.0803
0 | 0 | 156 | 71.98 | 104502.9 | 44.39 | 44.339
0 | 0 | 154.9 | 72.43 | 105147.2 | 44.72 | 44.6162
0 | 0 | 153.6 | 73.12 | 106159.2 | 45 | 45.0413
0 | 0 | 152.5 | 73.49 | 106686.5 | 45.39 | 45.2692
0 | 0 | 151.3 | 74.08 | 107541.7 | 45.58 | 45.6326
0 | 0 | 150 | 74.57 | 108261.8 | 45.92 | 45.9345
0 | 0 | 148.8 | 75.3 | 109314.9 | 46.43 | 46.3841
0 | 0 | 147.6 | 75.94 | 110243.1 | 46.74 | 46.7784
0 | 0 | 146.4 | 76.26 | 110713.4 | 46.93 | 46.9755
0 | 0 | 145.2 | 76.95 | 111717 | 47.41 | 47.4005
0 | 0 | 144 | 77.38 | 112335.9 | 47.63 | 47.6654
0 | 0 | 142.9 | 77.85 | 113024.2 | 47.94 | 47.9549
0 | 0 | 141.6 | 78.61 | 114129 | 48.41 | 48.4231
0 | 0 | 140.4 | 79.16 | 114925.9 | 48.74 | 48.7619
0 | 0 | 139.2 | 79.86 | 115941 | 49.23 | 49.1931
0 | 0 | 138 | 80.37 | 116678.8 | 49.53 | 49.5072
0 | 0 | 136.8 | 80.87 | 117402.7 | 49.8 | 49.8152
0 | 0 | 135.6 | 81.59 | 118450.8 | 50.16 | 50.2587
0 | 0 | 134.4 | 82.22 | 119364.2 | 50.52 | 50.6468
0 | 0 | 133.2 | 82.85 | 120274.3 | 50.96 | 51.0349
0 | 0 | 132 | 83.31 | 120955.4 | 51.35 | 51.3182
0 | 0 | 130.8 | 84.03 | 121996.3 | 51.7 | 51.7617
0 | 0 | 129.6 | 84.6 | 122823.8 | 52.08 | 52.1128
0 | 0 | 128.4 | 85.2 | 123696.6 | 52.51 | 52.4824
0 | 0 | 127.2 | 85.83 | 124612.6 | 52.84 | 52.8705
0 | 0 | 126 | 86.62 | 125750.4 | 53.29 | 53.3572
0 | 0 | 124.8 | 87.19 | 126574 | 53.67 | 53.7083
0 | 0 | 123.5 | 87.78 | 127436.6 | 54 | 54.0717
0 | 0 | 122.4 | 88.32 | 128217.1 | 54.51 | 54.4043
0 | 0 | 121.2 | 88.96 | 129154.5 | 54.91 | 54.7986
0 | 0 | 120 | 89.69 | 130214.3 | 55.28 | 55.2482
0 | 0 | 118.8 | 90.34 | 131156.9 | 55.65 | 55.6486
0 | 0 | 117.5 | 90.89 | 131954.6 | 56.02 | 55.9874
0 | 0 | 116.4 | 91.71 | 133147.3 | 56.58 | 56.4925
0 | 0 | 115.1 | 92.37 | 134096.2 | 56.79 | 56.8991
0 | 0 | 113.9 | 93.11 | 135175.7 | 57.37 | 57.3549
0 | 0 | 112.8 | 93.7 | 136037.7 | 57.81 | 57.7184
0 | 0 | 111.5 | 94.39 | 137038.7 | 58.14 | 58.1434
0 | 0 | 110.3 | 95.11 | 138084 | 58.66 | 58.5869
0 | 0 | 109.1 | 96.01 | 139387.5 | 59.17 | 59.1413
0 | 0 | 107.9 | 96.48 | 140063.3 | 59.49 | 59.4308
0 | 0 | 106.7 | 97.25 | 141186.5 | 59.97 | 59.9051
0 | 0 | 105.5 | 97.91 | 142150.4 | 60.31 | 60.3117
0 | 0 | 104.3 | 98.75 | 143364.1 | 60.76 | 60.8291
0 | 0 | 103.1 | 99.63 | 144647.5 | 61.34 | 61.3712
0 | 0 | 101.9 | 100.06 | 145267.2 | 61.68 | 61.6361
0 | 0 | 100.7 | 100.76 | 146281.7 | 62.14 | 62.0673
0 | 0 | 99.5 | 101.73 | 147693.4 | 62.67 | 62.6648
0 | 0 | 98.3 | 102.3 | 148512.1 | 63.05 | 63.0159
0 | 0 | 97.1 | 103.2 | 149829.4 | 63.55 | 63.5703
0 | 0 | 95.9 | 104.13 | 151174.4 | 64.07 | 64.1432
0 | 0 | 94.7 | 104.77 | 152106.9 | 64.53 | 64.5374
0 | 0 | 93.5 | 105.55 | 153232.8 | 65.16 | 65.0179
0 | 0 | 92.3 | 106.23 | 154219.3 | 65.52 | 65.4367
0 | 0 | 91.1 | 107.13 | 155529.2 | 65.9 | 65.9911
0 | 0 | 89.9 | 107.72 | 156385.1 | 66.19 | 66.3546
0 | 0 | 88.7 | 108.26 | 157168.7 | 66.65 | 66.6872
0 | 0 | 87.5 | 109.42 | 158853.3 | 67.37 | 67.4017
0 | 0 | 86.3 | 110.05 | 159766.8 | 67.77 | 67.7898
0 | 0 | 85 | 110.94 | 161062.1 | 68.3 | 68.3381
0 | 0 | 83.9 | 111.63 | 162068 | 68.89 | 68.7631
0 | 0 | 82.6 | 112.49 | 163318 | 69.2 | 69.2928
0 | 0 | 81.4 | 113.38 | 164603.3 | 69.75 | 69.8411
0 | 0 | 80.2 | 114.48 | 166202.4 | 70.5 | 70.5187
0 | 0 | 79 | 115.25 | 167316.9 | 71.03 | 70.993
0 | 0 | 77.8 | 116.09 | 168541 | 71.44 | 71.5104
0 | 0 | 76.6 | 116.67 | 169384.9 | 71.87 | 71.8677
0 | 0 | 75.4 | 117.76 | 170966.6 | 72.63 | 72.5391
0 | 0 | 74.2 | 118.42 | 171919.2 | 73.03 | 72.9457
0 | 0 | 73 | 119.3 | 173196.5 | 73.43 | 73.4877
0 | 0 | 71.8 | 120.17 | 174467.6 | 74.09 | 74.0237
0 | 0 | 70.6 | 121.18 | 175924.1 | 74.63 | 74.6458
0 | 0 | 69.3 | 121.99 | 177109.8 | 75.11 | 75.1448
0 | 0 | 68.1 | 122.86 | 178369.1 | 75.63 | 75.6807
0 | 0 | 66.9 | 123.78 | 179708.3 | 76.1 | 76.2474
0 | 0 | 65.7 | 124.62 | 180927.6 | 76.77 | 76.7648
0 | 0 | 64.5 | 125.44 | 182114.6 | 77.23 | 77.2699
0 | 0 | 63.3 | 126.34 | 183425.4 | 77.73 | 77.8243
0 | 0 | 62.1 | 126.96 | 184317.8 | 78.11 | 78.2062
0 | 0 | 60.9 | 128.13 | 186012 | 78.92 | 78.9269
0 | 0 | 59.7 | 128.79 | 186975.8 | 79.37 | 79.3335
0 | 0 | 58.5 | 129.77 | 188399.5 | 80.02 | 79.9372
0 | 0 | 57.3 | 130.82 | 189921.4 | 80.44 | 80.584
0 | 0 | 56.1 | 131.9 | 191485.6 | 81.24 | 81.2492
0 | 0 | 54.8 | 132.61 | 192524.4 | 81.73 | 81.6866
0 | 0 | 53.7 | 133.57 | 193922 | 82.2 | 82.2779
0 | 0 | 52.4 | 134.54 | 195317.9 | 82.85 | 82.8754
0 | 0 | 51.2 | 135.51 | 196726.7 | 83.25 | 83.473
0 | 0 | 50 | 136.31 | 197891 | 84.1 | 83.9658
0 | 0 | 48.8 | 137.3 | 199324.8 | 84.57 | 84.5756
0 | 0 | 47.6 | 138.25 | 200713.4 | 85.15 | 85.1608
0 | 0 | 46.4 | 139.31 | 202243.9 | 85.88 | 85.8137
0 | 0 | 45.2 | 139.97 | 203202 | 86.25 | 86.2203
0 | 0 | 44 | 141.06 | 204791.7 | 86.74 | 86.8917
0 | 0 | 42.8 | 142.19 | 206425.4 | 87.46 | 87.5878
0 | 0 | 41.6 | 142.99 | 207591.7 | 88.06 | 88.0806
0 | 0 | 40.4 | 144.01 | 209077.1 | 88.66 | 88.7089
0 | 0 | 39.1 | 145.02 | 210543.9 | 89.47 | 89.331
0 | 0 | 38 | 145.92 | 211849.2 | 89.86 | 89.8854
0 | 0 | 36.7 | 147.19 | 213685.9 | 90.61 | 90.6677
0 | 0 | 35.5 | 147.97 | 214813.6 | 91.13 | 91.1482
0 | 0 | 34.3 | 148.8 | 216027.1 | 91.82 | 91.6595
0 | 0 | 33.1 | 149.92 | 217657.2 | 92.18 | 92.3494
0 | 0 | 31.9 | 150.6 | 218639.2 | 92.62 | 92.7683
0 | 0 | 30.7 | 151.67 | 220198.5 | 93.5 | 93.4274
0 | 0 | 29.5 | 152.51 | 221412 | 94.17 | 93.9448
0 | 0 | 28.3 | 153.48 | 222821.9 | 94.62 | 94.5423
0 | 0 | 27.1 | 154.72 | 224620 | 95.27 | 95.3061
0 | 0 | 25.9 | 155.22 | 225347.9 | 95.64 | 95.6141
0 | 0 | 24.7 | 156.54 | 227262.3 | 96.33 | 96.4273
0 | 0 | 23.5 | 157.4 | 228516.8 | 96.88 | 96.957
0 | 0 | 22.3 | 158.45 | 230036.8 | 97.64 | 97.6038
0 | 0 | 21 | 159.16 | 231067.5 | 98.13 | 98.0411
0 | 0 | 19.9 | 160.06 | 232378 | 98.57 | 98.5955
0 | 0 | 18.6 | 160.73 | 233350.4 | 98.91 | 99.0083
0 | 0 | 17.4 | 161.31 | 234185.7 | 99.34 | 99.3655
0 | 0 | 16.2 | 162.05 | 235257.6 | 99.64 | 99.8214
0 | 0 | 15 | 162.11 | 235344 | 100 | 99.8583
0 | 0 | 13.8 | 162.34 | 235683 | 99.74 | 100
0 | 0 | 12.6 | 161.98 | 235156.1 | 99.83 | 99.7782
0 | 0 | 11.4 | 161 | 233739.5 | 98.95 | 99.1746
0 | 0 | 10.2 | 159.39 | 231402.5 | 98.23 | 98.1828
0 | 0 | 9 | 157.02 | 227962.5 | 96.52 | 96.7229
0 | 0 | 7.8 | 153.16 | 222362.6 | 94.42 | 94.3452
0 | 0 | 6.6 | 147.68 | 214400.2 | 91.04 | 90.9696
0 | 0 | 5.4 | 139.99 | 203240.6 | 86.23 | 86.2326
0 | 0 | 4.2 | 129.36 | 187799.6 | 79.6 | 79.6846
0 | 0 | 3 | 112.48 | 163299.6 | 69.43 | 69.2867
0 | 0 | 1.8 | 94.95 | 137843.9 | 58.54 | 58.4884
0 | 0 | 0.5 | 82.19 | 119327.1 | 50.67 | 50.6283
0 | 0 | -0.5 | 76.03 | 110382.9 | 46.77 | 46.8338