## 6 MeV ionization
Ionization | [5 empty] | Dose
Measurement time: | 07/09/2022 18:47:58 |  |  |  |  | Measurement time: | 07/09/2022 18:47:58
Radiation device: | Clinac3 Accelerator |  |  |  |  | Radiation device: | Clinac3 Accelerator
Energy: | 6 MeV |  |  |  |  | Energy: | 6 MeV
Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS |  |  |  |  | Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie |  |  |  |  | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 1000 mm |  |  |  |  | SSD: | 1000 mm
Field size: | 100 x 100 mm |  |  |  |  | Field size: | 100 x 100 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Beam |  |  |  |  | Scan type: | Beam
Scan mode: | Continuous |  |  |  |  | Scan mode: | Continuous
Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 200 | 2.73 | 1356.2 | 0.16 | 0 | 0 | 200 | 2.73 | 1356.2 | 0.16
0 | 0 | 199.3 | 2.24 | 1113.3 | 0.18 | 0 | 0 | 199.3 | 2.24 | 1113.3 | 0.19
0 | 0 | 198 | 2.08 | 1033.1 | 0.18 | 0 | 0 | 198 | 2.08 | 1033.1 | 0.18
0 | 0 | 196.8 | 2.25 | 1116.7 | 0.18 | 0 | 0 | 196.8 | 2.25 | 1116.7 | 0.18
0 | 0 | 195.6 | 2.15 | 1067.2 | 0.18 | 0 | 0 | 195.6 | 2.15 | 1067.2 | 0.18
0 | 0 | 194.4 | 2.18 | 1084.5 | 0.18 | 0 | 0 | 194.4 | 2.18 | 1084.5 | 0.19
0 | 0 | 193.1 | 2.23 | 1109.8 | 0.18 | 0 | 0 | 193.1 | 2.23 | 1109.8 | 0.18
0 | 0 | 192 | 2.22 | 1103.9 | 0.19 | 0 | 0 | 192 | 2.22 | 1103.9 | 0.19
0 | 0 | 190.8 | 2.15 | 1070.4 | 0.18 | 0 | 0 | 190.8 | 2.15 | 1070.4 | 0.18
0 | 0 | 189.6 | 2.35 | 1169.7 | 0.19 | 0 | 0 | 189.6 | 2.35 | 1169.7 | 0.19
0 | 0 | 188.4 | 2.45 | 1217.6 | 0.21 | 0 | 0 | 188.4 | 2.45 | 1217.6 | 0.21
0 | 0 | 187.1 | 2.29 | 1138.4 | 0.2 | 0 | 0 | 187.1 | 2.29 | 1138.4 | 0.2
0 | 0 | 185.9 | 2.64 | 1313.9 | 0.21 | 0 | 0 | 185.9 | 2.64 | 1313.9 | 0.21
0 | 0 | 184.7 | 2.4 | 1190.7 | 0.2 | 0 | 0 | 184.7 | 2.4 | 1190.7 | 0.2
0 | 0 | 183.5 | 2.45 | 1217.9 | 0.21 | 0 | 0 | 183.5 | 2.45 | 1217.9 | 0.21
0 | 0 | 182.3 | 2.38 | 1183.5 | 0.2 | 0 | 0 | 182.3 | 2.38 | 1183.5 | 0.2
0 | 0 | 181.1 | 2.49 | 1235.4 | 0.21 | 0 | 0 | 181.1 | 2.49 | 1235.4 | 0.21
0 | 0 | 179.9 | 2.56 | 1270.7 | 0.21 | 0 | 0 | 179.9 | 2.56 | 1270.7 | 0.21
0 | 0 | 178.7 | 2.62 | 1303.6 | 0.22 | 0 | 0 | 178.7 | 2.62 | 1303.6 | 0.22
0 | 0 | 177.5 | 2.52 | 1252.3 | 0.21 | 0 | 0 | 177.5 | 2.52 | 1252.3 | 0.21
0 | 0 | 176.3 | 2.58 | 1280.8 | 0.21 | 0 | 0 | 176.3 | 2.58 | 1280.8 | 0.21
0 | 0 | 175.1 | 2.48 | 1229.7 | 0.2 | 0 | 0 | 175.1 | 2.48 | 1229.7 | 0.21
0 | 0 | 173.9 | 2.5 | 1243.1 | 0.21 | 0 | 0 | 173.9 | 2.5 | 1243.1 | 0.21
0 | 0 | 172.7 | 2.68 | 1330.4 | 0.22 | 0 | 0 | 172.7 | 2.68 | 1330.4 | 0.22
0 | 0 | 171.5 | 2.74 | 1361 | 0.23 | 0 | 0 | 171.5 | 2.74 | 1361 | 0.23
0 | 0 | 170.3 | 2.53 | 1258.2 | 0.22 | 0 | 0 | 170.3 | 2.53 | 1258.2 | 0.22
0 | 0 | 169.1 | 2.7 | 1340.4 | 0.22 | 0 | 0 | 169.1 | 2.7 | 1340.4 | 0.22
0 | 0 | 167.9 | 2.64 | 1311.6 | 0.22 | 0 | 0 | 167.9 | 2.64 | 1311.6 | 0.22
0 | 0 | 166.7 | 2.73 | 1356 | 0.23 | 0 | 0 | 166.7 | 2.73 | 1356 | 0.23
0 | 0 | 165.5 | 2.88 | 1428.5 | 0.23 | 0 | 0 | 165.5 | 2.88 | 1428.5 | 0.24
0 | 0 | 164.3 | 2.82 | 1399 | 0.24 | 0 | 0 | 164.3 | 2.82 | 1399 | 0.24
0 | 0 | 163.1 | 2.74 | 1360 | 0.23 | 0 | 0 | 163.1 | 2.74 | 1360 | 0.23
0 | 0 | 161.9 | 2.8 | 1391.2 | 0.24 | 0 | 0 | 161.9 | 2.8 | 1391.2 | 0.24
0 | 0 | 160.7 | 2.99 | 1486.8 | 0.25 | 0 | 0 | 160.7 | 2.99 | 1486.8 | 0.25
0 | 0 | 159.5 | 3.04 | 1508.8 | 0.25 | 0 | 0 | 159.5 | 3.04 | 1508.8 | 0.25
0 | 0 | 158.3 | 2.77 | 1378.4 | 0.23 | 0 | 0 | 158.3 | 2.77 | 1378.4 | 0.24
0 | 0 | 157.1 | 2.86 | 1422.5 | 0.24 | 0 | 0 | 157.1 | 2.86 | 1422.5 | 0.25
0 | 0 | 155.9 | 2.94 | 1461.2 | 0.24 | 0 | 0 | 155.9 | 2.94 | 1461.2 | 0.24
0 | 0 | 154.7 | 2.95 | 1464.3 | 0.25 | 0 | 0 | 154.7 | 2.95 | 1464.3 | 0.26
0 | 0 | 153.5 | 2.87 | 1427.5 | 0.23 | 0 | 0 | 153.5 | 2.87 | 1427.5 | 0.23
0 | 0 | 152.3 | 2.99 | 1486.9 | 0.26 | 0 | 0 | 152.3 | 2.99 | 1486.9 | 0.26
0 | 0 | 151.1 | 3.02 | 1502.4 | 0.25 | 0 | 0 | 151.1 | 3.02 | 1502.4 | 0.25
0 | 0 | 149.9 | 3.1 | 1539.4 | 0.25 | 0 | 0 | 149.9 | 3.1 | 1539.4 | 0.25
0 | 0 | 148.7 | 3.04 | 1509.8 | 0.26 | 0 | 0 | 148.7 | 3.04 | 1509.8 | 0.27
0 | 0 | 147.5 | 2.93 | 1454.9 | 0.24 | 0 | 0 | 147.5 | 2.93 | 1454.9 | 0.24
0 | 0 | 146.2 | 3.08 | 1529.2 | 0.26 | 0 | 0 | 146.2 | 3.08 | 1529.2 | 0.27
0 | 0 | 145.1 | 3.2 | 1592 | 0.26 | 0 | 0 | 145.1 | 3.2 | 1592 | 0.27
0 | 0 | 143.9 | 3.1 | 1541.4 | 0.25 | 0 | 0 | 143.9 | 3.1 | 1541.4 | 0.25
0 | 0 | 142.7 | 3.03 | 1505.9 | 0.26 | 0 | 0 | 142.7 | 3.03 | 1505.9 | 0.26
0 | 0 | 141.5 | 3.21 | 1592.5 | 0.26 | 0 | 0 | 141.5 | 3.21 | 1592.5 | 0.27
0 | 0 | 140.3 | 3.23 | 1603.8 | 0.27 | 0 | 0 | 140.3 | 3.23 | 1603.8 | 0.28
0 | 0 | 139 | 3.24 | 1607.8 | 0.26 | 0 | 0 | 139 | 3.24 | 1607.8 | 0.27
0 | 0 | 137.8 | 2.93 | 1456.6 | 0.25 | 0 | 0 | 137.8 | 2.93 | 1456.6 | 0.25
0 | 0 | 136.6 | 3.17 | 1577 | 0.26 | 0 | 0 | 136.6 | 3.17 | 1577 | 0.26
0 | 0 | 135.4 | 3.12 | 1552.4 | 0.26 | 0 | 0 | 135.4 | 3.12 | 1552.4 | 0.27
0 | 0 | 134.2 | 3.38 | 1677 | 0.27 | 0 | 0 | 134.2 | 3.38 | 1677 | 0.28
0 | 0 | 133 | 3.02 | 1501.5 | 0.26 | 0 | 0 | 133 | 3.02 | 1501.5 | 0.26
0 | 0 | 131.8 | 3.25 | 1617.1 | 0.28 | 0 | 0 | 131.8 | 3.25 | 1617.1 | 0.28
0 | 0 | 130.6 | 3.36 | 1668.5 | 0.28 | 0 | 0 | 130.6 | 3.36 | 1668.5 | 0.29
0 | 0 | 129.4 | 3.47 | 1722 | 0.28 | 0 | 0 | 129.4 | 3.47 | 1722 | 0.29
0 | 0 | 128.2 | 3.38 | 1681.5 | 0.29 | 0 | 0 | 128.2 | 3.38 | 1681.5 | 0.29
0 | 0 | 127 | 3.71 | 1845.2 | 0.3 | 0 | 0 | 127 | 3.71 | 1845.2 | 0.31
0 | 0 | 125.8 | 3.46 | 1719 | 0.29 | 0 | 0 | 125.8 | 3.46 | 1719 | 0.3
0 | 0 | 124.6 | 3.54 | 1756.7 | 0.3 | 0 | 0 | 124.6 | 3.54 | 1756.7 | 0.3
0 | 0 | 123.4 | 3.56 | 1768.3 | 0.29 | 0 | 0 | 123.4 | 3.56 | 1768.3 | 0.29
0 | 0 | 122.2 | 3.6 | 1788.6 | 0.31 | 0 | 0 | 122.2 | 3.6 | 1788.6 | 0.32
0 | 0 | 121 | 3.57 | 1771.5 | 0.3 | 0 | 0 | 121 | 3.57 | 1771.5 | 0.3
0 | 0 | 119.8 | 3.58 | 1780.4 | 0.3 | 0 | 0 | 119.8 | 3.58 | 1780.4 | 0.31
0 | 0 | 118.6 | 3.78 | 1876.4 | 0.31 | 0 | 0 | 118.6 | 3.78 | 1876.4 | 0.31
0 | 0 | 117.4 | 3.66 | 1817.6 | 0.31 | 0 | 0 | 117.4 | 3.66 | 1817.6 | 0.32
0 | 0 | 116.2 | 3.79 | 1883.2 | 0.31 | 0 | 0 | 116.2 | 3.79 | 1883.2 | 0.32
0 | 0 | 115 | 3.76 | 1868.9 | 0.32 | 0 | 0 | 115 | 3.76 | 1868.9 | 0.33
0 | 0 | 113.8 | 3.69 | 1831.4 | 0.31 | 0 | 0 | 113.8 | 3.69 | 1831.4 | 0.32
0 | 0 | 112.6 | 3.82 | 1898.1 | 0.32 | 0 | 0 | 112.6 | 3.82 | 1898.1 | 0.32
0 | 0 | 111.4 | 4.09 | 2032.1 | 0.34 | 0 | 0 | 111.4 | 4.09 | 2032.1 | 0.34
0 | 0 | 110.2 | 3.78 | 1878.4 | 0.32 | 0 | 0 | 110.2 | 3.78 | 1878.4 | 0.33
0 | 0 | 109 | 4.17 | 2072.2 | 0.34 | 0 | 0 | 109 | 4.17 | 2072.2 | 0.35
0 | 0 | 107.8 | 3.93 | 1951.1 | 0.34 | 0 | 0 | 107.8 | 3.93 | 1951.1 | 0.34
0 | 0 | 106.6 | 4.23 | 2101.1 | 0.34 | 0 | 0 | 106.6 | 4.23 | 2101.1 | 0.35
0 | 0 | 105.4 | 4.12 | 2046.6 | 0.34 | 0 | 0 | 105.4 | 4.12 | 2046.6 | 0.35
0 | 0 | 104.2 | 4.04 | 2008.7 | 0.34 | 0 | 0 | 104.2 | 4.04 | 2008.7 | 0.35
0 | 0 | 103 | 4.12 | 2046.1 | 0.35 | 0 | 0 | 103 | 4.12 | 2046.1 | 0.36
0 | 0 | 101.8 | 4.33 | 2149.9 | 0.35 | 0 | 0 | 101.8 | 4.33 | 2149.9 | 0.36
0 | 0 | 100.6 | 4.19 | 2083.6 | 0.36 | 0 | 0 | 100.6 | 4.19 | 2083.6 | 0.37
0 | 0 | 99.4 | 4.11 | 2042.4 | 0.35 | 0 | 0 | 99.4 | 4.11 | 2042.4 | 0.36
0 | 0 | 98.2 | 4.57 | 2268.5 | 0.37 | 0 | 0 | 98.2 | 4.57 | 2268.5 | 0.38
0 | 0 | 97 | 4.34 | 2155.4 | 0.37 | 0 | 0 | 97 | 4.34 | 2155.4 | 0.38
0 | 0 | 95.8 | 4.42 | 2196.9 | 0.38 | 0 | 0 | 95.8 | 4.42 | 2196.9 | 0.39
0 | 0 | 94.6 | 4.34 | 2157.5 | 0.36 | 0 | 0 | 94.6 | 4.34 | 2157.5 | 0.37
0 | 0 | 93.4 | 4.54 | 2256.8 | 0.38 | 0 | 0 | 93.4 | 4.54 | 2256.8 | 0.39
0 | 0 | 92.1 | 4.74 | 2352.6 | 0.38 | 0 | 0 | 92.1 | 4.74 | 2352.6 | 0.39
0 | 0 | 91 | 4.83 | 2399.8 | 0.4 | 0 | 0 | 91 | 4.83 | 2399.8 | 0.41
0 | 0 | 89.7 | 4.82 | 2395.8 | 0.4 | 0 | 0 | 89.7 | 4.82 | 2395.8 | 0.42
0 | 0 | 88.6 | 4.55 | 2262.7 | 0.39 | 0 | 0 | 88.6 | 4.55 | 2262.7 | 0.4
0 | 0 | 87.3 | 4.93 | 2448.5 | 0.4 | 0 | 0 | 87.3 | 4.93 | 2448.5 | 0.42
0 | 0 | 86.1 | 4.7 | 2336 | 0.39 | 0 | 0 | 86.1 | 4.7 | 2336 | 0.41
0 | 0 | 84.9 | 4.79 | 2381.3 | 0.41 | 0 | 0 | 84.9 | 4.79 | 2381.3 | 0.42
0 | 0 | 83.7 | 4.9 | 2432.9 | 0.4 | 0 | 0 | 83.7 | 4.9 | 2432.9 | 0.41
0 | 0 | 82.5 | 4.77 | 2369.2 | 0.4 | 0 | 0 | 82.5 | 4.77 | 2369.2 | 0.42
0 | 0 | 81.3 | 4.9 | 2436.5 | 0.42 | 0 | 0 | 81.3 | 4.9 | 2436.5 | 0.43
0 | 0 | 80.1 | 5.29 | 2627.5 | 0.42 | 0 | 0 | 80.1 | 5.29 | 2627.5 | 0.44
0 | 0 | 78.9 | 4.98 | 2476.2 | 0.42 | 0 | 0 | 78.9 | 4.98 | 2476.2 | 0.44
0 | 0 | 77.7 | 5.02 | 2493.6 | 0.43 | 0 | 0 | 77.7 | 5.02 | 2493.6 | 0.44
0 | 0 | 76.5 | 5.42 | 2691.3 | 0.44 | 0 | 0 | 76.5 | 5.42 | 2691.3 | 0.45
0 | 0 | 75.3 | 5.13 | 2546.2 | 0.43 | 0 | 0 | 75.3 | 5.13 | 2546.2 | 0.45
0 | 0 | 74.1 | 5.26 | 2614.7 | 0.45 | 0 | 0 | 74.1 | 5.26 | 2614.7 | 0.46
0 | 0 | 72.9 | 5.55 | 2756.1 | 0.45 | 0 | 0 | 72.9 | 5.55 | 2756.1 | 0.47
0 | 0 | 71.7 | 5.35 | 2658.6 | 0.45 | 0 | 0 | 71.7 | 5.35 | 2658.6 | 0.47
0 | 0 | 70.5 | 5.3 | 2634 | 0.45 | 0 | 0 | 70.5 | 5.3 | 2634 | 0.47
0 | 0 | 69.2 | 5.48 | 2720.9 | 0.44 | 0 | 0 | 69.2 | 5.48 | 2720.9 | 0.46
0 | 0 | 68 | 5.51 | 2739.5 | 0.47 | 0 | 0 | 68 | 5.51 | 2739.5 | 0.49
0 | 0 | 66.8 | 5.46 | 2711.7 | 0.46 | 0 | 0 | 66.8 | 5.46 | 2711.7 | 0.48
0 | 0 | 65.6 | 5.87 | 2915.9 | 0.48 | 0 | 0 | 65.6 | 5.87 | 2915.9 | 0.49
0 | 0 | 64.4 | 5.56 | 2762 | 0.47 | 0 | 0 | 64.4 | 5.56 | 2762 | 0.48
0 | 0 | 63.2 | 5.39 | 2678.4 | 0.46 | 0 | 0 | 63.2 | 5.39 | 2678.4 | 0.47
0 | 0 | 62 | 5.92 | 2941 | 0.48 | 0 | 0 | 62 | 5.92 | 2941 | 0.5
0 | 0 | 60.8 | 5.67 | 2819.2 | 0.48 | 0 | 0 | 60.8 | 5.67 | 2819.2 | 0.5
0 | 0 | 59.6 | 6 | 2982.2 | 0.51 | 0 | 0 | 59.6 | 6 | 2982.2 | 0.53
0 | 0 | 58.4 | 6.14 | 3049.7 | 0.5 | 0 | 0 | 58.4 | 6.14 | 3049.7 | 0.52
0 | 0 | 57.2 | 5.86 | 2911.6 | 0.5 | 0 | 0 | 57.2 | 5.86 | 2911.6 | 0.51
0 | 0 | 56 | 6.07 | 3015.3 | 0.52 | 0 | 0 | 56 | 6.07 | 3015.3 | 0.53
0 | 0 | 54.8 | 6.23 | 3094.8 | 0.5 | 0 | 0 | 54.8 | 6.23 | 3094.8 | 0.52
0 | 0 | 53.6 | 6.16 | 3058.6 | 0.52 | 0 | 0 | 53.6 | 6.16 | 3058.6 | 0.54
0 | 0 | 52.4 | 6.17 | 3064.4 | 0.52 | 0 | 0 | 52.4 | 6.17 | 3064.4 | 0.54
0 | 0 | 51.2 | 6.52 | 3239.2 | 0.53 | 0 | 0 | 51.2 | 6.52 | 3239.2 | 0.55
0 | 0 | 50 | 6.26 | 3108.9 | 0.53 | 0 | 0 | 50 | 6.26 | 3108.9 | 0.55
0 | 0 | 48.8 | 6.23 | 3094.7 | 0.53 | 0 | 0 | 48.8 | 6.23 | 3094.7 | 0.55
0 | 0 | 47.6 | 6.8 | 3377.4 | 0.55 | 0 | 0 | 47.6 | 6.8 | 3377.4 | 0.57
0 | 0 | 46.4 | 6.61 | 3286.3 | 0.56 | 0 | 0 | 46.4 | 6.61 | 3286.3 | 0.58
0 | 0 | 45.2 | 6.43 | 3193.4 | 0.55 | 0 | 0 | 45.2 | 6.43 | 3193.4 | 0.57
0 | 0 | 44 | 6.76 | 3359.2 | 0.55 | 0 | 0 | 44 | 6.76 | 3359.2 | 0.57
0 | 0 | 42.7 | 6.64 | 3296.4 | 0.56 | 0 | 0 | 42.7 | 6.64 | 3296.4 | 0.58
0 | 0 | 41.5 | 6.86 | 3407.9 | 0.58 | 0 | 0 | 41.5 | 6.86 | 3407.9 | 0.6
0 | 0 | 40.3 | 7.1 | 3526.7 | 0.58 | 0 | 0 | 40.3 | 7.1 | 3526.7 | 0.6
0 | 0 | 39.1 | 7.09 | 3520.7 | 0.59 | 0 | 0 | 39.1 | 7.09 | 3520.7 | 0.61
0 | 0 | 37.9 | 7.08 | 3515.8 | 0.6 | 0 | 0 | 37.9 | 7.08 | 3515.8 | 0.62
0 | 0 | 36.7 | 7.02 | 3486.7 | 0.58 | 0 | 0 | 36.7 | 7.02 | 3486.7 | 0.61
0 | 0 | 35.5 | 7.62 | 3787.6 | 0.63 | 0 | 0 | 35.5 | 7.62 | 3787.6 | 0.65
0 | 0 | 34.3 | 8.65 | 4295 | 0.73 | 0 | 0 | 34.3 | 8.65 | 4295 | 0.76
0 | 0 | 33.1 | 12.53 | 6225.6 | 1.04 | 0 | 0 | 33.1 | 12.53 | 6225.6 | 1.09
0 | 0 | 31.9 | 23.82 | 11835.5 | 1.96 | 0 | 0 | 31.9 | 23.82 | 11835.5 | 2.04
0 | 0 | 30.7 | 47.34 | 23520.8 | 4.01 | 0 | 0 | 30.7 | 47.34 | 23520.8 | 4.18
0 | 0 | 29.5 | 86.88 | 43163.4 | 7.29 | 0 | 0 | 29.5 | 86.88 | 43163.4 | 7.59
0 | 0 | 28.3 | 153.58 | 76300.7 | 12.63 | 0 | 0 | 28.3 | 153.58 | 76300.7 | 13.15
0 | 0 | 27.1 | 234.32 | 116413.2 | 19.78 | 0 | 0 | 27.1 | 234.32 | 116413.2 | 20.58
0 | 0 | 25.9 | 336.46 | 167157.4 | 28.15 | 0 | 0 | 25.9 | 336.46 | 167157.4 | 29.27
0 | 0 | 24.7 | 472.55 | 234770.8 | 38.71 | 0 | 0 | 24.7 | 472.55 | 234770.8 | 40.2
0 | 0 | 23.5 | 565.8 | 281094.5 | 48.36 | 0 | 0 | 23.5 | 565.8 | 281094.5 | 50.07
0 | 0 | 22.3 | 701.07 | 348298.3 | 58.34 | 0 | 0 | 22.3 | 701.07 | 348298.3 | 60.21
0 | 0 | 21.1 | 831.94 | 413319.2 | 68.41 | 0 | 0 | 21.1 | 831.94 | 413319.2 | 70.38
0 | 0 | 19.9 | 910.74 | 452467.1 | 76.64 | 0 | 0 | 19.9 | 910.74 | 452467.1 | 78.6
0 | 0 | 18.6 | 1006.22 | 499902.9 | 84.42 | 0 | 0 | 18.6 | 1006.22 | 499902.9 | 86.19
0 | 0 | 17.4 | 1105.78 | 549366.8 | 90.22 | 0 | 0 | 17.4 | 1105.78 | 549366.8 | 91.76
0 | 0 | 16.2 | 1106.87 | 549907.9 | 94.96 | 0 | 0 | 16.2 | 1106.87 | 549907.9 | 96.21
0 | 0 | 15 | 1184.5 | 588476.7 | 98.04 | 0 | 0 | 15 | 1184.5 | 588476.7 | 98.92
0 | 0 | 13.8 | 1209.92 | 601104.5 | 99.54 | 0 | 0 | 13.8 | 1209.92 | 601104.5 | 100
0 | 0 | 12.6 | 1180.48 | 586475.5 | 100 | 0 | 0 | 12.6 | 1180.48 | 586475.5 | 100
0 | 0 | 11.4 | 1186.86 | 589646.2 | 99.27 | 0 | 0 | 11.4 | 1186.86 | 589646.2 | 98.86
0 | 0 | 10.2 | 1191.52 | 591963.2 | 98.04 | 0 | 0 | 10.2 | 1191.52 | 591963.2 | 97.26
0 | 0 | 9 | 1134.86 | 563813.3 | 96.11 | 0 | 0 | 9 | 1134.86 | 563813.3 | 94.97
0 | 0 | 7.8 | 1129.48 | 561140.4 | 93.95 | 0 | 0 | 7.8 | 1129.48 | 561140.4 | 92.48
0 | 0 | 6.6 | 1113.34 | 553121.3 | 91.76 | 0 | 0 | 6.6 | 1113.34 | 553121.3 | 89.98
0 | 0 | 5.4 | 1060.06 | 526648.8 | 89.42 | 0 | 0 | 5.4 | 1060.06 | 526648.8 | 87.37
0 | 0 | 4.2 | 1039.68 | 516523.7 | 87.24 | 0 | 0 | 4.2 | 1039.68 | 516523.7 | 84.95
0 | 0 | 3 | 1039.33 | 516352 | 85.03 | 0 | 0 | 3 | 1039.33 | 516352 | 82.57
0 | 0 | 1.8 | 982.36 | 488049.2 | 83.44 | 0 | 0 | 1.8 | 982.36 | 488049.2 | 80.91
0 | 0 | 0.6 | 983.85 | 488788.2 | 82.82 | 0 | 0 | 0.6 | 983.85 | 488788.2 | 80.2
0 | 0 | -0.5 | 1017.72 | 505615.6 | 83 | 0 | 0 | -0.5 | 1017.72 | 505615.6 | 80.36