## rendements DSP
Measurement time: | 07/09/2022 19:04:57 |  |  |  |  | Measurement time: | 29/09/2022 19:18:35
Radiation device: | Clinac3 Accelerator |  |  |  |  | Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV |  |  |  |  | Energy: | 12 MeV
Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS |  |  |  |  | Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie |  |  |  |  | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 1100 mm |  |  |  |  | SSD: | 1000 mm
Field size: | 100 x 100 mm |  |  |  |  | Field size: | 100 x 100 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Beam |  |  |  |  | Scan type: | Beam
Scan mode: | Continuous |  |  |  |  | Scan mode: | Continuous
Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 0.69 | 3849 | 1.22 | 0 | 0 | 150 | 1.24 | 7053.7 | 1.11
0 | 0 | 149.3 | 0.99 | 5475.5 | 1.14 | 0 | 0 | 149.1 | 1.1 | 6288.6 | 1.07
0 | 0 | 147.9 | 1.01 | 5578.3 | 1.14 | 0 | 0 | 147.6 | 1.25 | 7138.1 | 1.1
0 | 0 | 146.7 | 1.02 | 5656.5 | 1.16 | 0 | 0 | 146.6 | 1.14 | 6513.3 | 1.09
0 | 0 | 145.5 | 1.03 | 5687.8 | 1.16 | 0 | 0 | 145.3 | 1.12 | 6370 | 1.11
0 | 0 | 144.3 | 1.05 | 5829.9 | 1.18 | 0 | 0 | 144.1 | 1.24 | 7059.8 | 1.12
0 | 0 | 143.1 | 1.04 | 5794 | 1.18 | 0 | 0 | 142.9 | 1.2 | 6842.5 | 1.14
0 | 0 | 141.9 | 1.06 | 5903.3 | 1.2 | 0 | 0 | 141.7 | 1.26 | 7162.8 | 1.15
0 | 0 | 140.7 | 1.06 | 5895.8 | 1.19 | 0 | 0 | 140.5 | 1.17 | 6650.7 | 1.15
0 | 0 | 139.5 | 1.08 | 5984.5 | 1.21 | 0 | 0 | 139.3 | 1.17 | 6655.2 | 1.16
0 | 0 | 138.3 | 1.03 | 5698.3 | 1.21 | 0 | 0 | 138.1 | 1.32 | 7495.2 | 1.19
0 | 0 | 137.1 | 1.09 | 6018.2 | 1.21 | 0 | 0 | 136.9 | 1.23 | 6990.5 | 1.19
0 | 0 | 135.9 | 1.11 | 6172.4 | 1.25 | 0 | 0 | 135.7 | 1.23 | 7016.4 | 1.18
0 | 0 | 134.7 | 1.11 | 6172.5 | 1.24 | 0 | 0 | 134.5 | 1.33 | 7565.9 | 1.2
0 | 0 | 133.5 | 1.13 | 6246.1 | 1.26 | 0 | 0 | 133.3 | 1.22 | 6943.3 | 1.2
0 | 0 | 132.3 | 1.11 | 6182.6 | 1.25 | 0 | 0 | 132.1 | 1.33 | 7588.3 | 1.22
0 | 0 | 131.1 | 1.13 | 6261.9 | 1.26 | 0 | 0 | 130.9 | 1.3 | 7375.1 | 1.22
0 | 0 | 129.9 | 1.13 | 6247.9 | 1.27 | 0 | 0 | 129.7 | 1.27 | 7216.7 | 1.23
0 | 0 | 128.7 | 1.16 | 6411 | 1.29 | 0 | 0 | 128.5 | 1.38 | 7864 | 1.22
0 | 0 | 127.5 | 1.14 | 6327.7 | 1.28 | 0 | 0 | 127.3 | 1.28 | 7276.8 | 1.25
0 | 0 | 126.3 | 1.16 | 6415.5 | 1.29 | 0 | 0 | 126 | 1.29 | 7356.1 | 1.25
0 | 0 | 125.1 | 1.16 | 6457.5 | 1.3 | 0 | 0 | 124.9 | 1.3 | 7425.9 | 1.27
0 | 0 | 123.9 | 1.18 | 6530.4 | 1.32 | 0 | 0 | 123.6 | 1.42 | 8066.2 | 1.29
0 | 0 | 122.7 | 1.18 | 6541.6 | 1.32 | 0 | 0 | 122.4 | 1.38 | 7828.8 | 1.29
0 | 0 | 121.5 | 1.2 | 6628.1 | 1.34 | 0 | 0 | 121.3 | 1.34 | 7606.8 | 1.3
0 | 0 | 120.3 | 1.2 | 6642.2 | 1.35 | 0 | 0 | 120 | 1.42 | 8100 | 1.29
0 | 0 | 119.1 | 1.2 | 6636.4 | 1.35 | 0 | 0 | 118.9 | 1.38 | 7860.3 | 1.32
0 | 0 | 117.9 | 1.2 | 6638.7 | 1.36 | 0 | 0 | 117.6 | 1.36 | 7751.1 | 1.31
0 | 0 | 116.7 | 1.24 | 6867.2 | 1.36 | 0 | 0 | 116.4 | 1.44 | 8198.9 | 1.35
0 | 0 | 115.5 | 1.22 | 6749.6 | 1.4 | 0 | 0 | 115.2 | 1.37 | 7781.6 | 1.34
0 | 0 | 114.3 | 1.22 | 6740.7 | 1.39 | 0 | 0 | 114 | 1.5 | 8516.1 | 1.35
0 | 0 | 113.1 | 1.23 | 6824.8 | 1.41 | 0 | 0 | 112.8 | 1.39 | 7911.4 | 1.35
0 | 0 | 111.9 | 1.19 | 6612.6 | 1.42 | 0 | 0 | 111.6 | 1.44 | 8208.1 | 1.4
0 | 0 | 110.7 | 1.3 | 7203.7 | 1.42 | 0 | 0 | 110.4 | 1.49 | 8459.5 | 1.39
0 | 0 | 109.5 | 1.26 | 6964.8 | 1.43 | 0 | 0 | 109.2 | 1.45 | 8239.2 | 1.42
0 | 0 | 108.3 | 1.21 | 6721.4 | 1.45 | 0 | 0 | 108 | 1.54 | 8778.8 | 1.41
0 | 0 | 107.1 | 1.32 | 7339.1 | 1.47 | 0 | 0 | 106.8 | 1.48 | 8403.7 | 1.43
0 | 0 | 105.9 | 1.27 | 7038 | 1.46 | 0 | 0 | 105.6 | 1.48 | 8421.6 | 1.45
0 | 0 | 104.7 | 1.25 | 6906.6 | 1.48 | 0 | 0 | 104.4 | 1.56 | 8884 | 1.45
0 | 0 | 103.5 | 1.36 | 7565.4 | 1.5 | 0 | 0 | 103.2 | 1.52 | 8633.4 | 1.48
0 | 0 | 102.3 | 1.27 | 7045.1 | 1.52 | 0 | 0 | 102 | 1.62 | 9226.3 | 1.49
0 | 0 | 101.1 | 1.34 | 7433.5 | 1.53 | 0 | 0 | 100.8 | 1.53 | 8697.6 | 1.48
0 | 0 | 99.9 | 1.4 | 7774.5 | 1.54 | 0 | 0 | 99.6 | 1.52 | 8636.7 | 1.5
0 | 0 | 98.7 | 1.31 | 7254.5 | 1.55 | 0 | 0 | 98.4 | 1.66 | 9462.9 | 1.52
0 | 0 | 97.5 | 1.37 | 7617.5 | 1.56 | 0 | 0 | 97.2 | 1.57 | 8946.2 | 1.56
0 | 0 | 96.3 | 1.39 | 7689.5 | 1.58 | 0 | 0 | 95.9 | 1.68 | 9536.3 | 1.54
0 | 0 | 95.1 | 1.4 | 7742.8 | 1.6 | 0 | 0 | 94.8 | 1.58 | 8974.1 | 1.53
0 | 0 | 93.9 | 1.4 | 7774.7 | 1.6 | 0 | 0 | 93.5 | 1.6 | 9117.1 | 1.57
0 | 0 | 92.7 | 1.41 | 7792.2 | 1.6 | 0 | 0 | 92.3 | 1.73 | 9821.1 | 1.59
0 | 0 | 91.4 | 1.42 | 7874.6 | 1.63 | 0 | 0 | 91.2 | 1.63 | 9249.9 | 1.62
0 | 0 | 90.3 | 1.43 | 7943.2 | 1.63 | 0 | 0 | 89.9 | 1.73 | 9845.2 | 1.6
0 | 0 | 89.1 | 1.45 | 8065.8 | 1.65 | 0 | 0 | 88.7 | 1.65 | 9366.3 | 1.61
0 | 0 | 87.9 | 1.46 | 8076.6 | 1.67 | 0 | 0 | 87.5 | 1.67 | 9522.9 | 1.65
0 | 0 | 86.6 | 1.46 | 8086.7 | 1.65 | 0 | 0 | 86.3 | 1.82 | 10383 | 1.67
0 | 0 | 85.4 | 1.46 | 8108 | 1.67 | 0 | 0 | 85.1 | 1.67 | 9497.1 | 1.66
0 | 0 | 84.2 | 1.52 | 8414.3 | 1.72 | 0 | 0 | 83.9 | 1.81 | 10277.1 | 1.69
0 | 0 | 83 | 1.5 | 8332 | 1.7 | 0 | 0 | 82.7 | 1.76 | 9996.7 | 1.69
0 | 0 | 81.8 | 1.51 | 8358.3 | 1.72 | 0 | 0 | 81.5 | 1.73 | 9820.7 | 1.69
0 | 0 | 80.6 | 1.54 | 8543.3 | 1.75 | 0 | 0 | 80.3 | 1.84 | 10461.6 | 1.71
0 | 0 | 79.4 | 1.55 | 8606.4 | 1.76 | 0 | 0 | 79.1 | 1.76 | 10022.4 | 1.77
0 | 0 | 78.2 | 1.56 | 8649.1 | 1.75 | 0 | 0 | 77.9 | 1.83 | 10397.9 | 1.73
0 | 0 | 77 | 1.59 | 8841.7 | 1.79 | 0 | 0 | 76.7 | 1.8 | 10254.8 | 1.75
0 | 0 | 75.8 | 1.49 | 8237.4 | 1.79 | 0 | 0 | 75.4 | 1.82 | 10378.2 | 1.78
0 | 0 | 74.6 | 1.61 | 8944.2 | 1.81 | 0 | 0 | 74.2 | 1.91 | 10846.9 | 1.78
0 | 0 | 73.3 | 1.61 | 8949 | 1.84 | 0 | 0 | 73.1 | 1.83 | 10426.2 | 1.81
0 | 0 | 72.2 | 1.57 | 8681.4 | 1.85 | 0 | 0 | 71.8 | 1.86 | 10606.3 | 1.81
0 | 0 | 70.9 | 1.65 | 9158.7 | 1.87 | 0 | 0 | 70.6 | 1.88 | 10707 | 1.84
0 | 0 | 69.7 | 1.68 | 9306.8 | 1.91 | 0 | 0 | 69.4 | 1.96 | 11140 | 1.88
0 | 0 | 68.5 | 1.78 | 9892 | 2.02 | 0 | 0 | 68.2 | 2.11 | 12027 | 1.94
0 | 0 | 67.3 | 1.89 | 10503.9 | 2.24 | 0 | 0 | 67 | 2.06 | 11746.2 | 2.07
0 | 0 | 66.1 | 2.26 | 12541.6 | 2.58 | 0 | 0 | 65.8 | 2.37 | 13495.9 | 2.3
0 | 0 | 64.9 | 2.73 | 15152.4 | 3.11 | 0 | 0 | 64.6 | 2.78 | 15828.8 | 2.74
0 | 0 | 63.7 | 3.54 | 19632.7 | 4.03 | 0 | 0 | 63.4 | 3.63 | 20657 | 3.43
0 | 0 | 62.5 | 4.65 | 25786.7 | 5.33 | 0 | 0 | 62.2 | 4.9 | 27869.8 | 4.52
0 | 0 | 61.3 | 6.51 | 36127.4 | 7.32 | 0 | 0 | 61 | 6.45 | 36712.8 | 6.39
0 | 0 | 60.1 | 8.71 | 48326 | 9.84 | 0 | 0 | 59.7 | 8.24 | 46912 | 8.06
0 | 0 | 58.9 | 11.24 | 62340.8 | 12.79 | 0 | 0 | 58.6 | 11.14 | 63397.9 | 10.97
0 | 0 | 57.7 | 14.73 | 81707.5 | 16.62 | 0 | 0 | 57.3 | 15.39 | 87572 | 14.5
0 | 0 | 56.5 | 17.63 | 97794.3 | 20.8 | 0 | 0 | 56.1 | 19.73 | 112317 | 18.16
0 | 0 | 55.3 | 22.46 | 124575.8 | 25.63 | 0 | 0 | 54.9 | 23.56 | 134097.7 | 23.16
0 | 0 | 54.1 | 27.4 | 151935.8 | 30.85 | 0 | 0 | 53.7 | 28.38 | 161511.3 | 27.72
0 | 0 | 52.9 | 32.25 | 178874.8 | 36.07 | 0 | 0 | 52.5 | 33.19 | 188911.6 | 32.4
0 | 0 | 51.7 | 37.24 | 206539.4 | 41.87 | 0 | 0 | 51.3 | 41.7 | 237335.8 | 38.69
0 | 0 | 50.5 | 41.84 | 232048.3 | 47.24 | 0 | 0 | 50.1 | 48.06 | 273529.8 | 43.66
0 | 0 | 49.3 | 47.24 | 262004.5 | 52.81 | 0 | 0 | 48.9 | 50.91 | 289741.6 | 49.68
0 | 0 | 48.1 | 52.02 | 288479.6 | 58.57 | 0 | 0 | 47.7 | 57.53 | 327471.1 | 55.74
0 | 0 | 46.9 | 56.27 | 312094.1 | 63.64 | 0 | 0 | 46.5 | 62.01 | 352948.7 | 59.65
0 | 0 | 45.6 | 60.99 | 338264.3 | 68.96 | 0 | 0 | 45.3 | 72.67 | 413639.8 | 66.31
0 | 0 | 44.4 | 64.47 | 357543.3 | 73.64 | 0 | 0 | 44.1 | 77.28 | 439847.2 | 70.47
0 | 0 | 43.2 | 70.69 | 392034.8 | 77.87 | 0 | 0 | 42.9 | 79.14 | 450472.8 | 75.02
0 | 0 | 42 | 73.97 | 410239.2 | 82.1 | 0 | 0 | 41.7 | 82.92 | 471942.6 | 79.78
0 | 0 | 40.8 | 75.1 | 416501.6 | 85.4 | 0 | 0 | 40.5 | 86.96 | 494954.2 | 82.54
0 | 0 | 39.6 | 78.67 | 436269.3 | 88.37 | 0 | 0 | 39.2 | 94.94 | 540384.4 | 86.34
0 | 0 | 38.4 | 84.34 | 467751.4 | 91.27 | 0 | 0 | 38.1 | 98.6 | 561191.7 | 89.2
0 | 0 | 37.2 | 82.52 | 457644.2 | 93.29 | 0 | 0 | 36.8 | 96.86 | 551312.5 | 91.72
0 | 0 | 36 | 86.07 | 477332.1 | 95.24 | 0 | 0 | 35.6 | 99 | 563489 | 94.09
0 | 0 | 34.8 | 90.95 | 504426.5 | 96.86 | 0 | 0 | 34.4 | 102.01 | 580594.8 | 95.7
0 | 0 | 33.6 | 87.2 | 483606.1 | 97.99 | 0 | 0 | 33.2 | 108 | 614686 | 96.78
0 | 0 | 32.4 | 90.53 | 502051.6 | 98.75 | 0 | 0 | 32 | 111.5 | 634613.4 | 98.34
0 | 0 | 31.2 | 94.83 | 525892.8 | 99.43 | 0 | 0 | 30.8 | 103.76 | 590566.5 | 99.05
0 | 0 | 30 | 89.39 | 495772.4 | 99.63 | 0 | 0 | 29.6 | 113.27 | 644725.4 | 99.75
0 | 0 | 28.8 | 92.21 | 511386.3 | 100 | 0 | 0 | 28.4 | 107.98 | 614622.2 | 99.92
0 | 0 | 27.6 | 96.09 | 532918 | 99.93 | 0 | 0 | 27.2 | 111.85 | 636599.7 | 100
0 | 0 | 26.4 | 89.59 | 496848.2 | 99.56 | 0 | 0 | 26 | 113.89 | 648235.3 | 100
0 | 0 | 25.1 | 93.62 | 519196.4 | 99.39 | 0 | 0 | 24.8 | 105.99 | 603240 | 100
0 | 0 | 23.9 | 91.96 | 509977.5 | 98.65 | 0 | 0 | 23.6 | 112.96 | 642938 | 99.73
0 | 0 | 22.7 | 89.29 | 495202.9 | 98.14 | 0 | 0 | 22.3 | 106.71 | 607350.3 | 99.21
0 | 0 | 21.5 | 92.53 | 513161.9 | 97.93 | 0 | 0 | 21.1 | 111.46 | 634379.1 | 98.96
0 | 0 | 20.3 | 90.16 | 500003.2 | 97.17 | 0 | 0 | 19.9 | 112.95 | 642855.8 | 98.41
0 | 0 | 19.1 | 88.63 | 491511 | 96.56 | 0 | 0 | 18.7 | 104.1 | 592538.5 | 98.11
0 | 0 | 17.9 | 91.16 | 505569.6 | 96.27 | 0 | 0 | 17.5 | 111.43 | 634228.9 | 97.59
0 | 0 | 16.7 | 93.68 | 519566.5 | 95.76 | 0 | 0 | 16.3 | 105.04 | 597872.9 | 97.12
0 | 0 | 15.5 | 87.66 | 486164.5 | 94.91 | 0 | 0 | 15.1 | 111.25 | 633205.7 | 96.61
0 | 0 | 14.3 | 89.3 | 495254.2 | 94.49 | 0 | 0 | 13.9 | 111.55 | 634920.3 | 96.14
0 | 0 | 13.1 | 92.38 | 512331.3 | 93.82 | 0 | 0 | 12.7 | 102.78 | 584993.3 | 95.56
0 | 0 | 11.9 | 86.01 | 476994 | 93.26 | 0 | 0 | 11.5 | 107.88 | 613998.7 | 95.21
0 | 0 | 10.7 | 88.66 | 491702.6 | 92.84 | 0 | 0 | 10.3 | 106.39 | 605570.4 | 94.61
0 | 0 | 9.5 | 90.42 | 501444.4 | 92.08 | 0 | 0 | 9.1 | 110.57 | 629313.5 | 93.95
0 | 0 | 8.3 | 84.66 | 469542.9 | 91.49 | 0 | 0 | 7.9 | 110.18 | 627098.9 | 93.71
0 | 0 | 7.1 | 87.45 | 484978.8 | 90.61 | 0 | 0 | 6.7 | 102.13 | 581302.6 | 93.1
0 | 0 | 5.9 | 88.68 | 491794.1 | 90.1 | 0 | 0 | 5.4 | 105.35 | 599601.3 | 92.42
0 | 0 | 4.7 | 82.6 | 458076.9 | 89.26 | 0 | 0 | 4.2 | 102.76 | 584904.4 | 91.59
0 | 0 | 3.5 | 84.68 | 469621.5 | 88.15 | 0 | 0 | 3 | 108.8 | 619288.9 | 90.85
0 | 0 | 2.3 | 86.65 | 480565.8 | 87.17 | 0 | 0 | 1.8 | 100.12 | 569879.3 | 90.04
0 | 0 | 1.1 | 79.98 | 443587.7 | 86.65 | 0 | 0 | 0.6 | 98.99 | 563413.5 | 89.42
0 | 0 | -0.5 | 83.83 | 464911.6 | 86.72 | 0 | 0 | -0.5 | 105.85 | 602447 | 89.15
Measurement time: | 07/09/2022 19:06:57
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1050 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 1.28 | 7090.2 | 1.16
0 | 0 | 149.3 | 1.11 | 6139.9 | 1.13
0 | 0 | 147.9 | 1.11 | 6182.7 | 1.13
0 | 0 | 146.7 | 1.11 | 6137.6 | 1.13
0 | 0 | 145.5 | 1.1 | 6114.5 | 1.13
0 | 0 | 144.3 | 1.14 | 6312.3 | 1.17
0 | 0 | 143.1 | 1.12 | 6239.1 | 1.16
0 | 0 | 141.9 | 1.15 | 6355.4 | 1.17
0 | 0 | 140.7 | 1.16 | 6434.7 | 1.19
0 | 0 | 139.5 | 1.14 | 6348.1 | 1.18
0 | 0 | 138.3 | 1.16 | 6440.8 | 1.19
0 | 0 | 137.1 | 1.17 | 6514.5 | 1.2
0 | 0 | 135.9 | 1.19 | 6610.5 | 1.23
0 | 0 | 134.7 | 1.18 | 6531.8 | 1.21
0 | 0 | 133.5 | 1.19 | 6598.1 | 1.22
0 | 0 | 132.3 | 1.18 | 6564.7 | 1.22
0 | 0 | 131.1 | 1.19 | 6599.5 | 1.23
0 | 0 | 129.9 | 1.24 | 6893.8 | 1.26
0 | 0 | 128.7 | 1.21 | 6709.1 | 1.25
0 | 0 | 127.5 | 1.22 | 6767.2 | 1.27
0 | 0 | 126.3 | 1.25 | 6927.3 | 1.29
0 | 0 | 125.1 | 1.25 | 6931.7 | 1.29
0 | 0 | 123.9 | 1.25 | 6908.8 | 1.31
0 | 0 | 122.7 | 1.31 | 7242.1 | 1.31
0 | 0 | 121.5 | 1.25 | 6944.4 | 1.3
0 | 0 | 120.3 | 1.27 | 7045.6 | 1.31
0 | 0 | 119.1 | 1.34 | 7412.2 | 1.34
0 | 0 | 117.9 | 1.29 | 7147.2 | 1.32
0 | 0 | 116.7 | 1.29 | 7154.2 | 1.36
0 | 0 | 115.5 | 1.36 | 7563.7 | 1.36
0 | 0 | 114.3 | 1.29 | 7181.9 | 1.37
0 | 0 | 113.1 | 1.34 | 7433.4 | 1.38
0 | 0 | 111.9 | 1.4 | 7756.6 | 1.4
0 | 0 | 110.7 | 1.34 | 7405.3 | 1.41
0 | 0 | 109.5 | 1.39 | 7734.6 | 1.43
0 | 0 | 108.3 | 1.44 | 7988.6 | 1.44
0 | 0 | 107.1 | 1.39 | 7708.9 | 1.47
0 | 0 | 105.8 | 1.43 | 7905 | 1.46
0 | 0 | 104.7 | 1.49 | 8258 | 1.48
0 | 0 | 103.4 | 1.4 | 7765.8 | 1.5
0 | 0 | 102.2 | 1.46 | 8071 | 1.5
0 | 0 | 101 | 1.53 | 8473.6 | 1.52
0 | 0 | 99.8 | 1.43 | 7954.6 | 1.53
0 | 0 | 98.6 | 1.49 | 8269.1 | 1.54
0 | 0 | 97.4 | 1.56 | 8641.3 | 1.54
0 | 0 | 96.2 | 1.47 | 8130 | 1.56
0 | 0 | 95 | 1.54 | 8552.7 | 1.58
0 | 0 | 93.8 | 1.59 | 8834.7 | 1.58
0 | 0 | 92.6 | 1.52 | 8430.4 | 1.62
0 | 0 | 91.4 | 1.56 | 8666.1 | 1.61
0 | 0 | 90.2 | 1.66 | 9217.9 | 1.64
0 | 0 | 89 | 1.54 | 8523.9 | 1.64
0 | 0 | 87.8 | 1.64 | 9087.1 | 1.69
0 | 0 | 86.6 | 1.71 | 9459 | 1.67
0 | 0 | 85.4 | 1.55 | 8618.9 | 1.68
0 | 0 | 84.2 | 1.65 | 9167.7 | 1.7
0 | 0 | 83 | 1.73 | 9584.3 | 1.7
0 | 0 | 81.8 | 1.61 | 8901.8 | 1.73
0 | 0 | 80.6 | 1.7 | 9433.3 | 1.75
0 | 0 | 79.4 | 1.76 | 9773.8 | 1.73
0 | 0 | 78.1 | 1.65 | 9124.4 | 1.76
0 | 0 | 77 | 1.73 | 9577.3 | 1.77
0 | 0 | 75.8 | 1.83 | 10157.8 | 1.82
0 | 0 | 74.5 | 1.69 | 9391.8 | 1.81
0 | 0 | 73.3 | 1.77 | 9817.9 | 1.83
0 | 0 | 72.1 | 1.83 | 10140.8 | 1.82
0 | 0 | 70.9 | 1.76 | 9748.7 | 1.86
0 | 0 | 69.7 | 1.86 | 10307.6 | 1.92
0 | 0 | 68.5 | 2.02 | 11196.6 | 2.01
0 | 0 | 67.3 | 2.09 | 11592.3 | 2.24
0 | 0 | 66.1 | 2.47 | 13684.9 | 2.56
0 | 0 | 64.9 | 3.14 | 17401.3 | 3.18
0 | 0 | 63.7 | 3.81 | 21148.1 | 4.11
0 | 0 | 62.5 | 5.24 | 29068.4 | 5.43
0 | 0 | 61.3 | 7.35 | 40771.7 | 7.41
0 | 0 | 60.1 | 9.09 | 50428.5 | 9.85
0 | 0 | 58.9 | 12.4 | 68758.6 | 12.96
0 | 0 | 57.7 | 16.56 | 91835.1 | 16.75
0 | 0 | 56.5 | 19.85 | 110103.9 | 21.02
0 | 0 | 55.3 | 24.79 | 137495 | 25.74
0 | 0 | 54.1 | 30.85 | 171092 | 30.95
0 | 0 | 52.9 | 34.91 | 193630 | 36.25
0 | 0 | 51.7 | 40.43 | 224227.6 | 42.18
0 | 0 | 50.4 | 45.77 | 253823 | 47.62
0 | 0 | 49.2 | 52.03 | 288561.6 | 53.32
0 | 0 | 48 | 55.86 | 309775.2 | 58.87
0 | 0 | 46.8 | 63.41 | 351658.3 | 63.87
0 | 0 | 45.6 | 66.4 | 368237.4 | 69.18
0 | 0 | 44.4 | 72.66 | 402965.6 | 73.7
0 | 0 | 43.2 | 78.42 | 434896.5 | 78.07
0 | 0 | 42 | 80.32 | 445438.9 | 82.2
0 | 0 | 40.8 | 85.2 | 472519.4 | 85.55
0 | 0 | 39.6 | 89.43 | 495945.7 | 88.42
0 | 0 | 38.4 | 89.89 | 498527.4 | 91.39
0 | 0 | 37.2 | 94.02 | 521436.9 | 93.63
0 | 0 | 35.9 | 97.07 | 538334.9 | 95.38
0 | 0 | 34.8 | 97.34 | 539814.9 | 96.95
0 | 0 | 33.6 | 99.09 | 549544.8 | 98.22
0 | 0 | 32.4 | 101.53 | 563081.1 | 99.01
0 | 0 | 31.1 | 100.91 | 559633 | 99.52
0 | 0 | 29.9 | 102.24 | 566996.8 | 99.99
0 | 0 | 28.7 | 103.68 | 574988.3 | 100
0 | 0 | 27.5 | 102.09 | 566173.6 | 99.78
0 | 0 | 26.3 | 102.63 | 569169.6 | 99.84
0 | 0 | 25.1 | 103.36 | 573226.3 | 99.55
0 | 0 | 23.9 | 102.51 | 568506.9 | 99.01
0 | 0 | 22.7 | 101.72 | 564131.3 | 98.73
0 | 0 | 21.5 | 102.67 | 569404.4 | 98.12
0 | 0 | 20.3 | 100.87 | 559437.6 | 97.76
0 | 0 | 19.1 | 102.25 | 567072.4 | 97.23
0 | 0 | 17.9 | 100.85 | 559286.9 | 96.72
0 | 0 | 16.7 | 100.05 | 554874.5 | 96.18
0 | 0 | 15.5 | 99.49 | 551773.4 | 95.67
0 | 0 | 14.3 | 100.35 | 556532.7 | 95.02
0 | 0 | 13.1 | 98.22 | 544715.1 | 94.34
0 | 0 | 11.8 | 98.57 | 546680.8 | 93.88
0 | 0 | 10.6 | 98.05 | 543766.9 | 93.31
0 | 0 | 9.4 | 96.77 | 536653.5 | 92.61
0 | 0 | 8.2 | 96.85 | 537144.5 | 92.21
0 | 0 | 7 | 96.55 | 535477.7 | 91.53
0 | 0 | 5.8 | 95.08 | 527317.6 | 90.83
0 | 0 | 4.6 | 94.54 | 524293.8 | 90.01
0 | 0 | 3.4 | 93.93 | 520909.9 | 89.01
0 | 0 | 2.2 | 91.57 | 507819.5 | 87.91
0 | 0 | 1 | 91.38 | 506786 | 87.54
0 | 0 | -0.5 | 92.52 | 513103.9 | 87.7