## rendements taille de champ
Measurement time: | 07/09/2022 19:11:32 |  |  |  |  | Measurement time: | 29/09/2022 19:18:35
Radiation device: | Clinac3 Accelerator |  |  |  |  | Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV |  |  |  |  | Energy: | 12 MeV
Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS |  |  |  |  | Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie |  |  |  |  | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 1000 mm |  |  |  |  | SSD: | 1000 mm
Field size: | 60 x 60 mm |  |  |  |  | Field size: | 100 x 100 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Beam |  |  |  |  | Scan type: | Beam
Scan mode: | Continuous |  |  |  |  | Scan mode: | Continuous
Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 0.36 | 2018.7 | 0.57 | 0 | 0 | 150 | 1.24 | 7053.7 | 1.11
0 | 0 | 149.3 | 1.04 | 5746 | 0.98 | 0 | 0 | 149.1 | 1.1 | 6288.6 | 1.07
0 | 0 | 147.9 | 1.09 | 6059.1 | 1.01 | 0 | 0 | 147.6 | 1.25 | 7138.1 | 1.1
0 | 0 | 146.7 | 1.06 | 5854.3 | 1 | 0 | 0 | 146.6 | 1.14 | 6513.3 | 1.09
0 | 0 | 145.5 | 1.08 | 5999.7 | 1 | 0 | 0 | 145.3 | 1.12 | 6370 | 1.11
0 | 0 | 144.3 | 1.09 | 6049.6 | 1.02 | 0 | 0 | 144.1 | 1.24 | 7059.8 | 1.12
0 | 0 | 143.1 | 1.09 | 6047.4 | 1.02 | 0 | 0 | 142.9 | 1.2 | 6842.5 | 1.14
0 | 0 | 141.9 | 1.11 | 6168.9 | 1.04 | 0 | 0 | 141.7 | 1.26 | 7162.8 | 1.15
0 | 0 | 140.7 | 1.1 | 6083.7 | 1.03 | 0 | 0 | 140.5 | 1.17 | 6650.7 | 1.15
0 | 0 | 139.5 | 1.12 | 6185.4 | 1.05 | 0 | 0 | 139.3 | 1.17 | 6655.2 | 1.16
0 | 0 | 138.3 | 1.12 | 6200.2 | 1.05 | 0 | 0 | 138.1 | 1.32 | 7495.2 | 1.19
0 | 0 | 137.1 | 1.16 | 6419.3 | 1.08 | 0 | 0 | 136.9 | 1.23 | 6990.5 | 1.19
0 | 0 | 135.9 | 1.15 | 6379.1 | 1.08 | 0 | 0 | 135.7 | 1.23 | 7016.4 | 1.18
0 | 0 | 134.7 | 1.15 | 6397.7 | 1.08 | 0 | 0 | 134.5 | 1.33 | 7565.9 | 1.2
0 | 0 | 133.5 | 1.18 | 6529.9 | 1.1 | 0 | 0 | 133.3 | 1.22 | 6943.3 | 1.2
0 | 0 | 132.3 | 1.17 | 6488.1 | 1.09 | 0 | 0 | 132.1 | 1.33 | 7588.3 | 1.22
0 | 0 | 131.1 | 1.2 | 6677.3 | 1.11 | 0 | 0 | 130.9 | 1.3 | 7375.1 | 1.22
0 | 0 | 129.9 | 1.19 | 6609.6 | 1.12 | 0 | 0 | 129.7 | 1.27 | 7216.7 | 1.23
0 | 0 | 128.7 | 1.2 | 6675.2 | 1.12 | 0 | 0 | 128.5 | 1.38 | 7864 | 1.22
0 | 0 | 127.5 | 1.24 | 6901.4 | 1.15 | 0 | 0 | 127.3 | 1.28 | 7276.8 | 1.25
0 | 0 | 126.3 | 1.22 | 6752.2 | 1.14 | 0 | 0 | 126 | 1.29 | 7356.1 | 1.25
0 | 0 | 125.1 | 1.25 | 6910.4 | 1.16 | 0 | 0 | 124.9 | 1.3 | 7425.9 | 1.27
0 | 0 | 123.9 | 1.24 | 6893.2 | 1.16 | 0 | 0 | 123.6 | 1.42 | 8066.2 | 1.29
0 | 0 | 122.7 | 1.27 | 7016.5 | 1.18 | 0 | 0 | 122.4 | 1.38 | 7828.8 | 1.29
0 | 0 | 121.5 | 1.27 | 7034.3 | 1.19 | 0 | 0 | 121.3 | 1.34 | 7606.8 | 1.3
0 | 0 | 120.3 | 1.28 | 7096.2 | 1.19 | 0 | 0 | 120 | 1.42 | 8100 | 1.29
0 | 0 | 119.1 | 1.3 | 7226.7 | 1.22 | 0 | 0 | 118.9 | 1.38 | 7860.3 | 1.32
0 | 0 | 117.9 | 1.32 | 7314.4 | 1.23 | 0 | 0 | 117.6 | 1.36 | 7751.1 | 1.31
0 | 0 | 116.7 | 1.32 | 7293.9 | 1.23 | 0 | 0 | 116.4 | 1.44 | 8198.9 | 1.35
0 | 0 | 115.5 | 1.32 | 7312.7 | 1.24 | 0 | 0 | 115.2 | 1.37 | 7781.6 | 1.34
0 | 0 | 114.3 | 1.36 | 7524.5 | 1.27 | 0 | 0 | 114 | 1.5 | 8516.1 | 1.35
0 | 0 | 113.1 | 1.34 | 7442.3 | 1.27 | 0 | 0 | 112.8 | 1.39 | 7911.4 | 1.35
0 | 0 | 111.9 | 1.37 | 7574.9 | 1.28 | 0 | 0 | 111.6 | 1.44 | 8208.1 | 1.4
0 | 0 | 110.7 | 1.37 | 7598.7 | 1.28 | 0 | 0 | 110.4 | 1.49 | 8459.5 | 1.39
0 | 0 | 109.5 | 1.38 | 7653.4 | 1.31 | 0 | 0 | 109.2 | 1.45 | 8239.2 | 1.42
0 | 0 | 108.3 | 1.4 | 7770.5 | 1.32 | 0 | 0 | 108 | 1.54 | 8778.8 | 1.41
0 | 0 | 107.1 | 1.41 | 7834.9 | 1.33 | 0 | 0 | 106.8 | 1.48 | 8403.7 | 1.43
0 | 0 | 105.8 | 1.43 | 7938.4 | 1.35 | 0 | 0 | 105.6 | 1.48 | 8421.6 | 1.45
0 | 0 | 104.7 | 1.43 | 7951 | 1.36 | 0 | 0 | 104.4 | 1.56 | 8884 | 1.45
0 | 0 | 103.4 | 1.47 | 8127.4 | 1.36 | 0 | 0 | 103.2 | 1.52 | 8633.4 | 1.48
0 | 0 | 102.3 | 1.44 | 7976.5 | 1.36 | 0 | 0 | 102 | 1.62 | 9226.3 | 1.49
0 | 0 | 101.1 | 1.47 | 8164.5 | 1.39 | 0 | 0 | 100.8 | 1.53 | 8697.6 | 1.48
0 | 0 | 99.8 | 1.5 | 8301.6 | 1.4 | 0 | 0 | 99.6 | 1.52 | 8636.7 | 1.5
0 | 0 | 98.7 | 1.5 | 8343.7 | 1.43 | 0 | 0 | 98.4 | 1.66 | 9462.9 | 1.52
0 | 0 | 97.4 | 1.51 | 8357.6 | 1.43 | 0 | 0 | 97.2 | 1.57 | 8946.2 | 1.56
0 | 0 | 96.2 | 1.59 | 8822.4 | 1.45 | 0 | 0 | 95.9 | 1.68 | 9536.3 | 1.54
0 | 0 | 95 | 1.51 | 8378 | 1.45 | 0 | 0 | 94.8 | 1.58 | 8974.1 | 1.53
0 | 0 | 93.8 | 1.58 | 8758.4 | 1.46 | 0 | 0 | 93.5 | 1.6 | 9117.1 | 1.57
0 | 0 | 92.6 | 1.63 | 9014.7 | 1.49 | 0 | 0 | 92.3 | 1.73 | 9821.1 | 1.59
0 | 0 | 91.4 | 1.57 | 8697 | 1.51 | 0 | 0 | 91.2 | 1.63 | 9249.9 | 1.62
0 | 0 | 90.2 | 1.65 | 9124.2 | 1.52 | 0 | 0 | 89.9 | 1.73 | 9845.2 | 1.6
0 | 0 | 89 | 1.61 | 8936.4 | 1.54 | 0 | 0 | 88.7 | 1.65 | 9366.3 | 1.61
0 | 0 | 87.8 | 1.61 | 8947.3 | 1.55 | 0 | 0 | 87.5 | 1.67 | 9522.9 | 1.65
0 | 0 | 86.6 | 1.67 | 9277.2 | 1.57 | 0 | 0 | 86.3 | 1.82 | 10383 | 1.67
0 | 0 | 85.4 | 1.75 | 9679.3 | 1.58 | 0 | 0 | 85.1 | 1.67 | 9497.1 | 1.66
0 | 0 | 84.2 | 1.65 | 9152.9 | 1.58 | 0 | 0 | 83.9 | 1.81 | 10277.1 | 1.69
0 | 0 | 83 | 1.62 | 8973.8 | 1.6 | 0 | 0 | 82.7 | 1.76 | 9996.7 | 1.69
0 | 0 | 81.8 | 1.81 | 10018.4 | 1.62 | 0 | 0 | 81.5 | 1.73 | 9820.7 | 1.69
0 | 0 | 80.6 | 1.71 | 9464.6 | 1.65 | 0 | 0 | 80.3 | 1.84 | 10461.6 | 1.71
0 | 0 | 79.4 | 1.69 | 9373.1 | 1.65 | 0 | 0 | 79.1 | 1.76 | 10022.4 | 1.77
0 | 0 | 78.2 | 1.86 | 10302.6 | 1.68 | 0 | 0 | 77.9 | 1.83 | 10397.9 | 1.73
0 | 0 | 76.9 | 1.74 | 9657.6 | 1.7 | 0 | 0 | 76.7 | 1.8 | 10254.8 | 1.75
0 | 0 | 75.7 | 1.81 | 10062 | 1.7 | 0 | 0 | 75.4 | 1.82 | 10378.2 | 1.78
0 | 0 | 74.6 | 1.84 | 10183 | 1.72 | 0 | 0 | 74.2 | 1.91 | 10846.9 | 1.78
0 | 0 | 73.3 | 1.77 | 9819.5 | 1.74 | 0 | 0 | 73.1 | 1.83 | 10426.2 | 1.81
0 | 0 | 72.1 | 1.9 | 10519.8 | 1.77 | 0 | 0 | 71.8 | 1.86 | 10606.3 | 1.81
0 | 0 | 70.9 | 1.99 | 11023.2 | 1.79 | 0 | 0 | 70.6 | 1.88 | 10707 | 1.84
0 | 0 | 69.7 | 1.88 | 10420.9 | 1.85 | 0 | 0 | 69.4 | 1.96 | 11140 | 1.88
0 | 0 | 68.5 | 2.1 | 11639.8 | 1.97 | 0 | 0 | 68.2 | 2.11 | 12027 | 1.94
0 | 0 | 67.3 | 2.39 | 13272.7 | 2.17 | 0 | 0 | 67 | 2.06 | 11746.2 | 2.07
0 | 0 | 66.1 | 2.54 | 14106 | 2.5 | 0 | 0 | 65.8 | 2.37 | 13495.9 | 2.3
0 | 0 | 64.9 | 3.3 | 18320.1 | 3.11 | 0 | 0 | 64.6 | 2.78 | 15828.8 | 2.74
0 | 0 | 63.7 | 4.48 | 24867.3 | 4.04 | 0 | 0 | 63.4 | 3.63 | 20657 | 3.43
0 | 0 | 62.5 | 5.62 | 31174.4 | 5.52 | 0 | 0 | 62.2 | 4.9 | 27869.8 | 4.52
0 | 0 | 61.3 | 7.8 | 43244.5 | 7.33 | 0 | 0 | 61 | 6.45 | 36712.8 | 6.39
0 | 0 | 60.1 | 11.04 | 61245.1 | 9.93 | 0 | 0 | 59.7 | 8.24 | 46912 | 8.06
0 | 0 | 58.9 | 13.42 | 74451.4 | 13.19 | 0 | 0 | 58.6 | 11.14 | 63397.9 | 10.97
0 | 0 | 57.7 | 17.72 | 98259.3 | 16.7 | 0 | 0 | 57.3 | 15.39 | 87572 | 14.5
0 | 0 | 56.5 | 23.39 | 129709 | 21.03 | 0 | 0 | 56.1 | 19.73 | 112317 | 18.16
0 | 0 | 55.3 | 26.28 | 145763.3 | 25.79 | 0 | 0 | 54.9 | 23.56 | 134097.7 | 23.16
0 | 0 | 54.1 | 32.56 | 180582.3 | 30.44 | 0 | 0 | 53.7 | 28.38 | 161511.3 | 27.72
0 | 0 | 52.9 | 40.37 | 223875.5 | 35.98 | 0 | 0 | 52.5 | 33.19 | 188911.6 | 32.4
0 | 0 | 51.7 | 42.61 | 236290.5 | 41.69 | 0 | 0 | 51.3 | 41.7 | 237335.8 | 38.69
0 | 0 | 50.4 | 50.27 | 278791.9 | 46.67 | 0 | 0 | 50.1 | 48.06 | 273529.8 | 43.66
0 | 0 | 49.2 | 59.24 | 328545.2 | 52.48 | 0 | 0 | 48.9 | 50.91 | 289741.6 | 49.68
0 | 0 | 48 | 59.58 | 330448.1 | 57.58 | 0 | 0 | 47.7 | 57.53 | 327471.1 | 55.74
0 | 0 | 46.8 | 67.98 | 377033.8 | 62.55 | 0 | 0 | 46.5 | 62.01 | 352948.7 | 59.65
0 | 0 | 45.6 | 76.88 | 426385.4 | 67.54 | 0 | 0 | 45.3 | 72.67 | 413639.8 | 66.31
0 | 0 | 44.4 | 75.22 | 417173.5 | 72.05 | 0 | 0 | 44.1 | 77.28 | 439847.2 | 70.47
0 | 0 | 43.2 | 83.05 | 460573.8 | 76.25 | 0 | 0 | 42.9 | 79.14 | 450472.8 | 75.02
0 | 0 | 42 | 91.05 | 504955.4 | 79.85 | 0 | 0 | 41.7 | 82.92 | 471942.6 | 79.78
0 | 0 | 40.8 | 87.93 | 487661.4 | 83.49 | 0 | 0 | 40.5 | 86.96 | 494954.2 | 82.54
0 | 0 | 39.6 | 95.29 | 528476.6 | 86.62 | 0 | 0 | 39.2 | 94.94 | 540384.4 | 86.34
0 | 0 | 38.4 | 101.98 | 565594.4 | 89.13 | 0 | 0 | 38.1 | 98.6 | 561191.7 | 89.2
0 | 0 | 37.2 | 97.57 | 541119.1 | 91.7 | 0 | 0 | 36.8 | 96.86 | 551312.5 | 91.72
0 | 0 | 36 | 104.12 | 577412.9 | 93.58 | 0 | 0 | 35.6 | 99 | 563489 | 94.09
0 | 0 | 34.8 | 109.95 | 609762.3 | 95.35 | 0 | 0 | 34.4 | 102.01 | 580594.8 | 95.7
0 | 0 | 33.6 | 103.34 | 573102.8 | 96.73 | 0 | 0 | 33.2 | 108 | 614686 | 96.78
0 | 0 | 32.4 | 109.37 | 606556.5 | 97.8 | 0 | 0 | 32 | 111.5 | 634613.4 | 98.34
0 | 0 | 31.2 | 113.46 | 629253.8 | 98.64 | 0 | 0 | 30.8 | 103.76 | 590566.5 | 99.05
0 | 0 | 30 | 108.02 | 599051.7 | 99.21 | 0 | 0 | 29.6 | 113.27 | 644725.4 | 99.75
0 | 0 | 28.8 | 112.56 | 624259.4 | 99.69 | 0 | 0 | 28.4 | 107.98 | 614622.2 | 99.92
0 | 0 | 27.5 | 115.85 | 642478.4 | 99.98 | 0 | 0 | 27.2 | 111.85 | 636599.7 | 100
0 | 0 | 26.3 | 110.76 | 614280.3 | 100 | 0 | 0 | 26 | 113.89 | 648235.3 | 100
0 | 0 | 25.1 | 113.13 | 627398.1 | 99.97 | 0 | 0 | 24.8 | 105.99 | 603240 | 100
0 | 0 | 23.9 | 115.66 | 641442.6 | 99.61 | 0 | 0 | 23.6 | 112.96 | 642938 | 99.73
0 | 0 | 22.7 | 111.06 | 615946.3 | 99.21 | 0 | 0 | 22.3 | 106.71 | 607350.3 | 99.21
0 | 0 | 21.5 | 113.49 | 629417.4 | 99.18 | 0 | 0 | 21.1 | 111.46 | 634379.1 | 98.96
0 | 0 | 20.3 | 115.59 | 641077.4 | 98.61 | 0 | 0 | 19.9 | 112.95 | 642855.8 | 98.41
0 | 0 | 19.1 | 110.91 | 615108.3 | 98.49 | 0 | 0 | 18.7 | 104.1 | 592538.5 | 98.11
0 | 0 | 17.9 | 113.13 | 627420.7 | 98.14 | 0 | 0 | 17.5 | 111.43 | 634228.9 | 97.59
0 | 0 | 16.7 | 109.51 | 607346.9 | 97.56 | 0 | 0 | 16.3 | 105.04 | 597872.9 | 97.12
0 | 0 | 15.5 | 109.91 | 609523.2 | 97.34 | 0 | 0 | 15.1 | 111.25 | 633205.7 | 96.61
0 | 0 | 14.3 | 112.37 | 623177 | 96.97 | 0 | 0 | 13.9 | 111.55 | 634920.3 | 96.14
0 | 0 | 13.1 | 113.11 | 627309.1 | 96.43 | 0 | 0 | 12.7 | 102.78 | 584993.3 | 95.56
0 | 0 | 11.9 | 109.51 | 607325.7 | 96.08 | 0 | 0 | 11.5 | 107.88 | 613998.7 | 95.21
0 | 0 | 10.7 | 110.75 | 614217.7 | 95.53 | 0 | 0 | 10.3 | 106.39 | 605570.4 | 94.61
0 | 0 | 9.5 | 107.39 | 595584.6 | 95.04 | 0 | 0 | 9.1 | 110.57 | 629313.5 | 93.95
0 | 0 | 8.3 | 108.11 | 599579.4 | 94.43 | 0 | 0 | 7.9 | 110.18 | 627098.9 | 93.71
0 | 0 | 7.1 | 109.55 | 607574.5 | 93.94 | 0 | 0 | 6.7 | 102.13 | 581302.6 | 93.1
0 | 0 | 5.9 | 110.8 | 614484.4 | 93.26 | 0 | 0 | 5.4 | 105.35 | 599601.3 | 92.42
0 | 0 | 4.7 | 103.05 | 571481 | 92.49 | 0 | 0 | 4.2 | 102.76 | 584904.4 | 91.59
0 | 0 | 3.4 | 105.42 | 584630.3 | 91.68 | 0 | 0 | 3 | 108.8 | 619288.9 | 90.85
0 | 0 | 2.2 | 107.95 | 598674.7 | 90.7 | 0 | 0 | 1.8 | 100.12 | 569879.3 | 90.04
0 | 0 | 1 | 104.78 | 581111.2 | 90.07 | 0 | 0 | 0.6 | 98.99 | 563413.5 | 89.42
0 | 0 | -0.5 | 101.6 | 563480.2 | 90.41 | 0 | 0 | -0.5 | 105.85 | 602447 | 89.15
Measurement time: | 07/09/2022 19:15:56
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 150 x 150 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 1.53 | 8486.4 | 1.15
0 | 0 | 149.3 | 1.22 | 6771.3 | 1.14
0 | 0 | 147.9 | 1.27 | 7042.1 | 1.15
0 | 0 | 146.7 | 1.26 | 6965.5 | 1.16
0 | 0 | 145.5 | 1.29 | 7176.7 | 1.18
0 | 0 | 144.3 | 1.31 | 7239.4 | 1.19
0 | 0 | 143.1 | 1.31 | 7264.3 | 1.2
0 | 0 | 141.9 | 1.3 | 7224 | 1.2
0 | 0 | 140.7 | 1.33 | 7351.9 | 1.21
0 | 0 | 139.5 | 1.33 | 7387.7 | 1.22
0 | 0 | 138.3 | 1.31 | 7287.5 | 1.21
0 | 0 | 137.1 | 1.36 | 7562.6 | 1.24
0 | 0 | 135.9 | 1.32 | 7313.1 | 1.24
0 | 0 | 134.7 | 1.34 | 7426.4 | 1.25
0 | 0 | 133.5 | 1.37 | 7615.9 | 1.27
0 | 0 | 132.3 | 1.34 | 7453.5 | 1.26
0 | 0 | 131.1 | 1.38 | 7632.7 | 1.28
0 | 0 | 129.9 | 1.41 | 7816.3 | 1.3
0 | 0 | 128.7 | 1.39 | 7710.9 | 1.28
0 | 0 | 127.5 | 1.41 | 7804.3 | 1.29
0 | 0 | 126.3 | 1.42 | 7864.9 | 1.31
0 | 0 | 125.1 | 1.44 | 7982 | 1.32
0 | 0 | 123.9 | 1.44 | 7969.4 | 1.33
0 | 0 | 122.7 | 1.45 | 8063.3 | 1.34
0 | 0 | 121.5 | 1.46 | 8086.1 | 1.35
0 | 0 | 120.3 | 1.47 | 8125.6 | 1.36
0 | 0 | 119.1 | 1.49 | 8264.8 | 1.38
0 | 0 | 117.9 | 1.48 | 8187 | 1.38
0 | 0 | 116.7 | 1.48 | 8228.3 | 1.39
0 | 0 | 115.5 | 1.5 | 8319.3 | 1.41
0 | 0 | 114.3 | 1.54 | 8513.1 | 1.42
0 | 0 | 113.1 | 1.53 | 8472.8 | 1.42
0 | 0 | 111.9 | 1.53 | 8496.2 | 1.45
0 | 0 | 110.7 | 1.57 | 8701.1 | 1.45
0 | 0 | 109.5 | 1.55 | 8621.3 | 1.46
0 | 0 | 108.3 | 1.62 | 8970.8 | 1.49
0 | 0 | 107.1 | 1.59 | 8793.5 | 1.48
0 | 0 | 105.9 | 1.65 | 9138 | 1.51
0 | 0 | 104.7 | 1.63 | 9031.5 | 1.51
0 | 0 | 103.5 | 1.65 | 9138 | 1.54
0 | 0 | 102.3 | 1.66 | 9197.6 | 1.54
0 | 0 | 101.1 | 1.7 | 9411.5 | 1.57
0 | 0 | 99.9 | 1.71 | 9510.9 | 1.58
0 | 0 | 98.7 | 1.72 | 9548.5 | 1.59
0 | 0 | 97.5 | 1.73 | 9586.8 | 1.59
0 | 0 | 96.3 | 1.71 | 9479 | 1.57
0 | 0 | 95.1 | 1.75 | 9694.8 | 1.62
0 | 0 | 93.9 | 1.67 | 9282.5 | 1.63
0 | 0 | 92.6 | 1.76 | 9757.7 | 1.62
0 | 0 | 91.4 | 1.74 | 9644.3 | 1.63
0 | 0 | 90.2 | 1.8 | 9966.2 | 1.67
0 | 0 | 89 | 1.8 | 9972.6 | 1.66
0 | 0 | 87.8 | 1.77 | 9800.1 | 1.69
0 | 0 | 86.6 | 1.78 | 9895.7 | 1.7
0 | 0 | 85.4 | 1.85 | 10251.4 | 1.7
0 | 0 | 84.2 | 1.82 | 10106 | 1.72
0 | 0 | 83 | 1.79 | 9921.8 | 1.74
0 | 0 | 81.8 | 1.9 | 10537.4 | 1.74
0 | 0 | 80.6 | 1.88 | 10433.6 | 1.76
0 | 0 | 79.4 | 1.83 | 10175.2 | 1.77
0 | 0 | 78.2 | 1.96 | 10858 | 1.78
0 | 0 | 77 | 1.89 | 10475.4 | 1.81
0 | 0 | 75.7 | 1.95 | 10799.8 | 1.82
0 | 0 | 74.6 | 2.01 | 11168.8 | 1.82
0 | 0 | 73.3 | 1.93 | 10682.6 | 1.86
0 | 0 | 72.1 | 1.99 | 11031.9 | 1.87
0 | 0 | 70.9 | 2.11 | 11705.6 | 1.9
0 | 0 | 69.7 | 2.03 | 11243.7 | 1.96
0 | 0 | 68.5 | 2.22 | 12303.3 | 2.07
0 | 0 | 67.3 | 2.51 | 13941.2 | 2.26
0 | 0 | 66.1 | 2.69 | 14906.1 | 2.61
0 | 0 | 64.9 | 3.47 | 19266.1 | 3.23
0 | 0 | 63.7 | 4.67 | 25905.9 | 4.19
0 | 0 | 62.5 | 5.83 | 32324.7 | 5.63
0 | 0 | 61.3 | 8.17 | 45319.4 | 7.53
0 | 0 | 60.1 | 11.38 | 63096.4 | 10.18
0 | 0 | 58.9 | 13.75 | 76270.4 | 13.25
0 | 0 | 57.7 | 18.39 | 101993.9 | 16.96
0 | 0 | 56.5 | 22.75 | 126169.5 | 21.31
0 | 0 | 55.3 | 27.33 | 151556.8 | 26.09
0 | 0 | 54.1 | 33.91 | 188045.5 | 31.31
0 | 0 | 52.9 | 40.97 | 227227.4 | 36.45
0 | 0 | 51.7 | 44.37 | 246078.8 | 42.39
0 | 0 | 50.5 | 51.54 | 285808.8 | 47.93
0 | 0 | 49.2 | 59.49 | 329917.3 | 53.28
0 | 0 | 48 | 61.39 | 340476.5 | 59.01
0 | 0 | 46.8 | 69.05 | 382924.4 | 63.9
0 | 0 | 45.6 | 78.28 | 434151.8 | 69.11
0 | 0 | 44.4 | 77.46 | 429580.3 | 73.53
0 | 0 | 43.2 | 85.19 | 472455 | 77.77
0 | 0 | 42 | 93.52 | 518640.4 | 81.89
0 | 0 | 40.8 | 91 | 504673.8 | 85.39
0 | 0 | 39.6 | 98.64 | 547063.2 | 88.58
0 | 0 | 38.4 | 105.57 | 585504.8 | 90.95
0 | 0 | 37.2 | 99.56 | 552150.3 | 93.52
0 | 0 | 36 | 106.98 | 593323.2 | 95.2
0 | 0 | 34.8 | 112.09 | 621649.3 | 96.64
0 | 0 | 33.6 | 105.94 | 587551.4 | 98.03
0 | 0 | 32.4 | 112.31 | 622849 | 98.84
0 | 0 | 31.1 | 116.74 | 647426 | 99.3
0 | 0 | 29.9 | 108.54 | 601945.6 | 99.7
0 | 0 | 28.7 | 114.44 | 634665.4 | 99.83
0 | 0 | 27.5 | 117.23 | 650143.1 | 100
0 | 0 | 26.3 | 110.77 | 614315.5 | 99.81
0 | 0 | 25.1 | 113.84 | 631357.2 | 99.64
0 | 0 | 23.9 | 112.6 | 624471.2 | 99.39
0 | 0 | 22.7 | 110.57 | 613230.8 | 98.68
0 | 0 | 21.5 | 113.22 | 627898.2 | 98.48
0 | 0 | 20.3 | 116.7 | 647214.6 | 98.01
0 | 0 | 19.1 | 109.58 | 607724.2 | 97.45
0 | 0 | 17.9 | 112.59 | 624388.1 | 97.05
0 | 0 | 16.7 | 115.73 | 641838 | 96.39
0 | 0 | 15.5 | 108.38 | 601050.2 | 96.07
0 | 0 | 14.3 | 111.88 | 620487 | 95.51
0 | 0 | 13.1 | 114.39 | 634386.5 | 94.96
0 | 0 | 11.9 | 107.34 | 595280.2 | 94.52
0 | 0 | 10.7 | 110.34 | 611920.8 | 94.25
0 | 0 | 9.5 | 112.24 | 622472.4 | 93.61
0 | 0 | 8.3 | 105.65 | 585933 | 93.3
0 | 0 | 7 | 107.69 | 597254.2 | 92.64
0 | 0 | 5.9 | 111.05 | 615898.1 | 91.8
0 | 0 | 4.7 | 102.53 | 568605 | 91.24
0 | 0 | 3.4 | 104.96 | 582124.9 | 90.32
0 | 0 | 2.2 | 106.69 | 591694.5 | 89.39
0 | 0 | 1 | 103.03 | 571415.8 | 88.91
0 | 0 | -0.5 | 102.25 | 567082 | 89.05
Measurement time: | 07/09/2022 19:20:05
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 200 x 200 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Beam
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
0 | 0 | 150 | 1.56 | 8659 | 1.33
0 | 0 | 149.3 | 1.36 | 7548.3 | 1.28
0 | 0 | 147.9 | 1.34 | 7443.5 | 1.29
0 | 0 | 146.7 | 1.37 | 7609.6 | 1.29
0 | 0 | 145.5 | 1.39 | 7716.1 | 1.31
0 | 0 | 144.3 | 1.4 | 7791.9 | 1.32
0 | 0 | 143.1 | 1.41 | 7815.7 | 1.32
0 | 0 | 141.9 | 1.42 | 7861.9 | 1.33
0 | 0 | 140.7 | 1.45 | 8020.1 | 1.35
0 | 0 | 139.5 | 1.43 | 7956.9 | 1.35
0 | 0 | 138.3 | 1.42 | 7896.4 | 1.36
0 | 0 | 137.1 | 1.46 | 8103.8 | 1.37
0 | 0 | 135.9 | 1.47 | 8128.6 | 1.38
0 | 0 | 134.7 | 1.49 | 8287.8 | 1.4
0 | 0 | 133.5 | 1.48 | 8208.2 | 1.4
0 | 0 | 132.3 | 1.49 | 8281 | 1.4
0 | 0 | 131.1 | 1.52 | 8426.9 | 1.42
0 | 0 | 129.9 | 1.51 | 8368.9 | 1.43
0 | 0 | 128.7 | 1.55 | 8578.6 | 1.44
0 | 0 | 127.5 | 1.53 | 8504.7 | 1.44
0 | 0 | 126.3 | 1.56 | 8638.6 | 1.46
0 | 0 | 125.1 | 1.56 | 8674.3 | 1.47
0 | 0 | 123.9 | 1.57 | 8711.4 | 1.47
0 | 0 | 122.7 | 1.58 | 8742.9 | 1.47
0 | 0 | 121.5 | 1.58 | 8772.4 | 1.49
0 | 0 | 120.3 | 1.59 | 8808 | 1.49
0 | 0 | 119.1 | 1.61 | 8932.1 | 1.5
0 | 0 | 117.9 | 1.61 | 8907.3 | 1.51
0 | 0 | 116.7 | 1.62 | 8975.1 | 1.52
0 | 0 | 115.5 | 1.62 | 8995.9 | 1.54
0 | 0 | 114.3 | 1.67 | 9254.3 | 1.56
0 | 0 | 113.1 | 1.65 | 9150.3 | 1.56
0 | 0 | 111.9 | 1.67 | 9286 | 1.58
0 | 0 | 110.7 | 1.69 | 9369.8 | 1.6
0 | 0 | 109.5 | 1.7 | 9442.3 | 1.61
0 | 0 | 108.3 | 1.7 | 9434.2 | 1.61
0 | 0 | 107.1 | 1.7 | 9417.5 | 1.61
0 | 0 | 105.8 | 1.71 | 9492 | 1.63
0 | 0 | 104.7 | 1.75 | 9693 | 1.65
0 | 0 | 103.5 | 1.75 | 9687.8 | 1.66
0 | 0 | 102.3 | 1.76 | 9768 | 1.67
0 | 0 | 101.1 | 1.76 | 9785.2 | 1.68
0 | 0 | 99.9 | 1.78 | 9881.1 | 1.69
0 | 0 | 98.7 | 1.8 | 10001.1 | 1.71
0 | 0 | 97.5 | 1.81 | 10051.9 | 1.73
0 | 0 | 96.3 | 1.84 | 10207.8 | 1.75
0 | 0 | 95.1 | 1.83 | 10158.8 | 1.74
0 | 0 | 93.9 | 1.86 | 10319.5 | 1.77
0 | 0 | 92.6 | 1.85 | 10235.2 | 1.75
0 | 0 | 91.4 | 1.87 | 10378.6 | 1.78
0 | 0 | 90.2 | 1.89 | 10503.2 | 1.8
0 | 0 | 89 | 1.89 | 10461.1 | 1.81
0 | 0 | 87.8 | 1.92 | 10636.7 | 1.81
0 | 0 | 86.6 | 1.91 | 10586.1 | 1.83
0 | 0 | 85.4 | 1.97 | 10919.1 | 1.85
0 | 0 | 84.2 | 1.95 | 10789.3 | 1.85
0 | 0 | 83 | 1.96 | 10855 | 1.88
0 | 0 | 81.8 | 1.99 | 11033.8 | 1.88
0 | 0 | 80.6 | 2.04 | 11321.6 | 1.92
0 | 0 | 79.4 | 1.99 | 11041.1 | 1.92
0 | 0 | 78.2 | 2.01 | 11125.3 | 1.92
0 | 0 | 77 | 2.12 | 11736.5 | 1.97
0 | 0 | 75.8 | 2.02 | 11223.2 | 1.96
0 | 0 | 74.6 | 2.02 | 11196.9 | 1.94
0 | 0 | 73.4 | 2.16 | 11956.9 | 2
0 | 0 | 72.2 | 2.1 | 11619.6 | 2.01
0 | 0 | 71 | 2.15 | 11903 | 2.03
0 | 0 | 69.7 | 2.24 | 12426.3 | 2.09
0 | 0 | 68.5 | 2.24 | 12401 | 2.17
0 | 0 | 67.3 | 2.57 | 14274.8 | 2.42
0 | 0 | 66.1 | 2.94 | 16312.9 | 2.78
0 | 0 | 64.9 | 3.39 | 18825.5 | 3.36
0 | 0 | 63.7 | 4.68 | 25969.1 | 4.44
0 | 0 | 62.5 | 6.04 | 33471.4 | 5.65
0 | 0 | 61.3 | 7.87 | 43642.8 | 7.77
0 | 0 | 60.1 | 10.67 | 59188 | 10.21
0 | 0 | 58.9 | 14.07 | 78040.4 | 13.16
0 | 0 | 57.7 | 17.48 | 96950.3 | 17.25
0 | 0 | 56.5 | 22.12 | 122648.4 | 21.17
0 | 0 | 55.3 | 27.67 | 153467.4 | 25.89
0 | 0 | 54.1 | 32.42 | 179777 | 31.18
0 | 0 | 52.9 | 36.52 | 202561.4 | 36.06
0 | 0 | 51.7 | 45.26 | 251005.1 | 42.02
0 | 0 | 50.5 | 49.81 | 276229.8 | 47.59
0 | 0 | 49.3 | 53.69 | 297733 | 53.1
0 | 0 | 48 | 64.32 | 356698.2 | 58.83
0 | 0 | 46.9 | 68.14 | 377921.3 | 64.19
0 | 0 | 45.6 | 70.6 | 391565.4 | 68.84
0 | 0 | 44.4 | 82.15 | 455591.1 | 73.91
0 | 0 | 43.2 | 79.43 | 440493.3 | 78.04
0 | 0 | 42 | 88.02 | 488150.9 | 81.68
0 | 0 | 40.8 | 95.58 | 530076.6 | 85.33
0 | 0 | 39.6 | 91.23 | 505928.1 | 88.33
0 | 0 | 38.4 | 97.89 | 542906.6 | 90.91
0 | 0 | 37.2 | 105.15 | 583156.8 | 93.1
0 | 0 | 36 | 98.48 | 546138.2 | 95.03
0 | 0 | 34.8 | 104.65 | 580362.1 | 96.35
0 | 0 | 33.6 | 111.27 | 617074.8 | 97.64
0 | 0 | 32.4 | 106.68 | 591661.8 | 98.54
0 | 0 | 31.1 | 104 | 576769.4 | 99.29
0 | 0 | 30 | 114.03 | 632415.2 | 99.68
0 | 0 | 28.7 | 109.8 | 608915.1 | 99.85
0 | 0 | 27.5 | 106.57 | 591028.3 | 100
0 | 0 | 26.3 | 113.49 | 629421 | 99.85
0 | 0 | 25.1 | 108.24 | 600300.5 | 99.71
0 | 0 | 23.9 | 111.65 | 619218.5 | 99.38
0 | 0 | 22.7 | 115.07 | 638145.3 | 99.13
0 | 0 | 21.5 | 106.83 | 592483.7 | 98.81
0 | 0 | 20.3 | 110.17 | 610992.4 | 98.45
0 | 0 | 19.1 | 114.58 | 635463.7 | 98.08
0 | 0 | 17.9 | 106.48 | 590499.7 | 97.71
0 | 0 | 16.7 | 109.43 | 606892.4 | 97.18
0 | 0 | 15.5 | 114.4 | 634440.1 | 96.96
0 | 0 | 14.3 | 106.32 | 589620.3 | 96.48
0 | 0 | 13.1 | 108.44 | 601414.7 | 96.08
0 | 0 | 11.9 | 113.98 | 632143.4 | 95.75
0 | 0 | 10.7 | 104.48 | 579411.1 | 95.37
0 | 0 | 9.5 | 108.13 | 599675.3 | 94.7
0 | 0 | 8.3 | 111.98 | 621035.6 | 94.22
0 | 0 | 7 | 103.92 | 576327.4 | 93.9
0 | 0 | 5.9 | 107.63 | 596889.9 | 93.33
0 | 0 | 4.7 | 110.54 | 613026.7 | 92.41
0 | 0 | 3.4 | 101.49 | 562873.5 | 91.96
0 | 0 | 2.2 | 105.82 | 586867.2 | 91
0 | 0 | 1 | 107.99 | 598912.6 | 90.56
0 | 0 | -0.5 | 103.44 | 573652.5 | 90.65