## Profils énergie
Measurement time: | 02/09/2022 19:04:52
Radiation device: | Clinac3 Accelerator
Energy: | 6 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-150.7 | 0 | 13 | 0.47 | 918.3 | 0.34
-150 | 0 | 13 | 0.33 | 642.8 | 0.34
-148.6 | 0 | 13 | 0.36 | 703 | 0.35
-147.4 | 0 | 13 | 0.33 | 636.3 | 0.33
-146.2 | 0 | 13 | 0.33 | 634.3 | 0.33
-145 | 0 | 13 | 0.32 | 625.5 | 0.31
-143.8 | 0 | 13 | 0.39 | 750.9 | 0.39
-142.6 | 0 | 13 | 0.3 | 591.9 | 0.31
-141.4 | 0 | 13 | 0.37 | 718.2 | 0.36
-140.2 | 0 | 13 | 0.4 | 774.4 | 0.4
-139 | 0 | 13 | 0.4 | 783.1 | 0.39
-137.8 | 0 | 13 | 0.35 | 682.8 | 0.36
-136.6 | 0 | 13 | 0.39 | 756.7 | 0.39
-135.4 | 0 | 13 | 0.38 | 746.9 | 0.38
-134.2 | 0 | 13 | 0.43 | 847.1 | 0.42
-133 | 0 | 13 | 0.39 | 761.3 | 0.39
-131.8 | 0 | 13 | 0.37 | 724.5 | 0.37
-130.6 | 0 | 13 | 0.44 | 858.5 | 0.43
-129.4 | 0 | 13 | 0.37 | 723.4 | 0.38
-128.2 | 0 | 13 | 0.42 | 822.3 | 0.42
-127 | 0 | 13 | 0.43 | 838.3 | 0.43
-125.8 | 0 | 13 | 0.45 | 877.7 | 0.43
-124.6 | 0 | 13 | 0.43 | 839.5 | 0.42
-123.4 | 0 | 13 | 0.42 | 818.3 | 0.41
-122.2 | 0 | 13 | 0.41 | 790.7 | 0.41
-121 | 0 | 13 | 0.42 | 821.6 | 0.41
-119.8 | 0 | 13 | 0.41 | 791.5 | 0.4
-118.6 | 0 | 13 | 0.43 | 833.9 | 0.43
-117.4 | 0 | 13 | 0.43 | 837 | 0.42
-116.2 | 0 | 13 | 0.36 | 705.7 | 0.36
-115 | 0 | 13 | 0.35 | 683.5 | 0.36
-113.8 | 0 | 13 | 0.41 | 803.9 | 0.4
-112.6 | 0 | 13 | 0.36 | 704.1 | 0.37
-111.4 | 0 | 13 | 0.37 | 729.8 | 0.38
-110.2 | 0 | 13 | 0.37 | 726.6 | 0.36
-109 | 0 | 13 | 0.38 | 741.2 | 0.38
-107.8 | 0 | 13 | 0.32 | 629.3 | 0.33
-106.6 | 0 | 13 | 0.33 | 646.7 | 0.32
-105.4 | 0 | 13 | 0.32 | 633.4 | 0.33
-104.2 | 0 | 13 | 0.36 | 699.9 | 0.36
-102.9 | 0 | 13 | 0.34 | 672.1 | 0.34
-101.8 | 0 | 13 | 0.37 | 712.6 | 0.36
-100.5 | 0 | 13 | 0.35 | 672.9 | 0.35
-99.4 | 0 | 13 | 0.37 | 729 | 0.36
-98.1 | 0 | 13 | 0.38 | 748.5 | 0.39
-96.9 | 0 | 13 | 0.41 | 796.6 | 0.41
-95.7 | 0 | 13 | 0.43 | 845.7 | 0.42
-94.5 | 0 | 13 | 0.39 | 765.9 | 0.4
-93.3 | 0 | 13 | 0.35 | 688.5 | 0.36
-92.1 | 0 | 13 | 0.41 | 795.6 | 0.39
-90.9 | 0 | 13 | 0.4 | 781.3 | 0.39
-89.7 | 0 | 13 | 0.42 | 815.4 | 0.42
-88.5 | 0 | 13 | 0.45 | 874.1 | 0.44
-87.3 | 0 | 13 | 0.45 | 875.3 | 0.45
-86.1 | 0 | 13 | 0.44 | 865.4 | 0.45
-84.9 | 0 | 13 | 0.47 | 923 | 0.46
-83.7 | 0 | 13 | 0.52 | 1016.8 | 0.52
-82.5 | 0 | 13 | 0.5 | 971.9 | 0.5
-81.3 | 0 | 13 | 0.55 | 1073.5 | 0.54
-80.1 | 0 | 13 | 0.59 | 1144.3 | 0.59
-78.9 | 0 | 13 | 0.58 | 1132.8 | 0.6
-77.7 | 0 | 13 | 0.64 | 1252.3 | 0.63
-76.5 | 0 | 13 | 0.75 | 1460.7 | 0.75
-75.3 | 0 | 13 | 0.75 | 1453.9 | 0.75
-74.1 | 0 | 13 | 0.95 | 1847.3 | 0.92
-72.9 | 0 | 13 | 1.05 | 2048.7 | 1.06
-71.7 | 0 | 13 | 1.35 | 2626.8 | 1.35
-70.5 | 0 | 13 | 1.69 | 3288.5 | 1.65
-69.3 | 0 | 13 | 2.19 | 4267.6 | 2.18
-68.1 | 0 | 13 | 2.74 | 5339.1 | 2.76
-66.9 | 0 | 13 | 3.6 | 7017.9 | 3.61
-65.7 | 0 | 13 | 4.77 | 9303.6 | 4.65
-64.4 | 0 | 13 | 5.95 | 11601.4 | 6.05
-63.3 | 0 | 13 | 7.81 | 15231.7 | 7.77
-62 | 0 | 13 | 9.82 | 19151.3 | 9.91
-60.8 | 0 | 13 | 12.84 | 25037.5 | 12.64
-59.7 | 0 | 13 | 16.08 | 31353.6 | 15.92
-58.4 | 0 | 13 | 19.6 | 38201.8 | 19.86
-57.2 | 0 | 13 | 25.23 | 49190.3 | 24.65
-56 | 0 | 13 | 30.28 | 59023.2 | 30.24
-54.8 | 0 | 13 | 37.38 | 72867.5 | 36.61
-53.6 | 0 | 13 | 42.7 | 83251 | 43.38
-52.4 | 0 | 13 | 50.18 | 97827.5 | 50.29
-51.2 | 0 | 13 | 58.5 | 114042.3 | 57.46
-50 | 0 | 13 | 63.04 | 122896.1 | 64.05
-48.8 | 0 | 13 | 70.23 | 136913.1 | 70.12
-47.6 | 0 | 13 | 77.42 | 150931 | 75.52
-46.4 | 0 | 13 | 79.11 | 154230.1 | 79.79
-45.2 | 0 | 13 | 83.76 | 163293.3 | 83.79
-44 | 0 | 13 | 85.85 | 167357.8 | 86.41
-42.8 | 0 | 13 | 91.52 | 178414.3 | 89.15
-41.6 | 0 | 13 | 91.09 | 177580.7 | 91.2
-40.4 | 0 | 13 | 94.25 | 183737.8 | 92.35
-39.2 | 0 | 13 | 93.01 | 181327.6 | 93.83
-38 | 0 | 13 | 95.06 | 185320.9 | 95.01
-36.8 | 0 | 13 | 97.39 | 189862.7 | 95.93
-35.6 | 0 | 13 | 96.2 | 187551.4 | 96.35
-34.4 | 0 | 13 | 97.1 | 189302 | 97.1
-33.2 | 0 | 13 | 96.22 | 187576.7 | 97.44
-32 | 0 | 13 | 100.64 | 196189.1 | 97.67
-30.8 | 0 | 13 | 97.97 | 190984.6 | 98.16
-29.6 | 0 | 13 | 100.98 | 196868 | 98.26
-28.4 | 0 | 13 | 96.38 | 187896.3 | 98.66
-27.2 | 0 | 13 | 99.39 | 193765.1 | 98.82
-26 | 0 | 13 | 101.21 | 197315 | 98.95
-24.8 | 0 | 13 | 97.59 | 190249.3 | 99.04
-23.6 | 0 | 13 | 98.79 | 192600.8 | 98.99
-22.4 | 0 | 13 | 102.05 | 198939.2 | 99.37
-21.2 | 0 | 13 | 98.09 | 191217.6 | 99.21
-20 | 0 | 13 | 99.49 | 193946.4 | 99.32
-18.8 | 0 | 13 | 101.93 | 198708 | 99.48
-17.6 | 0 | 13 | 98.1 | 191241.2 | 99.51
-16.3 | 0 | 13 | 98.99 | 192972.1 | 99.61
-15.1 | 0 | 13 | 102.52 | 199868.8 | 99.3
-14 | 0 | 13 | 98.17 | 191390.3 | 99.5
-12.7 | 0 | 13 | 98.75 | 192513.6 | 99.64
-11.5 | 0 | 13 | 102.99 | 200783.7 | 99.54
-10.3 | 0 | 13 | 97.49 | 190058.4 | 99.44
-9.1 | 0 | 13 | 99.89 | 194741.8 | 99.7
-7.9 | 0 | 13 | 102.02 | 198894 | 99.63
-6.7 | 0 | 13 | 99.35 | 193689.9 | 99.68
-5.5 | 0 | 13 | 99.56 | 194098.9 | 99.9
-4.3 | 0 | 13 | 102.4 | 199624.9 | 99.84
-3.1 | 0 | 13 | 97.15 | 189402.7 | 99.7
-1.9 | 0 | 13 | 101.25 | 197393 | 99.78
-0.7 | 0 | 13 | 101.75 | 198361 | 99.93
0.5 | 0 | 13 | 98.08 | 191204 | 99.74
1.7 | 0 | 13 | 101.23 | 197351.1 | 99.9
2.9 | 0 | 13 | 100.32 | 195582.3 | 100
4.1 | 0 | 13 | 98.83 | 192674.7 | 99.89
5.3 | 0 | 13 | 101.46 | 197796 | 99.94
6.5 | 0 | 13 | 99.94 | 194825.4 | 100
7.7 | 0 | 13 | 98.71 | 192441.7 | 99.94
8.9 | 0 | 13 | 101.5 | 197876.8 | 99.95
10.1 | 0 | 13 | 101.15 | 197188.8 | 99.82
11.3 | 0 | 13 | 98.98 | 192968.5 | 99.92
12.5 | 0 | 13 | 101.17 | 197231 | 99.99
13.7 | 0 | 13 | 100.05 | 195044.6 | 99.84
14.9 | 0 | 13 | 98.24 | 191529 | 99.63
16.1 | 0 | 13 | 106.67 | 207945.9 | 99.84
17.3 | 0 | 13 | 99.69 | 194351.8 | 99.52
18.5 | 0 | 13 | 98 | 191050.2 | 99.33
19.7 | 0 | 13 | 102.4 | 199629.3 | 99.63
20.9 | 0 | 13 | 98.42 | 191865.1 | 99.47
22.1 | 0 | 13 | 99.49 | 193957.4 | 99.41
23.4 | 0 | 13 | 102.25 | 199335.8 | 99.61
24.5 | 0 | 13 | 99.48 | 193940 | 99.35
25.8 | 0 | 13 | 97.58 | 190228 | 99.13
27 | 0 | 13 | 102.12 | 199086.2 | 99.2
28.2 | 0 | 13 | 97.78 | 190618.5 | 98.84
29.4 | 0 | 13 | 97.47 | 190019.6 | 98.66
30.6 | 0 | 13 | 101.56 | 197982.7 | 98.41
31.8 | 0 | 13 | 97.4 | 189888.6 | 98.05
33 | 0 | 13 | 96.32 | 187779.9 | 97.59
34.2 | 0 | 13 | 99.57 | 194107 | 96.97
35.4 | 0 | 13 | 94.73 | 184669.2 | 96.17
36.6 | 0 | 13 | 93.87 | 182995.2 | 95.17
37.8 | 0 | 13 | 96.35 | 187840.8 | 93.69
39 | 0 | 13 | 91.49 | 178356.7 | 92.14
40.2 | 0 | 13 | 88.82 | 173153.3 | 89.91
41.4 | 0 | 13 | 90.19 | 175824.1 | 87.52
42.6 | 0 | 13 | 83.9 | 163570.9 | 83.97
43.8 | 0 | 13 | 78.74 | 153509.3 | 80.31
45 | 0 | 13 | 78.3 | 152639.8 | 75.67
46.2 | 0 | 13 | 69.67 | 135825.9 | 69.93
47.4 | 0 | 13 | 62.84 | 122505.2 | 63.84
48.6 | 0 | 13 | 61.33 | 119570.6 | 57.21
49.8 | 0 | 13 | 49.33 | 96162.9 | 49.82
51 | 0 | 13 | 42.58 | 83013.8 | 42.93
52.2 | 0 | 13 | 37.05 | 72232 | 35.95
53.4 | 0 | 13 | 29.49 | 57489.3 | 29.65
54.6 | 0 | 13 | 24.02 | 46833 | 24.14
55.8 | 0 | 13 | 19.86 | 38712.8 | 19.34
57 | 0 | 13 | 15.06 | 29360.1 | 15.41
58.3 | 0 | 13 | 12.16 | 23705.6 | 12.1
59.4 | 0 | 13 | 9.62 | 18761.4 | 9.56
60.6 | 0 | 13 | 7.44 | 14509.3 | 7.52
61.9 | 0 | 13 | 6 | 11687.7 | 5.88
63.1 | 0 | 13 | 4.63 | 9018.2 | 4.6
64.3 | 0 | 13 | 3.48 | 6778.1 | 3.53
65.5 | 0 | 13 | 2.85 | 5557.3 | 2.79
66.7 | 0 | 13 | 2.16 | 4209.2 | 2.17
67.9 | 0 | 13 | 1.69 | 3302.7 | 1.72
69.1 | 0 | 13 | 1.38 | 2685.9 | 1.35
70.3 | 0 | 13 | 1.07 | 2091.4 | 1.07
71.5 | 0 | 13 | 0.89 | 1739.3 | 0.9
72.7 | 0 | 13 | 0.74 | 1434.8 | 0.71
73.9 | 0 | 13 | 0.6 | 1161.3 | 0.6
75.1 | 0 | 13 | 0.55 | 1064.1 | 0.55
76.3 | 0 | 13 | 0.52 | 1020.7 | 0.51
77.5 | 0 | 13 | 0.44 | 850.1 | 0.44
78.7 | 0 | 13 | 0.41 | 807.5 | 0.4
79.9 | 0 | 13 | 0.36 | 701.3 | 0.37
81.1 | 0 | 13 | 0.35 | 683.8 | 0.35
82.3 | 0 | 13 | 0.35 | 675.4 | 0.34
83.5 | 0 | 13 | 0.31 | 605 | 0.3
84.7 | 0 | 13 | 0.27 | 523.9 | 0.27
85.9 | 0 | 13 | 0.25 | 479.4 | 0.25
87.1 | 0 | 13 | 0.29 | 564.5 | 0.28
88.3 | 0 | 13 | 0.25 | 483.8 | 0.25
89.5 | 0 | 13 | 0.26 | 500 | 0.26
90.7 | 0 | 13 | 0.25 | 479 | 0.24
91.9 | 0 | 13 | 0.25 | 485.4 | 0.25
93.1 | 0 | 13 | 0.2 | 387.8 | 0.19
94.3 | 0 | 13 | 0.25 | 496.9 | 0.26
95.6 | 0 | 13 | 0.23 | 446.1 | 0.23
96.8 | 0 | 13 | 0.22 | 430.6 | 0.22
98 | 0 | 13 | 0.27 | 533.7 | 0.27
99.2 | 0 | 13 | 0.23 | 439.4 | 0.23
100.4 | 0 | 13 | 0.2 | 389.5 | 0.2
101.6 | 0 | 13 | 0.21 | 410.7 | 0.2
102.8 | 0 | 13 | 0.18 | 354.8 | 0.18
104 | 0 | 13 | 0.22 | 427.6 | 0.22
105.2 | 0 | 13 | 0.23 | 443.7 | 0.22
106.4 | 0 | 13 | 0.15 | 294.6 | 0.15
107.6 | 0 | 13 | 0.21 | 417.3 | 0.22
108.8 | 0 | 13 | 0.23 | 451.1 | 0.22
110 | 0 | 13 | 0.2 | 387.6 | 0.2
111.2 | 0 | 13 | 0.17 | 336.8 | 0.17
112.4 | 0 | 13 | 0.19 | 364.4 | 0.18
113.6 | 0 | 13 | 0.19 | 371.3 | 0.19
114.8 | 0 | 13 | 0.18 | 346 | 0.18
116 | 0 | 13 | 0.16 | 308.8 | 0.15
117.2 | 0 | 13 | 0.17 | 338.6 | 0.18
118.4 | 0 | 13 | 0.15 | 298.7 | 0.15
119.6 | 0 | 13 | 0.14 | 278.6 | 0.14
120.8 | 0 | 13 | 0.16 | 306.6 | 0.16
122 | 0 | 13 | 0.17 | 337.3 | 0.17
123.2 | 0 | 13 | 0.16 | 310 | 0.16
124.4 | 0 | 13 | 0.14 | 278.1 | 0.14
125.6 | 0 | 13 | 0.08 | 159.8 | 0.08
126.8 | 0 | 13 | 0.1 | 203.1 | 0.1
128 | 0 | 13 | 0.12 | 235 | 0.12
129.2 | 0 | 13 | 0.07 | 133.5 | 0.07
130.4 | 0 | 13 | 0.11 | 222.2 | 0.11
131.6 | 0 | 13 | 0.12 | 233.6 | 0.12
132.8 | 0 | 13 | 0.08 | 157.6 | 0.08
134 | 0 | 13 | 0.11 | 215.7 | 0.11
135.2 | 0 | 13 | 0.1 | 199.4 | 0.1
136.4 | 0 | 13 | 0.09 | 168 | 0.08
137.6 | 0 | 13 | 0.12 | 231.3 | 0.12
138.8 | 0 | 13 | 0.11 | 223.7 | 0.12
140 | 0 | 13 | 0.08 | 153.2 | 0.08
141.2 | 0 | 13 | 0.05 | 100.4 | 0.05
142.4 | 0 | 13 | 0.1 | 194.9 | 0.1
143.6 | 0 | 13 | 0.11 | 216.6 | 0.11
144.8 | 0 | 13 | 0.13 | 245.7 | 0.13
146 | 0 | 13 | 0.12 | 234.8 | 0.12
147.2 | 0 | 13 | 0.14 | 270.6 | 0.14
148.4 | 0 | 13 | 0.14 | 267.5 | 0.14
149.6 | 0 | 13 | 0.08 | 153.7 | 0.08
150.7 | 0 | 13 | 0.13 | 250.1 | 0.12
Measurement time: | 02/09/2022 19:06:58
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
151.4 | 0 | 28 | 0.17 | 326.3 | 0.15
150.7 | 0 | 28 | 0.15 | 293.8 | 0.15
149.3 | 0 | 28 | 0.18 | 359.3 | 0.17
148.1 | 0 | 28 | 0.14 | 264.6 | 0.14
146.9 | 0 | 28 | 0.19 | 367 | 0.18
145.7 | 0 | 28 | 0.17 | 330.8 | 0.16
144.5 | 0 | 28 | 0.2 | 394.3 | 0.19
143.3 | 0 | 28 | 0.18 | 350.1 | 0.17
142.1 | 0 | 28 | 0.2 | 387.6 | 0.19
140.9 | 0 | 28 | 0.2 | 397.2 | 0.19
139.7 | 0 | 28 | 0.2 | 390.5 | 0.19
138.5 | 0 | 28 | 0.18 | 342.8 | 0.17
137.3 | 0 | 28 | 0.17 | 321.9 | 0.16
136.1 | 0 | 28 | 0.2 | 395.7 | 0.19
134.9 | 0 | 28 | 0.17 | 339.8 | 0.17
133.7 | 0 | 28 | 0.21 | 419.1 | 0.21
132.5 | 0 | 28 | 0.2 | 393 | 0.19
131.3 | 0 | 28 | 0.2 | 385 | 0.19
130.1 | 0 | 28 | 0.19 | 367.3 | 0.18
128.9 | 0 | 28 | 0.22 | 430 | 0.21
127.7 | 0 | 28 | 0.22 | 419.4 | 0.21
126.5 | 0 | 28 | 0.23 | 443.1 | 0.22
125.3 | 0 | 28 | 0.2 | 386.9 | 0.19
124 | 0 | 28 | 0.2 | 389.7 | 0.19
122.9 | 0 | 28 | 0.23 | 458.1 | 0.22
121.6 | 0 | 28 | 0.25 | 490.3 | 0.25
120.4 | 0 | 28 | 0.22 | 423.3 | 0.21
119.3 | 0 | 28 | 0.24 | 469.9 | 0.23
118.1 | 0 | 28 | 0.21 | 401 | 0.2
116.9 | 0 | 28 | 0.23 | 440.6 | 0.22
115.6 | 0 | 28 | 0.26 | 512 | 0.25
114.5 | 0 | 28 | 0.24 | 468.7 | 0.23
113.3 | 0 | 28 | 0.27 | 533 | 0.26
112 | 0 | 28 | 0.26 | 509.3 | 0.25
110.9 | 0 | 28 | 0.28 | 546.4 | 0.27
109.6 | 0 | 28 | 0.29 | 563.5 | 0.28
108.5 | 0 | 28 | 0.28 | 537.3 | 0.26
107.3 | 0 | 28 | 0.29 | 565.6 | 0.28
106 | 0 | 28 | 0.31 | 601.5 | 0.3
104.9 | 0 | 28 | 0.27 | 524.1 | 0.27
103.6 | 0 | 28 | 0.36 | 693 | 0.34
102.4 | 0 | 28 | 0.33 | 644.4 | 0.32
101.2 | 0 | 28 | 0.37 | 725.8 | 0.36
100 | 0 | 28 | 0.38 | 737.9 | 0.36
98.8 | 0 | 28 | 0.37 | 722.9 | 0.36
97.6 | 0 | 28 | 0.44 | 854.2 | 0.42
96.4 | 0 | 28 | 0.44 | 854.7 | 0.42
95.2 | 0 | 28 | 0.46 | 903.8 | 0.44
94 | 0 | 28 | 0.52 | 1013.8 | 0.5
92.8 | 0 | 28 | 0.51 | 990.5 | 0.48
91.6 | 0 | 28 | 0.59 | 1140.7 | 0.56
90.4 | 0 | 28 | 0.61 | 1197.9 | 0.59
89.2 | 0 | 28 | 0.68 | 1326.2 | 0.66
88 | 0 | 28 | 0.76 | 1474.3 | 0.75
86.8 | 0 | 28 | 0.8 | 1555.6 | 0.79
85.6 | 0 | 28 | 0.93 | 1809.4 | 0.88
84.4 | 0 | 28 | 1.03 | 2015.1 | 1
83.2 | 0 | 28 | 1.12 | 2183.7 | 1.07
82 | 0 | 28 | 1.32 | 2576.6 | 1.25
80.8 | 0 | 28 | 1.4 | 2722.1 | 1.35
79.6 | 0 | 28 | 1.53 | 2983.5 | 1.46
78.4 | 0 | 28 | 1.7 | 3313.5 | 1.63
77.2 | 0 | 28 | 1.86 | 3632.4 | 1.83
75.9 | 0 | 28 | 2.15 | 4192.8 | 2.06
74.8 | 0 | 28 | 2.42 | 4712.9 | 2.29
73.5 | 0 | 28 | 2.7 | 5260.6 | 2.6
72.4 | 0 | 28 | 3.04 | 5928.8 | 2.93
71.2 | 0 | 28 | 3.48 | 6789.2 | 3.3
70 | 0 | 28 | 3.89 | 7578 | 3.79
68.8 | 0 | 28 | 4.42 | 8623.1 | 4.31
67.5 | 0 | 28 | 5.32 | 10363.8 | 4.98
66.3 | 0 | 28 | 5.96 | 11616.8 | 5.8
65.1 | 0 | 28 | 6.95 | 13544.7 | 6.75
63.9 | 0 | 28 | 8.03 | 15645.3 | 7.91
62.7 | 0 | 28 | 9.42 | 18355.2 | 9.28
61.5 | 0 | 28 | 11.51 | 22437.8 | 10.94
60.3 | 0 | 28 | 14.13 | 27552.9 | 13.15
59.1 | 0 | 28 | 15.87 | 30932.2 | 15.7
57.9 | 0 | 28 | 20.12 | 39233 | 18.9
56.7 | 0 | 28 | 24.47 | 47707.7 | 22.64
55.5 | 0 | 28 | 27.69 | 53990.9 | 27.4
54.3 | 0 | 28 | 32.34 | 63048.4 | 32.27
53.1 | 0 | 28 | 40.77 | 79489.9 | 37.87
51.9 | 0 | 28 | 44.45 | 86662.5 | 44.18
50.7 | 0 | 28 | 52.43 | 102216.8 | 50.28
49.5 | 0 | 28 | 61.87 | 120614.7 | 56.89
48.3 | 0 | 28 | 63.22 | 123240.3 | 62.8
47.1 | 0 | 28 | 71.26 | 138913.9 | 68.3
45.9 | 0 | 28 | 79.58 | 155148.3 | 73.68
44.7 | 0 | 28 | 78.05 | 152159.4 | 77.94
43.5 | 0 | 28 | 84.58 | 164890.4 | 81.94
42.3 | 0 | 28 | 91.54 | 178467.2 | 85.12
41.1 | 0 | 28 | 88.58 | 172696.1 | 87.48
39.9 | 0 | 28 | 92.99 | 181283.3 | 89.68
38.7 | 0 | 28 | 100.05 | 195038.6 | 91.37
37.5 | 0 | 28 | 93.37 | 182016.5 | 92.72
36.3 | 0 | 28 | 98.45 | 191927.4 | 93.83
35.1 | 0 | 28 | 103.09 | 200969.6 | 94.88
33.8 | 0 | 28 | 95.42 | 186030.1 | 95.64
32.7 | 0 | 28 | 100.91 | 196717.8 | 96.21
31.4 | 0 | 28 | 102.51 | 199843.8 | 96.77
30.3 | 0 | 28 | 99.35 | 193685.1 | 97.26
29 | 0 | 28 | 104.98 | 204663.3 | 97.56
27.8 | 0 | 28 | 100.78 | 196463.9 | 98.05
26.7 | 0 | 28 | 98.06 | 191171.2 | 98.09
25.4 | 0 | 28 | 102.17 | 199186.7 | 98.43
24.2 | 0 | 28 | 106.52 | 207660.5 | 98.64
23 | 0 | 28 | 97.89 | 190842.1 | 98.72
21.8 | 0 | 28 | 103.42 | 201609.3 | 98.85
20.6 | 0 | 28 | 107.64 | 209839.2 | 99.28
19.4 | 0 | 28 | 99.93 | 194812.2 | 99.09
18.2 | 0 | 28 | 104.35 | 203425.2 | 99.19
17 | 0 | 28 | 108.54 | 211599.8 | 99.51
15.8 | 0 | 28 | 99.87 | 194696.7 | 99.41
14.6 | 0 | 28 | 105.41 | 205487.6 | 99.66
13.4 | 0 | 28 | 103.72 | 202206.2 | 99.68
12.2 | 0 | 28 | 99.55 | 194073.1 | 99.67
11 | 0 | 28 | 104.39 | 203503.7 | 99.7
9.8 | 0 | 28 | 108.81 | 212123.8 | 99.92
8.6 | 0 | 28 | 100.77 | 196458.8 | 99.88
7.4 | 0 | 28 | 103.95 | 202649.7 | 99.83
6.2 | 0 | 28 | 108.23 | 211000.7 | 100
5 | 0 | 28 | 101.79 | 198433.6 | 99.67
3.8 | 0 | 28 | 103.56 | 201893.1 | 99.88
2.6 | 0 | 28 | 106.91 | 208416.4 | 99.91
1.4 | 0 | 28 | 101.23 | 197346.3 | 99.68
0.2 | 0 | 28 | 104.66 | 204040.5 | 99.87
-1 | 0 | 28 | 104.8 | 204311 | 99.83
-2.2 | 0 | 28 | 105.13 | 204950.5 | 99.64
-3.4 | 0 | 28 | 100.55 | 196019.8 | 99.76
-4.6 | 0 | 28 | 106.03 | 206705.6 | 99.66
-5.8 | 0 | 28 | 100.57 | 196053.4 | 99.56
-7 | 0 | 28 | 105.21 | 205102.5 | 99.58
-8.2 | 0 | 28 | 107.59 | 209747.2 | 99.83
-9.4 | 0 | 28 | 101.98 | 198806.1 | 99.57
-10.6 | 0 | 28 | 103.54 | 201858.1 | 99.75
-11.8 | 0 | 28 | 106.27 | 207174.9 | 99.74
-13 | 0 | 28 | 101.01 | 196915.4 | 99.76
-14.2 | 0 | 28 | 102.74 | 200287 | 99.63
-15.4 | 0 | 28 | 106.39 | 207398.9 | 99.51
-16.6 | 0 | 28 | 102.85 | 200513.4 | 99.48
-17.8 | 0 | 28 | 102.24 | 199321.3 | 99.47
-19 | 0 | 28 | 104.63 | 203972.3 | 99.3
-20.2 | 0 | 28 | 101.99 | 198830.3 | 99.19
-21.5 | 0 | 28 | 103.44 | 201647.7 | 99.03
-22.7 | 0 | 28 | 104.63 | 203976.5 | 99.11
-23.9 | 0 | 28 | 103.45 | 201676 | 98.88
-25.1 | 0 | 28 | 103.7 | 202163.5 | 98.72
-26.3 | 0 | 28 | 101.16 | 197204.4 | 98.48
-27.5 | 0 | 28 | 101.05 | 197003.7 | 98.18
-28.7 | 0 | 28 | 100.24 | 195416.4 | 98.08
-29.9 | 0 | 28 | 103.37 | 201522.3 | 97.94
-31.1 | 0 | 28 | 101.21 | 197317.7 | 97.43
-32.3 | 0 | 28 | 101.01 | 196928.6 | 97.18
-33.5 | 0 | 28 | 102.63 | 200083.2 | 96.71
-34.7 | 0 | 28 | 99.33 | 193644.9 | 96.23
-35.9 | 0 | 28 | 100.06 | 195068.6 | 95.7
-37.1 | 0 | 28 | 99.27 | 193518.1 | 95.18
-38.3 | 0 | 28 | 98.34 | 191717.7 | 94.39
-39.5 | 0 | 28 | 95.6 | 186371.4 | 93.24
-40.7 | 0 | 28 | 96.7 | 188515.7 | 92.22
-41.9 | 0 | 28 | 93.32 | 181937.1 | 90.53
-43.1 | 0 | 28 | 92.34 | 180013.4 | 88.73
-44.3 | 0 | 28 | 90.33 | 176099.9 | 86.42
-45.5 | 0 | 28 | 86.12 | 167891.3 | 83.59
-46.7 | 0 | 28 | 84.66 | 165052.3 | 80.15
-47.9 | 0 | 28 | 79.63 | 155243.7 | 76.2
-49.1 | 0 | 28 | 73.97 | 144207 | 71.22
-50.3 | 0 | 28 | 68.57 | 133680.8 | 65.77
-51.5 | 0 | 28 | 61.32 | 119552.5 | 59.97
-52.7 | 0 | 28 | 55.22 | 107653.4 | 53.53
-53.9 | 0 | 28 | 48.22 | 93997.8 | 47.31
-55.1 | 0 | 28 | 42.1 | 82075.2 | 41.23
-56.3 | 0 | 28 | 36.54 | 71230.9 | 35.04
-57.5 | 0 | 28 | 31.07 | 60573.7 | 30.04
-58.7 | 0 | 28 | 26.56 | 51780.1 | 25.22
-59.9 | 0 | 28 | 21.84 | 42578.3 | 21.19
-61.1 | 0 | 28 | 18.5 | 36064.2 | 17.7
-62.3 | 0 | 28 | 15.72 | 30655.5 | 14.84
-63.5 | 0 | 28 | 12.91 | 25175.9 | 12.52
-64.8 | 0 | 28 | 11.22 | 21866 | 10.55
-65.9 | 0 | 28 | 9.21 | 17954.3 | 8.94
-67.2 | 0 | 28 | 7.93 | 15453 | 7.56
-68.4 | 0 | 28 | 6.56 | 12792.1 | 6.48
-69.6 | 0 | 28 | 5.74 | 11187.3 | 5.57
-70.8 | 0 | 28 | 5.07 | 9880.6 | 4.78
-72 | 0 | 28 | 4.25 | 8290.1 | 4.15
-73.2 | 0 | 28 | 3.75 | 7304.8 | 3.58
-74.4 | 0 | 28 | 3.22 | 6280.2 | 3.08
-75.6 | 0 | 28 | 2.84 | 5537.5 | 2.71
-76.8 | 0 | 28 | 2.46 | 4801.7 | 2.33
-78 | 0 | 28 | 2.14 | 4178.9 | 2.06
-79.2 | 0 | 28 | 1.83 | 3571.6 | 1.76
-80.4 | 0 | 28 | 1.67 | 3261.8 | 1.61
-81.6 | 0 | 28 | 1.54 | 3005.6 | 1.45
-82.8 | 0 | 28 | 1.28 | 2489.7 | 1.28
-84 | 0 | 28 | 1.17 | 2278.3 | 1.14
-85.2 | 0 | 28 | 1.11 | 2157.2 | 1.03
-86.4 | 0 | 28 | 0.93 | 1805.8 | 0.93
-87.6 | 0 | 28 | 0.89 | 1743.2 | 0.85
-88.8 | 0 | 28 | 0.83 | 1617.7 | 0.8
-90 | 0 | 28 | 0.74 | 1446.2 | 0.71
-91.2 | 0 | 28 | 0.64 | 1251.6 | 0.61
-92.4 | 0 | 28 | 0.59 | 1157.6 | 0.57
-93.6 | 0 | 28 | 0.59 | 1155.9 | 0.56
-94.8 | 0 | 28 | 0.57 | 1108 | 0.54
-96 | 0 | 28 | 0.54 | 1050.8 | 0.51
-97.2 | 0 | 28 | 0.51 | 996.8 | 0.49
-98.4 | 0 | 28 | 0.46 | 888.6 | 0.44
-99.6 | 0 | 28 | 0.43 | 835.5 | 0.41
-100.9 | 0 | 28 | 0.42 | 816.2 | 0.41
-102 | 0 | 28 | 0.38 | 737.8 | 0.36
-103.3 | 0 | 28 | 0.42 | 825.4 | 0.39
-104.5 | 0 | 28 | 0.37 | 714.6 | 0.36
-105.6 | 0 | 28 | 0.38 | 739 | 0.36
-106.9 | 0 | 28 | 0.35 | 673.9 | 0.34
-108 | 0 | 28 | 0.33 | 641.6 | 0.31
-109.2 | 0 | 28 | 0.31 | 594.8 | 0.3
-110.4 | 0 | 28 | 0.3 | 594 | 0.3
-111.6 | 0 | 28 | 0.32 | 627.1 | 0.32
-112.9 | 0 | 28 | 0.3 | 594.5 | 0.29
-114 | 0 | 28 | 0.34 | 662.5 | 0.32
-115.3 | 0 | 28 | 0.26 | 502.5 | 0.25
-116.5 | 0 | 28 | 0.32 | 621.2 | 0.3
-117.7 | 0 | 28 | 0.29 | 568.5 | 0.28
-118.9 | 0 | 28 | 0.29 | 574.3 | 0.29
-120.1 | 0 | 28 | 0.29 | 565.3 | 0.27
-121.3 | 0 | 28 | 0.29 | 557.1 | 0.27
-122.5 | 0 | 28 | 0.31 | 602.3 | 0.3
-123.7 | 0 | 28 | 0.28 | 555.2 | 0.27
-124.9 | 0 | 28 | 0.27 | 519 | 0.25
-126.1 | 0 | 28 | 0.28 | 553.1 | 0.27
-127.3 | 0 | 28 | 0.27 | 525.6 | 0.26
-128.5 | 0 | 28 | 0.3 | 585.2 | 0.28
-129.7 | 0 | 28 | 0.25 | 497.1 | 0.25
-130.9 | 0 | 28 | 0.27 | 523 | 0.26
-132.1 | 0 | 28 | 0.26 | 515.1 | 0.25
-133.3 | 0 | 28 | 0.31 | 601.5 | 0.3
-134.5 | 0 | 28 | 0.3 | 588 | 0.29
-135.7 | 0 | 28 | 0.29 | 562.7 | 0.27
-136.9 | 0 | 28 | 0.29 | 572.6 | 0.28
-138.1 | 0 | 28 | 0.33 | 650.3 | 0.32
-139.3 | 0 | 28 | 0.32 | 631 | 0.31
-140.5 | 0 | 28 | 0.3 | 585.4 | 0.29
-141.7 | 0 | 28 | 0.3 | 589.4 | 0.28
-142.9 | 0 | 28 | 0.31 | 606.4 | 0.3
-144.1 | 0 | 28 | 0.3 | 580.8 | 0.28
-145.3 | 0 | 28 | 0.32 | 623.1 | 0.31
-146.5 | 0 | 28 | 0.31 | 599.6 | 0.29
-147.7 | 0 | 28 | 0.31 | 610 | 0.3
-148.9 | 0 | 28 | 0.3 | 585.4 | 0.29
-150.1 | 0 | 28 | 0.29 | 570.8 | 0.28
-151.4 | 0 | 28 | 0.26 | 504.5 | 0.25
Measurement time: | 02/09/2022 19:09:02
Radiation device: | Clinac3 Accelerator
Energy: | 18 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1000 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-151.1 | 0 | 21 | 0.26 | 512.4 | 0.31
-150.3 | 0 | 21 | 0.79 | 1542.5 | 0.79
-148.9 | 0 | 21 | 0.8 | 1557.6 | 0.8
-147.8 | 0 | 21 | 0.8 | 1554.9 | 0.8
-146.6 | 0 | 21 | 0.81 | 1574.6 | 0.78
-145.3 | 0 | 21 | 0.85 | 1665.3 | 0.86
-144.1 | 0 | 21 | 0.85 | 1665.3 | 0.86
-142.9 | 0 | 21 | 0.83 | 1609.6 | 0.8
-141.7 | 0 | 21 | 0.79 | 1548.3 | 0.79
-140.5 | 0 | 21 | 0.85 | 1657.5 | 0.82
-139.3 | 0 | 21 | 0.8 | 1564.8 | 0.81
-138.1 | 0 | 21 | 0.83 | 1608.5 | 0.83
-136.9 | 0 | 21 | 0.85 | 1660.3 | 0.83
-135.7 | 0 | 21 | 0.84 | 1632.6 | 0.84
-134.5 | 0 | 21 | 0.77 | 1504.6 | 0.78
-133.3 | 0 | 21 | 0.87 | 1696.4 | 0.84
-132.1 | 0 | 21 | 0.79 | 1549.5 | 0.79
-130.9 | 0 | 21 | 0.74 | 1443.4 | 0.75
-129.7 | 0 | 21 | 0.8 | 1552.7 | 0.78
-128.5 | 0 | 21 | 0.75 | 1471.1 | 0.74
-127.3 | 0 | 21 | 0.77 | 1509.5 | 0.77
-126.1 | 0 | 21 | 0.81 | 1574.5 | 0.79
-124.9 | 0 | 21 | 0.8 | 1553.8 | 0.78
-123.7 | 0 | 21 | 0.83 | 1625.2 | 0.83
-122.5 | 0 | 21 | 0.85 | 1665.9 | 0.85
-121.3 | 0 | 21 | 0.85 | 1662.3 | 0.82
-120.1 | 0 | 21 | 0.82 | 1593.2 | 0.82
-118.9 | 0 | 21 | 0.84 | 1629.5 | 0.83
-117.7 | 0 | 21 | 0.87 | 1702.3 | 0.84
-116.5 | 0 | 21 | 0.84 | 1630.8 | 0.84
-115.3 | 0 | 21 | 0.82 | 1604.5 | 0.82
-114.1 | 0 | 21 | 0.89 | 1728.4 | 0.85
-112.9 | 0 | 21 | 0.89 | 1732 | 0.89
-111.7 | 0 | 21 | 0.88 | 1714.5 | 0.88
-110.5 | 0 | 21 | 0.92 | 1784.8 | 0.88
-109.3 | 0 | 21 | 0.85 | 1661.6 | 0.85
-108.1 | 0 | 21 | 0.92 | 1796.8 | 0.92
-106.9 | 0 | 21 | 0.93 | 1813.1 | 0.9
-105.7 | 0 | 21 | 0.94 | 1838.1 | 0.95
-104.5 | 0 | 21 | 1 | 1942.4 | 1
-103.2 | 0 | 21 | 1.01 | 1978.1 | 0.98
-102.1 | 0 | 21 | 1.06 | 2062 | 1.06
-100.8 | 0 | 21 | 1.06 | 2074.7 | 1.07
-99.6 | 0 | 21 | 1.15 | 2247.2 | 1.11
-98.4 | 0 | 21 | 1.12 | 2192.6 | 1.13
-97.2 | 0 | 21 | 1.16 | 2262.8 | 1.16
-96 | 0 | 21 | 1.3 | 2537.1 | 1.25
-94.8 | 0 | 21 | 1.31 | 2553.6 | 1.32
-93.6 | 0 | 21 | 1.39 | 2707.1 | 1.4
-92.4 | 0 | 21 | 1.55 | 3022.4 | 1.49
-91.2 | 0 | 21 | 1.58 | 3078 | 1.59
-90 | 0 | 21 | 1.69 | 3298.9 | 1.68
-88.8 | 0 | 21 | 1.85 | 3603.1 | 1.76
-87.6 | 0 | 21 | 1.89 | 3687.4 | 1.89
-86.4 | 0 | 21 | 1.96 | 3818.2 | 1.96
-85.2 | 0 | 21 | 2.19 | 4265.4 | 2.08
-84 | 0 | 21 | 2.11 | 4113.5 | 2.11
-82.8 | 0 | 21 | 2.24 | 4361 | 2.24
-81.6 | 0 | 21 | 2.46 | 4797.4 | 2.35
-80.4 | 0 | 21 | 2.42 | 4716.5 | 2.4
-79.2 | 0 | 21 | 2.44 | 4761.8 | 2.46
-78 | 0 | 21 | 2.7 | 5271.5 | 2.59
-76.8 | 0 | 21 | 2.67 | 5211.4 | 2.68
-75.6 | 0 | 21 | 2.81 | 5480 | 2.81
-74.4 | 0 | 21 | 3.12 | 6079.1 | 2.98
-73.2 | 0 | 21 | 3.17 | 6174.8 | 3.18
-72 | 0 | 21 | 3.33 | 6495.8 | 3.33
-70.8 | 0 | 21 | 3.67 | 7163.4 | 3.51
-69.6 | 0 | 21 | 3.73 | 7281.3 | 3.74
-68.4 | 0 | 21 | 4.05 | 7893.5 | 4.02
-67.2 | 0 | 21 | 4.44 | 8652.3 | 4.29
-66 | 0 | 21 | 4.61 | 8979.8 | 4.59
-64.7 | 0 | 21 | 5.01 | 9761 | 5.03
-63.6 | 0 | 21 | 5.81 | 11321.3 | 5.57
-62.4 | 0 | 21 | 6.2 | 12095.2 | 6.2
-61.2 | 0 | 21 | 7.11 | 13856.6 | 7.12
-59.9 | 0 | 21 | 8.99 | 17520.8 | 8.63
-58.7 | 0 | 21 | 10.99 | 21420.4 | 11.04
-57.5 | 0 | 21 | 14.87 | 28989.2 | 14.85
-56.3 | 0 | 21 | 21.84 | 42567.8 | 20.85
-55.1 | 0 | 21 | 29.44 | 57399.3 | 29.43
-53.9 | 0 | 21 | 39.99 | 77959.1 | 39.88
-52.7 | 0 | 21 | 54.83 | 106884.1 | 52.21
-51.5 | 0 | 21 | 64.63 | 125995.3 | 64.4
-50.3 | 0 | 21 | 74.91 | 146027.6 | 74.81
-49.1 | 0 | 21 | 86.94 | 169487.8 | 82.89
-47.9 | 0 | 21 | 88.64 | 172797.3 | 88.33
-46.7 | 0 | 21 | 92.21 | 179770.6 | 91.94
-45.5 | 0 | 21 | 98.79 | 192598.6 | 94.2
-44.3 | 0 | 21 | 95.05 | 185291.5 | 95.28
-43.1 | 0 | 21 | 96.03 | 187219.8 | 96.08
-41.9 | 0 | 21 | 101.6 | 198064.7 | 96.9
-40.7 | 0 | 21 | 96.68 | 188469.7 | 97.19
-39.5 | 0 | 21 | 96.13 | 187404.7 | 97.53
-38.3 | 0 | 21 | 102.26 | 199359.7 | 97.88
-37.1 | 0 | 21 | 96.77 | 188653.6 | 97.97
-35.9 | 0 | 21 | 97.67 | 190410 | 98.06
-34.7 | 0 | 21 | 101.83 | 198525.5 | 98.34
-33.5 | 0 | 21 | 97.87 | 190794 | 98.35
-32.3 | 0 | 21 | 98.03 | 191108 | 98.35
-31.1 | 0 | 21 | 102.42 | 199668.6 | 98.65
-29.9 | 0 | 21 | 98.66 | 192344.8 | 98.75
-28.7 | 0 | 21 | 98.46 | 191955.4 | 98.85
-27.5 | 0 | 21 | 103.48 | 201742.3 | 98.88
-26.3 | 0 | 21 | 99.02 | 193035.1 | 98.99
-25.1 | 0 | 21 | 98.93 | 192868.6 | 99.01
-23.9 | 0 | 21 | 103.86 | 202476.9 | 99.21
-22.7 | 0 | 21 | 99 | 193005.1 | 99.12
-21.5 | 0 | 21 | 99.44 | 193865.3 | 99.28
-20.3 | 0 | 21 | 103.7 | 202164.7 | 99.37
-19.1 | 0 | 21 | 99.22 | 193432.5 | 99.42
-17.9 | 0 | 21 | 98.93 | 192859.6 | 99.43
-16.7 | 0 | 21 | 104.11 | 202956.6 | 99.58
-15.5 | 0 | 21 | 98.73 | 192465.8 | 99.37
-14.3 | 0 | 21 | 99.75 | 194471.5 | 99.63
-13 | 0 | 21 | 103.35 | 201473.6 | 99.66
-11.8 | 0 | 21 | 99.24 | 193472.3 | 99.37
-10.6 | 0 | 21 | 99.57 | 194115.6 | 99.6
-9.4 | 0 | 21 | 103.66 | 202082.4 | 99.78
-8.2 | 0 | 21 | 99.67 | 194316.2 | 99.75
-7 | 0 | 21 | 99.16 | 193319.4 | 99.58
-5.8 | 0 | 21 | 103.66 | 202088.9 | 99.79
-4.6 | 0 | 21 | 99.43 | 193838.5 | 99.73
-3.4 | 0 | 21 | 99.58 | 194122.5 | 99.68
-2.2 | 0 | 21 | 103.28 | 201340.1 | 99.68
-1 | 0 | 21 | 100.67 | 196261.7 | 99.82
0.2 | 0 | 21 | 99.11 | 193208.7 | 99.94
1.4 | 0 | 21 | 104.94 | 204585.4 | 99.85
2.6 | 0 | 21 | 99.7 | 194374.5 | 99.73
3.8 | 0 | 21 | 101.09 | 197084.7 | 99.78
5 | 0 | 21 | 103.36 | 201493.1 | 99.81
6.2 | 0 | 21 | 97.59 | 190247.4 | 100
7.4 | 0 | 21 | 101.94 | 198729.2 | 99.9
8.6 | 0 | 21 | 102.49 | 199807.9 | 99.88
9.8 | 0 | 21 | 99.98 | 194917.1 | 99.79
11 | 0 | 21 | 100.11 | 195165.7 | 99.83
12.2 | 0 | 21 | 102.38 | 199589.3 | 99.77
13.4 | 0 | 21 | 99.77 | 194500.1 | 99.53
14.6 | 0 | 21 | 100.78 | 196465.9 | 99.63
15.8 | 0 | 21 | 102.91 | 200616.8 | 99.66
17 | 0 | 21 | 100.41 | 195742.2 | 99.42
18.2 | 0 | 21 | 100.25 | 195443.1 | 99.56
19.4 | 0 | 21 | 103.11 | 201021.1 | 99.56
20.6 | 0 | 21 | 99.43 | 193842.2 | 99.34
21.8 | 0 | 21 | 100.48 | 195882.4 | 99.37
23 | 0 | 21 | 102.99 | 200786.1 | 99.26
24.2 | 0 | 21 | 98.94 | 192876.6 | 99.27
25.4 | 0 | 21 | 100.37 | 195677.5 | 99.17
26.6 | 0 | 21 | 102 | 198857.5 | 99.16
27.8 | 0 | 21 | 98.14 | 191330.1 | 99.09
29 | 0 | 21 | 100.52 | 195966.4 | 98.93
30.2 | 0 | 21 | 100.32 | 195581.1 | 99.02
31.4 | 0 | 21 | 99.35 | 193688.6 | 98.63
32.6 | 0 | 21 | 99.24 | 193477.9 | 98.55
33.8 | 0 | 21 | 100.06 | 195068.5 | 98.51
35 | 0 | 21 | 98.7 | 192409.4 | 98.35
36.2 | 0 | 21 | 98.53 | 192077.7 | 98.1
37.4 | 0 | 21 | 100.36 | 195658 | 97.92
38.6 | 0 | 21 | 96.32 | 187778.9 | 97.52
39.8 | 0 | 21 | 98.54 | 192100.7 | 96.97
41 | 0 | 21 | 98.03 | 191117.3 | 96.55
42.2 | 0 | 21 | 95.07 | 185334.3 | 95.4
43.4 | 0 | 21 | 96.58 | 188282.5 | 93.72
44.6 | 0 | 21 | 92.69 | 180696.8 | 91.35
45.8 | 0 | 21 | 86.71 | 169042.5 | 86.8
47 | 0 | 21 | 82.77 | 161355.2 | 80.57
48.2 | 0 | 21 | 72.51 | 141355.4 | 71.63
49.4 | 0 | 21 | 60.66 | 118265.8 | 60.26
50.6 | 0 | 21 | 49.78 | 97055.2 | 48.06
51.9 | 0 | 21 | 36.2 | 70579.3 | 36
53.1 | 0 | 21 | 26.04 | 50767 | 25.99
54.3 | 0 | 21 | 19.39 | 37793.7 | 18.62
55.5 | 0 | 21 | 13.48 | 26279.2 | 13.46
56.7 | 0 | 21 | 10.19 | 19874.9 | 10.19
57.9 | 0 | 21 | 8.28 | 16145.7 | 8.27
59.1 | 0 | 21 | 7.35 | 14336.3 | 7.04
60.3 | 0 | 21 | 6.32 | 12313.4 | 6.29
61.5 | 0 | 21 | 5.97 | 11638.4 | 5.73
62.7 | 0 | 21 | 5.31 | 10353.9 | 5.3
63.9 | 0 | 21 | 4.93 | 9609.5 | 4.94
65.1 | 0 | 21 | 4.74 | 9238 | 4.56
66.3 | 0 | 21 | 4.22 | 8219.9 | 4.24
67.5 | 0 | 21 | 3.95 | 7700.2 | 3.96
68.7 | 0 | 21 | 3.84 | 7494.2 | 3.7
69.9 | 0 | 21 | 3.49 | 6794.9 | 3.47
71.1 | 0 | 21 | 3.29 | 6417.6 | 3.31
72.3 | 0 | 21 | 3.25 | 6344.8 | 3.17
73.5 | 0 | 21 | 3.05 | 5953.6 | 3.02
74.7 | 0 | 21 | 2.87 | 5590.8 | 2.87
75.9 | 0 | 21 | 2.84 | 5537.5 | 2.78
77.1 | 0 | 21 | 2.68 | 5231.7 | 2.65
78.3 | 0 | 21 | 2.55 | 4979.9 | 2.57
79.5 | 0 | 21 | 2.5 | 4865.3 | 2.43
80.7 | 0 | 21 | 2.38 | 4643.8 | 2.4
81.9 | 0 | 21 | 2.28 | 4443.4 | 2.26
83.1 | 0 | 21 | 2.17 | 4220.7 | 2.12
84.3 | 0 | 21 | 1.95 | 3809 | 1.97
85.6 | 0 | 21 | 1.85 | 3605.1 | 1.81
86.8 | 0 | 21 | 1.74 | 3394 | 1.71
88 | 0 | 21 | 1.58 | 3087.5 | 1.58
89.2 | 0 | 21 | 1.54 | 3011.9 | 1.5
90.4 | 0 | 21 | 1.47 | 2864.6 | 1.44
91.6 | 0 | 21 | 1.36 | 2647.9 | 1.35
92.8 | 0 | 21 | 1.31 | 2557.4 | 1.28
94 | 0 | 21 | 1.26 | 2463.7 | 1.25
95.2 | 0 | 21 | 1.17 | 2282.7 | 1.17
96.4 | 0 | 21 | 1.1 | 2154.1 | 1.07
97.6 | 0 | 21 | 1.11 | 2163.4 | 1.1
98.8 | 0 | 21 | 1.01 | 1969.5 | 1.01
100 | 0 | 21 | 1.04 | 2023.2 | 1
101.2 | 0 | 21 | 1.01 | 1966.4 | 1
102.4 | 0 | 21 | 0.92 | 1799.1 | 0.92
103.6 | 0 | 21 | 0.9 | 1761.6 | 0.87
104.8 | 0 | 21 | 0.88 | 1706.1 | 0.87
106 | 0 | 21 | 0.89 | 1729.6 | 0.86
107.2 | 0 | 21 | 0.85 | 1653.7 | 0.85
108.4 | 0 | 21 | 0.85 | 1664.8 | 0.85
109.6 | 0 | 21 | 0.86 | 1681.7 | 0.87
110.8 | 0 | 21 | 0.84 | 1642.9 | 0.81
112 | 0 | 21 | 0.8 | 1553.7 | 0.79
113.2 | 0 | 21 | 0.79 | 1545.1 | 0.79
114.4 | 0 | 21 | 0.8 | 1564.3 | 0.77
115.6 | 0 | 21 | 0.72 | 1413.2 | 0.72
116.8 | 0 | 21 | 0.7 | 1367.2 | 0.71
118 | 0 | 21 | 0.8 | 1551.6 | 0.77
119.2 | 0 | 21 | 0.71 | 1379.6 | 0.71
120.4 | 0 | 21 | 0.68 | 1325 | 0.68
121.6 | 0 | 21 | 0.72 | 1411.3 | 0.7
122.8 | 0 | 21 | 0.7 | 1359 | 0.7
124 | 0 | 21 | 0.7 | 1359.9 | 0.7
125.2 | 0 | 21 | 0.74 | 1435.9 | 0.71
126.4 | 0 | 21 | 0.66 | 1282.1 | 0.66
127.6 | 0 | 21 | 0.65 | 1258.2 | 0.65
128.8 | 0 | 21 | 0.66 | 1290.3 | 0.64
130 | 0 | 21 | 0.61 | 1197.4 | 0.62
131.2 | 0 | 21 | 0.63 | 1235.4 | 0.64
132.4 | 0 | 21 | 0.68 | 1329.4 | 0.65
133.6 | 0 | 21 | 0.62 | 1214.4 | 0.62
134.8 | 0 | 21 | 0.61 | 1186.8 | 0.61
136 | 0 | 21 | 0.7 | 1355.2 | 0.67
137.2 | 0 | 21 | 0.65 | 1271.6 | 0.65
138.4 | 0 | 21 | 0.65 | 1270.8 | 0.65
139.6 | 0 | 21 | 0.67 | 1309.6 | 0.64
140.8 | 0 | 21 | 0.62 | 1204.7 | 0.61
142 | 0 | 21 | 0.63 | 1218.7 | 0.63
143.3 | 0 | 21 | 0.64 | 1239.2 | 0.61
144.4 | 0 | 21 | 0.61 | 1189.9 | 0.61
145.7 | 0 | 21 | 0.62 | 1200.1 | 0.61
146.9 | 0 | 21 | 0.61 | 1182.2 | 0.58
148.1 | 0 | 21 | 0.57 | 1110.3 | 0.57
149.3 | 0 | 21 | 0.6 | 1170 | 0.6
151.1 | 0 | 21 | 0.62 | 1200.5 | 0.6