## Profils DSP
Measurement time: | 07/09/2022 17:35:28 |  |  |  |  | Measurement time: | 02/09/2022 20:44:47
Radiation device: | Clinac3 Accelerator |  |  |  |  | Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV |  |  |  |  | Energy: | 12 MeV
Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F |  |  |  |  | Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie |  |  |  |  | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 1050 mm |  |  |  |  | SSD: | 1000 mm
Field size: | 100 x 100 mm |  |  |  |  | Field size: | 100 x 100 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | Continuous |  |  |  |  | Scan mode: | Continuous
Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
153.9 | 0 | 28 | 0.28 | 506.7 | 0.19 | -151.4 | 0 | 28 | 1.76 | 2581.1 | 1.57
153.2 | 0 | 28 | 0.09 | 161.1 | 0.09 | -150.6 | 0 | 28 | 1.72 | 2526.5 | 1.28
151.8 | 0 | 28 | 0.11 | 196.8 | 0.11 | -149.3 | 0 | 28 | 1.71 | 2502 | 1.26
150.6 | 0 | 28 | 0.05 | 90.8 | 0.05 | -148.1 | 0 | 28 | 1.73 | 2530.3 | 1.22
149.4 | 0 | 28 | 0.1 | 185.9 | 0.1 | -146.9 | 0 | 28 | 1.73 | 2530.4 | 1.31
148.2 | 0 | 28 | 0.08 | 139.3 | 0.08 | -145.7 | 0 | 28 | 1.71 | 2505.3 | 1.25
147 | 0 | 28 | 0.06 | 111.7 | 0.06 | -144.5 | 0 | 28 | 1.75 | 2566 | 1.23
145.8 | 0 | 28 | 0.11 | 194.2 | 0.11 | -143.3 | 0 | 28 | 1.83 | 2679.8 | 1.4
144.6 | 0 | 28 | 0.1 | 175.6 | 0.1 | -142.1 | 0 | 28 | 1.76 | 2571.7 | 1.34
143.4 | 0 | 28 | 0.08 | 146 | 0.08 | -140.9 | 0 | 28 | 1.7 | 2495.2 | 1.19
142.2 | 0 | 28 | 0.1 | 182.2 | 0.1 | -139.7 | 0 | 28 | 1.77 | 2588.1 | 1.34
141 | 0 | 28 | 0.11 | 196.8 | 0.11 | -138.5 | 0 | 28 | 1.77 | 2589.5 | 1.36
139.8 | 0 | 28 | 0.14 | 260.2 | 0.14 | -137.3 | 0 | 28 | 1.67 | 2450.5 | 1.16
138.6 | 0 | 28 | 0.1 | 177.2 | 0.1 | -136.1 | 0 | 28 | 1.72 | 2524.8 | 1.32
137.4 | 0 | 28 | 0.11 | 200.7 | 0.11 | -134.9 | 0 | 28 | 1.77 | 2591.1 | 1.35
136.2 | 0 | 28 | 0.13 | 233.7 | 0.13 | -133.7 | 0 | 28 | 1.74 | 2547.9 | 1.22
135 | 0 | 28 | 0.12 | 209.5 | 0.11 | -132.5 | 0 | 28 | 1.74 | 2549.7 | 1.31
133.8 | 0 | 28 | 0.15 | 268.7 | 0.15 | -131.3 | 0 | 28 | 1.72 | 2519.9 | 1.26
132.6 | 0 | 28 | 0.17 | 306.1 | 0.17 | -130.1 | 0 | 28 | 1.78 | 2606.6 | 1.25
131.4 | 0 | 28 | 0.19 | 338.5 | 0.19 | -128.9 | 0 | 28 | 1.77 | 2589.6 | 1.35
130.2 | 0 | 28 | 0.15 | 276.4 | 0.15 | -127.7 | 0 | 28 | 1.79 | 2627.4 | 1.31
129 | 0 | 28 | 0.17 | 303.6 | 0.16 | -126.5 | 0 | 28 | 1.73 | 2535.8 | 1.22
127.8 | 0 | 28 | 0.14 | 250.6 | 0.14 | -125.3 | 0 | 28 | 1.74 | 2544.3 | 1.34
126.6 | 0 | 28 | 0.19 | 353.5 | 0.2 | -124.1 | 0 | 28 | 1.69 | 2482.2 | 1.25
125.4 | 0 | 28 | 0.2 | 369.5 | 0.2 | -122.9 | 0 | 28 | 1.73 | 2533.2 | 1.23
124.2 | 0 | 28 | 0.13 | 242.8 | 0.14 | -121.7 | 0 | 28 | 1.74 | 2551.5 | 1.35
123 | 0 | 28 | 0.23 | 414.5 | 0.23 | -120.5 | 0 | 28 | 1.73 | 2538.4 | 1.24
121.8 | 0 | 28 | 0.21 | 383.6 | 0.21 | -119.3 | 0 | 28 | 1.73 | 2540.7 | 1.3
120.6 | 0 | 28 | 0.22 | 409.4 | 0.22 | -118.1 | 0 | 28 | 1.71 | 2504.9 | 1.26
119.4 | 0 | 28 | 0.22 | 403.6 | 0.23 | -116.9 | 0 | 28 | 1.71 | 2509.6 | 1.32
118.2 | 0 | 28 | 0.27 | 495.8 | 0.27 | -115.7 | 0 | 28 | 1.69 | 2473.8 | 1.21
117 | 0 | 28 | 0.3 | 551 | 0.3 | -114.5 | 0 | 28 | 1.8 | 2631.3 | 1.35
115.8 | 0 | 28 | 0.26 | 469.6 | 0.26 | -113.3 | 0 | 28 | 1.68 | 2465.4 | 1.28
114.6 | 0 | 28 | 0.31 | 566.6 | 0.31 | -112.1 | 0 | 28 | 1.67 | 2441 | 1.2
113.4 | 0 | 28 | 0.29 | 529.5 | 0.29 | -110.9 | 0 | 28 | 1.68 | 2456.4 | 1.25
112.2 | 0 | 28 | 0.3 | 552.1 | 0.3 | -109.6 | 0 | 28 | 1.71 | 2504.8 | 1.3
111 | 0 | 28 | 0.33 | 593.4 | 0.32 | -108.4 | 0 | 28 | 1.62 | 2373.4 | 1.18
109.8 | 0 | 28 | 0.33 | 609.6 | 0.33 | -107.2 | 0 | 28 | 1.6 | 2349.4 | 1.2
108.6 | 0 | 28 | 0.35 | 643.3 | 0.37 | -106 | 0 | 28 | 1.63 | 2381.6 | 1.23
107.4 | 0 | 28 | 0.37 | 665.1 | 0.36 | -104.9 | 0 | 28 | 1.58 | 2307.3 | 1.14
106.2 | 0 | 28 | 0.3 | 544.5 | 0.3 | -103.6 | 0 | 28 | 1.57 | 2299.4 | 1.17
105 | 0 | 28 | 0.39 | 701.1 | 0.38 | -102.4 | 0 | 28 | 1.6 | 2341.7 | 1.17
103.8 | 0 | 28 | 0.42 | 771.8 | 0.43 | -101.2 | 0 | 28 | 1.54 | 2249.7 | 1.1
102.6 | 0 | 28 | 0.43 | 784.6 | 0.42 | -100 | 0 | 28 | 1.59 | 2329.3 | 1.18
101.4 | 0 | 28 | 0.44 | 796.9 | 0.43 | -98.8 | 0 | 28 | 1.48 | 2166 | 1.07
100.2 | 0 | 28 | 0.45 | 819.5 | 0.45 | -97.6 | 0 | 28 | 1.42 | 2082.6 | 1.01
99 | 0 | 28 | 0.51 | 931.5 | 0.5 | -96.4 | 0 | 28 | 1.42 | 2075.7 | 1.06
97.7 | 0 | 28 | 0.54 | 983.1 | 0.57 | -95.2 | 0 | 28 | 1.42 | 2079.8 | 1.04
96.6 | 0 | 28 | 0.63 | 1153.7 | 0.63 | -94 | 0 | 28 | 1.42 | 2079.3 | 1.06
95.4 | 0 | 28 | 0.62 | 1124.1 | 0.61 | -92.8 | 0 | 28 | 1.3 | 1906.7 | 0.97
94.1 | 0 | 28 | 0.69 | 1256.5 | 0.7 | -91.6 | 0 | 28 | 1.22 | 1793.7 | 0.89
93 | 0 | 28 | 0.76 | 1380.8 | 0.75 | -90.4 | 0 | 28 | 1.13 | 1651.3 | 0.81
91.7 | 0 | 28 | 0.82 | 1485.7 | 0.82 | -89.2 | 0 | 28 | 1.07 | 1565.8 | 0.79
90.6 | 0 | 28 | 0.89 | 1613.5 | 0.91 | -88 | 0 | 28 | 0.9 | 1319.6 | 0.65
89.3 | 0 | 28 | 1.01 | 1833.8 | 1.02 | -86.8 | 0 | 28 | 0.8 | 1164.9 | 0.57
88.1 | 0 | 28 | 1.1 | 2008.3 | 1.13 | -85.6 | 0 | 28 | 0.71 | 1046.1 | 0.53
87 | 0 | 28 | 1.21 | 2197.7 | 1.19 | -84.4 | 0 | 28 | 0.67 | 978.5 | 0.51
85.7 | 0 | 28 | 1.32 | 2406.9 | 1.34 | -83.2 | 0 | 28 | 0.44 | 650.2 | 0.32
84.5 | 0 | 28 | 1.46 | 2666.7 | 1.45 | -82 | 0 | 28 | 0.24 | 345.7 | 0.17
83.3 | 0 | 28 | 1.64 | 2994.4 | 1.64 | -80.8 | 0 | 28 | 0.01 | 15.2 | 0.01
82.1 | 0 | 28 | 1.78 | 3242.3 | 1.77 | -79.6 | 0 | 28 | 0.21 | 307.1 | 0.15
80.9 | 0 | 28 | 2.07 | 3772.8 | 2.03 | -78.4 | 0 | 28 | 0.54 | 787.5 | 0.4
79.7 | 0 | 28 | 2.27 | 4125.9 | 2.27 | -77.2 | 0 | 28 | 0.86 | 1255 | 0.63
78.5 | 0 | 28 | 2.57 | 4675.2 | 2.54 | -76 | 0 | 28 | 1.27 | 1860.9 | 0.93
77.3 | 0 | 28 | 2.8 | 5092.6 | 2.82 | -74.8 | 0 | 28 | 1.74 | 2554.7 | 1.29
76.1 | 0 | 28 | 3.16 | 5752.9 | 3.15 | -73.6 | 0 | 28 | 2.33 | 3418.8 | 1.74
74.9 | 0 | 28 | 3.61 | 6575.7 | 3.63 | -72.4 | 0 | 28 | 3.01 | 4415.3 | 2.22
73.7 | 0 | 28 | 4.03 | 7341.6 | 4.06 | -71.2 | 0 | 28 | 3.69 | 5404.3 | 2.74
72.5 | 0 | 28 | 4.73 | 8602.3 | 4.68 | -70 | 0 | 28 | 4.73 | 6928.3 | 3.51
71.3 | 0 | 28 | 5.44 | 9906 | 5.46 | -68.8 | 0 | 28 | 5.76 | 8443.7 | 4.28
70.1 | 0 | 28 | 6.36 | 11576.6 | 6.3 | -67.6 | 0 | 28 | 7.23 | 10583.4 | 5.36
68.9 | 0 | 28 | 7.42 | 13512.2 | 7.3 | -66.4 | 0 | 28 | 8.86 | 12980.4 | 6.57
67.7 | 0 | 28 | 8.61 | 15677.8 | 8.61 | -65.2 | 0 | 28 | 10.97 | 16075.8 | 8.08
66.5 | 0 | 28 | 10.25 | 18651.1 | 10.05 | -63.9 | 0 | 28 | 13.35 | 19553 | 9.84
65.3 | 0 | 28 | 11.95 | 21755.4 | 11.84 | -62.8 | 0 | 28 | 16.1 | 23577.8 | 11.89
64.1 | 0 | 28 | 14.04 | 25563.2 | 14.07 | -61.5 | 0 | 28 | 20.37 | 29837.9 | 14.75
62.9 | 0 | 28 | 16.54 | 30113.3 | 16.54 | -60.3 | 0 | 28 | 24.25 | 35515.2 | 17.91
61.7 | 0 | 28 | 19.38 | 35277.7 | 19.45 | -59.1 | 0 | 28 | 29.9 | 43803.2 | 21.86
60.5 | 0 | 28 | 22.94 | 41765.8 | 23.05 | -57.9 | 0 | 28 | 36.28 | 53146.8 | 26.52
59.3 | 0 | 28 | 26.99 | 49134.2 | 27.27 | -56.7 | 0 | 28 | 43.08 | 63106.3 | 31.8
58.1 | 0 | 28 | 31.99 | 58231.8 | 31.97 | -55.5 | 0 | 28 | 52.41 | 76767.7 | 37.85
56.9 | 0 | 28 | 37.38 | 68045.1 | 37.07 | -54.3 | 0 | 28 | 58 | 84962.2 | 44.08
55.7 | 0 | 28 | 42.64 | 77617.7 | 42.8 | -53.1 | 0 | 28 | 69.25 | 101444.7 | 50.93
54.5 | 0 | 28 | 48.36 | 88030 | 48.28 | -51.9 | 0 | 28 | 79.02 | 115743.6 | 57.53
53.3 | 0 | 28 | 52.99 | 96456.4 | 53.94 | -50.7 | 0 | 28 | 86.8 | 127141.9 | 63.92
52 | 0 | 28 | 59.66 | 108610.4 | 59.72 | -49.5 | 0 | 28 | 95.39 | 139734.3 | 69.8
50.9 | 0 | 28 | 64.45 | 117322.7 | 64.95 | -48.3 | 0 | 28 | 100.92 | 147834.9 | 74.86
49.6 | 0 | 28 | 72.18 | 131393.8 | 70.05 | -47.1 | 0 | 28 | 107.55 | 157547.6 | 79.45
48.5 | 0 | 28 | 73.38 | 133569.9 | 74.25 | -45.9 | 0 | 28 | 112.24 | 164407.3 | 83.08
47.2 | 0 | 28 | 77.31 | 140728.7 | 78.19 | -44.7 | 0 | 28 | 115.96 | 169859.7 | 85.93
46 | 0 | 28 | 83.71 | 152389.7 | 81.81 | -43.5 | 0 | 28 | 119.52 | 175068.4 | 88.54
44.8 | 0 | 28 | 84.34 | 153522.2 | 84.51 | -42.3 | 0 | 28 | 121.97 | 178662.5 | 90.3
43.6 | 0 | 28 | 84.81 | 154377.5 | 87.17 | -41.1 | 0 | 28 | 124.96 | 183044 | 91.99
42.5 | 0 | 28 | 89.23 | 162428.4 | 89.36 | -39.9 | 0 | 28 | 128.95 | 188887.2 | 93.34
41.2 | 0 | 28 | 89.96 | 163762.6 | 90.88 | -38.7 | 0 | 28 | 126.97 | 185986.1 | 94.27
40 | 0 | 28 | 90.72 | 165151.9 | 92.61 | -37.5 | 0 | 28 | 131.39 | 192459.4 | 95.2
38.8 | 0 | 28 | 94.51 | 172040.1 | 93.69 | -36.3 | 0 | 28 | 131.67 | 192877.3 | 95.76
37.6 | 0 | 28 | 93.32 | 169877 | 94.84 | -35.1 | 0 | 28 | 130.81 | 191611.9 | 96.4
36.4 | 0 | 28 | 93.98 | 171083.3 | 95.61 | -33.9 | 0 | 28 | 126.7 | 185596.5 | 96.49
35.2 | 0 | 28 | 99.33 | 180820.9 | 96.32 | -32.7 | 0 | 28 | 132.69 | 194362.4 | 97.32
34 | 0 | 28 | 95.96 | 174675.1 | 96.84 | -31.5 | 0 | 28 | 135.43 | 198380.1 | 97.64
32.8 | 0 | 28 | 97.44 | 177381.5 | 97.21 | -30.3 | 0 | 28 | 131.52 | 192648.6 | 97.77
31.6 | 0 | 28 | 102.13 | 185917.2 | 97.97 | -29.1 | 0 | 28 | 133.63 | 195742.4 | 98.17
30.4 | 0 | 28 | 96.17 | 175067.9 | 98.18 | -27.9 | 0 | 28 | 133.33 | 195301.7 | 98.24
29.2 | 0 | 28 | 100.02 | 182080.8 | 98.54 | -26.7 | 0 | 28 | 132.89 | 194654.7 | 98.46
28 | 0 | 28 | 101.18 | 184188.4 | 98.63 | -25.5 | 0 | 28 | 136.53 | 199995.8 | 98.82
26.8 | 0 | 28 | 95.11 | 173139.8 | 98.89 | -24.3 | 0 | 28 | 133.36 | 195342.4 | 98.81
25.6 | 0 | 28 | 100.9 | 183674.4 | 99.09 | -23.1 | 0 | 28 | 130.51 | 191170.9 | 98.88
24.4 | 0 | 28 | 100.48 | 182913.4 | 99.3 | -21.9 | 0 | 28 | 134.22 | 196604.4 | 99.1
23.2 | 0 | 28 | 95.43 | 173724.7 | 98.95 | -20.7 | 0 | 28 | 134.82 | 197493.1 | 99.41
22 | 0 | 28 | 99.9 | 181849.5 | 99.49 | -19.5 | 0 | 28 | 134.3 | 196732 | 99.4
20.8 | 0 | 28 | 102.91 | 187334.3 | 99.55 | -18.3 | 0 | 28 | 135.67 | 198726.8 | 99.51
19.6 | 0 | 28 | 96.35 | 175382.6 | 99.51 | -17 | 0 | 28 | 135.52 | 198509.6 | 99.43
18.4 | 0 | 28 | 98.96 | 180135.9 | 99.4 | -15.8 | 0 | 28 | 136.51 | 199966.6 | 99.47
17.2 | 0 | 28 | 104.8 | 190777.6 | 99.78 | -14.6 | 0 | 28 | 129.76 | 190072 | 99.58
16 | 0 | 28 | 97.46 | 177405.6 | 99.59 | -13.4 | 0 | 28 | 137.47 | 201362.9 | 99.62
14.8 | 0 | 28 | 100.18 | 182370.8 | 99.81 | -12.2 | 0 | 28 | 136.58 | 200059.9 | 99.55
13.6 | 0 | 28 | 104.29 | 189846.3 | 99.81 | -11 | 0 | 28 | 130.46 | 191098 | 99.64
12.4 | 0 | 28 | 97.86 | 178140.9 | 99.88 | -9.8 | 0 | 28 | 137.42 | 201295 | 99.76
11.2 | 0 | 28 | 100.43 | 182809.7 | 99.82 | -8.6 | 0 | 28 | 137.13 | 200874.9 | 99.73
10 | 0 | 28 | 102.25 | 186139.7 | 100 | -7.4 | 0 | 28 | 136.26 | 199592 | 99.84
8.8 | 0 | 28 | 96.26 | 175225.6 | 99.68 | -6.2 | 0 | 28 | 137.02 | 200711.1 | 99.93
7.6 | 0 | 28 | 102.36 | 186335.9 | 99.79 | -5 | 0 | 28 | 137.34 | 201171.9 | 99.68
6.4 | 0 | 28 | 99.86 | 181780.8 | 99.94 | -3.8 | 0 | 28 | 130.05 | 190500.8 | 99.71
5.2 | 0 | 28 | 95.92 | 174609.2 | 99.73 | -2.6 | 0 | 28 | 137.63 | 201597.9 | 99.95
4 | 0 | 28 | 100.71 | 183323.6 | 99.89 | -1.4 | 0 | 28 | 136.49 | 199935.9 | 99.75
2.8 | 0 | 28 | 105.02 | 191181.6 | 99.92 | -0.2 | 0 | 28 | 136.8 | 200386.7 | 99.87
1.5 | 0 | 28 | 96.3 | 175301 | 99.66 | 1 | 0 | 28 | 136.64 | 200154.9 | 99.78
0.4 | 0 | 28 | 96.36 | 175401.1 | 99.76 | 2.2 | 0 | 28 | 135.82 | 198954.9 | 99.87
-0.8 | 0 | 28 | 105.37 | 191818.7 | 99.75 | 3.4 | 0 | 28 | 135.96 | 199159.9 | 100
-2 | 0 | 28 | 95.81 | 174399.7 | 99.8 | 4.6 | 0 | 28 | 135.68 | 198749.5 | 99.88
-3.2 | 0 | 28 | 97.39 | 177278.1 | 99.82 | 5.8 | 0 | 28 | 135.01 | 197769.5 | 99.87
-4.5 | 0 | 28 | 104.05 | 189401.1 | 99.9 | 7 | 0 | 28 | 135.8 | 198923.4 | 99.87
-5.6 | 0 | 28 | 96.1 | 174944.8 | 99.47 | 8.2 | 0 | 28 | 134.75 | 197379.3 | 99.77
-6.9 | 0 | 28 | 99.51 | 181134.7 | 99.55 | 9.4 | 0 | 28 | 134.72 | 197337.5 | 99.71
-8.1 | 0 | 28 | 102.9 | 187307.2 | 99.6 | 10.6 | 0 | 28 | 134.49 | 196997.1 | 99.79
-9.2 | 0 | 28 | 98.35 | 179038.4 | 99.44 | 11.8 | 0 | 28 | 139.41 | 204215.9 | 99.69
-10.5 | 0 | 28 | 97.08 | 176712.3 | 99.41 | 13 | 0 | 28 | 133.72 | 195882.7 | 99.41
-11.6 | 0 | 28 | 102.03 | 185727.2 | 99.49 | 14.2 | 0 | 28 | 127.98 | 187471.3 | 99.3
-12.9 | 0 | 28 | 98.62 | 179518.4 | 99.07 | 15.4 | 0 | 28 | 139.75 | 204704.5 | 99.56
-14.1 | 0 | 28 | 95.2 | 173298.9 | 99.26 | 16.6 | 0 | 28 | 133.45 | 195479.2 | 99.24
-15.3 | 0 | 28 | 103.46 | 188342.5 | 99.3 | 17.8 | 0 | 28 | 127.96 | 187434.1 | 99.22
-16.5 | 0 | 28 | 95.05 | 173020.8 | 99.38 | 19 | 0 | 28 | 139.58 | 204460 | 99.43
-17.7 | 0 | 28 | 99.44 | 181016.1 | 99.05 | 20.2 | 0 | 28 | 132.73 | 194429.6 | 98.97
-18.9 | 0 | 28 | 104.47 | 190181.4 | 99.13 | 21.4 | 0 | 28 | 131.15 | 192112.8 | 99.01
-20.1 | 0 | 28 | 95.6 | 174034.6 | 98.89 | 22.6 | 0 | 28 | 133.38 | 195381 | 98.87
-21.3 | 0 | 28 | 100.38 | 182731 | 98.78 | 23.8 | 0 | 28 | 132.62 | 194262.1 | 98.59
-22.5 | 0 | 28 | 104.03 | 189368 | 98.91 | 25 | 0 | 28 | 128.1 | 187643.9 | 98.41
-23.7 | 0 | 28 | 95.47 | 173783.9 | 98.68 | 26.2 | 0 | 28 | 136.81 | 200405.7 | 98.31
-24.9 | 0 | 28 | 98.7 | 179673.3 | 98.41 | 27.4 | 0 | 28 | 127.93 | 187396.9 | 97.77
-26.1 | 0 | 28 | 101.47 | 184717.4 | 98.3 | 28.6 | 0 | 28 | 131.63 | 192811.3 | 97.8
-27.3 | 0 | 28 | 94.54 | 172097.4 | 98.1 | 29.8 | 0 | 28 | 135.84 | 198979.3 | 97.56
-28.5 | 0 | 28 | 97.03 | 176633.6 | 97.95 | 31 | 0 | 28 | 130.28 | 190834.3 | 96.98
-29.7 | 0 | 28 | 99.37 | 180882.4 | 97.71 | 32.2 | 0 | 28 | 125.7 | 184131.4 | 96.36
-30.9 | 0 | 28 | 95.19 | 173288.3 | 97.49 | 33.4 | 0 | 28 | 133.01 | 194831.8 | 95.97
-32.1 | 0 | 28 | 96.51 | 175681.2 | 97.15 | 34.6 | 0 | 28 | 129.35 | 189476.7 | 95.17
-33.3 | 0 | 28 | 99.26 | 180696.2 | 96.91 | 35.8 | 0 | 28 | 123.04 | 180229.1 | 94.18
-34.5 | 0 | 28 | 94.66 | 172316.9 | 96.36 | 37 | 0 | 28 | 128.87 | 188770.2 | 93.27
-35.7 | 0 | 28 | 97.18 | 176902.9 | 95.85 | 38.2 | 0 | 28 | 120.54 | 176573.1 | 91.7
-36.9 | 0 | 28 | 97.68 | 177821.5 | 95.48 | 39.4 | 0 | 28 | 120.95 | 177175.5 | 90.58
-38.1 | 0 | 28 | 93.89 | 170913.2 | 94.52 | 40.6 | 0 | 28 | 123.67 | 181155.8 | 88.59
-39.3 | 0 | 28 | 96.18 | 175075.1 | 93.93 | 41.8 | 0 | 28 | 112.19 | 164335.3 | 85.84
-40.5 | 0 | 28 | 96.43 | 175535.7 | 93.07 | 43 | 0 | 28 | 112.45 | 164723.3 | 83.13
-41.7 | 0 | 28 | 89.61 | 163120.4 | 91.63 | 44.2 | 0 | 28 | 107.69 | 157750.8 | 79.2
-42.9 | 0 | 28 | 89.96 | 163762.1 | 90.39 | 45.4 | 0 | 28 | 102.23 | 149752.8 | 74.93
-44.1 | 0 | 28 | 90.75 | 165192.4 | 88.93 | 46.7 | 0 | 28 | 94.94 | 139073.2 | 69.99
-45.3 | 0 | 28 | 85.52 | 155685.8 | 86.79 | 47.9 | 0 | 28 | 87.18 | 127706.4 | 64.33
-46.5 | 0 | 28 | 84.8 | 154372.9 | 84.41 | 49 | 0 | 28 | 79.03 | 115758.9 | 58
-47.7 | 0 | 28 | 80.99 | 147424.6 | 81.66 | 50.3 | 0 | 28 | 70.06 | 102621.9 | 51.41
-48.9 | 0 | 28 | 77.19 | 140520.9 | 78.2 | 51.5 | 0 | 28 | 61.22 | 89682 | 44.74
-50.1 | 0 | 28 | 74.9 | 136341.6 | 74.22 | 52.7 | 0 | 28 | 51.94 | 76086.8 | 38.05
-51.3 | 0 | 28 | 72.12 | 131292.3 | 69.87 | 53.9 | 0 | 28 | 44.1 | 64595 | 32.13
-52.5 | 0 | 28 | 63.97 | 116450 | 64.76 | 55.1 | 0 | 28 | 36.25 | 53099.6 | 26.51
-53.7 | 0 | 28 | 60.41 | 109967.2 | 59.68 | 56.3 | 0 | 28 | 29.9 | 43795.1 | 21.79
-54.9 | 0 | 28 | 55.28 | 100629.6 | 53.9 | 57.5 | 0 | 28 | 24.34 | 35657.7 | 17.76
-56.2 | 0 | 28 | 48 | 87381.5 | 48.17 | 58.7 | 0 | 28 | 19.78 | 28970.7 | 14.44
-57.4 | 0 | 28 | 43.21 | 78663.8 | 42.71 | 59.9 | 0 | 28 | 16.01 | 23455.8 | 11.7
-58.6 | 0 | 28 | 37.85 | 68894 | 37.09 | 61.1 | 0 | 28 | 12.25 | 17948 | 9.36
-59.8 | 0 | 28 | 31.57 | 57462.9 | 32.17 | 62.3 | 0 | 28 | 10.65 | 15602.3 | 7.77
-61 | 0 | 28 | 27.36 | 49798.4 | 27.55 | 63.5 | 0 | 28 | 8.69 | 12735.1 | 6.31
-62.2 | 0 | 28 | 23.42 | 42630.8 | 23.5 | 64.7 | 0 | 28 | 6.54 | 9573.4 | 5.03
-63.4 | 0 | 28 | 19.73 | 35906.7 | 20.01 | 65.9 | 0 | 28 | 5.73 | 8387 | 4.17
-64.6 | 0 | 28 | 17.33 | 31547.8 | 16.91 | 67.1 | 0 | 28 | 4.58 | 6711.7 | 3.35
-65.8 | 0 | 28 | 14.52 | 26430.3 | 14.52 | 68.3 | 0 | 28 | 3.7 | 5420.7 | 2.71
-67 | 0 | 28 | 12.11 | 22042.9 | 12.13 | 69.5 | 0 | 28 | 3 | 4394.6 | 2.2
-68.2 | 0 | 28 | 10.28 | 18712.7 | 10.4 | 70.7 | 0 | 28 | 2.32 | 3398 | 1.71
-69.4 | 0 | 28 | 8.9 | 16199.8 | 8.87 | 71.9 | 0 | 28 | 1.85 | 2708.1 | 1.36
-70.6 | 0 | 28 | 7.4 | 13467.1 | 7.5 | 73.1 | 0 | 28 | 1.34 | 1956.5 | 0.99
-71.8 | 0 | 28 | 6.5 | 11833.6 | 6.47 | 74.3 | 0 | 28 | 0.95 | 1385.7 | 0.7
-73 | 0 | 28 | 5.54 | 10079.5 | 5.56 | 75.5 | 0 | 28 | 0.71 | 1040.4 | 0.53
-74.2 | 0 | 28 | 4.78 | 8693.7 | 4.8 | 76.7 | 0 | 28 | 0.36 | 521.2 | 0.26
-75.4 | 0 | 28 | 4.16 | 7578.3 | 4.14 | 77.9 | 0 | 28 | 0.02 | 33.1 | 0.02
-76.6 | 0 | 28 | 3.57 | 6503.2 | 3.54 | 79.1 | 0 | 28 | 0.11 | 158.2 | 0.08
-77.8 | 0 | 28 | 3.15 | 5733.3 | 3.11 | 80.3 | 0 | 28 | 0.32 | 475.5 | 0.24
-79 | 0 | 28 | 2.73 | 4972.6 | 2.71 | 81.5 | 0 | 28 | 0.53 | 778 | 0.4
-80.2 | 0 | 28 | 2.41 | 4389.4 | 2.38 | 82.7 | 0 | 28 | 0.66 | 968.7 | 0.49
-81.4 | 0 | 28 | 2.09 | 3808 | 2.09 | 84 | 0 | 28 | 0.77 | 1121.3 | 0.56
-82.6 | 0 | 28 | 1.87 | 3395.9 | 1.88 | 85.2 | 0 | 28 | 1.01 | 1483.7 | 0.76
-83.8 | 0 | 28 | 1.66 | 3020.1 | 1.66 | 86.3 | 0 | 28 | 1.08 | 1578.4 | 0.8
-85 | 0 | 28 | 1.47 | 2681.8 | 1.48 | 87.6 | 0 | 28 | 1.12 | 1642.9 | 0.81
-86.2 | 0 | 28 | 1.32 | 2411.1 | 1.33 | 88.8 | 0 | 28 | 1.25 | 1825.8 | 0.93
-87.4 | 0 | 28 | 1.21 | 2204 | 1.22 | 90 | 0 | 28 | 1.36 | 1993.8 | 1.01
-88.6 | 0 | 28 | 1.12 | 2041.4 | 1.11 | 91.2 | 0 | 28 | 1.4 | 2047.8 | 1.01
-89.8 | 0 | 28 | 1.01 | 1843.4 | 1.02 | 92.4 | 0 | 28 | 1.5 | 2197.4 | 1.13
-91 | 0 | 28 | 0.97 | 1756.7 | 0.94 | 93.6 | 0 | 28 | 1.51 | 2215.4 | 1.1
-92.2 | 0 | 28 | 0.81 | 1479.8 | 0.81 | 94.8 | 0 | 28 | 1.59 | 2323.3 | 1.14
-93.4 | 0 | 28 | 0.8 | 1458.1 | 0.8 | 96 | 0 | 28 | 1.58 | 2309 | 1.17
-94.7 | 0 | 28 | 0.74 | 1347.4 | 0.71 | 97.2 | 0 | 28 | 1.67 | 2445.3 | 1.2
-95.8 | 0 | 28 | 0.66 | 1194 | 0.67 | 98.4 | 0 | 28 | 1.68 | 2454.9 | 1.21
-97.1 | 0 | 28 | 0.6 | 1084.6 | 0.6 | 99.6 | 0 | 28 | 1.7 | 2486.5 | 1.25
-98.3 | 0 | 28 | 0.59 | 1074.2 | 0.59 | 100.8 | 0 | 28 | 1.72 | 2526.8 | 1.32
-99.5 | 0 | 28 | 0.51 | 926.9 | 0.51 | 102 | 0 | 28 | 1.72 | 2516.9 | 1.23
-100.7 | 0 | 28 | 0.48 | 870.6 | 0.49 | 103.2 | 0 | 28 | 1.74 | 2552.4 | 1.3
-101.8 | 0 | 28 | 0.48 | 878.3 | 0.5 | 104.4 | 0 | 28 | 1.78 | 2601.1 | 1.36
-103.1 | 0 | 28 | 0.46 | 838.4 | 0.45 | 105.6 | 0 | 28 | 1.75 | 2556.5 | 1.24
-104.3 | 0 | 28 | 0.41 | 743.5 | 0.42 | 106.8 | 0 | 28 | 1.8 | 2636 | 1.36
-105.5 | 0 | 28 | 0.42 | 763 | 0.41 | 108 | 0 | 28 | 1.79 | 2628.3 | 1.32
-106.7 | 0 | 28 | 0.4 | 722.1 | 0.4 | 109.2 | 0 | 28 | 1.77 | 2592.7 | 1.25
-107.9 | 0 | 28 | 0.36 | 658.5 | 0.37 | 110.4 | 0 | 28 | 1.81 | 2647.1 | 1.38
-109.1 | 0 | 28 | 0.37 | 677.6 | 0.37 | 111.6 | 0 | 28 | 1.84 | 2689.2 | 1.35
-110.3 | 0 | 28 | 0.3 | 548.9 | 0.3 | 112.8 | 0 | 28 | 1.84 | 2696.7 | 1.32
-111.5 | 0 | 28 | 0.29 | 520.6 | 0.29 | 114 | 0 | 28 | 1.82 | 2659.2 | 1.4
-112.7 | 0 | 28 | 0.31 | 560.6 | 0.3 | 115.2 | 0 | 28 | 1.86 | 2719.4 | 1.37
-113.9 | 0 | 28 | 0.33 | 599 | 0.33 | 116.4 | 0 | 28 | 1.79 | 2623.8 | 1.27
-115.1 | 0 | 28 | 0.31 | 564 | 0.31 | 117.6 | 0 | 28 | 1.9 | 2778.8 | 1.46
-116.3 | 0 | 28 | 0.27 | 496.7 | 0.27 | 118.8 | 0 | 28 | 1.88 | 2750.3 | 1.39
-117.5 | 0 | 28 | 0.25 | 461.6 | 0.25 | 120 | 0 | 28 | 1.83 | 2683.6 | 1.29
-118.7 | 0 | 28 | 0.26 | 479.4 | 0.27 | 121.3 | 0 | 28 | 1.85 | 2713.4 | 1.43
-119.9 | 0 | 28 | 0.28 | 508.2 | 0.27 | 122.4 | 0 | 28 | 1.85 | 2703.6 | 1.36
-121.1 | 0 | 28 | 0.21 | 385.8 | 0.21 | 123.6 | 0 | 28 | 1.84 | 2693.1 | 1.3
-122.3 | 0 | 28 | 0.23 | 426.2 | 0.24 | 124.9 | 0 | 28 | 1.97 | 2887.6 | 1.53
-123.5 | 0 | 28 | 0.2 | 372.7 | 0.21 | 126 | 0 | 28 | 1.89 | 2773.6 | 1.38
-124.7 | 0 | 28 | 0.19 | 351.8 | 0.19 | 127.2 | 0 | 28 | 1.89 | 2762.5 | 1.33
-125.9 | 0 | 28 | 0.17 | 307.5 | 0.17 | 128.4 | 0 | 28 | 1.9 | 2779.7 | 1.44
-127.1 | 0 | 28 | 0.13 | 242.3 | 0.13 | 129.7 | 0 | 28 | 1.91 | 2802.2 | 1.39
-128.3 | 0 | 28 | 0.2 | 356.7 | 0.19 | 130.8 | 0 | 28 | 1.93 | 2833.1 | 1.42
-129.5 | 0 | 28 | 0.18 | 332.4 | 0.18 | 132.1 | 0 | 28 | 1.92 | 2808.4 | 1.46
-130.7 | 0 | 28 | 0.13 | 244.5 | 0.13 | 133.3 | 0 | 28 | 1.96 | 2877.7 | 1.43
-131.9 | 0 | 28 | 0.2 | 369.7 | 0.2 | 134.5 | 0 | 28 | 1.96 | 2864.8 | 1.38
-133.1 | 0 | 28 | 0.2 | 360.3 | 0.2 | 135.7 | 0 | 28 | 1.96 | 2865 | 1.49
-134.3 | 0 | 28 | 0.12 | 224.8 | 0.12 | 136.9 | 0 | 28 | 1.94 | 2849 | 1.42
-135.5 | 0 | 28 | 0.16 | 289.1 | 0.16 | 138.1 | 0 | 28 | 1.86 | 2721.4 | 1.31
-136.7 | 0 | 28 | 0.14 | 258.6 | 0.14 | 139.3 | 0 | 28 | 1.96 | 2877.7 | 1.49
-137.9 | 0 | 28 | 0.15 | 279.5 | 0.15 | 140.5 | 0 | 28 | 1.96 | 2867.1 | 1.5
-139.1 | 0 | 28 | 0.17 | 308.1 | 0.17 | 141.7 | 0 | 28 | 1.91 | 2802.6 | 1.34
-140.3 | 0 | 28 | 0.13 | 236.1 | 0.13 | 142.9 | 0 | 28 | 2.01 | 2942.1 | 1.55
-141.5 | 0 | 28 | 0.15 | 268.4 | 0.15 | 144.1 | 0 | 28 | 1.91 | 2799.4 | 1.41
-142.7 | 0 | 28 | 0.2 | 364.8 | 0.2 | 145.3 | 0 | 28 | 1.92 | 2817.3 | 1.36
-143.9 | 0 | 28 | 0.15 | 278.6 | 0.15 | 146.5 | 0 | 28 | 2.01 | 2937.9 | 1.56
-145.1 | 0 | 28 | 0.18 | 330.7 | 0.18 | 147.7 | 0 | 28 | 1.95 | 2861.5 | 1.44
-146.3 | 0 | 28 | 0.19 | 352.7 | 0.19 | 148.9 | 0 | 28 | 2.02 | 2958.4 | 1.44
-147.5 | 0 | 28 | 0.16 | 294.7 | 0.16 | 150.1 | 0 | 28 | 1.97 | 2880.6 | 1.46
-148.7 | 0 | 28 | 0.16 | 290.9 | 0.16 | 151.4 | 0 | 28 | 1.99 | 2922.2 | 1.49
-149.9 | 0 | 28 | 0.18 | 319.5 | 0.17
-151.1 | 0 | 28 | 0.15 | 274.5 | 0.15
-152.3 | 0 | 28 | 0.24 | 429.2 | 0.23
-153.9 | 0 | 28 | 0.13 | 228.3 | 0.12
Measurement time: | 07/09/2022 17:39:43
Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV
Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | CC 13 - F
Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm
SSD: | 1100 mm
Field size: | 100 x 100 mm
Gantry angle: | 0
Measurement medium: | Water
Scan type: | Inline
Scan mode: | Continuous
Wedge: | none
Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio
-156.4 | 0 | 28 | 0.29 | 531.3 | 0.49
-155.7 | 0 | 28 | 0.09 | 165.4 | 0.1
-154.3 | 0 | 28 | 0.15 | 279 | 0.17
-153.2 | 0 | 28 | 0.11 | 199.9 | 0.12
-151.9 | 0 | 28 | 0.14 | 253.6 | 0.15
-150.7 | 0 | 28 | 0.12 | 225 | 0.14
-149.5 | 0 | 28 | 0.1 | 189.5 | 0.12
-148.3 | 0 | 28 | 0.08 | 147 | 0.09
-147.1 | 0 | 28 | 0.1 | 189.2 | 0.12
-145.9 | 0 | 28 | 0.16 | 283.1 | 0.17
-144.7 | 0 | 28 | 0.15 | 277 | 0.17
-143.5 | 0 | 28 | 0.16 | 294.9 | 0.18
-142.3 | 0 | 28 | 0.18 | 321.1 | 0.2
-141.1 | 0 | 28 | 0.21 | 378.3 | 0.23
-139.9 | 0 | 28 | 0.14 | 246.6 | 0.15
-138.7 | 0 | 28 | 0.2 | 368.1 | 0.22
-137.5 | 0 | 28 | 0.2 | 356.1 | 0.21
-136.3 | 0 | 28 | 0.22 | 396.9 | 0.25
-135.1 | 0 | 28 | 0.24 | 431.2 | 0.26
-133.9 | 0 | 28 | 0.22 | 396.2 | 0.23
-132.7 | 0 | 28 | 0.22 | 403.5 | 0.25
-131.5 | 0 | 28 | 0.24 | 433.7 | 0.26
-130.3 | 0 | 28 | 0.27 | 489 | 0.29
-129.1 | 0 | 28 | 0.19 | 344.4 | 0.22
-127.9 | 0 | 28 | 0.24 | 443.8 | 0.27
-126.7 | 0 | 28 | 0.28 | 511.9 | 0.3
-125.5 | 0 | 28 | 0.21 | 388.9 | 0.25
-124.3 | 0 | 28 | 0.26 | 472.9 | 0.29
-123.1 | 0 | 28 | 0.27 | 499.6 | 0.29
-121.9 | 0 | 28 | 0.27 | 484.4 | 0.31
-120.7 | 0 | 28 | 0.31 | 560.6 | 0.34
-119.5 | 0 | 28 | 0.35 | 641.9 | 0.37
-118.3 | 0 | 28 | 0.31 | 562.2 | 0.36
-117.1 | 0 | 28 | 0.29 | 534 | 0.32
-115.9 | 0 | 28 | 0.35 | 632.8 | 0.36
-114.7 | 0 | 28 | 0.35 | 632.9 | 0.4
-113.5 | 0 | 28 | 0.37 | 666.6 | 0.4
-112.3 | 0 | 28 | 0.43 | 789.2 | 0.46
-111.1 | 0 | 28 | 0.36 | 647.8 | 0.41
-109.9 | 0 | 28 | 0.41 | 755 | 0.46
-108.7 | 0 | 28 | 0.42 | 772.4 | 0.46
-107.5 | 0 | 28 | 0.46 | 831.6 | 0.51
-106.3 | 0 | 28 | 0.45 | 827.5 | 0.52
-105.1 | 0 | 28 | 0.56 | 1017.7 | 0.6
-103.9 | 0 | 28 | 0.55 | 998.7 | 0.63
-102.7 | 0 | 28 | 0.6 | 1098.3 | 0.66
-101.5 | 0 | 28 | 0.62 | 1134.5 | 0.66
-100.3 | 0 | 28 | 0.61 | 1116.5 | 0.71
-99.1 | 0 | 28 | 0.77 | 1394.3 | 0.84
-97.9 | 0 | 28 | 0.87 | 1590.4 | 0.92
-96.7 | 0 | 28 | 0.81 | 1475.1 | 0.93
-95.5 | 0 | 28 | 0.92 | 1677.4 | 1.02
-94.3 | 0 | 28 | 1.06 | 1926.2 | 1.12
-93.1 | 0 | 28 | 1.05 | 1904.6 | 1.22
-91.8 | 0 | 28 | 1.23 | 2230 | 1.32
-90.6 | 0 | 28 | 1.3 | 2372.3 | 1.46
-89.4 | 0 | 28 | 1.42 | 2580.3 | 1.59
-88.2 | 0 | 28 | 1.51 | 2741.2 | 1.75
-87 | 0 | 28 | 1.86 | 3377 | 1.98
-85.8 | 0 | 28 | 1.98 | 3605.4 | 2.21
-84.6 | 0 | 28 | 2.15 | 3910.2 | 2.48
-83.4 | 0 | 28 | 2.69 | 4903.5 | 2.87
-82.2 | 0 | 28 | 2.84 | 5172.1 | 3.22
-81 | 0 | 28 | 3.43 | 6248.7 | 3.75
-79.8 | 0 | 28 | 3.95 | 7190.4 | 4.27
-78.6 | 0 | 28 | 4.36 | 7932.7 | 4.94
-77.4 | 0 | 28 | 5.11 | 9305.2 | 5.69
-76.2 | 0 | 28 | 6.08 | 11069 | 6.6
-75 | 0 | 28 | 6.71 | 12223.2 | 7.71
-73.8 | 0 | 28 | 7.97 | 14516.7 | 8.89
-72.6 | 0 | 28 | 9.68 | 17620.6 | 10.44
-71.4 | 0 | 28 | 10.63 | 19358.9 | 12.17
-70.2 | 0 | 28 | 12.65 | 23018.4 | 14.05
-69 | 0 | 28 | 15.11 | 27497.3 | 16.47
-67.8 | 0 | 28 | 17.11 | 31152.1 | 19.13
-66.6 | 0 | 28 | 20.01 | 36433.6 | 22.18
-65.4 | 0 | 28 | 23.72 | 43172.3 | 25.58
-64.2 | 0 | 28 | 26.3 | 47876.2 | 29.54
-63 | 0 | 28 | 30.79 | 56053.9 | 33.67
-61.8 | 0 | 28 | 35.56 | 64739.9 | 38.34
-60.6 | 0 | 28 | 38.03 | 69219.2 | 43.1
-59.4 | 0 | 28 | 43.56 | 79294.1 | 48.02
-58.2 | 0 | 28 | 48.9 | 89013.7 | 53.29
-57 | 0 | 28 | 51.25 | 93295.7 | 58.1
-55.8 | 0 | 28 | 56.45 | 102764.5 | 62.98
-54.6 | 0 | 28 | 61.69 | 112303.7 | 67.71
-53.4 | 0 | 28 | 63.69 | 115945.2 | 71.88
-52.2 | 0 | 28 | 68.17 | 124089.3 | 75.44
-51 | 0 | 28 | 72.5 | 131974 | 78.99
-49.8 | 0 | 28 | 73.7 | 134168.4 | 82.06
-48.6 | 0 | 28 | 76.24 | 138793.1 | 84.22
-47.4 | 0 | 28 | 76.94 | 140054.2 | 86.71
-46.1 | 0 | 28 | 79.57 | 144845.5 | 88.5
-44.9 | 0 | 28 | 82.32 | 149859.4 | 90.27
-43.8 | 0 | 28 | 80.35 | 146264.9 | 91.51
-42.6 | 0 | 28 | 82.84 | 150792.8 | 92.7
-41.3 | 0 | 28 | 84.61 | 154012.9 | 93.56
-40.1 | 0 | 28 | 86.36 | 157197.4 | 94.34
-38.9 | 0 | 28 | 84.03 | 152969.6 | 95.13
-37.7 | 0 | 28 | 85 | 154727.7 | 95.4
-36.5 | 0 | 28 | 89.78 | 163429.3 | 96.04
-35.3 | 0 | 28 | 86.1 | 156739.7 | 96.4
-34.1 | 0 | 28 | 86.61 | 157663.6 | 96.66
-32.9 | 0 | 28 | 87.59 | 159449.5 | 97.2
-31.7 | 0 | 28 | 87.5 | 159279.9 | 97.54
-30.5 | 0 | 28 | 87.26 | 158837.3 | 97.71
-29.3 | 0 | 28 | 89.66 | 163207.4 | 98.01
-28.1 | 0 | 28 | 87.91 | 160033.6 | 98.19
-26.9 | 0 | 28 | 89.52 | 162963.7 | 98.39
-25.7 | 0 | 28 | 90.08 | 163972.2 | 98.49
-24.5 | 0 | 28 | 88.73 | 161515.2 | 98.64
-23.3 | 0 | 28 | 88.99 | 161991.9 | 98.35
-22.1 | 0 | 28 | 89.63 | 163157.2 | 98.76
-20.9 | 0 | 28 | 88.11 | 160399.7 | 99.09
-19.7 | 0 | 28 | 88.39 | 160909.6 | 98.93
-18.5 | 0 | 28 | 89.4 | 162739.4 | 99.12
-17.3 | 0 | 28 | 88.69 | 161441.4 | 99.29
-16.1 | 0 | 28 | 90.15 | 164104.4 | 99.09
-14.9 | 0 | 28 | 90.6 | 164933 | 99.4
-13.7 | 0 | 28 | 89.57 | 163058.1 | 99.58
-12.5 | 0 | 28 | 90.47 | 164695.2 | 99.32
-11.3 | 0 | 28 | 91.06 | 165766.1 | 99.41
-10.1 | 0 | 28 | 90.78 | 165251.2 | 99.58
-8.9 | 0 | 28 | 86.37 | 157229.5 | 99.4
-7.7 | 0 | 28 | 89.84 | 163548.3 | 99.54
-6.5 | 0 | 28 | 89.21 | 162399.2 | 99.72
-5.3 | 0 | 28 | 88.05 | 160276.9 | 99.54
-4.1 | 0 | 28 | 90.5 | 164751 | 99.65
-2.9 | 0 | 28 | 88.91 | 161854.7 | 99.75
-1.7 | 0 | 28 | 90.32 | 164422.8 | 99.63
-0.5 | 0 | 28 | 91.72 | 166956.3 | 99.76
0.7 | 0 | 28 | 89.8 | 163465.8 | 99.72
1.9 | 0 | 28 | 90.87 | 165412.8 | 99.77
3.1 | 0 | 28 | 91.78 | 167065.1 | 99.91
4.3 | 0 | 28 | 89.62 | 163133 | 99.73
5.5 | 0 | 28 | 92 | 167474.1 | 99.61
6.7 | 0 | 28 | 89.49 | 162894.8 | 99.87
8 | 0 | 28 | 90.15 | 164103.6 | 100
9.2 | 0 | 28 | 90.54 | 164806.4 | 99.58
10.4 | 0 | 28 | 86.13 | 156779.5 | 99.81
11.6 | 0 | 28 | 90.83 | 165352 | 99.75
12.7 | 0 | 28 | 87.08 | 158520.6 | 99.7
14 | 0 | 28 | 89.57 | 163051.4 | 99.62
15.2 | 0 | 28 | 91.93 | 167336.4 | 99.86
16.4 | 0 | 28 | 89.67 | 163234.1 | 99.58
17.6 | 0 | 28 | 87.38 | 159068.1 | 99.52
18.8 | 0 | 28 | 88.71 | 161476.6 | 99.49
20 | 0 | 28 | 90.07 | 163966 | 99.31
21.2 | 0 | 28 | 90.35 | 164470.5 | 99.32
22.4 | 0 | 28 | 90.6 | 164921.9 | 99.25
23.6 | 0 | 28 | 85.68 | 155962.4 | 99.06
24.8 | 0 | 28 | 89.67 | 163231.8 | 98.96
26 | 0 | 28 | 89.37 | 162687 | 98.85
27.2 | 0 | 28 | 89.45 | 162822.2 | 98.88
28.4 | 0 | 28 | 88.77 | 161587.1 | 98.32
29.6 | 0 | 28 | 88.15 | 160473.3 | 98.19
30.8 | 0 | 28 | 87.88 | 159971.6 | 98.06
32 | 0 | 28 | 88.11 | 160393.9 | 97.62
33.2 | 0 | 28 | 87.12 | 158598.3 | 97.34
34.4 | 0 | 28 | 86.83 | 158056.7 | 96.55
35.6 | 0 | 28 | 87.4 | 159101.4 | 96.17
36.8 | 0 | 28 | 86.24 | 156979.6 | 95.52
38 | 0 | 28 | 85.96 | 156476.9 | 94.63
39.2 | 0 | 28 | 85.68 | 155971 | 93.94
40.4 | 0 | 28 | 84.11 | 153107.7 | 92.94
41.6 | 0 | 28 | 83.51 | 152025.7 | 91.76
42.8 | 0 | 28 | 82.15 | 149549.8 | 90.34
44 | 0 | 28 | 80.8 | 147084.6 | 88.61
45.2 | 0 | 28 | 79.27 | 144299.8 | 86.57
46.4 | 0 | 28 | 72.75 | 132426.2 | 84.3
47.6 | 0 | 28 | 74.01 | 134727.2 | 81.63
48.8 | 0 | 28 | 70.97 | 129195.1 | 78.41
50 | 0 | 28 | 67.46 | 122804.9 | 75.07
51.2 | 0 | 28 | 63.94 | 116394.9 | 71.09
52.4 | 0 | 28 | 59.83 | 108911.1 | 66.67
53.6 | 0 | 28 | 55.23 | 100540.7 | 62.34
54.8 | 0 | 28 | 53.43 | 97259.7 | 57.4
56 | 0 | 28 | 46.13 | 83966.8 | 52.12
57.2 | 0 | 28 | 42.04 | 76530.3 | 47.32
58.5 | 0 | 28 | 39.07 | 71120.8 | 42.24
59.7 | 0 | 28 | 31.88 | 58038.7 | 37.15
60.9 | 0 | 28 | 30.09 | 54780.3 | 32.9
62.1 | 0 | 28 | 25.92 | 47180 | 28.6
63.3 | 0 | 28 | 22.49 | 40937.7 | 24.69
64.5 | 0 | 28 | 19.5 | 35496 | 21.37
65.7 | 0 | 28 | 16.64 | 30299 | 18.33
66.9 | 0 | 28 | 14.43 | 26268.9 | 15.79
68.1 | 0 | 28 | 12.25 | 22304.2 | 13.51
69.3 | 0 | 28 | 10.62 | 19323.9 | 11.58
70.5 | 0 | 28 | 9.11 | 16582.7 | 10.02
71.7 | 0 | 28 | 7.78 | 14169.8 | 8.59
72.9 | 0 | 28 | 6.77 | 12319.3 | 7.45
74.1 | 0 | 28 | 5.76 | 10492.7 | 6.36
75.3 | 0 | 28 | 4.99 | 9082.4 | 5.54
76.5 | 0 | 28 | 4.28 | 7791 | 4.8
77.7 | 0 | 28 | 3.72 | 6780.1 | 4.19
78.9 | 0 | 28 | 3.28 | 5967.8 | 3.69
80.1 | 0 | 28 | 2.96 | 5386 | 3.23
81.3 | 0 | 28 | 2.52 | 4584.5 | 2.84
82.5 | 0 | 28 | 2.18 | 3960.6 | 2.49
83.7 | 0 | 28 | 2.02 | 3684.9 | 2.22
84.9 | 0 | 28 | 1.76 | 3198.5 | 1.98
86.1 | 0 | 28 | 1.53 | 2783 | 1.76
87.3 | 0 | 28 | 1.47 | 2676.3 | 1.61
88.5 | 0 | 28 | 1.28 | 2330.7 | 1.45
89.7 | 0 | 28 | 1.16 | 2109.3 | 1.28
90.9 | 0 | 28 | 1.08 | 1961.6 | 1.19
92.1 | 0 | 28 | 0.99 | 1804.3 | 1.1
93.3 | 0 | 28 | 0.85 | 1548.5 | 0.96
94.5 | 0 | 28 | 0.85 | 1543.1 | 0.93
95.7 | 0 | 28 | 0.73 | 1326.4 | 0.82
96.9 | 0 | 28 | 0.66 | 1197.1 | 0.74
98.1 | 0 | 28 | 0.6 | 1096.9 | 0.67
99.3 | 0 | 28 | 0.6 | 1084 | 0.65
100.5 | 0 | 28 | 0.54 | 977 | 0.59
101.8 | 0 | 28 | 0.48 | 879.7 | 0.53
103 | 0 | 28 | 0.46 | 831.6 | 0.51
104.2 | 0 | 28 | 0.44 | 796.9 | 0.49
105.4 | 0 | 28 | 0.38 | 686.9 | 0.42
106.6 | 0 | 28 | 0.35 | 631.8 | 0.39
107.7 | 0 | 28 | 0.36 | 661.8 | 0.39
109 | 0 | 28 | 0.33 | 601.5 | 0.36
110.2 | 0 | 28 | 0.36 | 656.4 | 0.4
111.4 | 0 | 28 | 0.34 | 625.2 | 0.38
112.6 | 0 | 28 | 0.29 | 519.7 | 0.31
113.8 | 0 | 28 | 0.27 | 494.1 | 0.3
115 | 0 | 28 | 0.22 | 392.8 | 0.24
116.2 | 0 | 28 | 0.18 | 326.1 | 0.2
117.4 | 0 | 28 | 0.22 | 403.3 | 0.25
118.6 | 0 | 28 | 0.2 | 361.8 | 0.22
119.8 | 0 | 28 | 0.18 | 323 | 0.2
121 | 0 | 28 | 0.18 | 329.4 | 0.2
122.2 | 0 | 28 | 0.15 | 272.8 | 0.16
123.4 | 0 | 28 | 0.15 | 275.3 | 0.17
124.6 | 0 | 28 | 0.19 | 352.7 | 0.21
125.8 | 0 | 28 | 0.19 | 348.6 | 0.21
127 | 0 | 28 | 0.16 | 288.3 | 0.17
128.2 | 0 | 28 | 0.2 | 355.8 | 0.22
129.4 | 0 | 28 | 0.15 | 265.9 | 0.16
130.6 | 0 | 28 | 0.14 | 262.6 | 0.16
131.8 | 0 | 28 | 0.12 | 225.3 | 0.14
133 | 0 | 28 | 0.09 | 159.6 | 0.1
134.2 | 0 | 28 | 0.11 | 191.8 | 0.12
135.4 | 0 | 28 | 0.14 | 262.4 | 0.16
136.6 | 0 | 28 | 0.14 | 261.6 | 0.16
137.8 | 0 | 28 | 0.1 | 177.9 | 0.11
139 | 0 | 28 | 0.09 | 160.2 | 0.1
140.2 | 0 | 28 | 0.05 | 94.7 | 0.06
141.4 | 0 | 28 | 0.02 | 36.5 | 0.02
142.6 | 0 | 28 | 0.04 | 77.6 | 0.05
143.8 | 0 | 28 | 0.06 | 100.3 | 0.06
145 | 0 | 28 | 0 | 0.6 | 0
146.2 | 0 | 28 | 0.03 | 55.3 | 0.03
147.4 | 0 | 28 | 0.02 | 43 | 0.03
148.6 | 0 | 28 | 0.05 | 90.9 | 0.06
149.8 | 0 | 28 | 0.03 | 50.5 | 0.03
151 | 0 | 28 | 0.03 | 60 | 0.04
152.2 | 0 | 28 | 0.02 | 36.1 | 0.02
153.4 | 0 | 28 | 0.03 | 56.2 | 0.03
154.6 | 0 | 28 | 0.02 | 39.4 | 0.02
156.4 | 0 | 28 | 0.02 | 27.6 | 0.02