## Feuil1
Measurement time: | 07/09/2022 18:27:16 |  |  |  |  | Measurement time: | 02/09/2022 20:44:47 |  |  |  |  | Measurement time: | 29/09/2022 19:19:17
Radiation device: | Clinac3 Accelerator |  |  |  |  | Radiation device: | Clinac3 Accelerator |  |  |  |  | Radiation device: | Clinac3 Accelerator
Energy: | 12 MeV |  |  |  |  | Energy: | 12 MeV |  |  |  |  | Energy: | 12 MeV
Controller: | CCU |  |  |  |  | Controller: | CCU |  |  |  |  | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41) |  |  |  |  | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | PTW31014 perpendiculaire |  |  |  |  | Field detector: | CC 13 - F |  |  |  |  | Field detector: | ROOS
Reference detector: | CC 13 - R pas offset energie |  |  |  |  | Reference detector: | CC 13 - R pas offset energie |  |  |  |  | Reference detector: | CC 13 - R pas offset energie
SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm |  |  |  |  | SAD: | 1000 mm
SSD: | 1000 mm |  |  |  |  | SSD: | 1000 mm |  |  |  |  | SSD: | 1000 mm
Field size: | 100 x 100 mm |  |  |  |  | Field size: | 100 x 100 mm |  |  |  |  | Field size: | 100 x 100 mm
Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0 |  |  |  |  | Gantry angle: | 0
Measurement medium: | Water |  |  |  |  | Measurement medium: | Water |  |  |  |  | Measurement medium: | Water
Scan type: | Inline |  |  |  |  | Scan type: | Inline |  |  |  |  | Scan type: | Inline
Scan mode: | Continuous |  |  |  |  | Scan mode: | Continuous |  |  |  |  | Scan mode: | Continuous
Wedge: | none |  |  |  |  | Wedge: | none |  |  |  |  | Wedge: | none
Points [mm]: | [5 empty] | Points [mm]: | [5 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Inline | Crossline | Depth | Normalized field | Current field | Ratio
-151.4 | 0 | 28 | 4.88 | 1002.3 | 2.99 | -151.4 | 0 | 28 | 1.76 | 2581.1 | 1.57 | -151.4 | 0 | 28 | 0.3 | 1705.9 | 0.31
-150.7 | 0 | 28 | 0.59 | 120.8 | 0.51 | -150.6 | 0 | 28 | 1.72 | 2526.5 | 1.28 | -150.4 | 0 | 28 | 0.3 | 1719.2 | 0.28
-149.3 | 0 | 28 | 0.48 | 98.7 | 0.41 | -149.3 | 0 | 28 | 1.71 | 2502 | 1.26 | -149 | 0 | 28 | 0.31 | 1737.6 | 0.27
-148.1 | 0 | 28 | 0.45 | 92.1 | 0.39 | -148.1 | 0 | 28 | 1.73 | 2530.3 | 1.22 | -147.8 | 0 | 28 | 0.29 | 1672.7 | 0.27
-146.9 | 0 | 28 | 0.73 | 150.1 | 0.64 | -146.9 | 0 | 28 | 1.73 | 2530.4 | 1.31 | -146.7 | 0 | 28 | 0.31 | 1789 | 0.28
-145.7 | 0 | 28 | 0.61 | 124.9 | 0.53 | -145.7 | 0 | 28 | 1.71 | 2505.3 | 1.25 | -145.4 | 0 | 28 | 0.31 | 1784.4 | 0.28
-144.5 | 0 | 28 | 1.2 | 246.3 | 1.05 | -144.5 | 0 | 28 | 1.75 | 2566 | 1.23 | -144.2 | 0 | 28 | 0.31 | 1754.7 | 0.29
-143.3 | 0 | 28 | 1.04 | 214.5 | 0.91 | -143.3 | 0 | 28 | 1.83 | 2679.8 | 1.4 | -143 | 0 | 28 | 0.35 | 1999.2 | 0.3
-142.1 | 0 | 28 | 0.68 | 140.1 | 0.59 | -142.1 | 0 | 28 | 1.76 | 2571.7 | 1.34 | -141.8 | 0 | 28 | 0.31 | 1770.8 | 0.29
-140.9 | 0 | 28 | 0.89 | 183 | 0.77 | -140.9 | 0 | 28 | 1.7 | 2495.2 | 1.19 | -140.6 | 0 | 28 | 0.34 | 1910.2 | 0.29
-139.7 | 0 | 28 | 0.84 | 173 | 0.73 | -139.7 | 0 | 28 | 1.77 | 2588.1 | 1.34 | -139.4 | 0 | 28 | 0.33 | 1853.9 | 0.29
-138.5 | 0 | 28 | 0.97 | 199.5 | 0.83 | -138.5 | 0 | 28 | 1.77 | 2589.5 | 1.36 | -138.2 | 0 | 28 | 0.29 | 1665.8 | 0.28
-137.3 | 0 | 28 | 0.82 | 168.4 | 0.71 | -137.3 | 0 | 28 | 1.67 | 2450.5 | 1.16 | -137 | 0 | 28 | 0.3 | 1724.4 | 0.26
-136.1 | 0 | 28 | 0.86 | 176.4 | 0.74 | -136.1 | 0 | 28 | 1.72 | 2524.8 | 1.32 | -135.8 | 0 | 28 | 0.32 | 1820.3 | 0.3
-134.9 | 0 | 28 | 0.91 | 186.6 | 0.78 | -134.9 | 0 | 28 | 1.77 | 2591.1 | 1.35 | -134.6 | 0 | 28 | 0.33 | 1906.2 | 0.3
-133.7 | 0 | 28 | 0.46 | 94.3 | 0.39 | -133.7 | 0 | 28 | 1.74 | 2547.9 | 1.22 | -133.4 | 0 | 28 | 0.31 | 1739.5 | 0.28
-132.5 | 0 | 28 | 0.88 | 180.1 | 0.75 | -132.5 | 0 | 28 | 1.74 | 2549.7 | 1.31 | -132.2 | 0 | 28 | 0.31 | 1763.4 | 0.29
-131.3 | 0 | 28 | 1.08 | 221.2 | 0.97 | -131.3 | 0 | 28 | 1.72 | 2519.9 | 1.26 | -131 | 0 | 28 | 0.32 | 1841.3 | 0.28
-130.1 | 0 | 28 | 0.82 | 168.3 | 0.71 | -130.1 | 0 | 28 | 1.78 | 2606.6 | 1.25 | -129.8 | 0 | 28 | 0.31 | 1770.8 | 0.29
-128.9 | 0 | 28 | 1.1 | 225.9 | 0.94 | -128.9 | 0 | 28 | 1.77 | 2589.6 | 1.35 | -128.6 | 0 | 28 | 0.31 | 1786.3 | 0.28
-127.7 | 0 | 28 | 1.06 | 217.3 | 0.92 | -127.7 | 0 | 28 | 1.79 | 2627.4 | 1.31 | -127.4 | 0 | 28 | 0.33 | 1860.2 | 0.29
-126.5 | 0 | 28 | 1.11 | 229.1 | 0.97 | -126.5 | 0 | 28 | 1.73 | 2535.8 | 1.22 | -126.2 | 0 | 28 | 0.31 | 1750.2 | 0.29
-125.3 | 0 | 28 | 0.98 | 201.5 | 0.86 | -125.3 | 0 | 28 | 1.74 | 2544.3 | 1.34 | -125 | 0 | 28 | 0.33 | 1882.2 | 0.28
-124.1 | 0 | 28 | 1.05 | 214.8 | 0.91 | -124.1 | 0 | 28 | 1.69 | 2482.2 | 1.25 | -123.8 | 0 | 28 | 0.3 | 1714 | 0.28
-122.9 | 0 | 28 | 0.9 | 184.1 | 0.77 | -122.9 | 0 | 28 | 1.73 | 2533.2 | 1.23 | -122.6 | 0 | 28 | 0.33 | 1853.8 | 0.3
-121.7 | 0 | 28 | 0.93 | 191.8 | 0.82 | -121.7 | 0 | 28 | 1.74 | 2551.5 | 1.35 | -121.4 | 0 | 28 | 0.32 | 1802.6 | 0.29
-120.5 | 0 | 28 | 0.85 | 173.9 | 0.73 | -120.5 | 0 | 28 | 1.73 | 2538.4 | 1.24 | -120.2 | 0 | 28 | 0.32 | 1830.7 | 0.29
-119.3 | 0 | 28 | 1.33 | 272.6 | 1.13 | -119.3 | 0 | 28 | 1.73 | 2540.7 | 1.3 | -119 | 0 | 28 | 0.34 | 1913.4 | 0.29
-118.1 | 0 | 28 | 1.1 | 225.8 | 0.95 | -118.1 | 0 | 28 | 1.71 | 2504.9 | 1.26 | -117.7 | 0 | 28 | 0.33 | 1893.1 | 0.32
-116.9 | 0 | 28 | 0.78 | 159.8 | 0.66 | -116.9 | 0 | 28 | 1.71 | 2509.6 | 1.32 | -116.6 | 0 | 28 | 0.32 | 1848.2 | 0.3
-115.7 | 0 | 28 | 0.72 | 148.6 | 0.63 | -115.7 | 0 | 28 | 1.69 | 2473.8 | 1.21 | -115.4 | 0 | 28 | 0.32 | 1837.2 | 0.29
-114.5 | 0 | 28 | 1.31 | 268.6 | 1.12 | -114.5 | 0 | 28 | 1.8 | 2631.3 | 1.35 | -114.2 | 0 | 28 | 0.35 | 1992.6 | 0.32
-113.3 | 0 | 28 | 0.83 | 170.3 | 0.75 | -113.3 | 0 | 28 | 1.68 | 2465.4 | 1.28 | -113 | 0 | 28 | 0.36 | 2054.5 | 0.31
-112 | 0 | 28 | 1.09 | 224.8 | 0.94 | -112.1 | 0 | 28 | 1.67 | 2441 | 1.2 | -111.7 | 0 | 28 | 0.34 | 1949.4 | 0.32
-110.9 | 0 | 28 | 1.36 | 280.5 | 1.17 | -110.9 | 0 | 28 | 1.68 | 2456.4 | 1.25 | -110.6 | 0 | 28 | 0.36 | 2076.6 | 0.33
-109.6 | 0 | 28 | 1.05 | 216.3 | 0.92 | -109.6 | 0 | 28 | 1.71 | 2504.8 | 1.3 | -109.3 | 0 | 28 | 0.37 | 2086.1 | 0.34
-108.5 | 0 | 28 | 0.9 | 185.2 | 0.78 | -108.4 | 0 | 28 | 1.62 | 2373.4 | 1.18 | -108.2 | 0 | 28 | 0.37 | 2132.1 | 0.34
-107.3 | 0 | 28 | 1.09 | 224.8 | 0.95 | -107.2 | 0 | 28 | 1.6 | 2349.4 | 1.2 | -107 | 0 | 28 | 0.43 | 2430.2 | 0.37
-106 | 0 | 28 | 1.25 | 257.5 | 1.11 | -106 | 0 | 28 | 1.63 | 2381.6 | 1.23 | -105.7 | 0 | 28 | 0.4 | 2278.2 | 0.38
-104.9 | 0 | 28 | 0.91 | 186.4 | 0.79 | -104.9 | 0 | 28 | 1.58 | 2307.3 | 1.14 | -104.6 | 0 | 28 | 0.44 | 2498.8 | 0.4
-103.6 | 0 | 28 | 0.88 | 180.4 | 0.77 | -103.6 | 0 | 28 | 1.57 | 2299.4 | 1.17 | -103.3 | 0 | 28 | 0.43 | 2455.4 | 0.4
-102.4 | 0 | 28 | 1.11 | 228.1 | 0.98 | -102.4 | 0 | 28 | 1.6 | 2341.7 | 1.17 | -102.1 | 0 | 28 | 0.47 | 2671.2 | 0.42
-101.2 | 0 | 28 | 1.21 | 248 | 1.05 | -101.2 | 0 | 28 | 1.54 | 2249.7 | 1.1 | -100.9 | 0 | 28 | 0.49 | 2766.6 | 0.43
-100 | 0 | 28 | 1.13 | 232.9 | 1 | -100 | 0 | 28 | 1.59 | 2329.3 | 1.18 | -99.7 | 0 | 28 | 0.49 | 2815.8 | 0.46
-98.8 | 0 | 28 | 0.96 | 196.5 | 0.84 | -98.8 | 0 | 28 | 1.48 | 2166 | 1.07 | -98.5 | 0 | 28 | 0.53 | 3031.4 | 0.5
-97.6 | 0 | 28 | 1.36 | 278.9 | 1.2 | -97.6 | 0 | 28 | 1.42 | 2082.6 | 1.01 | -97.3 | 0 | 28 | 0.58 | 3310.9 | 0.53
-96.4 | 0 | 28 | 0.92 | 188.9 | 0.8 | -96.4 | 0 | 28 | 1.42 | 2075.7 | 1.06 | -96.1 | 0 | 28 | 0.63 | 3607.3 | 0.57
-95.2 | 0 | 28 | 1.08 | 221.3 | 0.94 | -95.2 | 0 | 28 | 1.42 | 2079.8 | 1.04 | -94.9 | 0 | 28 | 0.71 | 4039.1 | 0.61
-94 | 0 | 28 | 1.43 | 294.6 | 1.25 | -94 | 0 | 28 | 1.42 | 2079.3 | 1.06 | -93.7 | 0 | 28 | 0.77 | 4387.6 | 0.68
-92.8 | 0 | 28 | 1.19 | 244.3 | 1.03 | -92.8 | 0 | 28 | 1.3 | 1906.7 | 0.97 | -92.5 | 0 | 28 | 0.82 | 4666.2 | 0.76
-91.6 | 0 | 28 | 1.13 | 233.1 | 0.97 | -91.6 | 0 | 28 | 1.22 | 1793.7 | 0.89 | -91.3 | 0 | 28 | 0.88 | 4984.4 | 0.81
-90.4 | 0 | 28 | 1.49 | 305.9 | 1.28 | -90.4 | 0 | 28 | 1.13 | 1651.3 | 0.81 | -90.1 | 0 | 28 | 1.01 | 5759.6 | 0.91
-89.2 | 0 | 28 | 1.45 | 297.4 | 1.26 | -89.2 | 0 | 28 | 1.07 | 1565.8 | 0.79 | -88.9 | 0 | 28 | 1.17 | 6679 | 1.01
-88 | 0 | 28 | 1.88 | 385.7 | 1.62 | -88 | 0 | 28 | 0.9 | 1319.6 | 0.65 | -87.7 | 0 | 28 | 1.19 | 6761.9 | 1.12
-86.8 | 0 | 28 | 1.72 | 354.1 | 1.49 | -86.8 | 0 | 28 | 0.8 | 1164.9 | 0.57 | -86.5 | 0 | 28 | 1.45 | 8242.1 | 1.26
-85.6 | 0 | 28 | 1.75 | 359.8 | 1.53 | -85.6 | 0 | 28 | 0.71 | 1046.1 | 0.53 | -85.3 | 0 | 28 | 1.47 | 8357.5 | 1.38
-84.4 | 0 | 28 | 1.68 | 344.4 | 1.46 | -84.4 | 0 | 28 | 0.67 | 978.5 | 0.51 | -84.1 | 0 | 28 | 1.82 | 10376.7 | 1.57
-83.2 | 0 | 28 | 2.35 | 482.3 | 2.02 | -83.2 | 0 | 28 | 0.44 | 650.2 | 0.32 | -82.9 | 0 | 28 | 1.95 | 11123.9 | 1.75
-82 | 0 | 28 | 2.3 | 472.9 | 2.02 | -82 | 0 | 28 | 0.24 | 345.7 | 0.17 | -81.7 | 0 | 28 | 2.07 | 11754.7 | 1.96
-80.8 | 0 | 28 | 2.12 | 435.5 | 1.86 | -80.8 | 0 | 28 | 0.01 | 15.2 | 0.01 | -80.5 | 0 | 28 | 2.53 | 14404 | 2.2
-79.6 | 0 | 28 | 2.67 | 548.3 | 2.32 | -79.6 | 0 | 28 | 0.21 | 307.1 | 0.15 | -79.3 | 0 | 28 | 2.58 | 14693 | 2.45
-78.4 | 0 | 28 | 2.83 | 581.1 | 2.49 | -78.4 | 0 | 28 | 0.54 | 787.5 | 0.4 | -78.1 | 0 | 28 | 3.17 | 18052.7 | 2.8
-77.2 | 0 | 28 | 3.04 | 624.3 | 2.68 | -77.2 | 0 | 28 | 0.86 | 1255 | 0.63 | -76.9 | 0 | 28 | 3.56 | 20260.6 | 3.1
-76 | 0 | 28 | 3.28 | 674 | 2.81 | -76 | 0 | 28 | 1.27 | 1860.9 | 0.93 | -75.6 | 0 | 28 | 3.71 | 21096.7 | 3.52
-74.8 | 0 | 28 | 3.6 | 739.3 | 3.2 | -74.8 | 0 | 28 | 1.74 | 2554.7 | 1.29 | -74.5 | 0 | 28 | 4.52 | 25712.8 | 3.93
-73.6 | 0 | 28 | 4.4 | 903.7 | 3.79 | -73.6 | 0 | 28 | 2.33 | 3418.8 | 1.74 | -73.2 | 0 | 28 | 4.74 | 26967.2 | 4.45
-72.4 | 0 | 28 | 4.5 | 925.3 | 3.8 | -72.4 | 0 | 28 | 3.01 | 4415.3 | 2.22 | -72 | 0 | 28 | 5.82 | 33135 | 5.04
-71.2 | 0 | 28 | 5.35 | 1100.3 | 4.74 | -71.2 | 0 | 28 | 3.69 | 5404.3 | 2.74 | -70.9 | 0 | 28 | 6.54 | 37248.1 | 5.82
-70 | 0 | 28 | 6.32 | 1298.4 | 5.38 | -70 | 0 | 28 | 4.73 | 6928.3 | 3.51 | -69.6 | 0 | 28 | 6.94 | 39496.3 | 6.6
-68.8 | 0 | 28 | 6.96 | 1429.8 | 5.9 | -68.8 | 0 | 28 | 5.76 | 8443.7 | 4.28 | -68.4 | 0 | 28 | 8.76 | 49831.9 | 7.63
-67.6 | 0 | 28 | 7.66 | 1574.9 | 6.74 | -67.6 | 0 | 28 | 7.23 | 10583.4 | 5.36 | -67.2 | 0 | 28 | 9.42 | 53602.8 | 8.79
-66.4 | 0 | 28 | 9.31 | 1913.7 | 8.07 | -66.4 | 0 | 28 | 8.86 | 12980.4 | 6.57 | -66 | 0 | 28 | 11.82 | 67267.1 | 10.28
-65.2 | 0 | 28 | 10.37 | 2131 | 9.05 | -65.2 | 0 | 28 | 10.97 | 16075.8 | 8.08 | -64.8 | 0 | 28 | 13.5 | 76828.5 | 11.96
-64 | 0 | 28 | 11.97 | 2460.3 | 10.68 | -63.9 | 0 | 28 | 13.35 | 19553 | 9.84 | -63.6 | 0 | 28 | 15.45 | 87938.5 | 14.08
-62.8 | 0 | 28 | 14.67 | 3014.7 | 12.54 | -62.8 | 0 | 28 | 16.1 | 23577.8 | 11.89 | -62.4 | 0 | 28 | 18.16 | 103381.4 | 16.47
-61.6 | 0 | 28 | 16.93 | 3479.6 | 14.71 | -61.5 | 0 | 28 | 20.37 | 29837.9 | 14.75 | -61.2 | 0 | 28 | 20.48 | 116591.3 | 19.22
-60.4 | 0 | 28 | 19.53 | 4012.4 | 17.33 | -60.3 | 0 | 28 | 24.25 | 35515.2 | 17.91 | -60 | 0 | 28 | 25.58 | 145616.3 | 22.55
-59.2 | 0 | 28 | 24.16 | 4963.6 | 20.73 | -59.1 | 0 | 28 | 29.9 | 43803.2 | 21.86 | -58.8 | 0 | 28 | 30.36 | 172804.4 | 26.18
-58 | 0 | 28 | 29.75 | 6113.8 | 24.7 | -57.9 | 0 | 28 | 36.28 | 53146.8 | 26.52 | -57.6 | 0 | 28 | 31.89 | 181502.5 | 30.38
-56.8 | 0 | 28 | 32.8 | 6740.1 | 29.56 | -56.7 | 0 | 28 | 43.08 | 63106.3 | 31.8 | -56.4 | 0 | 28 | 40.16 | 228607 | 34.68
-55.5 | 0 | 28 | 40.69 | 8360.6 | 34.95 | -55.5 | 0 | 28 | 52.41 | 76767.7 | 37.85 | -55.2 | 0 | 28 | 41.68 | 237216.3 | 39.4
-54.3 | 0 | 28 | 48.98 | 10064.6 | 40.79 | -54.3 | 0 | 28 | 58 | 84962.2 | 44.08 | -54 | 0 | 28 | 50.67 | 288383.6 | 44.69
-53.1 | 0 | 28 | 52.41 | 10768.9 | 46.95 | -53.1 | 0 | 28 | 69.25 | 101444.7 | 50.93 | -52.8 | 0 | 28 | 57.52 | 327409.9 | 49.78
-51.9 | 0 | 28 | 61.79 | 12696.5 | 53.34 | -51.9 | 0 | 28 | 79.02 | 115743.6 | 57.53 | -51.6 | 0 | 28 | 58 | 330098.5 | 55.03
-50.7 | 0 | 28 | 70.84 | 14555.8 | 59.58 | -50.7 | 0 | 28 | 86.8 | 127141.9 | 63.92 | -50.4 | 0 | 28 | 69.54 | 395820.8 | 60.11
-49.5 | 0 | 28 | 72.75 | 14949.4 | 65.71 | -49.5 | 0 | 28 | 95.39 | 139734.3 | 69.8 | -49.2 | 0 | 28 | 68.67 | 390829.2 | 64.8
-48.4 | 0 | 28 | 81.1 | 16663.8 | 71.04 | -48.3 | 0 | 28 | 100.92 | 147834.9 | 74.86 | -48 | 0 | 28 | 79.84 | 454447.4 | 69.49
-47.1 | 0 | 28 | 90.75 | 18647.3 | 76.07 | -47.1 | 0 | 28 | 107.55 | 157547.6 | 79.45 | -46.7 | 0 | 28 | 85.26 | 485260.7 | 74.04
-45.9 | 0 | 28 | 88.63 | 18212.1 | 80.35 | -45.9 | 0 | 28 | 112.24 | 164407.3 | 83.08 | -45.5 | 0 | 28 | 82.72 | 470833.8 | 77.75
-44.7 | 0 | 28 | 95.02 | 19524.9 | 83.69 | -44.7 | 0 | 28 | 115.96 | 169859.7 | 85.93 | -44.4 | 0 | 28 | 88.76 | 505192.9 | 80.95
-43.5 | 0 | 28 | 103.43 | 21252.3 | 86.71 | -43.5 | 0 | 28 | 119.52 | 175068.4 | 88.54 | -43.1 | 0 | 28 | 92.39 | 525852 | 83.94
-42.3 | 0 | 28 | 98.15 | 20168 | 88.73 | -42.3 | 0 | 28 | 121.97 | 178662.5 | 90.3 | -42 | 0 | 28 | 93.5 | 532154.8 | 86.13
-41.1 | 0 | 28 | 104.73 | 21521.1 | 90.66 | -41.1 | 0 | 28 | 124.96 | 183044 | 91.99 | -40.7 | 0 | 28 | 102.31 | 582304.3 | 88.6
-39.9 | 0 | 28 | 108.77 | 22349.9 | 91.89 | -39.9 | 0 | 28 | 128.95 | 188887.2 | 93.34 | -39.5 | 0 | 28 | 94.72 | 539144.5 | 90.23
-38.7 | 0 | 28 | 103.53 | 21273.8 | 93.12 | -38.7 | 0 | 28 | 126.97 | 185986.1 | 94.27 | -38.3 | 0 | 28 | 100.22 | 570434.4 | 91.62
-37.5 | 0 | 28 | 108.46 | 22285.6 | 94.16 | -37.5 | 0 | 28 | 131.39 | 192459.4 | 95.2 | -37.1 | 0 | 28 | 99.94 | 568850.2 | 92.58
-36.3 | 0 | 28 | 112.96 | 23211 | 94.73 | -36.3 | 0 | 28 | 131.67 | 192877.3 | 95.76 | -35.9 | 0 | 28 | 104.54 | 595022.5 | 93.45
-35.1 | 0 | 28 | 105.15 | 21607.1 | 95.52 | -35.1 | 0 | 28 | 130.81 | 191611.9 | 96.4 | -34.7 | 0 | 28 | 109.29 | 622050.2 | 94.52
-33.9 | 0 | 28 | 110.91 | 22789.8 | 96.44 | -33.9 | 0 | 28 | 126.7 | 185596.5 | 96.49 | -33.5 | 0 | 28 | 101.3 | 576550 | 95.29
-32.7 | 0 | 28 | 116.29 | 23896.4 | 96.88 | -32.7 | 0 | 28 | 132.69 | 194362.4 | 97.32 | -32.3 | 0 | 28 | 103.63 | 589824.3 | 95.82
-31.5 | 0 | 28 | 108.24 | 22240.4 | 97.08 | -31.5 | 0 | 28 | 135.43 | 198380.1 | 97.64 | -31.1 | 0 | 28 | 103.54 | 589327.6 | 96.24
-30.3 | 0 | 28 | 112.85 | 23188.5 | 97.42 | -30.3 | 0 | 28 | 131.52 | 192648.6 | 97.77 | -29.9 | 0 | 28 | 109.04 | 620619.3 | 96.55
-29.1 | 0 | 28 | 113.66 | 23354.1 | 97.35 | -29.1 | 0 | 28 | 133.63 | 195742.4 | 98.17 | -28.7 | 0 | 28 | 110.89 | 631172.1 | 96.55
-27.9 | 0 | 28 | 108.64 | 22322.9 | 97.75 | -27.9 | 0 | 28 | 133.33 | 195301.7 | 98.24 | -27.5 | 0 | 28 | 104.14 | 592743.7 | 97.42
-26.7 | 0 | 28 | 114.38 | 23503 | 98.52 | -26.7 | 0 | 28 | 132.89 | 194654.7 | 98.46 | -26.3 | 0 | 28 | 104.8 | 596499.7 | 97.61
-25.5 | 0 | 28 | 117.35 | 24113.8 | 98.77 | -25.5 | 0 | 28 | 136.53 | 199995.8 | 98.82 | -25.1 | 0 | 28 | 106.25 | 604771.9 | 98.01
-24.3 | 0 | 28 | 109.35 | 22468.7 | 98.77 | -24.3 | 0 | 28 | 133.36 | 195342.4 | 98.81 | -23.9 | 0 | 28 | 109.46 | 623001.9 | 98.06
-23.1 | 0 | 28 | 112.26 | 23067.3 | 98.46 | -23.1 | 0 | 28 | 130.51 | 191170.9 | 98.88 | -22.7 | 0 | 28 | 113.62 | 646702.6 | 97.87
-21.9 | 0 | 28 | 116.82 | 24005.3 | 98.82 | -21.9 | 0 | 28 | 134.22 | 196604.4 | 99.1 | -21.5 | 0 | 28 | 103.99 | 591873.4 | 98.41
-20.7 | 0 | 28 | 113.23 | 23267.5 | 99.07 | -20.7 | 0 | 28 | 134.82 | 197493.1 | 99.41 | -20.3 | 0 | 28 | 106.36 | 605354.8 | 98.72
-19.5 | 0 | 28 | 111.91 | 22995.6 | 99.56 | -19.5 | 0 | 28 | 134.3 | 196732 | 99.4 | -19 | 0 | 28 | 105.6 | 601066.9 | 98.89
-18.3 | 0 | 28 | 113.92 | 23407.5 | 98.9 | -18.3 | 0 | 28 | 135.67 | 198726.8 | 99.51 | -17.8 | 0 | 28 | 110.54 | 629148.4 | 98.67
-17.1 | 0 | 28 | 114.05 | 23435.7 | 99.63 | -17 | 0 | 28 | 135.52 | 198509.6 | 99.43 | -16.6 | 0 | 28 | 114.66 | 652594.2 | 98.94
-15.9 | 0 | 28 | 112.62 | 23140.6 | 99.66 | -15.8 | 0 | 28 | 136.51 | 199966.6 | 99.47 | -15.4 | 0 | 28 | 104.13 | 592691 | 98.99
-14.7 | 0 | 28 | 113.35 | 23290.7 | 99.24 | -14.6 | 0 | 28 | 129.76 | 190072 | 99.58 | -14.2 | 0 | 28 | 113.85 | 647982.4 | 99.23
-13.5 | 0 | 28 | 112.95 | 23208.7 | 99.39 | -13.4 | 0 | 28 | 137.47 | 201362.9 | 99.62 | -13 | 0 | 28 | 105.45 | 600217.5 | 99.38
-12.3 | 0 | 28 | 109.77 | 22556.1 | 99.47 | -12.2 | 0 | 28 | 136.58 | 200059.9 | 99.55 | -11.8 | 0 | 28 | 114.4 | 651129.2 | 99.26
-11.1 | 0 | 28 | 118.23 | 24294.8 | 99.46 | -11 | 0 | 28 | 130.46 | 191098 | 99.64 | -10.6 | 0 | 28 | 110.47 | 628782.9 | 99.25
-9.9 | 0 | 28 | 109.47 | 22494.6 | 99.63 | -9.8 | 0 | 28 | 137.42 | 201295 | 99.76 | -9.4 | 0 | 28 | 112.03 | 637640.6 | 99.44
-8.7 | 0 | 28 | 115.72 | 23778.8 | 100 | -8.6 | 0 | 28 | 137.13 | 200874.9 | 99.73 | -8.2 | 0 | 28 | 106.4 | 605627 | 99.56
-7.5 | 0 | 28 | 120.8 | 24822.3 | 99.89 | -7.4 | 0 | 28 | 136.26 | 199592 | 99.84 | -7 | 0 | 28 | 111.89 | 636867.3 | 99.6
-6.2 | 0 | 28 | 110.94 | 22797 | 99.74 | -6.2 | 0 | 28 | 137.02 | 200711.1 | 99.93 | -5.8 | 0 | 28 | 105.45 | 600174.8 | 99.7
-5.1 | 0 | 28 | 115.45 | 23721.9 | 99.8 | -5 | 0 | 28 | 137.34 | 201171.9 | 99.68 | -4.6 | 0 | 28 | 111.38 | 633943.1 | 99.67
-3.8 | 0 | 28 | 119.85 | 24627.6 | 99.94 | -3.8 | 0 | 28 | 130.05 | 190500.8 | 99.71 | -3.4 | 0 | 28 | 109.93 | 625699.1 | 99.76
-2.6 | 0 | 28 | 109.87 | 22577.2 | 99.66 | -2.6 | 0 | 28 | 137.63 | 201597.9 | 99.95 | -2.2 | 0 | 28 | 108.23 | 616004.2 | 99.71
-1.4 | 0 | 28 | 114.42 | 23511.9 | 99.93 | -1.4 | 0 | 28 | 136.49 | 199935.9 | 99.75 | -1 | 0 | 28 | 113.43 | 645626 | 99.78
-0.2 | 0 | 28 | 119.36 | 24525.5 | 99.84 | -0.2 | 0 | 28 | 136.8 | 200386.7 | 99.87 | 0.2 | 0 | 28 | 105.41 | 599969.6 | 99.5
1 | 0 | 28 | 109.16 | 22431.1 | 99.5 | 1 | 0 | 28 | 136.64 | 200154.9 | 99.78 | 1.4 | 0 | 28 | 112.25 | 638925.2 | 99.74
2.2 | 0 | 28 | 114.11 | 23447.6 | 99.79 | 2.2 | 0 | 28 | 135.82 | 198954.9 | 99.87 | 2.6 | 0 | 28 | 103.96 | 591721.8 | 99.88
3.4 | 0 | 28 | 119.88 | 24633.6 | 99.8 | 3.4 | 0 | 28 | 135.96 | 199159.9 | 100 | 3.8 | 0 | 28 | 114.65 | 652582.8 | 100
4.6 | 0 | 28 | 111.04 | 22816.7 | 99.29 | 4.6 | 0 | 28 | 135.68 | 198749.5 | 99.88 | 5 | 0 | 28 | 111.91 | 636970.7 | 99.85
5.8 | 0 | 28 | 115.72 | 23777.8 | 99.71 | 5.8 | 0 | 28 | 135.01 | 197769.5 | 99.87 | 6.2 | 0 | 28 | 105.25 | 599043 | 99.62
7 | 0 | 28 | 119.12 | 24477.1 | 99.5 | 7 | 0 | 28 | 135.8 | 198923.4 | 99.87 | 7.4 | 0 | 28 | 112.19 | 638570.8 | 99.72
8.2 | 0 | 28 | 112.15 | 23044.9 | 99.44 | 8.2 | 0 | 28 | 134.75 | 197379.3 | 99.77 | 8.7 | 0 | 28 | 106.12 | 604024.4 | 99.8
9.4 | 0 | 28 | 116.2 | 23876.8 | 99.76 | 9.4 | 0 | 28 | 134.72 | 197337.5 | 99.71 | 9.8 | 0 | 28 | 110.47 | 628790.6 | 99.81
10.6 | 0 | 28 | 118.91 | 24433.3 | 99.47 | 10.6 | 0 | 28 | 134.49 | 196997.1 | 99.79 | 11.1 | 0 | 28 | 103.98 | 591852 | 99.69
11.8 | 0 | 28 | 110.02 | 22607.5 | 99.57 | 11.8 | 0 | 28 | 139.41 | 204215.9 | 99.69 | 12.2 | 0 | 28 | 114.55 | 652015.9 | 99.19
13 | 0 | 28 | 113.55 | 23331.9 | 99.73 | 13 | 0 | 28 | 133.72 | 195882.7 | 99.41 | 13.5 | 0 | 28 | 112.53 | 640493.2 | 99.7
14.2 | 0 | 28 | 118.04 | 24254.5 | 99.15 | 14.2 | 0 | 28 | 127.98 | 187471.3 | 99.3 | 14.7 | 0 | 28 | 105.54 | 600695.3 | 99.66
15.4 | 0 | 28 | 111.68 | 22947.4 | 99.68 | 15.4 | 0 | 28 | 139.75 | 204704.5 | 99.56 | 15.8 | 0 | 28 | 114.24 | 650240.9 | 99.5
16.6 | 0 | 28 | 114.19 | 23463.8 | 99.59 | 16.6 | 0 | 28 | 133.45 | 195479.2 | 99.24 | 17.1 | 0 | 28 | 106.04 | 603571.3 | 99.47
17.8 | 0 | 28 | 117.3 | 24104 | 99.61 | 17.8 | 0 | 28 | 127.96 | 187434.1 | 99.22 | 18.3 | 0 | 28 | 113.97 | 648715.2 | 98.96
19 | 0 | 28 | 111.79 | 22971.7 | 99.13 | 19 | 0 | 28 | 139.58 | 204460 | 99.43 | 19.5 | 0 | 28 | 111.67 | 635619.7 | 99.06
20.2 | 0 | 28 | 113.39 | 23299.5 | 99.29 | 20.2 | 0 | 28 | 132.73 | 194429.6 | 98.97 | 20.7 | 0 | 28 | 106.44 | 605812 | 99.06
21.4 | 0 | 28 | 116.54 | 23945.8 | 99.14 | 21.4 | 0 | 28 | 131.15 | 192112.8 | 99.01 | 21.9 | 0 | 28 | 107.64 | 612654.3 | 99
22.6 | 0 | 28 | 111.36 | 22882.8 | 98.65 | 22.6 | 0 | 28 | 133.38 | 195381 | 98.87 | 23.1 | 0 | 28 | 110.21 | 627285.7 | 98.83
23.8 | 0 | 28 | 114.53 | 23533.1 | 99.03 | 23.8 | 0 | 28 | 132.62 | 194262.1 | 98.59 | 24.3 | 0 | 28 | 113.01 | 643204.2 | 98.39
25 | 0 | 28 | 114.55 | 23538.5 | 98.42 | 25 | 0 | 28 | 128.1 | 187643.9 | 98.41 | 25.5 | 0 | 28 | 111.69 | 635716.9 | 97.97
26.2 | 0 | 28 | 111.24 | 22858.4 | 98.07 | 26.2 | 0 | 28 | 136.81 | 200405.7 | 98.31 | 26.7 | 0 | 28 | 98.77 | 562170.3 | 97.97
27.4 | 0 | 28 | 113.38 | 23297 | 97.89 | 27.4 | 0 | 28 | 127.93 | 187396.9 | 97.77 | 27.9 | 0 | 28 | 111.11 | 632384.3 | 97.77
28.6 | 0 | 28 | 112.62 | 23141.3 | 97.62 | 28.6 | 0 | 28 | 131.63 | 192811.3 | 97.8 | 29.1 | 0 | 28 | 102.71 | 584578 | 97.36
29.8 | 0 | 28 | 110.07 | 22616.9 | 97.21 | 29.8 | 0 | 28 | 135.84 | 198979.3 | 97.56 | 30.3 | 0 | 28 | 110.53 | 629112.5 | 96.59
31 | 0 | 28 | 112.55 | 23126.6 | 97.69 | 31 | 0 | 28 | 130.28 | 190834.3 | 96.98 | 31.5 | 0 | 28 | 104.56 | 595103.5 | 96.32
32.2 | 0 | 28 | 113.94 | 23412.7 | 96.89 | 32.2 | 0 | 28 | 125.7 | 184131.4 | 96.36 | 32.7 | 0 | 28 | 102.25 | 581959.6 | 95.68
33.4 | 0 | 28 | 110.64 | 22734.7 | 96.03 | 33.4 | 0 | 28 | 133.01 | 194831.8 | 95.97 | 33.9 | 0 | 28 | 108.84 | 619507.2 | 94.91
34.6 | 0 | 28 | 111.42 | 22894 | 95.76 | 34.6 | 0 | 28 | 129.35 | 189476.7 | 95.17 | 35.1 | 0 | 28 | 99.86 | 568401.6 | 93.92
35.8 | 0 | 28 | 105.62 | 21703.8 | 94.64 | 35.8 | 0 | 28 | 123.04 | 180229.1 | 94.18 | 36.3 | 0 | 28 | 105.55 | 600747.1 | 92.82
37 | 0 | 28 | 107.71 | 22132.2 | 94.18 | 37 | 0 | 28 | 128.87 | 188770.2 | 93.27 | 37.5 | 0 | 28 | 102.41 | 582871.6 | 91.37
38.2 | 0 | 28 | 106.97 | 21980.7 | 92.8 | 38.2 | 0 | 28 | 120.54 | 176573.1 | 91.7 | 38.7 | 0 | 28 | 92.8 | 528199.8 | 89.89
39.4 | 0 | 28 | 108.12 | 22217.6 | 91.68 | 39.4 | 0 | 28 | 120.95 | 177175.5 | 90.58 | 39.9 | 0 | 28 | 100.64 | 572813.9 | 87.97
40.6 | 0 | 28 | 104.4 | 21452.1 | 90.42 | 40.6 | 0 | 28 | 123.67 | 181155.8 | 88.59 | 41.1 | 0 | 28 | 90.14 | 513029.6 | 85.69
41.8 | 0 | 28 | 97.04 | 19939.1 | 88.02 | 41.8 | 0 | 28 | 112.19 | 164335.3 | 85.84 | 42.3 | 0 | 28 | 92.51 | 526565.3 | 82.95
43 | 0 | 28 | 100.56 | 20664.1 | 85.68 | 43 | 0 | 28 | 112.45 | 164723.3 | 83.13 | 43.5 | 0 | 28 | 88.35 | 502871.1 | 79.56
44.2 | 0 | 28 | 93.99 | 19314 | 82.7 | 44.2 | 0 | 28 | 107.69 | 157750.8 | 79.2 | 44.8 | 0 | 28 | 80.41 | 457662.8 | 75.83
45.4 | 0 | 28 | 90.6 | 18616.2 | 78.98 | 45.4 | 0 | 28 | 102.23 | 149752.8 | 74.93 | 46 | 0 | 28 | 82.08 | 467166.9 | 72.05
46.6 | 0 | 28 | 86.31 | 17734.6 | 74.54 | 46.7 | 0 | 28 | 94.94 | 139073.2 | 69.99 | 47.2 | 0 | 28 | 73.63 | 419080.1 | 67.87
47.8 | 0 | 28 | 78.3 | 16089.2 | 68.86 | 47.9 | 0 | 28 | 87.18 | 127706.4 | 64.33 | 48.3 | 0 | 28 | 72.55 | 412939.7 | 62.64
49 | 0 | 28 | 71.64 | 14720 | 63.48 | 49 | 0 | 28 | 79.03 | 115758.9 | 58 | 49.6 | 0 | 28 | 63.92 | 363801.4 | 57.84
50.3 | 0 | 28 | 64.89 | 13333.3 | 57.21 | 50.3 | 0 | 28 | 70.06 | 102621.9 | 51.41 | 50.8 | 0 | 28 | 59.94 | 341178.1 | 52.25
51.5 | 0 | 28 | 56.81 | 11672.9 | 50.56 | 51.5 | 0 | 28 | 61.22 | 89682 | 44.74 | 52 | 0 | 28 | 50.64 | 288211.5 | 47.46
52.6 | 0 | 28 | 52.6 | 10807.9 | 44.23 | 52.7 | 0 | 28 | 51.94 | 76086.8 | 38.05 | 53.2 | 0 | 28 | 47.01 | 267589.3 | 42.1
53.9 | 0 | 28 | 43.13 | 8861.4 | 37.84 | 53.9 | 0 | 28 | 44.1 | 64595 | 32.13 | 54.4 | 0 | 28 | 40.17 | 228656 | 36.99
55.1 | 0 | 28 | 36.79 | 7560 | 31.94 | 55.1 | 0 | 28 | 36.25 | 53099.6 | 26.51 | 55.6 | 0 | 28 | 36.9 | 210052.1 | 32.41
56.3 | 0 | 28 | 30.42 | 6249.8 | 26.81 | 56.3 | 0 | 28 | 29.9 | 43795.1 | 21.79 | 56.8 | 0 | 28 | 30.35 | 172730.9 | 27.84
57.5 | 0 | 28 | 26.4 | 5425.2 | 22.61 | 57.5 | 0 | 28 | 24.34 | 35657.7 | 17.76 | 58 | 0 | 28 | 25.81 | 146905.8 | 24.05
58.6 | 0 | 28 | 21.45 | 4408.3 | 18.8 | 58.7 | 0 | 28 | 19.78 | 28970.7 | 14.44 | 59.2 | 0 | 28 | 22.84 | 129980.5 | 20.49
59.9 | 0 | 28 | 18.29 | 3758.7 | 15.78 | 59.9 | 0 | 28 | 16.01 | 23455.8 | 11.7 | 60.4 | 0 | 28 | 18.9 | 107583.2 | 17.55
61.1 | 0 | 28 | 15.4 | 3163.5 | 13.15 | 61.1 | 0 | 28 | 12.25 | 17948 | 9.36 | 61.6 | 0 | 28 | 16.11 | 91712.6 | 15.03
62.3 | 0 | 28 | 12.62 | 2593 | 11 | 62.3 | 0 | 28 | 10.65 | 15602.3 | 7.77 | 62.8 | 0 | 28 | 14.78 | 84116.1 | 12.72
63.5 | 0 | 28 | 10.69 | 2197 | 9.17 | 63.5 | 0 | 28 | 8.69 | 12735.1 | 6.31 | 64 | 0 | 28 | 11.78 | 67030.5 | 11.02
64.7 | 0 | 28 | 9.1 | 1869.8 | 7.8 | 64.7 | 0 | 28 | 6.54 | 9573.4 | 5.03 | 65.2 | 0 | 28 | 10.79 | 61416.8 | 9.42
65.9 | 0 | 28 | 7.8 | 1602.7 | 6.73 | 65.9 | 0 | 28 | 5.73 | 8387 | 4.17 | 66.4 | 0 | 28 | 8.51 | 48454.8 | 8.1
67.1 | 0 | 28 | 6.95 | 1428 | 5.94 | 67.1 | 0 | 28 | 4.58 | 6711.7 | 3.35 | 67.6 | 0 | 28 | 7.59 | 43185.6 | 7.03
68.3 | 0 | 28 | 5.93 | 1218.5 | 5.35 | 68.3 | 0 | 28 | 3.7 | 5420.7 | 2.71 | 68.8 | 0 | 28 | 7.03 | 40021.6 | 6.08
69.5 | 0 | 28 | 5 | 1027.8 | 4.32 | 69.5 | 0 | 28 | 3 | 4394.6 | 2.2 | 70 | 0 | 28 | 5.9 | 33561.8 | 5.34
70.7 | 0 | 28 | 4.14 | 850.5 | 3.58 | 70.7 | 0 | 28 | 2.32 | 3398 | 1.71 | 71.3 | 0 | 28 | 5.33 | 30358.7 | 4.66
71.9 | 0 | 28 | 3.81 | 782 | 3.29 | 71.9 | 0 | 28 | 1.85 | 2708.1 | 1.36 | 72.5 | 0 | 28 | 4.38 | 24949.5 | 4.09
73.1 | 0 | 28 | 3.61 | 742.4 | 3.13 | 73.1 | 0 | 28 | 1.34 | 1956.5 | 0.99 | 73.7 | 0 | 28 | 3.91 | 22271.2 | 3.63
74.3 | 0 | 28 | 2.95 | 605.6 | 2.55 | 74.3 | 0 | 28 | 0.95 | 1385.7 | 0.7 | 74.9 | 0 | 28 | 3.61 | 20529.3 | 3.21
75.5 | 0 | 28 | 2.47 | 507.9 | 2.16 | 75.5 | 0 | 28 | 0.71 | 1040.4 | 0.53 | 76.1 | 0 | 28 | 3.28 | 18661.4 | 2.83
76.7 | 0 | 28 | 2.17 | 446.1 | 1.88 | 76.7 | 0 | 28 | 0.36 | 521.2 | 0.26 | 77.3 | 0 | 28 | 2.72 | 15456.3 | 2.55
77.9 | 0 | 28 | 2.1 | 431.6 | 1.85 | 77.9 | 0 | 28 | 0.02 | 33.1 | 0.02 | 78.5 | 0 | 28 | 2.58 | 14689.8 | 2.24
79.1 | 0 | 28 | 2.04 | 418.7 | 1.76 | 79.1 | 0 | 28 | 0.11 | 158.2 | 0.08 | 79.7 | 0 | 28 | 2.18 | 12398.1 | 2.02
80.3 | 0 | 28 | 1.66 | 341.7 | 1.46 | 80.3 | 0 | 28 | 0.32 | 475.5 | 0.24 | 80.9 | 0 | 28 | 1.9 | 10828.5 | 1.77
81.5 | 0 | 28 | 1.54 | 316.5 | 1.36 | 81.5 | 0 | 28 | 0.53 | 778 | 0.4 | 82.1 | 0 | 28 | 1.79 | 10193 | 1.59
82.7 | 0 | 28 | 1.47 | 301.1 | 1.25 | 82.7 | 0 | 28 | 0.66 | 968.7 | 0.49 | 83.3 | 0 | 28 | 1.49 | 8469.2 | 1.41
83.9 | 0 | 28 | 1.4 | 287.9 | 1.22 | 84 | 0 | 28 | 0.77 | 1121.3 | 0.56 | 84.5 | 0 | 28 | 1.47 | 8358 | 1.26
85.1 | 0 | 28 | 1.57 | 322.6 | 1.35 | 85.2 | 0 | 28 | 1.01 | 1483.7 | 0.76 | 85.7 | 0 | 28 | 1.2 | 6813.2 | 1.11
86.3 | 0 | 28 | 1.55 | 319.1 | 1.34 | 86.3 | 0 | 28 | 1.08 | 1578.4 | 0.8 | 86.9 | 0 | 28 | 1.07 | 6109.7 | 0.99
87.5 | 0 | 28 | 0.97 | 198.6 | 0.84 | 87.6 | 0 | 28 | 1.12 | 1642.9 | 0.81 | 88.1 | 0 | 28 | 1.01 | 5758.3 | 0.89
88.7 | 0 | 28 | 0.93 | 191.6 | 0.79 | 88.8 | 0 | 28 | 1.25 | 1825.8 | 0.93 | 89.3 | 0 | 28 | 0.85 | 4814.6 | 0.79
89.9 | 0 | 28 | 0.85 | 175.4 | 0.74 | 90 | 0 | 28 | 1.36 | 1993.8 | 1.01 | 90.5 | 0 | 28 | 0.83 | 4702.1 | 0.71
91.1 | 0 | 28 | 1.15 | 235.3 | 0.98 | 91.2 | 0 | 28 | 1.4 | 2047.8 | 1.01 | 91.7 | 0 | 28 | 0.7 | 3984.6 | 0.65
92.3 | 0 | 28 | 1.26 | 259.7 | 1.13 | 92.4 | 0 | 28 | 1.5 | 2197.4 | 1.13 | 92.9 | 0 | 28 | 0.64 | 3641.1 | 0.59
93.5 | 0 | 28 | 1.06 | 217.6 | 0.92 | 93.6 | 0 | 28 | 1.51 | 2215.4 | 1.1 | 94.1 | 0 | 28 | 0.58 | 3311.8 | 0.52
94.7 | 0 | 28 | 0.84 | 173 | 0.73 | 94.8 | 0 | 28 | 1.59 | 2323.3 | 1.14 | 95.3 | 0 | 28 | 0.53 | 3014.8 | 0.5
95.9 | 0 | 28 | 1.08 | 221.3 | 0.94 | 96 | 0 | 28 | 1.58 | 2309 | 1.17 | 96.5 | 0 | 28 | 0.52 | 2955.6 | 0.45
97.1 | 0 | 28 | 0.56 | 114.8 | 0.49 | 97.2 | 0 | 28 | 1.67 | 2445.3 | 1.2 | 97.7 | 0 | 28 | 0.45 | 2584.8 | 0.42
98.3 | 0 | 28 | 0.26 | 52.6 | 0.22 | 98.4 | 0 | 28 | 1.68 | 2454.9 | 1.21 | 98.9 | 0 | 28 | 0.42 | 2392.4 | 0.39
99.5 | 0 | 28 | 0.52 | 106.1 | 0.45 | 99.6 | 0 | 28 | 1.7 | 2486.5 | 1.25 | 100.1 | 0 | 28 | 0.44 | 2512.3 | 0.39
100.7 | 0 | 28 | 0.34 | 70.5 | 0.3 | 100.8 | 0 | 28 | 1.72 | 2526.8 | 1.32 | 101.4 | 0 | 28 | 0.41 | 2321.9 | 0.37
101.9 | 0 | 28 | 0.43 | 88.4 | 0.38 | 102 | 0 | 28 | 1.72 | 2516.9 | 1.23 | 102.5 | 0 | 28 | 0.37 | 2095.7 | 0.34
103.2 | 0 | 28 | 0.75 | 153.3 | 0.66 | 103.2 | 0 | 28 | 1.74 | 2552.4 | 1.3 | 103.8 | 0 | 28 | 0.38 | 2172.6 | 0.33
104.4 | 0 | 28 | 0.63 | 130.1 | 0.55 | 104.4 | 0 | 28 | 1.78 | 2601.1 | 1.36 | 105 | 0 | 28 | 0.33 | 1868.7 | 0.31
105.6 | 0 | 28 | 0.34 | 69.4 | 0.29 | 105.6 | 0 | 28 | 1.75 | 2556.5 | 1.24 | 106.2 | 0 | 28 | 0.35 | 2001.8 | 0.31
106.8 | 0 | 28 | 0.4 | 82.9 | 0.35 | 106.8 | 0 | 28 | 1.8 | 2636 | 1.36 | 107.4 | 0 | 28 | 0.33 | 1902.8 | 0.31
108 | 0 | 28 | 0.8 | 165 | 0.69 | 108 | 0 | 28 | 1.79 | 2628.3 | 1.32 | 108.6 | 0 | 28 | 0.33 | 1900.5 | 0.3
109.2 | 0 | 28 | 0.54 | 110.7 | 0.46 | 109.2 | 0 | 28 | 1.77 | 2592.7 | 1.25 | 109.8 | 0 | 28 | 0.32 | 1824.5 | 0.28
110.4 | 0 | 28 | 0.55 | 113.9 | 0.48 | 110.4 | 0 | 28 | 1.81 | 2647.1 | 1.38 | 111 | 0 | 28 | 0.3 | 1733.4 | 0.29
111.6 | 0 | 28 | 0.4 | 83 | 0.35 | 111.6 | 0 | 28 | 1.84 | 2689.2 | 1.35 | 112.2 | 0 | 28 | 0.31 | 1748.9 | 0.27
112.8 | 0 | 28 | 0.51 | 104.3 | 0.44 | 112.8 | 0 | 28 | 1.84 | 2696.7 | 1.32 | 113.4 | 0 | 28 | 0.28 | 1585.2 | 0.26
114 | 0 | 28 | 0.69 | 142.1 | 0.59 | 114 | 0 | 28 | 1.82 | 2659.2 | 1.4 | 114.6 | 0 | 28 | 0.28 | 1620.5 | 0.26
115.2 | 0 | 28 | 0.38 | 77.7 | 0.34 | 115.2 | 0 | 28 | 1.86 | 2719.4 | 1.37 | 115.8 | 0 | 28 | 0.32 | 1810.1 | 0.27
116.4 | 0 | 28 | 0.38 | 77.8 | 0.33 | 116.4 | 0 | 28 | 1.79 | 2623.8 | 1.27 | 117 | 0 | 28 | 0.27 | 1559.9 | 0.25
117.6 | 0 | 28 | 0.19 | 39.2 | 0.17 | 117.6 | 0 | 28 | 1.9 | 2778.8 | 1.46 | 118.2 | 0 | 28 | 0.28 | 1598.7 | 0.24
118.8 | 0 | 28 | 0.47 | 97.2 | 0.41 | 118.8 | 0 | 28 | 1.88 | 2750.3 | 1.39 | 119.4 | 0 | 28 | 0.28 | 1582.1 | 0.26
120 | 0 | 28 | 0.39 | 80 | 0.33 | 120 | 0 | 28 | 1.83 | 2683.6 | 1.29 | 120.6 | 0 | 28 | 0.28 | 1604.3 | 0.26
121.2 | 0 | 28 | 0.35 | 72.5 | 0.31 | 121.3 | 0 | 28 | 1.85 | 2713.4 | 1.43 | 121.8 | 0 | 28 | 0.26 | 1497.8 | 0.24
122.4 | 0 | 28 | 0.09 | 17.7 | 0.08 | 122.4 | 0 | 28 | 1.85 | 2703.6 | 1.36 | 123 | 0 | 28 | 0.25 | 1416.7 | 0.23
123.6 | 0 | 28 | 0.6 | 122.7 | 0.5 | 123.6 | 0 | 28 | 1.84 | 2693.1 | 1.3 | 124.2 | 0 | 28 | 0.27 | 1556.2 | 0.23
124.8 | 0 | 28 | 0.51 | 104.8 | 0.46 | 124.9 | 0 | 28 | 1.97 | 2887.6 | 1.53 | 125.4 | 0 | 28 | 0.24 | 1364.9 | 0.23
126 | 0 | 28 | 0.52 | 106.9 | 0.45 | 126 | 0 | 28 | 1.89 | 2773.6 | 1.38 | 126.6 | 0 | 28 | 0.25 | 1409.2 | 0.23
127.2 | 0 | 28 | 0.51 | 104.6 | 0.44 | 127.2 | 0 | 28 | 1.89 | 2762.5 | 1.33 | 127.8 | 0 | 28 | 0.27 | 1510.8 | 0.24
128.4 | 0 | 28 | 0.45 | 92.8 | 0.39 | 128.4 | 0 | 28 | 1.9 | 2779.7 | 1.44 | 129 | 0 | 28 | 0.23 | 1322.2 | 0.21
129.6 | 0 | 28 | 0.2 | 40.7 | 0.17 | 129.7 | 0 | 28 | 1.91 | 2802.2 | 1.39 | 130.2 | 0 | 28 | 0.26 | 1472 | 0.22
130.8 | 0 | 28 | 0.1 | 19.6 | 0.08 | 130.8 | 0 | 28 | 1.93 | 2833.1 | 1.42 | 131.4 | 0 | 28 | 0.24 | 1347.1 | 0.22
132 | 0 | 28 | 0.16 | 33 | 0.14 | 132.1 | 0 | 28 | 1.92 | 2808.4 | 1.46 | 132.6 | 0 | 28 | 0.25 | 1399.7 | 0.22
133.2 | 0 | 28 | 0.6 | 124 | 0.52 | 133.3 | 0 | 28 | 1.96 | 2877.7 | 1.43 | 133.8 | 0 | 28 | 0.24 | 1370.2 | 0.22
134.4 | 0 | 28 | 0.44 | 89.4 | 0.38 | 134.5 | 0 | 28 | 1.96 | 2864.8 | 1.38 | 135 | 0 | 28 | 0.24 | 1353.6 | 0.22
135.6 | 0 | 28 | 0.01 | 2.5 | 0.01 | 135.7 | 0 | 28 | 1.96 | 2865 | 1.49 | 136.2 | 0 | 28 | 0.23 | 1320.4 | 0.2
136.8 | 0 | 28 | 0.39 | 80.3 | 0.34 | 136.9 | 0 | 28 | 1.94 | 2849 | 1.42 | 137.5 | 0 | 28 | 0.23 | 1296.3 | 0.22
138 | 0 | 28 | 0.01 | 2.2 | 0.01 | 138.1 | 0 | 28 | 1.86 | 2721.4 | 1.31 | 138.7 | 0 | 28 | 0.21 | 1178.7 | 0.19
139.2 | 0 | 28 | 0.16 | 32.7 | 0.14 | 139.3 | 0 | 28 | 1.96 | 2877.7 | 1.49 | 139.8 | 0 | 28 | 0.23 | 1310.5 | 0.21
140.4 | 0 | 28 | 0.56 | 114.1 | 0.49 | 140.5 | 0 | 28 | 1.96 | 2867.1 | 1.5 | 141.1 | 0 | 28 | 0.24 | 1388 | 0.21
141.6 | 0 | 28 | 0.64 | 130.9 | 0.54 | 141.7 | 0 | 28 | 1.91 | 2802.6 | 1.34 | 142.3 | 0 | 28 | 0.23 | 1286.4 | 0.2
142.8 | 0 | 28 | 0.27 | 54.7 | 0.24 | 142.9 | 0 | 28 | 2.01 | 2942.1 | 1.55 | 143.5 | 0 | 28 | 0.21 | 1191.6 | 0.2
144 | 0 | 28 | 0.23 | 48.1 | 0.2 | 144.1 | 0 | 28 | 1.91 | 2799.4 | 1.41 | 144.7 | 0 | 28 | 0.22 | 1226.6 | 0.19
145.2 | 0 | 28 | 0.33 | 68.6 | 0.29 | 145.3 | 0 | 28 | 1.92 | 2817.3 | 1.36 | 145.9 | 0 | 28 | 0.2 | 1125.9 | 0.18
146.4 | 0 | 28 | 0.52 | 106.8 | 0.45 | 146.5 | 0 | 28 | 2.01 | 2937.9 | 1.56 | 147.1 | 0 | 28 | 0.22 | 1254.8 | 0.2
147.6 | 0 | 28 | 0.54 | 111.2 | 0.47 | 147.7 | 0 | 28 | 1.95 | 2861.5 | 1.44 | 148.3 | 0 | 28 | 0.23 | 1287.5 | 0.2
148.8 | 0 | 28 | 0.58 | 118.4 | 0.5 | 148.9 | 0 | 28 | 2.02 | 2958.4 | 1.44 | 149.5 | 0 | 28 | 0.22 | 1236.5 | 0.21
150 | 0 | 28 | 0 | 1 | 0 | 150.1 | 0 | 28 | 1.97 | 2880.6 | 1.46 | 151.4 | 0 | 28 | 0.23 | 1314.4 | 0.21
151.4 | 0 | 28 | 0.5 | 101.7 | 0.43 | 151.4 | 0 | 28 | 1.99 | 2922.2 | 1.49