## 0,5x0,5
Measurement time: | 12/04/2023 18:06:08 |  |  | Taille de champ : |  |  | Measurement time: | 27/05/2023 10:11:08 |  |  | Taille de champ : |  |  | Measurement time: | 27/05/2023 12:25:45 |  |  |  | Taille de champ :
Radiation device: | Novalis TB Accelerator |  |  | 5.1717 |  |  | Radiation device: | Novalis TB Accelerator |  |  | 5.2881 |  |  | Radiation device: | Novalis TB Accelerator |  |  |  | 5.4433
Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF)
Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | PTW60019 Diamant |  | Côté G |  | Côté D |  | Field detector: | Razor Nano |  |  | Côté G |  | Côté D
Reference detector: | Stealth CI transmission |  | 35.1504 | 2.5 | 68.7576 | -2 | Reference detector: | Stealth CI transmission |  | 49.1348 | -2.5 | 59.8416 | 2 | Reference detector: | Stealth CI transmission |  |  | 49.4588 | -2.4 | 69.6025 | 2
SAD: | 1000 mm |  | 58.9098 | 2 | 46.3321 | -2.5 | SAD: | 1000 mm |  | 66.4601 | -2 | 40.6592 | 2.5 | SAD: | 1000 mm |  |  | 62.3624 | -2 | 53.1045 | 2.5
SSD: | 900 mm |  | 50 | 2.1875 | 50 | -2.4182 | SSD: | 900 mm |  | 50 | -2.475 | 50 | 2.2565 | SSD: | 900 mm |  |  | 50 | -2.3832 | 50 | 2.5941
Field size: | 5 x 5 mm | [5 empty] | Field size: | 5 x 5 mm | [5 empty] | Field size: | 5 x 5 mm
Gantry angle: | 0 |  |  | 4.6057 |  |  | Gantry angle: | 0 |  |  | 4.7316 |  |  | Gantry angle: | 0 |  |  |  | 4.9773
Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water
Scan type: | Inline | [5 empty] | Scan type: | Inline | [5 empty] | Scan type: | Inline
Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep
Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
5 | 0 | 100 | 97.71 | 34144.8 | 3.97 | 3.964 | -10 | 0 | 100 | 54.94 | 713 | 0.95 | 0.9469 | -10 | 0 | 100 | 23.3 | 135.8 | 1.18 | 1.1846
4.5 | 0 | 100 | 140.14 | 48973.5 | 5.7 | 5.6854 | -9.5 | 0 | 100 | 58.36 | 757.5 | 1.01 | 1.0058 | -9.4 | 0 | 100 | 27.27 | 158.9 | 1.39 | 1.3864
4 | 0 | 100 | 194.29 | 67896 | 7.9 | 7.8822 | -9 | 0 | 100 | 67.74 | 879.2 | 1.17 | 1.1675 | -8.9 | 0 | 100 | 39.6 | 230.8 | 2.02 | 2.0132
3.5 | 0 | 100 | 308.49 | 107801.2 | 12.55 | 12.5152 | -8.4 | 0 | 100 | 83.09 | 1078.4 | 1.44 | 1.4321 | -8.4 | 0 | 100 | 9.34 | 54.4 | 0.48 | 0.47484
3 | 0 | 100 | 499.06 | 174398.1 | 20.26 | 20.2465 | -8 | 0 | 100 | 96.87 | 1257.2 | 1.68 | 1.6696 | -8 | 0 | 100 | 36.96 | 215.4 | 1.88 | 1.879
2.5 | 0 | 100 | 866.43 | 302775.5 | 35.21 | 35.1504 | -7.5 | 0 | 100 | 116.72 | 1514.9 | 2.03 | 2.0117 | -7.5 | 0 | 100 | 67.52 | 393.5 | 3.45 | 3.4327
2 | 0 | 100 | 1452.08 | 507428.6 | 59.05 | 58.9098 | -7 | 0 | 100 | 134.31 | 1743.2 | 2.33 | 2.3149 | -7 | 0 | 100 | 77.93 | 454.2 | 3.98 | 3.9619
1.5 | 0 | 100 | 2003.44 | 700104.6 | 81.38 | 81.2781 | -6.6 | 0 | 100 | 158.31 | 2054.7 | 2.75 | 2.7285 | -6.5 | 0 | 100 | 110.63 | 644.7 | 5.63 | 5.6244
1 | 0 | 100 | 2275.87 | 795304.4 | 92.42 | 92.3304 | -6 | 0 | 100 | 206.94 | 2685.9 | 3.59 | 3.5667 | -6 | 0 | 100 | 133.58 | 778.5 | 6.8 | 6.7912
0.6 | 0 | 100 | 2407.5 | 841303.4 | 97.85 | 97.6705 | -5.5 | 0 | 100 | 260.45 | 3380.3 | 4.52 | 4.4889 | -5.5 | 0 | 100 | 168.55 | 982.3 | 8.59 | 8.569
-0.1 | 0 | 100 | 2464.92 | 861366.9 | 100.27 | 100 | -5 | 0 | 100 | 362.46 | 4704.3 | 6.3 | 6.2471 | -5 | 0 | 100 | 202.38 | 1179.5 | 10.36 | 10.2889
-0.5 | 0 | 100 | 2439.73 | 852564.1 | 99.25 | 98.9781 | -4.5 | 0 | 100 | 499.22 | 6479.3 | 8.65 | 8.6042 | -4.5 | 0 | 100 | 257.98 | 1503.5 | 13.18 | 13.1156
-1 | 0 | 100 | 2354.74 | 822866.6 | 95.73 | 95.5301 | -4.1 | 0 | 100 | 733.22 | 9516.4 | 12.69 | 12.6372 | -4 | 0 | 100 | 341.85 | 1992.3 | 17.49 | 17.3795
-1.5 | 0 | 100 | 2133.14 | 745428.1 | 86.74 | 86.5399 | -3.4 | 0 | 100 | 1329.75 | 17258.7 | 23.1 | 22.9185 | -3.5 | 0 | 100 | 462.27 | 2694.1 | 23.67 | 23.5016
-2 | 0 | 100 | 1694.82 | 592255.7 | 68.88 | 68.7576 | -3 | 0 | 100 | 1917.02 | 24880.8 | 33.25 | 33.0402 | -3 | 0 | 100 | 646.48 | 3767.7 | 33.03 | 32.8668
-2.5 | 0 | 100 | 1142.05 | 399091.4 | 46.39 | 46.3321 | -2.5 | 0 | 100 | 2850.84 | 37000.9 | 49.32 | 49.1348 | -2.4 | 0 | 100 | 972.84 | 5669.7 | 49.49 | 49.4588
-3 | 0 | 100 | 662.79 | 231612.8 | 26.97 | 26.8889 | -2 | 0 | 100 | 3856.07 | 50047.6 | 66.9 | 66.4601 | -2 | 0 | 100 | 1226.65 | 7148.9 | 62.34 | 62.3624
-3.5 | 0 | 100 | 408.1 | 142610.5 | 16.59 | 16.5563 | -1.5 | 0 | 100 | 4777.49 | 62006.6 | 82.48 | 82.341 | -1.5 | 0 | 100 | 1542.3 | 8988.6 | 78.31 | 78.4099
-4 | 0 | 100 | 266.3 | 93057.5 | 10.84 | 10.8036 | -1 | 0 | 100 | 5432.54 | 70508.5 | 93.71 | 93.6309 | -1 | 0 | 100 | 1772.41 | 10329.6 | 90.31 | 90.1086
-4.5 | 0 | 100 | 175.73 | 61408.8 | 7.14 | 7.1292 | -0.5 | 0 | 100 | 5742.58 | 74532.4 | 99.12 | 98.9745 | -0.6 | 0 | 100 | 1922.48 | 11204.2 | 98.06 | 97.7381
-5 | 0 | 100 | 122.57 | 42830.6 | 4.99 | 4.9726 | 0 | 0 | 100 | 5802.08 | 75304.7 | 100.04 | 100 | -0.1 | 0 | 100 | 1966.97 | 11463.6 | 100.09 | 100
[7 empty] | 0.6 | 0 | 100 | 5644.8 | 73263.4 | 97.56 | 97.2892 | 0.5 | 0 | 100 | 1943.51 | 11326.8 | 99.18 | 98.8073
[7 empty] | 1 | 0 | 100 | 5287.35 | 68624.1 | 91.58 | 91.1285 | 1 | 0 | 100 | 1864.16 | 10864.3 | 94.52 | 94.7732
Measurement time: | 12/04/2023 18:06:57 | [5 empty] | 1.5 | 0 | 100 | 4471.65 | 58037.2 | 77.52 | 77.0698 | 1.5 | 0 | 100 | 1669.65 | 9730.7 | 85.05 | 84.8844
Radiation device: | Novalis TB Accelerator | [5 empty] | 2 | 0 | 100 | 3472.06 | 45063.6 | 59.96 | 59.8416 | 2 | 0 | 100 | 1369.06 | 7978.9 | 69.73 | 69.6025
Energy: | 6 MV (FFF) | [5 empty] | 2.5 | 0 | 100 | 2359.08 | 30618.3 | 40.7 | 40.6592 | 2.5 | 0 | 100 | 1044.55 | 6087.7 | 53.28 | 53.1045
Controller: | CCU | [5 empty] | 3 | 0 | 100 | 1521.5 | 19747.4 | 26.29 | 26.2234 | 2.9 | 0 | 100 | 735.46 | 4286.3 | 37.52 | 37.3905
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | 3.5 | 0 | 100 | 890.33 | 11555.5 | 15.36 | 15.345 | 3.4 | 0 | 100 | 504.2 | 2938.5 | 25.69 | 25.6333
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | 4 | 0 | 100 | 564.82 | 7330.7 | 9.78 | 9.7348 | 4.1 | 0 | 100 | 316.78 | 1846.2 | 16.16 | 16.105
Reference detector: | Stealth CI transmission |  | 45.1159 | 3 | 67.8397 | -2.5 | 4.5 | 0 | 100 | 401.14 | 5206.4 | 6.93 | 6.9137 | 4.5 | 0 | 100 | 253.98 | 1480.2 | 12.93 | 12.9122
SAD: | 1000 mm |  | 67.6325 | 2.4 | 47.4426 | -3 | 5 | 0 | 100 | 291.24 | 3780 | 5.03 | 5.0196 | 5 | 0 | 100 | 199.65 | 1163.6 | 10.18 | 10.1501
SSD: | 900 mm |  | 50 | 2.8699 | 50 | -2.9373 | 5.4 | 0 | 100 | 221.85 | 2879.4 | 3.83 | 3.8236 | 5.6 | 0 | 100 | 165.69 | 965.6 | 8.45 | 8.4236
Field size: | 5 x 5 mm | [5 empty] | 6 | 0 | 100 | 164.01 | 2128.7 | 2.84 | 2.8267 | 6 | 0 | 100 | 146.94 | 856.4 | 7.49 | 7.4704
Gantry angle: | 0 |  |  | 5.8072 |  |  | 6.5 | 0 | 100 | 131.5 | 1706.7 | 2.28 | 2.2664 | 6.5 | 0 | 100 | 122.02 | 711.1 | 6.23 | 6.2034
Measurement medium: | Water | [5 empty] | 7 | 0 | 100 | 103.39 | 1341.9 | 1.79 | 1.7819 | 7 | 0 | 100 | 108.46 | 632.1 | 5.54 | 5.5141
Scan type: | Crossline | [5 empty] | 7.5 | 0 | 100 | 84.03 | 1090.6 | 1.45 | 1.4483 | 7.5 | 0 | 100 | 106.72 | 622 | 5.46 | 5.4256
Scan mode: | StepByStep | [5 empty] | 8 | 0 | 100 | 71.77 | 931.5 | 1.24 | 1.237 | 8.1 | 0 | 100 | 87.62 | 510.6 | 4.49 | 4.4546
Wedge: | none | [5 empty] | 8.5 | 0 | 100 | 63.14 | 819.5 | 1.09 | 1.0882 | 8.5 | 0 | 100 | 100.73 | 587 | 5.14 | 5.1211
[7 empty] | 9 | 0 | 100 | 54.7 | 710 | 0.94 | 0.94277 | 9 | 0 | 100 | 91.84 | 535.2 | 4.69 | 4.6691
Points [mm]: | [6 empty] | 9.5 | 0 | 100 | 51.78 | 672 | 0.89 | 0.89244 | 9.5 | 0 | 100 | 93.77 | 546.5 | 4.78 | 4.7672
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | 10 | 0 | 100 | 41.03 | 532.6 | 0.71 | 0.70716 | 10 | 0 | 100 | 78.07 | 455 | 3.97 | 3.969
0 | 5 | 100 | 155.02 | 54172.8 | 6.3 | 6.2982
0 | 4.5 | 100 | 257.82 | 90094.1 | 10.49 | 10.4748
0 | 4 | 100 | 420.45 | 146926.6 | 17.11 | 17.0822 | Measurement time: | 27/05/2023 10:12:41 | [5 empty] | Measurement time: | 27/05/2023 12:27:28
0 | 3.5 | 100 | 693.21 | 242243.6 | 28.19 | 28.164 | Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator
0 | 3 | 100 | 1110.45 | 388048.1 | 45.16 | 45.1159 | Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF)
0 | 2.4 | 100 | 1664.66 | 581714.5 | 67.72 | 67.6325 | Controller: | CCU | [5 empty] | Controller: | CCU
0 | 1.9 | 100 | 2023.11 | 706976.4 | 82.3 | 82.1958 | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 1.4 | 100 | 2250.34 | 786383.1 | 91.46 | 91.4278 | Field detector: | PTW60019 Diamant |  | Côté G |  | Côté D |  | Field detector: | Razor Nano |  | Côté G |  | Côté D
0 | 1 | 100 | 2363.39 | 825886.7 | 95.95 | 96.0209 | Reference detector: | Stealth CI transmission |  | 39.0553 | -3 | 57.0673 | 3 | Reference detector: | Stealth CI transmission |  | 49.0731 | 3 | 64.9751 | -2.5
0 | 0.5 | 100 | 2428.48 | 848633.3 | 98.7 | 98.6654 | SAD: | 1000 mm |  | 54.4022 | -2.6 | 38.9817 | 3.5 | SAD: | 1000 mm |  | 66.4863 | 2.5 | 49.3608 | -3
0 | 0 | 100 | 2461.33 | 860112.3 | 100.06 | 100 | SSD: | 900 mm |  | 50 | -2.7147 | 50 | 3.1954 | SSD: | 900 mm |  | 50 | 2.9734 | 50 | -2.9795
0 | -0.5 | 100 | 2458.56 | 859144.7 | 99.84 | 99.8875 | Field size: | 5 x 5 mm | [5 empty] | Field size: | 5 x 5 mm
0 | -1 | 100 | 2416.3 | 844378.3 | 98.19 | 98.1705 | Gantry angle: | 0 |  |  | 5.9101 |  |  | Gantry angle: | 0 |  |  | 5.9529
0 | -1.5 | 100 | 2301.1 | 804120.1 | 93.6 | 93.4901 | Measurement medium: | Water | [5 empty] | Measurement medium: | Water
0 | -2 | 100 | 2074.6 | 724970.8 | 84.37 | 84.2878 | Scan type: | Crossline | [5 empty] | Scan type: | Crossline
0 | -2.5 | 100 | 1669.76 | 583497.5 | 67.9 | 67.8397 | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep
0 | -3 | 100 | 1167.72 | 408061.8 | 47.51 | 47.4426 | Wedge: | none | [5 empty] | Wedge: | none
0 | -3.5 | 100 | 775.27 | 270919.4 | 31.54 | 31.498
0 | -4 | 100 | 502.35 | 175547.5 | 20.43 | 20.4097 | Points [mm]: | [6 empty] | Points [mm]:
0 | -4.5 | 100 | 322.33 | 112637.9 | 13.11 | 13.0958 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | -5 | 100 | 210.22 | 73462 | 8.55 | 8.5409 | 0 | -10 | 100 | 47.36 | 614.6 | 0.81 | 0.81641 | 0 | 10 | 100 | 75.67 | 441 | 3.84 | 3.84
[7 empty] | 0 | -9.5 | 100 | 56.05 | 727.5 | 0.96 | 0.96622 | 0 | 9.5 | 100 | 77.9 | 454 | 3.96 | 3.9532
[7 empty] | 0 | -9 | 100 | 63.16 | 819.8 | 1.08 | 1.0888 | 0 | 9 | 100 | 78.12 | 455.3 | 3.99 | 3.9643
[7 empty] | 0 | -8.5 | 100 | 70.65 | 917 | 1.21 | 1.2179 | 0 | 8.5 | 100 | 74.62 | 434.9 | 3.8 | 3.7867
[7 empty] | 0 | -8 | 100 | 85.68 | 1112.1 | 1.47 | 1.477 | 0 | 7.9 | 100 | 100.19 | 583.9 | 5.09 | 5.0843
[7 empty] | 0 | -7.4 | 100 | 112.71 | 1462.9 | 1.93 | 1.9429 | 0 | 7.5 | 100 | 97.9 | 570.6 | 4.98 | 4.9681
[7 empty] | 0 | -7 | 100 | 139.68 | 1812.9 | 2.39 | 2.4079 | 0 | 7 | 100 | 112.6 | 656.2 | 5.73 | 5.7141
[7 empty] | 0 | -6.5 | 100 | 181.19 | 2351.7 | 3.11 | 3.1234 | 0 | 6.5 | 100 | 119.43 | 696.1 | 6.08 | 6.0607
[7 empty] | 0 | -6 | 100 | 229.39 | 2977.2 | 3.93 | 3.9543 | 0 | 6 | 100 | 146.75 | 855.2 | 7.45 | 7.4471
[7 empty] | 0 | -5.5 | 100 | 326.53 | 4238 | 5.6 | 5.6289 | 0 | 5.6 | 100 | 189.09 | 1102 | 9.66 | 9.5957
[7 empty] | 0 | -5 | 100 | 462.81 | 6006.8 | 7.92 | 7.9781 | 0 | 5 | 100 | 246.84 | 1438.6 | 12.58 | 12.5263
[7 empty] | 0 | -4.5 | 100 | 680.4 | 8830.9 | 11.67 | 11.7291 | 0 | 4.5 | 100 | 325.98 | 1899.8 | 16.6 | 16.5424
[7 empty] | 0 | -4 | 100 | 1035.67 | 13441.9 | 17.81 | 17.8534 | 0 | 4 | 100 | 469.56 | 2736.6 | 23.87 | 23.8286
[7 empty] | 0 | -3.5 | 100 | 1525.9 | 19804.6 | 26.15 | 26.3042 | 0 | 3.6 | 100 | 692.01 | 4033 | 35.37 | 35.1173
[7 empty] | 0 | -3 | 100 | 2265.59 | 29405 | 38.92 | 39.0553 | 0 | 3 | 100 | 967.02 | 5635.8 | 49.32 | 49.0731
[7 empty] | 0 | -2.6 | 100 | 3155.86 | 40959.6 | 54.15 | 54.4022 | 0 | 2.5 | 100 | 1310.16 | 7635.6 | 66.68 | 66.4863
[7 empty] | 0 | -2 | 100 | 4157.88 | 53964.7 | 71.33 | 71.6755 | 0 | 2 | 100 | 1556.34 | 9070.4 | 79.11 | 78.9792
[7 empty] | 0 | -1.5 | 100 | 4896.35 | 63549.3 | 84.29 | 84.4056 | 0 | 1.4 | 100 | 1750.19 | 10200.1 | 88.92 | 88.8164
[7 empty] | 0 | -1 | 100 | 5431.68 | 70497.4 | 93.34 | 93.6338 | 0 | 1 | 100 | 1887.94 | 11002.9 | 95.96 | 95.8068
[7 empty] | 0 | -0.5 | 100 | 5701.08 | 73993.9 | 97.77 | 98.2779 | 0 | 0.5 | 100 | 1953.79 | 11386.7 | 99.12 | 99.1485
[7 empty] | 0 | 0 | 100 | 5800.98 | 75290.4 | 100.07 | 100 | 0 | -0.1 | 100 | 1970.57 | 11484.5 | 100.11 | 100
[7 empty] | 0 | 0.4 | 100 | 5781.65 | 75039.5 | 99.4 | 99.6668 | 0 | -0.5 | 100 | 1960.1 | 11423.5 | 99.79 | 99.4687
[7 empty] | 0 | 0.9 | 100 | 5632.02 | 73097.5 | 96.47 | 97.0874 | 0 | -1 | 100 | 1910.58 | 11134.9 | 97.13 | 96.9557
[7 empty] | 0 | 1.4 | 100 | 5375.74 | 69771.3 | 92.33 | 92.6695 | 0 | -1.5 | 100 | 1781.32 | 10381.6 | 90.54 | 90.3962
[7 empty] | 0 | 2 | 100 | 4877.15 | 63300.1 | 83.81 | 84.0746 | 0 | -2 | 100 | 1575.68 | 9183.1 | 79.66 | 79.9606
[7 empty] | 0 | 2.5 | 100 | 4090.82 | 53094.4 | 70.05 | 70.5195 | 0 | -2.5 | 100 | 1280.38 | 7462.1 | 64.94 | 64.9751
[7 empty] | 0 | 3 | 100 | 3310.46 | 42966.1 | 56.91 | 57.0673 | 0 | -3 | 100 | 972.69 | 5668.8 | 49.33 | 49.3608
[7 empty] | 0 | 3.5 | 100 | 2261.32 | 29349.5 | 38.91 | 38.9817 | 0 | -3.4 | 100 | 729.57 | 4252 | 37.14 | 37.0233
[7 empty] | 0 | 4 | 100 | 1530.51 | 19864.3 | 26.21 | 26.3836 | 0 | -3.9 | 100 | 527.41 | 3073.8 | 26.85 | 26.7643
[7 empty] | 0 | 4.6 | 100 | 869.64 | 11287 | 14.92 | 14.9913 | 0 | -4.5 | 100 | 357.37 | 2082.7 | 18.21 | 18.1354
[7 empty] | 0 | 5.1 | 100 | 526.86 | 6838 | 9.03 | 9.0823 | 0 | -4.9 | 100 | 274.4 | 1599.2 | 13.99 | 13.9249
[7 empty] | 0 | 5.5 | 100 | 377.07 | 4893.9 | 6.49 | 6.5001 | 0 | -5.5 | 100 | 199.57 | 1163.1 | 10.23 | 10.1275
[7 empty] | 0 | 5.9 | 100 | 277.16 | 3597.2 | 4.77 | 4.7778 | 0 | -6 | 100 | 164.18 | 956.8 | 8.37 | 8.3316
[7 empty] | 0 | 6.5 | 100 | 197.43 | 2562.4 | 3.39 | 3.4034 | 0 | -6.4 | 100 | 136.73 | 796.9 | 6.97 | 6.9386
[7 empty] | 0 | 6.9 | 100 | 146.27 | 1898.4 | 2.51 | 2.5215 | 0 | -7 | 100 | 118.98 | 693.4 | 6.06 | 6.0378
[7 empty] | 0 | 7.4 | 100 | 121.74 | 1580 | 2.09 | 2.0986 | 0 | -7.5 | 100 | 115.03 | 670.4 | 5.85 | 5.8374
[7 empty] | 0 | 7.9 | 100 | 96.84 | 1256.9 | 1.66 | 1.6694 | 0 | -8 | 100 | 105.07 | 612.3 | 5.34 | 5.332
[7 empty] | 0 | 8.5 | 100 | 80.78 | 1048.4 | 1.39 | 1.3925 | 0 | -8.5 | 100 | 103.98 | 606 | 5.3 | 5.2766
[7 empty] | 0 | 9 | 100 | 66.2 | 859.2 | 1.13 | 1.1412 | 0 | -9 | 100 | 85.52 | 498.4 | 4.35 | 4.3399
[7 empty] | 0 | 9.5 | 100 | 60.41 | 784.1 | 1.03 | 1.0414 | 0 | -9.5 | 100 | 88.69 | 516.9 | 4.52 | 4.5007
[7 empty] | 0 | 10 | 100 | 54.14 | 702.7 | 0.93 | 0.93329 | 0 | -10 | 100 | 79.93 | 465.8 | 4.07 | 4.0562

## 1x1
Measurement time: | 12/04/2023 18:28:42 | [5 empty] | Measurement time: | 12/04/2023 18:35:42 | [5 empty] | Measurement time: | 12/04/2023 18:39:08 | [5 empty] | Measurement time: | 12/04/2023 19:58:03 | [5 empty] | Measurement time: | 27/05/2023 10:34:11 | [5 empty] | Measurement time: | 27/05/2023 12:42:01
Radiation device: | Novalis TB Accelerator |  |  | Taille de champ | p.120 |  | Radiation device: | Novalis TB Accelerator |  |  | Taille de champ |  |  | Radiation device: | Novalis TB Accelerator |  |  | Taille de champ |  |  | Radiation device: | Novalis TB Accelerator |  | Taille de champ |  |  |  | Radiation device: | Novalis TB Accelerator |  |  | Taille de champ |  |  | Radiation device: | Novalis TB Accelerator |  |  | Taille de champ
Energy: | 6 MV (FFF) |  |  | 10.0772 |  |  | Energy: | 6 MV (FFF) |  |  | 10.0432 |  |  | Energy: | 6 MV |  |  | 10.0626 |  |  | Energy: | 6 MV (FFF) |  | 9.6866 |  |  |  | Energy: | 6 MV (FFF) |  |  | 10.0613 |  |  | Energy: | 6 MV (FFF) |  |  | 10.1858
Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | PTW60019 Diamant | [5 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission |  | 30.9153 | -5.5 | 51.9784 | 4.5 | Reference detector: | Stealth CI transmission |  | 30.5721 | 5 | 50.3784 | -5 | Reference detector: | Stealth CI transmission |  | 34.7497 | -5.5 | 62.5495 | 4 | Reference detector: | <none> |  | 29.3672 | 5.1 | 51.43 | -5 | Reference detector: | Stealth CI transmission | [5 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  | 50.4796 | -5 | 32.0843 | 5 | SAD: | 1000 mm |  | 66.5019 | 4 | 20.9868 | -5.9 | SAD: | 1000 mm |  | 56.2477 | -5 | 40.0432 | 4.6 | SAD: | 1000 mm |  | 61.3254 | 4.5 | 28.2511 | -5.4 | SAD: | 1000 mm |  | Côté G |  | Côté D |  | SAD: | 1000 mm |  | Côté G |  | Côté D
SSD: | 900 mm |  | 50 | -5.0123 | 50 | 4.5497 | SSD: | 900 mm |  | 50 | 4.4593 | 50 | -5.0116 | SSD: | 900 mm |  | 50 | -5.1453 | 50 | 4.3346 | SSD: | 900 mm |  | 50 | 4.7126 | 50 | -5.0247 | SSD: | 900 mm |  | 49.1745 | 4.6 | 64.4831 | -4.5 | SSD: | 900 mm |  | 42.3742 | -5 | 62.154 | 4.5
Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm |  | 64.1631 | 4.1 | 48.5138 | -5 | Field size: | 10 x 10 mm |  | 56.1622 | -4.5 | 47.9396 | 5
Gantry angle: | 0 |  |  | 9.562 |  |  | Gantry angle: | 0 |  |  | 9.4709 |  |  | Gantry angle: | 0 |  |  | 9.4799 |  |  | Gantry angle: | 0 | Cône 10 mm |  | 9.7373 |  |  | Gantry angle: | 0 |  | 50 | 4.5725 | 50 | -4.9535 | Gantry angle: | 0 |  | 50 | -4.7235 | 50 | 4.9275
Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water
Scan type: | Inline | [5 empty] | Scan type: | Inline | [5 empty] | Scan type: | Inline | [5 empty] | Scan type: | Inline | [5 empty] | Scan type: | Inline |  |  | 9.5259 |  |  | Scan type: | Inline |  |  | 9.651
Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | 0,1 cm |  |  |  |  | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep
Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
-10 | 0 | 100 | 2.24 | 23220.6 | 2.24 | 2.2382 | 10 | 0 | 100 | 1.77 | 18383.9 | 1.78 | 1.7859 | -10 | 0 | 100 | 0.34 | 3551.6 | 1.38 | 1.2243 | 10 | 0 | 100 | 2.08 | 21021.6 | 0 | 2.0727 | 20 | 0 | 100 | 63.26 | 534.4 | 0.58 | 0.5773 | -20 | 0 | 100 | 68.08 | 396.8 | 2.71 | 2.7106
-9.5 | 0 | 100 | 2.81 | 29134.9 | 2.81 | 2.8078 | 9 | 0 | 100 | 2.71 | 28073.6 | 2.72 | 2.7343 | -9.5 | 0 | 100 | 0.55 | 5677.7 | 2.21 | 1.9806 | 9.5 | 0 | 100 | 2.41 | 24354.7 | 0 | 2.4016 | 19.5 | 0 | 100 | 64.68 | 546.3 | 0.59 | 0.59026 | -19.5 | 0 | 100 | 68.68 | 400.3 | 2.73 | 2.7345
-8.9 | 0 | 100 | 3.55 | 36812.3 | 3.55 | 3.5472 | 8 | 0 | 100 | 4.2 | 43558.2 | 4.22 | 4.2377 | -9 | 0 | 100 | 0.8 | 8282.9 | 3.23 | 2.8808 | 9 | 0 | 100 | 2.83 | 28669.4 | 0 | 2.8201 | 19 | 0 | 100 | 71.45 | 603.5 | 0.65 | 0.65204 | -18.9 | 0 | 100 | 86.72 | 505.4 | 3.45 | 3.4528
-8.5 | 0 | 100 | 4.44 | 46032.2 | 4.44 | 4.4365 | 7 | 0 | 100 | 6.88 | 71303.6 | 6.91 | 6.9418 | -8.5 | 0 | 100 | 1.15 | 11867.7 | 4.62 | 4.1412 | 8.5 | 0 | 100 | 3.35 | 33941.4 | 0 | 3.3383 | 18.5 | 0 | 100 | 67.96 | 574.1 | 0.62 | 0.62019 | -18.5 | 0 | 100 | 100.72 | 587 | 4.01 | 4.0102
-8 | 0 | 100 | 5.67 | 58807.9 | 5.68 | 5.6655 | 6 | 0 | 100 | 12.86 | 133278.1 | 12.91 | 12.9755 | -8 | 0 | 100 | 1.57 | 16311.2 | 6.35 | 5.6536 | 8 | 0 | 100 | 4.05 | 40970.2 | 0 | 4.0359 | 18 | 0 | 100 | 78.11 | 659.8 | 0.71 | 0.71282 | -18 | 0 | 100 | 97.29 | 567 | 3.87 | 3.8736
-7.5 | 0 | 100 | 7.28 | 75471.7 | 7.29 | 7.2742 | 5 | 0 | 100 | 30.3 | 313993.6 | 30.41 | 30.5721 | -7.5 | 0 | 100 | 2.19 | 22670.9 | 8.83 | 7.8862 | 7.6 | 0 | 100 | 4.91 | 49654.4 | 0 | 4.8929 | 17.5 | 0 | 100 | 80.81 | 682.6 | 0.74 | 0.73746 | -17.5 | 0 | 100 | 89.89 | 523.9 | 3.58 | 3.579
-7 | 0 | 100 | 9.78 | 101327.9 | 9.78 | 9.7722 | 4 | 0 | 100 | 65.91 | 683095.6 | 66.14 | 66.5019 | -7 | 0 | 100 | 2.98 | 30891.1 | 12.03 | 10.731 | 7 | 0 | 100 | 6.42 | 65004.7 | 0 | 6.3976 | 17 | 0 | 100 | 82.8 | 699.4 | 0.76 | 0.75562 | -17 | 0 | 100 | 92.52 | 539.2 | 3.68 | 3.6837
-6.5 | 0 | 100 | 13.43 | 139196.1 | 13.44 | 13.4193 | 3 | 0 | 100 | 88 | 912065.8 | 88.31 | 88.7902 | -6.6 | 0 | 100 | 4.25 | 44065 | 17.15 | 15.3043 | 6.5 | 0 | 100 | 8.27 | 83677 | 0 | 8.2412 | 16.5 | 0 | 100 | 91.61 | 773.8 | 0.84 | 0.83602 | -16.5 | 0 | 100 | 99.41 | 579.4 | 3.96 | 3.958
-5.9 | 0 | 100 | 20.42 | 211640.3 | 20.44 | 20.4037 | 2 | 0 | 100 | 95.15 | 986163.3 | 95.52 | 96.0044 | -6 | 0 | 100 | 6.4 | 66341.8 | 25.87 | 23.0465 | 6.1 | 0 | 100 | 11.06 | 111900.2 | 0 | 11.0214 | 16 | 0 | 100 | 88.95 | 751.4 | 0.81 | 0.81175 | -16 | 0 | 100 | 99.82 | 581.7 | 3.97 | 3.9744
-5.5 | 0 | 100 | 30.94 | 320728.2 | 30.98 | 30.9153 | 1 | 0 | 100 | 98.08 | 1016582.5 | 98.46 | 98.9608 | -5.5 | 0 | 100 | 9.65 | 100055.6 | 38.95 | 34.7497 | 5.5 | 0 | 100 | 17.15 | 173530.5 | 0 | 17.0902 | 15.5 | 0 | 100 | 97.15 | 820.6 | 0.89 | 0.88658 | -15.6 | 0 | 100 | 87.61 | 510.6 | 3.48 | 3.4882
-5 | 0 | 100 | 50.52 | 523631.7 | 50.58 | 50.4796 | 0 | 0 | 100 | 99.11 | 1027179.5 | 99.49 | 100 | -5 | 0 | 100 | 15.62 | 161850.1 | 62.96 | 56.2477 | 5.1 | 0 | 100 | 29.47 | 298161.4 | 0 | 29.3672 | 15 | 0 | 100 | 98.78 | 834.4 | 0.9 | 0.90145 | -14.9 | 0 | 100 | 68.46 | 399 | 2.72 | 2.7258
-4.5 | 0 | 100 | 70.45 | 730144.2 | 70.51 | 70.3937 | -1 | 0 | 100 | 98.93 | 1025349.3 | 99.36 | 99.8184 | -4.5 | 0 | 100 | 20.78 | 215368.5 | 83.97 | 74.829 | 4.5 | 0 | 100 | 61.54 | 622620 | 0 | 61.3254 | 14.5 | 0 | 100 | 105.25 | 889.1 | 0.96 | 0.9605 | -14.4 | 0 | 100 | 68.27 | 397.9 | 2.72 | 2.7182
-4 | 0 | 100 | 84.37 | 874466.9 | 84.41 | 84.3026 | -2 | 0 | 100 | 97.44 | 1009891.6 | 97.81 | 98.315 | -4 | 0 | 100 | 23.58 | 244412 | 95.27 | 84.9118 | 4 | 0 | 100 | 80.5 | 814487.6 | 0 | 80.2192 | 14 | 0 | 100 | 106.81 | 902.2 | 0.97 | 0.97473 | -14.1 | 0 | 100 | 65.13 | 379.6 | 2.59 | 2.5932
-3.5 | 0 | 100 | 90.47 | 937651 | 90.61 | 90.3977 | -3 | 0 | 100 | 93.95 | 973793.2 | 94.32 | 94.7937 | -3.5 | 0 | 100 | 25.26 | 261765.8 | 101.93 | 90.9615 | 3.5 | 0 | 100 | 87.7 | 887289.9 | 0 | 87.3941 | 13.5 | 0 | 100 | 122.44 | 1034.3 | 1.12 | 1.1174 | -13.4 | 0 | 100 | 72.68 | 423.6 | 2.89 | 2.8938
-3 | 0 | 100 | 94.39 | 978284.7 | 94.52 | 94.3145 | -4 | 0 | 100 | 83.59 | 866420.1 | 83.96 | 84.3406 | -3 | 0 | 100 | 26.29 | 272445.3 | 106.12 | 94.6705 | 3 | 0 | 100 | 92.19 | 932793.8 | 0 | 91.8685 | 13 | 0 | 100 | 124.8 | 1054.2 | 1.14 | 1.1389 | -13 | 0 | 100 | 75.06 | 437.4 | 2.99 | 2.9885
-2.5 | 0 | 100 | 96.56 | 1000806.1 | 96.67 | 96.4828 | -5 | 0 | 100 | 49.93 | 517509.1 | 50.12 | 50.3784 | -2.5 | 0 | 100 | 26.92 | 278971.3 | 108.63 | 96.9391 | 2.5 | 0 | 100 | 94.94 | 960575.2 | 0 | 94.6089 | 12.5 | 0 | 100 | 137.17 | 1158.6 | 1.25 | 1.2518 | -12.6 | 0 | 100 | 71.7 | 417.9 | 2.85 | 2.8548
-2 | 0 | 100 | 98.26 | 1018395.3 | 98.35 | 98.1815 | -5.9 | 0 | 100 | 20.8 | 215631.1 | 20.88 | 20.9868 | -2 | 0 | 100 | 27.36 | 283605.8 | 110.43 | 98.5236 | 2.1 | 0 | 100 | 96.7 | 978414.8 | 0 | 96.3627 | 12 | 0 | 100 | 161.03 | 1360.3 | 1.47 | 1.4695 | -11.9 | 0 | 100 | 71.13 | 414.6 | 2.83 | 2.8321
-1.5 | 0 | 100 | 99.17 | 1027821.6 | 99.26 | 99.0907 | -7 | 0 | 100 | 9.99 | 103510.4 | 10.03 | 10.0797 | -1.5 | 0 | 100 | 27.52 | 285203.6 | 111.08 | 99.0997 | 1.5 | 0 | 100 | 98.3 | 994599.5 | 0 | 97.9571 | 11.5 | 0 | 100 | 177.43 | 1498.7 | 1.62 | 1.6192 | -11.5 | 0 | 100 | 67.82 | 395.2 | 2.7 | 2.7003
-1 | 0 | 100 | 99.69 | 1033220 | 99.81 | 99.6103 | -8 | 0 | 100 | 5.77 | 59838.5 | 5.8 | 5.8218 | -1 | 0 | 100 | 27.75 | 287628.6 | 112 | 99.928 | 1 | 0 | 100 | 99.02 | 1001844.1 | 0 | 98.6746 | 11 | 0 | 100 | 192.47 | 1625.8 | 1.76 | 1.7565 | -11 | 0 | 100 | 69.82 | 406.9 | 2.78 | 2.7799
-0.4 | 0 | 100 | 100.08 | 1037245.6 | 100.17 | 100 | -9 | 0 | 100 | 3.58 | 37092.4 | 3.59 | 3.6121 | -0.6 | 0 | 100 | 27.77 | 287860.7 | 112.2 | 100 | 0.5 | 0 | 100 | 99.66 | 1008313.5 | 0 | 99.3124 | 10.5 | 0 | 100 | 230.15 | 1944.1 | 2.1 | 2.1003 | -10.5 | 0 | 100 | 64.5 | 375.9 | 2.57 | 2.5681
0 | 0 | 100 | 99.94 | 1035803.1 | 100.05 | 99.8601 | -10 | 0 | 100 | 2.38 | 24667.3 | 2.39 | 2.4014 | -0.1 | 0 | 100 | 27.66 | 286654.5 | 111.73 | 99.6039 | 0 | 0 | 100 | 99.98 | 1011554.7 | 0 | 99.6313 | 10 | 0 | 100 | 254.87 | 2152.9 | 2.33 | 2.3259 | -10 | 0 | 100 | 85.06 | 495.7 | 3.39 | 3.3867
0.5 | 0 | 100 | 99.41 | 1030325.9 | 99.5 | 99.3305 | [7 empty] | 0.5 | 0 | 100 | 27.58 | 285876.2 | 111.35 | 99.3158 | -0.5 | 0 | 100 | 100.35 | 1015283.7 | 0 | 100 | 9.5 | 0 | 100 | 299.51 | 2530 | 2.74 | 2.7333 | -9.5 | 0 | 100 | 101.12 | 589.3 | 4.03 | 4.0261
1 | 0 | 100 | 99.05 | 1026603.6 | 99.16 | 98.9708 | [7 empty] | 1 | 0 | 100 | 27.4 | 284036.8 | 110.65 | 98.6676 | -0.9 | 0 | 100 | 99.6 | 1007688.9 | 0 | 99.2526 | 9 | 0 | 100 | 354.01 | 2990.3 | 3.23 | 3.2306 | -9 | 0 | 100 | 115.62 | 673.9 | 4.6 | 4.6034
1.5 | 0 | 100 | 97.91 | 1014791.9 | 98.11 | 97.8317 | [7 empty] | 1.5 | 0 | 100 | 27.15 | 281404.3 | 109.53 | 97.7674 | -1.5 | 0 | 100 | 99.2 | 1003647.2 | 0 | 98.854 | 8.5 | 0 | 100 | 429.12 | 3624.9 | 3.92 | 3.9161 | -8.5 | 0 | 100 | 138.34 | 806.2 | 5.51 | 5.508
2 | 0 | 100 | 96.17 | 996789.1 | 96.3 | 96.0931 | [7 empty] | 1.9 | 0 | 100 | 26.44 | 274023 | 106.77 | 95.2107 | -2.1 | 0 | 100 | 97.82 | 989755.4 | 0 | 97.4788 | 8 | 0 | 100 | 524.09 | 4427 | 4.79 | 4.7828 | -8 | 0 | 100 | 176.72 | 1029.9 | 7.04 | 7.0362
2.5 | 0 | 100 | 93.43 | 968307.9 | 93.5 | 93.3553 | [7 empty] | 2.6 | 0 | 100 | 25.31 | 262284.8 | 102.08 | 91.1415 | -2.5 | 0 | 100 | 96.47 | 976078.7 | 0 | 96.1335 | 7.5 | 0 | 100 | 660.51 | 5579.4 | 6.03 | 6.0277 | -7.5 | 0 | 100 | 217.48 | 1267.5 | 8.65 | 8.659
3 | 0 | 100 | 89.17 | 924157.1 | 89.28 | 89.0987 | [7 empty] | 3 | 0 | 100 | 23.78 | 246517.9 | 96.06 | 85.632 | -3 | 0 | 100 | 94.39 | 955021.1 | 0 | 94.0608 | 7 | 0 | 100 | 874.68 | 7388.5 | 7.99 | 7.9822 | -7.1 | 0 | 100 | 279.75 | 1630.4 | 11.15 | 11.1383
3.5 | 0 | 100 | 82.73 | 857417.4 | 82.82 | 82.6639 | [7 empty] | 3.5 | 0 | 100 | 21.58 | 223673.3 | 87.03 | 77.7098 | -3.4 | 0 | 100 | 91.64 | 927204 | 0 | 91.3204 | 6.5 | 0 | 100 | 1165.61 | 9846 | 10.66 | 10.6372 | -6.5 | 0 | 100 | 381.89 | 2225.7 | 15.2 | 15.205
4 | 0 | 100 | 70.87 | 734548.8 | 70.96 | 70.8133 | [7 empty] | 4 | 0 | 100 | 17.37 | 180072.1 | 70.14 | 62.5495 | -4 | 0 | 100 | 86.23 | 872471.4 | 0 | 85.9292 | 6 | 0 | 100 | 1597.97 | 13498.2 | 14.61 | 14.5829 | -6 | 0 | 100 | 537.68 | 3133.6 | 21.41 | 21.4079
4.5 | 0 | 100 | 52.02 | 539187.9 | 52.1 | 51.9784 | [7 empty] | 4.6 | 0 | 100 | 11.12 | 115299 | 44.95 | 40.0432 | -4.5 | 0 | 100 | 77.03 | 779381.5 | 0 | 76.7613 | 5.5 | 0 | 100 | 2427.12 | 20502.1 | 22.17 | 22.1496 | -5.4 | 0 | 100 | 829.84 | 4836.3 | 33.08 | 33.0403
5 | 0 | 100 | 32.11 | 332774.7 | 32.15 | 32.0843 | [7 empty] | 5 | 0 | 100 | 7.58 | 78542.2 | 30.57 | 27.2956 | -5 | 0 | 100 | 51.61 | 522224.6 | 0 | 51.43 | 5 | 0 | 100 | 3772.34 | 31865.3 | 34.46 | 34.4259 | -5 | 0 | 100 | 1064.27 | 6202.6 | 42.34 | 42.3742
5.5 | 0 | 100 | 20.66 | 214162.9 | 20.69 | 20.6435 | [7 empty] | 5.5 | 0 | 100 | 4.95 | 51330.2 | 19.98 | 17.825 | -5.4 | 0 | 100 | 28.35 | 286861.8 | 0 | 28.2511 | 4.6 | 0 | 100 | 5388.47 | 45516.9 | 49.25 | 49.1745 | -4.5 | 0 | 100 | 1410.57 | 8220.8 | 56.2 | 56.1622
6 | 0 | 100 | 13.29 | 137747.3 | 13.31 | 13.2794 | [7 empty] | 6 | 0 | 100 | 3.36 | 34785.7 | 13.55 | 12.0994 | -5.9 | 0 | 100 | 16.44 | 166380.5 | 0 | 16.3827 | 4.1 | 0 | 100 | 7030.9 | 59390.6 | 64.27 | 64.1631 | -3.9 | 0 | 100 | 1844.77 | 10751.4 | 73.45 | 73.45
6.5 | 0 | 100 | 9.52 | 98680.8 | 9.53 | 9.5124 | [7 empty] | 6.5 | 0 | 100 | 2.31 | 23896.9 | 9.3 | 8.3183 | -6.5 | 0 | 100 | 11.19 | 113182.6 | 0 | 11.151 | 3.5 | 0 | 100 | 8644.18 | 73018.2 | 78.97 | 78.8857 | -3.5 | 0 | 100 | 2088.23 | 12170.3 | 83.19 | 83.1434
7 | 0 | 100 | 7.04 | 72998.1 | 7.05 | 7.0344 | [7 empty] | 6.9 | 0 | 100 | 1.71 | 17711.5 | 6.9 | 6.1577 | -7 | 0 | 100 | 8.09 | 81868.3 | 0 | 8.0618 | 3 | 0 | 100 | 9572.27 | 80857.8 | 87.47 | 87.3553 | -3 | 0 | 100 | 2249.88 | 13112.3 | 89.56 | 89.5796
7.4 | 0 | 100 | 5.49 | 56952.5 | 5.5 | 5.4856 | [7 empty] | 7.5 | 0 | 100 | 1.2 | 12434.1 | 4.84 | 4.3212 | -7.6 | 0 | 100 | 6.03 | 60997.7 | 0 | 6.009 | 2.6 | 0 | 100 | 10165.77 | 85871.1 | 92.86 | 92.7715 | -2.5 | 0 | 100 | 2356.71 | 13734.9 | 93.88 | 93.833
8 | 0 | 100 | 4.29 | 44483.7 | 4.3 | 4.2866 | [7 empty] | 8 | 0 | 100 | 0.81 | 8351.5 | 3.25 | 2.9168 | -8 | 0 | 100 | 4.89 | 49456.3 | 0 | 4.8729 | 2 | 0 | 100 | 10535.72 | 88996.2 | 96.36 | 96.1476 | -2.1 | 0 | 100 | 2422.55 | 14118.7 | 96.47 | 96.4545
8.5 | 0 | 100 | 3.44 | 35604.4 | 3.44 | 3.4373 | [7 empty] | 8.5 | 0 | 100 | 0.57 | 5906.6 | 2.3 | 2.0526 | -8.5 | 0 | 100 | 4.01 | 40607.1 | 0 | 3.996 | 1.5 | 0 | 100 | 10759.03 | 90882.5 | 98.43 | 98.1855 | -1.5 | 0 | 100 | 2473.1 | 14413.3 | 98.51 | 98.4671
9 | 0 | 100 | 2.71 | 28130.3 | 2.72 | 2.7078 | [7 empty] | 9 | 0 | 100 | 0.37 | 3851.9 | 1.5 | 1.3324 | -9 | 0 | 100 | 3.32 | 33610.5 | 0 | 3.3084 | 1 | 0 | 100 | 10868.21 | 91804.7 | 99.24 | 99.1819 | -1.1 | 0 | 100 | 2484.57 | 14480.1 | 98.96 | 98.9238
9.5 | 0 | 100 | 2.18 | 22588.5 | 2.18 | 2.1783 | [7 empty] | 9.5 | 0 | 100 | 0.22 | 2231.3 | 0.87 | 0.79222 | -9.5 | 0 | 100 | 2.84 | 28699.8 | 0 | 2.8301 | 0.6 | 0 | 100 | 10894.97 | 92030.8 | 99.55 | 99.4261 | -0.5 | 0 | 100 | 2491.96 | 14523.2 | 99.34 | 99.218
10 | 0 | 100 | 1.76 | 18254.1 | 1.76 | 1.7586 | [7 empty] | 10 | 0 | 100 | 0.08 | 815.2 | 0.32 | 0.28808 | -10 | 0 | 100 | 2.45 | 24823.8 | 0 | 2.4415 | -0.1 | 0 | 100 | 10956.36 | 92549.4 | 100.07 | 99.9863 | -0.1 | 0 | 100 | 2511.6 | 14637.6 | 100.03 | 100
[28 empty] | -0.5 | 0 | 100 | 10957.86 | 92562 | 100.01 | 100 | 0.5 | 0 | 100 | 2503.23 | 14588.9 | 99.74 | 99.6667
[28 empty] | -1 | 0 | 100 | 10903.44 | 92102.3 | 99.58 | 99.5034 | 1 | 0 | 100 | 2502.61 | 14585.3 | 99.59 | 99.6421
Measurement time: | 12/04/2023 18:30:13 | [5 empty] | Measurement time: | 12/04/2023 18:36:37 | [5 empty] | Measurement time: | 12/04/2023 18:40:43 | [5 empty] | Measurement time: | 12/04/2023 19:59:40 | [5 empty] | -1.5 | 0 | 100 | 10821.78 | 91412.6 | 98.9 | 98.7582 | 1.4 | 0 | 100 | 2481.04 | 14459.6 | 98.74 | 98.7832
Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator | [5 empty] | -2 | 0 | 100 | 10709.65 | 90465.4 | 97.82 | 97.7349 | 2 | 0 | 100 | 2431.28 | 14169.5 | 96.92 | 96.802
Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV | [5 empty] | Energy: | 6 MV (FFF) | [5 empty] | -2.5 | 0 | 100 | 10512.66 | 88801.4 | 96.14 | 95.9372 | 2.4 | 0 | 100 | 2373.33 | 13831.8 | 94.53 | 94.4947
Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | -2.9 | 0 | 100 | 10213.84 | 86277.2 | 93.29 | 93.2102 | 2.9 | 0 | 100 | 2277.98 | 13276.1 | 90.74 | 90.6984
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | -3.5 | 0 | 100 | 9658.49 | 81586.1 | 88.19 | 88.1421 | 3.5 | 0 | 100 | 2092.99 | 12198 | 83.35 | 83.3329
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | SFD parallele |  | Côté G |  | Côté D |  | -4 | 0 | 100 | 8519.2 | 71962.4 | 77.85 | 77.7451 | 4 | 0 | 100 | 1867.59 | 10884.3 | 74.41 | 74.3586
Reference detector: | Stealth CI transmission |  | 43.327 | -5.5 | 62.7695 | 5 | Reference detector: | Stealth CI transmission |  | 26.8986 | 6 | 61.1791 | -5 | Reference detector: | Stealth CI transmission |  | 43.9885 | -5.5 | 61.6271 | 5 | Reference detector: | <none> |  | 38.7161 | -5 | 70.6129 | 4.5 | -4.5 | 0 | 100 | 7065.97 | 59686.9 | 64.54 | 64.4831 | 4.5 | 0 | 100 | 1561.06 | 9097.9 | 62.16 | 62.154
SAD: | 1000 mm |  | 61.7467 | -5 | 45.8137 | 5.4 | SAD: | 1000 mm |  | 60.5218 | 5 | 31.3379 | -5.9 | SAD: | 1000 mm |  | 61.4831 | -4.9 | 46.6163 | 5.5 | SAD: | 1000 mm |  | 63.1237 | -4.5 | 42.5557 | 5 | -5 | 0 | 100 | 5316.07 | 44905.3 | 48.51 | 48.5138 | 5 | 0 | 100 | 1204.05 | 7017.2 | 47.95 | 47.9396
SSD: | 900 mm |  | 50 | -5.3189 | 50 | 5.3012 | SSD: | 900 mm |  | 50 | 5.3129 | 50 | -5.3372 | SSD: | 900 mm |  | 50 | -5.2938 | 50 | 5.3873 | SSD: | 900 mm |  | 50 | -4.7688 | 50 | 4.8673 | -5.5 | 0 | 100 | 3759.7 | 31758.5 | 34.33 | 34.3105 | 5.5 | 0 | 100 | 851.93 | 4965.1 | 33.94 | 33.9198
Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm | [5 empty] | Field size: | 10 x 10 mm | [5 empty] | -6 | 0 | 100 | 2394.97 | 20230.5 | 21.87 | 21.8562 | 6 | 0 | 100 | 597.48 | 3482.1 | 23.82 | 23.7888
Gantry angle: | 0 |  |  | 10.6201 |  |  | Gantry angle: | 0 |  |  | 10.6501 |  |  | Gantry angle: | 0 |  |  | 10.6811 |  |  | Gantry angle: | 0 | Cône 10 mm |  | 9.6362 |  |  | -6.4 | 0 | 100 | 1625.12 | 13727.5 | 14.85 | 14.8306 | 6.4 | 0 | 100 | 451.95 | 2633.9 | 17.99 | 17.9945
Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | -7 | 0 | 100 | 1106.42 | 9346 | 10.11 | 10.097 | 7 | 0 | 100 | 331.31 | 1930.9 | 13.2 | 13.1912
Scan type: | Crossline | [5 empty] | Scan type: | Crossline | [5 empty] | Scan type: | Crossline | [5 empty] | Scan type: | Crossline | [5 empty] | -7.5 | 0 | 100 | 824.88 | 6967.8 | 7.53 | 7.5277 | 7.4 | 0 | 100 | 278.72 | 1624.4 | 11.09 | 11.0973
Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | 0,1 cm |  |  |  |  | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | [5 empty] | -8 | 0 | 100 | 638.47 | 5393.2 | 5.84 | 5.8266 | 7.9 | 0 | 100 | 227.72 | 1327.1 | 9.08 | 9.0667
Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | -8.5 | 0 | 100 | 503.4 | 4252.3 | 4.6 | 4.594 | 8.5 | 0 | 100 | 200.23 | 1166.9 | 7.98 | 7.9722
[28 empty] | -8.9 | 0 | 100 | 418.02 | 3531.1 | 3.82 | 3.8148 | 9 | 0 | 100 | 174.38 | 1016.3 | 6.95 | 6.943
Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | -9.5 | 0 | 100 | 336.97 | 2846.4 | 3.08 | 3.0751 | 9.5 | 0 | 100 | 168.3 | 980.8 | 6.71 | 6.7009
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | -10 | 0 | 100 | 287.8 | 2431.1 | 2.63 | 2.6264 | 10 | 0 | 100 | 128.87 | 751.1 | 5.13 | 5.131
0 | -10 | 100 | 2.55 | 26458.6 | 2.56 | 2.5569 | 0 | 10 | 100 | 2.4 | 24885.2 | 2.41 | 2.4269 | 0 | -10 | 100 | 0.37 | 3843.1 | 1.5 | 1.3319 | 0 | -10 | 100 | 2.24 | 22637.2 | 0 | 2.2398 | -10.4 | 0 | 100 | 250.41 | 2115.2 | 2.29 | 2.2852 | 10.5 | 0 | 100 | 141.43 | 824.3 | 5.63 | 5.6311
0 | -9.5 | 100 | 3.22 | 33400.9 | 3.23 | 3.2287 | 0 | 9 | 100 | 3.8 | 39376.5 | 3.82 | 3.8427 | 0 | -9.6 | 100 | 0.59 | 6131.6 | 2.39 | 2.1238 | 0 | -9.5 | 100 | 2.67 | 26990 | 0 | 2.6697 | -11 | 0 | 100 | 210.03 | 1774.1 | 1.92 | 1.9167 | 11 | 0 | 100 | 117.41 | 684.3 | 4.68 | 4.6747
0 | -9 | 100 | 3.97 | 41189.2 | 3.98 | 3.9807 | 0 | 8.1 | 100 | 6.17 | 63977.9 | 6.2 | 6.2393 | 0 | -9 | 100 | 0.86 | 8874.8 | 3.46 | 3.0958 | 0 | -9 | 100 | 3.1 | 31379.2 | 0 | 3.0997 | -11.5 | 0 | 100 | 181.01 | 1529 | 1.65 | 1.6519 | 11.6 | 0 | 100 | 106.33 | 619.7 | 4.24 | 4.2336
0 | -8.5 | 100 | 5.28 | 54759.5 | 5.29 | 5.2943 | 0 | 7.1 | 100 | 11.49 | 119091.2 | 11.54 | 11.619 | 0 | -8.5 | 100 | 1.3 | 13454.4 | 5.24 | 4.6796 | 0 | -8.5 | 100 | 3.7 | 37450.4 | 0 | 3.6996 | -12 | 0 | 100 | 164.6 | 1390.4 | 1.51 | 1.5021 | 12 | 0 | 100 | 104.41 | 608.5 | 4.17 | 4.1571
0 | -8 | 100 | 7.17 | 74319 | 7.18 | 7.1894 | 0 | 6 | 100 | 26.6 | 275692.9 | 26.72 | 26.8986 | 0 | -8 | 100 | 1.83 | 19002.7 | 7.41 | 6.5875 | 0 | -7.9 | 100 | 4.75 | 48051.9 | 0 | 4.7495 | -12.5 | 0 | 100 | 153.91 | 1300.1 | 1.41 | 1.4046 | 12.5 | 0 | 100 | 101.18 | 589.7 | 4.03 | 4.0285
0 | -7.5 | 100 | 9.63 | 99804 | 9.65 | 9.6561 | 0 | 5 | 100 | 59.85 | 620296.5 | 60.07 | 60.5218 | 0 | -7.5 | 100 | 2.63 | 27299 | 10.62 | 9.4672 | 0 | -7.5 | 100 | 5.59 | 56554.4 | 0 | 5.5894 | -12.9 | 0 | 100 | 140.96 | 1190.7 | 1.28 | 1.2864 | 13 | 0 | 100 | 105.82 | 616.7 | 4.21 | 4.2133
0 | -7 | 100 | 13.33 | 138150.2 | 13.35 | 13.3661 | 0 | 4 | 100 | 83.11 | 861434.2 | 83.45 | 84.0429 | 0 | -7.1 | 100 | 3.54 | 36731.7 | 14.31 | 12.743 | 0 | -7 | 100 | 7.14 | 72245.7 | 0 | 7.1393 | -13.5 | 0 | 100 | 130.4 | 1101.5 | 1.19 | 1.19 | 13.4 | 0 | 100 | 121.43 | 707.7 | 4.84 | 4.8348
0 | -6.4 | 100 | 20.27 | 210139.7 | 20.31 | 20.3249 | 0 | 3 | 100 | 92.29 | 956592.1 | 92.64 | 93.3259 | 0 | -6.5 | 100 | 5.56 | 57589.2 | 22.4 | 20.0144 | 0 | -6.4 | 100 | 10.04 | 101547.8 | 0 | 10.039 | -14 | 0 | 100 | 115.96 | 979.6 | 1.06 | 1.0582 | 14 | 0 | 100 | 101.43 | 591.2 | 4.04 | 4.0385
0 | -6 | 100 | 29.43 | 305029.8 | 29.5 | 29.5097 | 0 | 2.1 | 100 | 96.87 | 1003967 | 97.33 | 97.9573 | 0 | -6 | 100 | 8.38 | 86815.9 | 33.79 | 30.1656 | 0 | -5.9 | 100 | 14.01 | 141680 | 0 | 14.0086 | -14.4 | 0 | 100 | 113.16 | 955.8 | 1.03 | 1.0327 | 14.4 | 0 | 100 | 97.33 | 567.2 | 3.88 | 3.8752
0 | -5.5 | 100 | 43.21 | 447869.2 | 43.28 | 43.327 | 0 | 1.1 | 100 | 98.55 | 1021391.3 | 98.96 | 99.6562 | 0 | -5.5 | 100 | 12.22 | 126641.7 | 49.28 | 43.9885 | 0 | -5.5 | 100 | 22.09 | 223492.8 | 0 | 22.0878 | -15 | 0 | 100 | 109.32 | 923.4 | 1 | 0.99764 | 15 | 0 | 100 | 87.88 | 512.2 | 3.5 | 3.499
0 | -5 | 100 | 61.58 | 638248.4 | 61.71 | 61.7467 | 0 | 0.1 | 100 | 98.89 | 1024948.3 | 99.26 | 100 | 0 | -4.9 | 100 | 17.08 | 177044.2 | 69 | 61.4831 | 0 | -5 | 100 | 38.72 | 391739.9 | 0 | 38.7161 | -15.5 | 0 | 100 | 102.12 | 862.7 | 0.93 | 0.93193 | 15.5 | 0 | 100 | 94.38 | 550.1 | 3.76 | 3.7578
0 | -4.5 | 100 | 75.85 | 786119.9 | 75.92 | 76.0553 | 0 | -0.9 | 100 | 98.75 | 1023489.6 | 99.21 | 99.8584 | 0 | -4.5 | 100 | 20.8 | 215533.6 | 83.94 | 74.874 | 0 | -4.5 | 100 | 63.13 | 638587.5 | 0 | 63.1237 | -16 | 0 | 100 | 92.74 | 783.4 | 0.85 | 0.84633 | 15.9 | 0 | 100 | 97.39 | 567.6 | 3.88 | 3.8776
0 | -4.1 | 100 | 86.35 | 894976.4 | 86.46 | 86.5838 | 0 | -2 | 100 | 97.56 | 1011147.2 | 97.94 | 98.6551 | 0 | -3.9 | 100 | 23.72 | 245828.6 | 95.71 | 85.3852 | 0 | -4 | 100 | 82.12 | 830757.5 | 0 | 82.1118 | -16.4 | 0 | 100 | 94.81 | 800.8 | 0.87 | 0.86522 | 16.4 | 0 | 100 | 98.76 | 575.6 | 3.93 | 3.9322
0 | -3.6 | 100 | 92.01 | 953651.3 | 92.2 | 92.2591 | 0 | -3 | 100 | 95.08 | 985431.7 | 95.48 | 96.1472 | 0 | -3.5 | 100 | 25.27 | 261888.7 | 102.04 | 90.9647 | 0 | -3.5 | 100 | 89.22 | 902581.1 | 0 | 89.2111 | -17 | 0 | 100 | 86.83 | 733.5 | 0.79 | 0.7924 | 17 | 0 | 100 | 82.29 | 479.6 | 3.27 | 3.2764
0 | -3 | 100 | 95.17 | 986387.3 | 95.36 | 95.4277 | 0 | -4 | 100 | 86.11 | 892522.9 | 86.47 | 87.0765 | 0 | -3.1 | 100 | 26.26 | 272131.3 | 105.89 | 94.5284 | 0 | -3 | 100 | 92.62 | 936941.4 | 0 | 92.6107 | -17.5 | 0 | 100 | 75.67 | 639.2 | 0.69 | 0.69055 | 17.5 | 0 | 100 | 83.51 | 486.7 | 3.33 | 3.325
0 | -2.6 | 100 | 96.71 | 1002404.2 | 96.89 | 96.9718 | 0 | -5 | 100 | 60.5 | 627095.6 | 60.77 | 61.1791 | 0 | -2.5 | 100 | 26.74 | 277143.4 | 107.9 | 96.2563 | 0 | -2.5 | 100 | 95.26 | 963698.3 | 0 | 95.2505 | -18 | 0 | 100 | 82.82 | 699.6 | 0.76 | 0.7558 | 18 | 0 | 100 | 84.12 | 490.3 | 3.35 | 3.3493
0 | -2 | 100 | 98.04 | 1016119.6 | 98.25 | 98.3054 | 0 | -5.9 | 100 | 30.99 | 321157.7 | 31.12 | 31.3379 | 0 | -2 | 100 | 27.12 | 281074.7 | 109.45 | 97.6242 | 0 | -2 | 100 | 97.18 | 983065.7 | 0 | 97.1703 | -18.5 | 0 | 100 | 78.71 | 664.9 | 0.72 | 0.7183 | 18.4 | 0 | 100 | 72 | 419.6 | 2.86 | 2.8667
0 | -1.4 | 100 | 99.1 | 1027173 | 99.3 | 99.3683 | 0 | -7 | 100 | 13.93 | 144372.8 | 14 | 14.0864 | 0 | -1.5 | 100 | 27.46 | 284609.4 | 110.83 | 98.8481 | 0 | -1.5 | 100 | 98.58 | 997206.5 | 0 | 98.5701 | -19 | 0 | 100 | 71.71 | 605.8 | 0.66 | 0.65442 | 18.9 | 0 | 100 | 81.19 | 473.2 | 3.24 | 3.2326
0 | -1.1 | 100 | 99.27 | 1028878.4 | 99.45 | 99.5388 | 0 | -8 | 100 | 7.22 | 74859.9 | 7.26 | 7.301 | 0 | -1 | 100 | 27.56 | 285643.3 | 111.15 | 99.2081 | 0 | -1 | 100 | 99.08 | 1002351 | 0 | 99.0701 | -19.5 | 0 | 100 | 74.84 | 632.2 | 0.68 | 0.68298 | 19.5 | 0 | 100 | 80.67 | 470.2 | 3.22 | 3.2119
0 | -0.6 | 100 | 99.66 | 1032900.6 | 99.87 | 99.9298 | 0 | -8.9 | 100 | 4.24 | 43955.8 | 4.26 | 4.2876 | 0 | -0.5 | 100 | 27.78 | 287889 | 112.25 | 100 | 0 | -0.5 | 100 | 99.87 | 1010311.7 | 0 | 99.86 | -20 | 0 | 100 | 75.1 | 634.4 | 0.69 | 0.68535 | 20 | 0 | 100 | 90.01 | 524.6 | 3.59 | 3.5838
0 | 0 | 100 | 99.73 | 1033632.5 | 99.94 | 100 | 0 | -10 | 100 | 2.69 | 27846.1 | 2.7 | 2.7202 | 0 | 0 | 100 | 27.65 | 286548.4 | 111.58 | 99.532 | 0 | 0 | 100 | 100.01 | 1011696.6 | 0 | 100
0 | 0.5 | 100 | 99.66 | 1032965.2 | 99.84 | 99.9298 | [7 empty] | 0 | 0.5 | 100 | 27.67 | 286759 | 111.71 | 99.604 | 0 | 0.5 | 100 | 99.86 | 1010195.3 | 0 | 99.85
0 | 0.9 | 100 | 99.41 | 1030332 | 99.6 | 99.6791 | [7 empty] | 0 | 1 | 100 | 27.64 | 286462.8 | 111.58 | 99.496 | 0 | 1 | 100 | 99.46 | 1006122.9 | 0 | 99.4501 | Measurement time: | 27/05/2023 10:37:20 | [5 empty] | Measurement time: | 27/05/2023 12:45:21
0 | 1.4 | 100 | 98.7 | 1023011.3 | 98.92 | 98.9672 | [7 empty] | 0 | 1.5 | 100 | 27.36 | 283581.3 | 110.43 | 98.4881 | 0 | 1.4 | 100 | 98.95 | 1001030.1 | 0 | 98.9401 | Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator
0 | 2 | 100 | 97.66 | 1012214.8 | 97.83 | 97.9244 | [7 empty] | 0 | 1.9 | 100 | 27.01 | 279930.7 | 108.99 | 97.2282 | 0 | 2 | 100 | 97.76 | 988933.7 | 0 | 97.7502 | Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF)
0 | 2.4 | 100 | 95.99 | 994904.5 | 96.2 | 96.2499 | [7 empty] | 0 | 2.5 | 100 | 26.5 | 274634.5 | 106.84 | 95.3924 | 0 | 2.6 | 100 | 95.33 | 964374.2 | 0 | 95.3205 | Controller: | CCU | [5 empty] | Controller: | CCU
0 | 3 | 100 | 93.47 | 968799.1 | 93.68 | 93.7231 | [7 empty] | 0 | 3 | 100 | 25.68 | 266158 | 103.61 | 92.4406 | 0 | 2.9 | 100 | 93.57 | 946587.6 | 0 | 93.5606 | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 3.5 | 100 | 89.49 | 927488.6 | 89.64 | 89.7323 | [7 empty] | 0 | 3.5 | 100 | 24.67 | 255705.9 | 99.58 | 88.8049 | 0 | 3.4 | 100 | 90.8 | 918552.3 | 0 | 90.7909 | Field detector: | PTW60019 Diamant | [5 empty] | Field detector: | Razor Nano
0 | 4 | 100 | 83.71 | 867614.6 | 83.93 | 83.9366 | [7 empty] | 0 | 4 | 100 | 23.14 | 239822.6 | 93.39 | 83.2973 | 0 | 4 | 100 | 84.54 | 855241.6 | 0 | 84.5315 | Reference detector: | Stealth CI transmission | [5 empty] | Reference detector: | Stealth CI transmission
0 | 4.4 | 100 | 75.95 | 787173.4 | 76.05 | 76.1556 | [7 empty] | 0 | 4.5 | 100 | 20.82 | 215776.9 | 84 | 74.946 | 0 | 4.5 | 100 | 70.62 | 714373.3 | 0 | 70.6129 | SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm
0 | 5 | 100 | 62.6 | 648801.1 | 62.75 | 62.7695 | [7 empty] | 0 | 5 | 100 | 17.12 | 177485.4 | 69.15 | 61.6271 | 0 | 5 | 100 | 42.56 | 430561.6 | 0 | 42.5557 | SSD: | 900 mm |  | Côté G |  | Côté D |  | SSD: | 900 mm |  | Côté G |  | Côté D
0 | 5.4 | 100 | 45.69 | 473582.2 | 45.81 | 45.8137 | [7 empty] | 0 | 5.5 | 100 | 12.95 | 134226.8 | 52.31 | 46.6163 | 0 | 5.5 | 100 | 21.77 | 220232.7 | 0 | 21.7678 | Field size: | 10 x 10 mm |  | 36.5611 | 6 | 52.6066 | -5 | Field size: | 10 x 10 mm |  | 46.3184 | -5.5 | 59.7825 | 5
0 | 6 | 100 | 28.23 | 292540.2 | 28.29 | 28.3064 | [7 empty] | 0 | 6 | 100 | 8.33 | 86306.5 | 33.62 | 29.9856 | 0 | 6 | 100 | 13.5 | 136551.9 | 0 | 13.4987 | Gantry angle: | 0 |  | 50.9719 | 5.5 | 38.5852 | -5.5 | Gantry angle: | 0 |  | 60.4763 | -5 | 46.9189 | 5.5
0 | 6.5 | 100 | 17.79 | 184337.6 | 17.82 | 17.8382 | [7 empty] | 0 | 6.4 | 100 | 5.54 | 57468.5 | 22.4 | 19.9424 | 0 | 6.4 | 100 | 9.63 | 97407.6 | 0 | 9.629 | Measurement medium: | Water |  | 50 | 5.5337 | 50 | -5.093 | Measurement medium: | Water |  | 50 | -5.37 | 50 | 5.3802
0 | 7 | 100 | 12.2 | 126413.8 | 12.22 | 12.233 | [7 empty] | 0 | 6.9 | 100 | 3.68 | 38145.3 | 14.86 | 13.2469 | 0 | 7 | 100 | 7.15 | 72288.5 | 0 | 7.1493 | Scan type: | Crossline | [5 empty] | Scan type: | Crossline
0 | 7.5 | 100 | 8.76 | 90768.4 | 8.78 | 8.7837 | [7 empty] | 0 | 7.4 | 100 | 2.51 | 26037.1 | 10.13 | 9.0353 | 0 | 7.5 | 100 | 5.63 | 56969.8 | 0 | 5.6294 | Scan mode: | StepByStep |  |  | 10.6267 |  |  | Scan mode: | StepByStep |  |  | 10.7502
0 | 7.9 | 100 | 6.52 | 67526.4 | 6.53 | 6.5377 | [7 empty] | 0 | 8 | 100 | 1.72 | 17820.4 | 6.95 | 6.1915 | 0 | 7.9 | 100 | 4.59 | 46407 | 0 | 4.5895 | Wedge: | none | [5 empty] | Wedge: | none
0 | 8.4 | 100 | 5.02 | 52025.8 | 5.03 | 5.0336 | [7 empty] | 0 | 8.4 | 100 | 1.24 | 12803.6 | 4.99 | 4.4636 | 0 | 8.4 | 100 | 3.75 | 37979.7 | 0 | 3.7496
0 | 8.9 | 100 | 3.85 | 39936.5 | 3.86 | 3.8604 | [7 empty] | 0 | 9 | 100 | 0.86 | 8929.2 | 3.48 | 3.0958 | 0 | 9 | 100 | 3.05 | 30805.9 | 0 | 3.0497 | Points [mm]: | [6 empty] | Points [mm]:
0 | 9.4 | 100 | 3.06 | 31711 | 3.07 | 3.0683 | [7 empty] | 0 | 9.4 | 100 | 0.59 | 6096.5 | 2.37 | 2.1238 | 0 | 9.5 | 100 | 2.56 | 25915.6 | 0 | 2.5597 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | 10 | 100 | 2.43 | 25216.2 | 2.44 | 2.4366 | [7 empty] | 0 | 10 | 100 | 0.37 | 3826.7 | 1.49 | 1.3319 | 0 | 10 | 100 | 2.19 | 22175.6 | 0 | 2.1898 | 0 | 20 | 100 | 72.57 | 613 | 0.66 | 0.66156 | 0 | -20 | 100 | 89.11 | 519.3 | 3.56 | 3.5553
[28 empty] | 0 | 19.6 | 100 | 71.96 | 607.9 | 0.66 | 0.656 | 0 | -19.5 | 100 | 89.85 | 523.7 | 3.59 | 3.5849
[28 empty] | 0 | 19 | 100 | 82.01 | 692.7 | 0.75 | 0.74762 | 0 | -19 | 100 | 81.15 | 473 | 3.24 | 3.2378
[28 empty] | 0 | 18.4 | 100 | 81.59 | 689.2 | 0.75 | 0.74379 | 0 | -18.5 | 100 | 88.33 | 514.8 | 3.53 | 3.5242
[28 empty] | 0 | 17.9 | 100 | 84.62 | 714.8 | 0.77 | 0.77141 | 0 | -18 | 100 | 85.07 | 495.8 | 3.4 | 3.3942
[28 empty] | 0 | 17.4 | 100 | 86.7 | 732.3 | 0.79 | 0.79037 | 0 | -17.5 | 100 | 83.92 | 489.1 | 3.36 | 3.3483
[28 empty] | 0 | 17 | 100 | 94.4 | 797.4 | 0.86 | 0.86057 | 0 | -17 | 100 | 89.87 | 523.8 | 3.59 | 3.5857
[28 empty] | 0 | 16.4 | 100 | 103.76 | 876.4 | 0.95 | 0.9459 | 0 | -16.5 | 100 | 92.05 | 536.5 | 3.67 | 3.6726
[28 empty] | 0 | 15.9 | 100 | 106.34 | 898.3 | 0.97 | 0.96942 | 0 | -16 | 100 | 84.08 | 490 | 3.36 | 3.3547
[28 empty] | 0 | 15.4 | 100 | 106.65 | 900.9 | 0.97 | 0.97224 | 0 | -15.5 | 100 | 89.78 | 523.2 | 3.58 | 3.5821
[28 empty] | 0 | 15 | 100 | 114.11 | 963.9 | 1.04 | 1.0402 | 0 | -14.9 | 100 | 90.27 | 526.1 | 3.6 | 3.6016
[28 empty] | 0 | 14.5 | 100 | 125.42 | 1059.4 | 1.14 | 1.1434 | 0 | -14.5 | 100 | 92.36 | 538.3 | 3.69 | 3.685
[28 empty] | 0 | 13.9 | 100 | 137.57 | 1162.1 | 1.25 | 1.2541 | 0 | -14 | 100 | 88.69 | 516.9 | 3.54 | 3.5386
[28 empty] | 0 | 13.5 | 100 | 144.8 | 1223.1 | 1.32 | 1.32 | 0 | -13.5 | 100 | 93.05 | 542.3 | 3.72 | 3.7125
[28 empty] | 0 | 13 | 100 | 153.63 | 1297.7 | 1.4 | 1.4005 | 0 | -13 | 100 | 106.07 | 618.2 | 4.23 | 4.232
[28 empty] | 0 | 12.4 | 100 | 171.96 | 1452.6 | 1.57 | 1.5676 | 0 | -12.5 | 100 | 102.87 | 599.5 | 4.11 | 4.1043
[28 empty] | 0 | 11.9 | 100 | 191.17 | 1614.8 | 1.74 | 1.7427 | 0 | -12 | 100 | 113.01 | 658.6 | 4.51 | 4.5089
[28 empty] | 0 | 11.4 | 100 | 213.55 | 1803.9 | 1.95 | 1.9468 | 0 | -11.5 | 100 | 120.44 | 701.9 | 4.81 | 4.8054
[28 empty] | 0 | 11 | 100 | 246.11 | 2078.9 | 2.25 | 2.2436 | 0 | -11 | 100 | 124.93 | 728.1 | 4.99 | 4.9845
[28 empty] | 0 | 10.5 | 100 | 288.1 | 2433.6 | 2.63 | 2.6264 | 0 | -10.5 | 100 | 138.69 | 808.3 | 5.54 | 5.5335
[28 empty] | 0 | 10 | 100 | 350.66 | 2962 | 3.2 | 3.1967 | 0 | -10 | 100 | 154.4 | 899.8 | 6.17 | 6.1603
[28 empty] | 0 | 9.4 | 100 | 423.92 | 3580.9 | 3.87 | 3.8645 | 0 | -9.5 | 100 | 166.22 | 968.7 | 6.64 | 6.6319
[28 empty] | 0 | 9 | 100 | 511.73 | 4322.7 | 4.66 | 4.665 | 0 | -9 | 100 | 192.1 | 1119.6 | 7.68 | 7.6645
[28 empty] | 0 | 8.5 | 100 | 651.87 | 5506.4 | 5.94 | 5.9426 | 0 | -8.5 | 100 | 222.62 | 1297.5 | 8.89 | 8.8822
[28 empty] | 0 | 8 | 100 | 852.27 | 7199.2 | 7.77 | 7.7695 | 0 | -8 | 100 | 255.26 | 1487.7 | 10.2 | 10.1845
[28 empty] | 0 | 7.5 | 100 | 1190.31 | 10054.7 | 10.85 | 10.8511 | 0 | -7.5 | 100 | 348.43 | 2030.6 | 13.91 | 13.9018
[28 empty] | 0 | 7 | 100 | 1736.82 | 14671.1 | 15.83 | 15.8332 | 0 | -7 | 100 | 450.47 | 2625.3 | 17.98 | 17.973
[28 empty] | 0 | 6.5 | 100 | 2618.92 | 22122.3 | 23.86 | 23.8746 | 0 | -6.5 | 100 | 625.27 | 3644.1 | 24.97 | 24.9472
[28 empty] | 0 | 6 | 100 | 4010.56 | 33877.5 | 36.57 | 36.5611 | 0 | -6 | 100 | 867.36 | 5055 | 34.68 | 34.6062
[28 empty] | 0 | 5.5 | 100 | 5591.35 | 47230.7 | 51 | 50.9719 | 0 | -5.5 | 100 | 1160.91 | 6765.8 | 46.36 | 46.3184
[28 empty] | 0 | 5 | 100 | 7271.25 | 61420.9 | 66.29 | 66.2862 | 0 | -5 | 100 | 1515.76 | 8833.9 | 60.66 | 60.4763
[28 empty] | 0 | 4.5 | 100 | 8413.67 | 71071 | 76.87 | 76.7007 | 0 | -4.5 | 100 | 1848.04 | 10770.4 | 73.84 | 73.7337
[28 empty] | 0 | 4 | 100 | 9348.24 | 78965.4 | 85.24 | 85.2204 | 0 | -4 | 100 | 2075.08 | 12093.6 | 82.92 | 82.7922
[28 empty] | 0 | 3.5 | 100 | 9903.18 | 83653.1 | 90.32 | 90.2794 | 0 | -3.6 | 100 | 2273.46 | 13249.7 | 90.8 | 90.7073
[28 empty] | 0 | 3 | 100 | 10266.42 | 86721.4 | 93.72 | 93.5908 | 0 | -3 | 100 | 2365.07 | 13783.7 | 94.45 | 94.3624
[28 empty] | 0 | 2.5 | 100 | 10584.09 | 89404.8 | 96.56 | 96.4867 | 0 | -2.6 | 100 | 2409.69 | 14043.7 | 96.28 | 96.1426
[28 empty] | 0 | 2 | 100 | 10745.41 | 90767.4 | 98.06 | 97.9573 | 0 | -2 | 100 | 2451.99 | 14290.2 | 97.91 | 97.8303
[28 empty] | 0 | 1.6 | 100 | 10834.82 | 91522.7 | 98.98 | 98.7724 | 0 | -1.5 | 100 | 2485.58 | 14486 | 99.26 | 99.1705
[28 empty] | 0 | 1 | 100 | 10899.04 | 92065.2 | 99.43 | 99.3579 | 0 | -1 | 100 | 2497.45 | 14555.2 | 99.76 | 99.6441
[28 empty] | 0 | 0.5 | 100 | 10947.68 | 92476.1 | 99.9 | 99.8013 | 0 | -0.6 | 100 | 2500.06 | 14570.4 | 99.89 | 99.7482
[28 empty] | 0 | 0 | 100 | 10969.48 | 92660.1 | 100 | 100 | 0 | -0.1 | 100 | 2506.37 | 14607.2 | 100.04 | 100
[28 empty] | 0 | -0.5 | 100 | 10936.56 | 92382.1 | 99.85 | 99.6999 | 0 | 0.5 | 100 | 2502.2 | 14582.9 | 99.83 | 99.8336
[28 empty] | 0 | -1 | 100 | 10891.17 | 91998.7 | 99.31 | 99.2861 | 0 | 1 | 100 | 2489.33 | 14507.8 | 99.37 | 99.3201
[28 empty] | 0 | -1.5 | 100 | 10802.05 | 91245.9 | 98.46 | 98.4737 | 0 | 1.6 | 100 | 2465.95 | 14371.6 | 98.46 | 98.3873
[28 empty] | 0 | -2 | 100 | 10726.6 | 90608.5 | 97.84 | 97.7859 | 0 | 2 | 100 | 2427.05 | 14144.9 | 96.96 | 96.8353
[28 empty] | 0 | -2.5 | 100 | 10568.52 | 89273.2 | 96.26 | 96.3448 | 0 | 2.5 | 100 | 2382.18 | 13883.4 | 95.26 | 95.045
[28 empty] | 0 | -3 | 100 | 10238.92 | 86489.1 | 93.3 | 93.3401 | 0 | 3 | 100 | 2303.84 | 13426.8 | 91.99 | 91.9194
[28 empty] | 0 | -3.5 | 100 | 9600.94 | 81100 | 87.53 | 87.5241 | 0 | 3.5 | 100 | 2200.7 | 12825.7 | 87.91 | 87.8043
[28 empty] | 0 | -4 | 100 | 8793.4 | 74278.7 | 80.17 | 80.1624 | 0 | 4 | 100 | 2035.92 | 11865.4 | 81.43 | 81.2298
[28 empty] | 0 | -4.5 | 100 | 7413.78 | 62624.8 | 67.55 | 67.5855 | 0 | 4.5 | 100 | 1797.66 | 10476.8 | 71.86 | 71.7236
[28 empty] | 0 | -5 | 100 | 5770.67 | 48745.4 | 52.58 | 52.6066 | 0 | 5 | 100 | 1498.37 | 8732.5 | 59.85 | 59.7825
[28 empty] | 0 | -5.5 | 100 | 4232.6 | 35753.2 | 38.64 | 38.5852 | 0 | 5.5 | 100 | 1175.96 | 6853.5 | 46.96 | 46.9189
[28 empty] | 0 | -6 | 100 | 2968.08 | 25071.7 | 27.06 | 27.0576 | 0 | 5.9 | 100 | 865.44 | 5043.8 | 34.58 | 34.5296
[28 empty] | 0 | -6.5 | 100 | 2063.8 | 17433.1 | 18.82 | 18.814 | 0 | 6.5 | 100 | 607.52 | 3540.6 | 24.25 | 24.239
[28 empty] | 0 | -6.9 | 100 | 1418.62 | 11983.2 | 12.92 | 12.9324 | 0 | 7.1 | 100 | 439.16 | 2559.4 | 17.55 | 17.5218
[28 empty] | 0 | -7.5 | 100 | 981.39 | 8289.9 | 8.95 | 8.9465 | 0 | 7.5 | 100 | 335.14 | 1953.2 | 13.39 | 13.3715
[28 empty] | 0 | -8.1 | 100 | 729.32 | 6160.6 | 6.65 | 6.6486 | 0 | 8 | 100 | 268.64 | 1565.6 | 10.73 | 10.7183
[28 empty] | 0 | -8.5 | 100 | 571.03 | 4823.5 | 5.21 | 5.2056 | 0 | 8.5 | 100 | 235.84 | 1374.5 | 9.42 | 9.4096
[28 empty] | 0 | -8.9 | 100 | 455.51 | 3847.8 | 4.15 | 4.1525 | 0 | 9 | 100 | 190.78 | 1111.9 | 7.62 | 7.6118
[28 empty] | 0 | -9.4 | 100 | 377.52 | 3188.9 | 3.44 | 3.4415 | 0 | 9.4 | 100 | 158.66 | 924.7 | 6.34 | 6.3303
[28 empty] | 0 | -10 | 100 | 312.61 | 2640.7 | 2.85 | 2.8498 | 0 | 10 | 100 | 142.92 | 833 | 5.7 | 5.7023
[28 empty] | 0 | -10.4 | 100 | 255.83 | 2161 | 2.34 | 2.3322 | 0 | 10.5 | 100 | 130.91 | 763 | 5.23 | 5.2231
[28 empty] | 0 | -11 | 100 | 222.57 | 1880.1 | 2.03 | 2.029 | 0 | 11 | 100 | 121.31 | 707 | 4.84 | 4.8401
[28 empty] | 0 | -11.5 | 100 | 188.51 | 1592.4 | 1.72 | 1.7185 | 0 | 11.4 | 100 | 122.25 | 712.5 | 4.89 | 4.8776
[28 empty] | 0 | -11.9 | 100 | 160.77 | 1358 | 1.47 | 1.4656 | 0 | 12 | 100 | 106.28 | 619.4 | 4.24 | 4.2404
[28 empty] | 0 | -12.5 | 100 | 151 | 1275.5 | 1.38 | 1.3765 | 0 | 12.5 | 100 | 98.57 | 574.5 | 3.94 | 3.9328
[28 empty] | 0 | -12.9 | 100 | 141.81 | 1197.9 | 1.29 | 1.2928 | 0 | 13 | 100 | 90.38 | 526.7 | 3.61 | 3.606
[28 empty] | 0 | -13.5 | 100 | 125.57 | 1060.7 | 1.15 | 1.1447 | 0 | 13.4 | 100 | 84.33 | 491.5 | 3.37 | 3.3646
[28 empty] | 0 | -14 | 100 | 121.94 | 1030.1 | 1.11 | 1.1116 | 0 | 13.9 | 100 | 81.69 | 476.1 | 3.27 | 3.2593
[28 empty] | 0 | -14.5 | 100 | 115.85 | 978.6 | 1.06 | 1.0561 | 0 | 14.5 | 100 | 85.7 | 499.5 | 3.42 | 3.4193
[28 empty] | 0 | -15 | 100 | 98.49 | 832 | 0.9 | 0.89785 | 0 | 15 | 100 | 82.21 | 479.1 | 3.29 | 3.28
[28 empty] | 0 | -15.5 | 100 | 97.28 | 821.7 | 0.89 | 0.88682 | 0 | 15.4 | 100 | 76.99 | 448.7 | 3.08 | 3.0718
[28 empty] | 0 | -15.9 | 100 | 95.28 | 804.8 | 0.87 | 0.86859 | 0 | 16 | 100 | 81.83 | 476.9 | 3.27 | 3.2649
[28 empty] | 0 | -16.5 | 100 | 97.04 | 819.7 | 0.89 | 0.88464 | 0 | 16.5 | 100 | 71.44 | 416.3 | 2.86 | 2.8503
[28 empty] | 0 | -17 | 100 | 82.85 | 699.8 | 0.76 | 0.75528 | 0 | 17 | 100 | 66.61 | 388.2 | 2.67 | 2.6576
[28 empty] | 0 | -17.5 | 100 | 79.47 | 671.3 | 0.73 | 0.72446 | 0 | 17.5 | 100 | 66.72 | 388.9 | 2.67 | 2.662
[28 empty] | 0 | -18 | 100 | 72.48 | 612.2 | 0.66 | 0.66074 | 0 | 17.9 | 100 | 72.25 | 421.1 | 2.89 | 2.8827
[28 empty] | 0 | -18.5 | 100 | 75.42 | 637.1 | 0.69 | 0.68754 | 0 | 18.5 | 100 | 60.72 | 353.9 | 2.43 | 2.4226
[28 empty] | 0 | -19 | 100 | 77.73 | 656.6 | 0.71 | 0.7086 | 0 | 19 | 100 | 88.23 | 514.2 | 3.53 | 3.5202
[28 empty] | 0 | -19.4 | 100 | 66.59 | 562.5 | 0.61 | 0.60705 | 0 | 19.4 | 100 | 72.11 | 420.2 | 2.88 | 2.8771
[28 empty] | 0 | -20 | 100 | 65.87 | 556.4 | 0.6 | 0.60048 | 0 | 20 | 100 | 71.69 | 417.8 | 2.87 | 2.8603

## 2x2
Measurement time: | 12/04/2023 18:56:46 | [5 empty] | Measurement time: | 27/05/2023 10:55:28 | [6 empty] | Measurement time: | 27/05/2023 13:03:36 |  |  | Taille de champ
Radiation device: | Novalis TB Accelerator |  |  | Taille de champ |  |  | Radiation device: | Novalis TB Accelerator |  |  |  | Taille de champ |  |  | Radiation device: | Novalis TB Accelerator |  |  | 19.9939
Energy: | 6 MV (FFF) |  |  | 19.9707 |  |  | Energy: | 6 MV (FFF) |  |  |  | 20.0606 |  |  | Energy: | 6 MV (FFF)
Controller: | CCU | [5 empty] | Controller: | CCU | [6 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | PTW60019 Diamant | [6 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission |  | 25.8506 | 10 | 53.562 | -9.9 | Reference detector: | Stealth CI transmission |  |  | Côté G |  | Côté D |  | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  | 58.3143 | 9.1 | 24.1631 | -10.9 | SAD: | 1000 mm |  |  | 46.7252 | -10.1 | 62.1529 | 9 | SAD: | 1000 mm
SSD: | 900 mm |  | 50 | 9.3305 | 50 | -10.0212 | SSD: | 900 mm |  |  | 74.6307 | -9.1 | 33.1231 | 10 | SSD: | 900 mm |  | Côté G |  | Côté D
Field size: | 20 x 20 mm | [5 empty] | Field size: | 20 x 20 mm |  |  | 50 | -9.9826 | 50 | 9.4186 | Field size: | 20 x 20 mm |  | 42.8522 | -10 | 66.4939 | 9
Gantry angle: | 0 |  |  | 19.3517 |  |  | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 |  | 68.202 | -9 | 41.6456 | 10
Measurement medium: | Water | [5 empty] | Measurement medium: | Water |  |  |  | 19.4013 |  |  | Measurement medium: | Water |  | 50 | -9.718 | 50 | 9.6638
Scan type: | Inline | [5 empty] | Scan type: | Inline | [6 empty] | Scan type: | Inline
Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep |  |  | 19.3818
Wedge: | none | [5 empty] | Wedge: | none | [6 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]: | [7 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
20 | 0 | 100 | 2.13 | 22053.6 | 1.94 | 1.9429 | -30 | 0 | 100 | 130.51 | 1102.5 | 1.08 | 1.0782 |  | -30 | 0 | 100 | 97.06 | 565.7 | 3.47 | 3.4524
18.9 | 0 | 100 | 2.36 | 24456.3 | 2.15 | 2.1527 | -29 | 0 | 100 | 154.35 | 1303.8 | 1.28 | 1.2751 |  | -29.1 | 0 | 100 | 98.07 | 571.6 | 3.51 | 3.4884
18 | 0 | 100 | 2.69 | 27836 | 2.45 | 2.4537 | -28 | 0 | 100 | 155.29 | 1311.7 | 1.28 | 1.2829 |  | -28.1 | 0 | 100 | 93.04 | 542.2 | 3.32 | 3.3095
17 | 0 | 100 | 3.09 | 32030.1 | 2.82 | 2.8186 | -27 | 0 | 100 | 168.83 | 1426.2 | 1.4 | 1.3947 |  | -27 | 0 | 100 | 97.4 | 567.7 | 3.48 | 3.4645
16 | 0 | 100 | 3.53 | 36619.3 | 3.22 | 3.2199 | -26 | 0 | 100 | 182.91 | 1545 | 1.51 | 1.511 |  | -26 | 0 | 100 | 118.35 | 689.7 | 4.23 | 4.2097
15 | 0 | 100 | 4.39 | 45541.6 | 4.01 | 4.0044 | -25 | 0 | 100 | 196.83 | 1662.6 | 1.63 | 1.626 |  | -25.1 | 0 | 100 | 116.51 | 679 | 4.16 | 4.1443
14 | 0 | 100 | 5.39 | 55826.9 | 4.91 | 4.9165 | -24 | 0 | 100 | 208.07 | 1757.6 | 1.72 | 1.7189 |  | -24.1 | 0 | 100 | 127.31 | 742 | 4.55 | 4.5284
13 | 0 | 100 | 7.04 | 72998.1 | 6.42 | 6.4216 | -23 | 0 | 100 | 217.29 | 1835.5 | 1.8 | 1.7951 |  | -23 | 0 | 100 | 136.11 | 793.3 | 4.86 | 4.8415
12 | 0 | 100 | 9.88 | 102436.1 | 9.01 | 9.0121 | -22 | 0 | 100 | 237.47 | 2006 | 1.96 | 1.9618 |  | -22.1 | 0 | 100 | 145.83 | 849.9 | 5.21 | 5.1872
11.1 | 0 | 100 | 14.71 | 152454.4 | 13.41 | 13.4179 | -21 | 0 | 100 | 255.7 | 2159.9 | 2.11 | 2.1124 |  | -21 | 0 | 100 | 138.1 | 804.8 | 4.94 | 4.9122
10 | 0 | 100 | 28.34 | 293760 | 25.84 | 25.8506 | -20 | 0 | 100 | 287.85 | 2431.5 | 2.38 | 2.378 |  | -20 | 0 | 100 | 127.25 | 741.6 | 4.55 | 4.5263
9.1 | 0 | 100 | 63.93 | 662650.2 | 58.3 | 58.3143 | -19 | 0 | 100 | 308.09 | 2602.5 | 2.55 | 2.5452 |  | -19.1 | 0 | 100 | 132.87 | 774.4 | 4.75 | 4.7262
8.1 | 0 | 100 | 89.92 | 931979.2 | 81.99 | 82.0213 | -18 | 0 | 100 | 349.48 | 2952.1 | 2.89 | 2.8871 |  | -18 | 0 | 100 | 135.24 | 788.2 | 4.84 | 4.8105
7 | 0 | 100 | 100.07 | 1037173 | 91.23 | 91.2798 | -17 | 0 | 100 | 395.66 | 3342.1 | 3.27 | 3.2686 |  | -17 | 0 | 100 | 134.41 | 783.4 | 4.8 | 4.781
6.1 | 0 | 100 | 104.37 | 1081755.1 | 95.11 | 95.202 | -16 | 0 | 100 | 471.45 | 3982.3 | 3.9 | 3.8947 |  | -16 | 0 | 100 | 144.86 | 844.2 | 5.18 | 5.1527
5.1 | 0 | 100 | 106.37 | 1102512.7 | 96.97 | 97.0264 | -15 | 0 | 100 | 565.45 | 4776.4 | 4.68 | 4.6713 |  | -15 | 0 | 100 | 173.25 | 1009.7 | 6.2 | 6.1625
4.1 | 0 | 100 | 107.9 | 1118337.4 | 98.36 | 98.422 | -14 | 0 | 100 | 716.66 | 6053.7 | 5.93 | 5.9204 |  | -14 | 0 | 100 | 234.86 | 1368.8 | 8.4 | 8.354
3.1 | 0 | 100 | 108.61 | 1125681.3 | 99.07 | 99.0696 | -13 | 0 | 100 | 970.4 | 8197.1 | 8.03 | 8.0166 |  | -13 | 0 | 100 | 294.84 | 1718.3 | 10.54 | 10.4875
2.1 | 0 | 100 | 109.19 | 1131732.7 | 99.58 | 99.5987 | -12 | 0 | 100 | 1456.6 | 12304 | 12.05 | 12.0332 |  | -12 | 0 | 100 | 415.45 | 2421.2 | 14.87 | 14.7777
1 | 0 | 100 | 109.62 | 1136126.6 | 100 | 99.9909 | -11 | 0 | 100 | 2793.83 | 23599.7 | 23.1 | 23.0803 |  | -11 | 0 | 100 | 659.5 | 3843.6 | 23.58 | 23.4586
0 | 0 | 100 | 109.63 | 1136307.4 | 100 | 100 | -10.1 | 0 | 100 | 5656 | 47776.7 | 46.78 | 46.7252 |  | -10 | 0 | 100 | 1204.72 | 7021.1 | 43.06 | 42.8522
-1 | 0 | 100 | 109.53 | 1135231.2 | 99.91 | 99.9088 | -9.1 | 0 | 100 | 9033.92 | 76310.3 | 74.76 | 74.6307 |  | -9 | 0 | 100 | 1917.39 | 11174.6 | 68.56 | 68.202
-2.1 | 0 | 100 | 109.62 | 1136140.4 | 99.97 | 99.9909 | -8 | 0 | 100 | 10792.29 | 91163.5 | 89.33 | 89.1569 |  | -8 | 0 | 100 | 2414.26 | 14070.4 | 86.31 | 85.8758
-3 | 0 | 100 | 109.13 | 1131054.8 | 99.54 | 99.5439 | -7.1 | 0 | 100 | 11380.42 | 96131.4 | 94.21 | 94.0155 |  | -7.1 | 0 | 100 | 2578.94 | 15030.1 | 92.24 | 91.7335
-4 | 0 | 100 | 108.83 | 1128016 | 99.25 | 99.2703 | -5.9 | 0 | 100 | 11690.25 | 98748.6 | 96.76 | 96.5751 |  | -6 | 0 | 100 | 2669.22 | 15556.3 | 95.45 | 94.9448
-5 | 0 | 100 | 107.78 | 1117130 | 98.28 | 98.3125 | -4.9 | 0 | 100 | 11856.62 | 100154 | 98.15 | 97.9495 |  | -5 | 0 | 100 | 2716.6 | 15832.4 | 97.11 | 96.6301
-6 | 0 | 100 | 106.26 | 1101322.2 | 96.94 | 96.926 | -4 | 0 | 100 | 11987.23 | 101257.2 | 99.17 | 99.0285 |  | -4.1 | 0 | 100 | 2732.65 | 15925.9 | 97.69 | 97.201
-7.1 | 0 | 100 | 103.64 | 1074132.1 | 94.52 | 94.5362 | -3.1 | 0 | 100 | 12033.42 | 101647.4 | 99.57 | 99.4101 |  | -3.1 | 0 | 100 | 2737.86 | 15956.3 | 97.93 | 97.3863
-8.1 | 0 | 100 | 99.41 | 1030372.7 | 90.68 | 90.6777 | -2 | 0 | 100 | 12065.59 | 101919.1 | 99.88 | 99.6758 |  | -2 | 0 | 100 | 2753.36 | 16046.6 | 98.43 | 97.9376
-9 | 0 | 100 | 89.75 | 930210.3 | 81.86 | 81.8663 | -1.1 | 0 | 100 | 12104.83 | 102250.6 | 100.18 | 100 |  | -1.1 | 0 | 100 | 2778.32 | 16192.1 | 99.34 | 98.8255
-9.9 | 0 | 100 | 58.72 | 608609.6 | 53.55 | 53.562 | 0 | 0 | 100 | 12087.29 | 102102.4 | 99.99 | 99.8551 |  | -0.1 | 0 | 100 | 2794.04 | 16283.7 | 99.96 | 99.3846
-10.9 | 0 | 100 | 26.49 | 274576.5 | 24.16 | 24.1631 | 1 | 0 | 100 | 12063.37 | 101900.4 | 99.83 | 99.6575 |  | 0.9 | 0 | 100 | 2811.34 | 16384.5 | 100.54 | 100
-11.9 | 0 | 100 | 14.65 | 151850.9 | 13.37 | 13.3631 | 2 | 0 | 100 | 12062.91 | 101896.5 | 99.79 | 99.6537 |  | 2 | 0 | 100 | 2800.76 | 16322.9 | 100.13 | 99.6237
-12.9 | 0 | 100 | 9.56 | 99098.5 | 8.72 | 8.7202 | 3 | 0 | 100 | 11977.76 | 101177.2 | 99.16 | 98.9503 |  | 2.9 | 0 | 100 | 2793.92 | 16283 | 99.93 | 99.3804
-14 | 0 | 100 | 6.85 | 71017.3 | 6.25 | 6.2483 | 4 | 0 | 100 | 11919.25 | 100683 | 98.69 | 98.4669 |  | 3.9 | 0 | 100 | 2779.82 | 16200.8 | 99.44 | 98.8788
-15 | 0 | 100 | 5.3 | 54925.3 | 4.84 | 4.8344 | 5 | 0 | 100 | 11801.6 | 99689.2 | 97.64 | 97.495 |  | 5 | 0 | 100 | 2740.64 | 15972.5 | 97.99 | 97.4852
-16 | 0 | 100 | 4.27 | 44292.6 | 3.9 | 3.8949 | 6 | 0 | 100 | 11556.19 | 97616.2 | 95.61 | 95.4676 |  | 5.9 | 0 | 100 | 2690.09 | 15677.9 | 96.24 | 95.6871
-17.1 | 0 | 100 | 3.45 | 35796 | 3.15 | 3.1469 | 7 | 0 | 100 | 11104.11 | 93797.5 | 91.9 | 91.7329 |  | 7 | 0 | 100 | 2594 | 15117.9 | 92.75 | 92.2692
-18 | 0 | 100 | 3 | 31086.6 | 2.74 | 2.7365 | 8 | 0 | 100 | 9977.81 | 84283.5 | 82.55 | 82.4283 |  | 8 | 0 | 100 | 2392.21 | 13941.9 | 85.54 | 85.0915
-19 | 0 | 100 | 2.67 | 27696.4 | 2.44 | 2.4355 | 9 | 0 | 100 | 7523.5 | 63551.7 | 62.27 | 62.1529 |  | 9 | 0 | 100 | 1869.37 | 10894.7 | 66.85 | 66.4939
-20 | 0 | 100 | 2.36 | 24477.1 | 2.16 | 2.1527 | 10 | 0 | 100 | 4009.49 | 33868.5 | 33.18 | 33.1231 |  | 10 | 0 | 100 | 1170.8 | 6823.4 | 41.89 | 41.6456
[7 empty] | 11 | 0 | 100 | 1893.7 | 15996.2 | 15.67 | 15.6442 |  | 11.1 | 0 | 100 | 600.28 | 3498.4 | 21.47 | 21.3521
[7 empty] | 12 | 0 | 100 | 1152.24 | 9733.1 | 9.53 | 9.5188 |  | 12 | 0 | 100 | 397.38 | 2316 | 14.21 | 14.1349
Measurement time: | 12/04/2023 18:59:35 | [5 empty] | 13.1 | 0 | 100 | 803.36 | 6786.1 | 6.65 | 6.6367 |  | 12.9 | 0 | 100 | 273.43 | 1593.6 | 9.77 | 9.726
Radiation device: | Novalis TB Accelerator | [5 empty] | 14 | 0 | 100 | 606.85 | 5126.2 | 5.02 | 5.0133 |  | 14 | 0 | 100 | 223.7 | 1303.7 | 8 | 7.9571
Energy: | 6 MV (FFF) | [5 empty] | 15 | 0 | 100 | 493.33 | 4167.2 | 4.08 | 4.0755 |  | 15.1 | 0 | 100 | 176.31 | 1027.5 | 6.3 | 6.2714
Controller: | CCU | [5 empty] | 16 | 0 | 100 | 415.37 | 3508.7 | 3.44 | 3.4314 |  | 16 | 0 | 100 | 158.86 | 925.8 | 5.68 | 5.6507
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | 17 | 0 | 100 | 352.88 | 2980.8 | 2.92 | 2.9152 |  | 17.1 | 0 | 100 | 138.78 | 808.8 | 4.96 | 4.9364
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | 18 | 0 | 100 | 306.94 | 2592.8 | 2.54 | 2.5357 |  | 18 | 0 | 100 | 124.98 | 728.4 | 4.47 | 4.4456
Reference detector: | Stealth CI transmission |  | 26.4271 | 11.1 | 64.1968 | -9.9 | 19 | 0 | 100 | 285.56 | 2412.2 | 2.36 | 2.3591 |  | 19 | 0 | 100 | 125.84 | 733.4 | 4.5 | 4.4762
SAD: | 1000 mm |  | 56.3255 | 10 | 34.4285 | -10.9 | 20 | 0 | 100 | 260.04 | 2196.6 | 2.15 | 2.1482 |  | 19.9 | 0 | 100 | 119.29 | 695.2 | 4.27 | 4.2432
SSD: | 900 mm |  | 50 | 10.2327 | 50 | -10.3769 | 21.1 | 0 | 100 | 229.25 | 1936.5 | 1.9 | 1.8939 |  | 21 | 0 | 100 | 113.27 | 660.1 | 4.05 | 4.029
Field size: | 20 x 20 mm | [5 empty] | 21.9 | 0 | 100 | 220.06 | 1858.9 | 1.82 | 1.818 |  | 22 | 0 | 100 | 111.16 | 647.8 | 3.97 | 3.954
Gantry angle: | 0 |  |  | 20.6096 |  |  | 23 | 0 | 100 | 209.83 | 1772.4 | 1.74 | 1.7334 |  | 23 | 0 | 100 | 94.93 | 553.3 | 3.4 | 3.3767
Measurement medium: | Water | [5 empty] | 23.9 | 0 | 100 | 195.49 | 1651.3 | 1.62 | 1.615 |  | 24 | 0 | 100 | 91.01 | 530.4 | 3.26 | 3.2372
Scan type: | Crossline | [5 empty] | 25 | 0 | 100 | 182.35 | 1540.3 | 1.51 | 1.5064 |  | 25 | 0 | 100 | 92.88 | 541.3 | 3.32 | 3.3038
Scan mode: | StepByStep | [5 empty] | 26 | 0 | 100 | 173.65 | 1466.8 | 1.44 | 1.4346 |  | 26 | 0 | 100 | 92.81 | 540.9 | 3.32 | 3.3013
Wedge: | none | [5 empty] | 27.1 | 0 | 100 | 165 | 1393.8 | 1.37 | 1.3631 |  | 26.9 | 0 | 100 | 101.31 | 590.4 | 3.62 | 3.6036
[7 empty] | 28 | 0 | 100 | 162.47 | 1372.4 | 1.34 | 1.3422 |  | 27.9 | 0 | 100 | 93.03 | 542.2 | 3.33 | 3.3091
Points [mm]: | [6 empty] | 28.9 | 0 | 100 | 166.75 | 1408.5 | 1.38 | 1.3775 |  | 29 | 0 | 100 | 93.62 | 545.6 | 3.35 | 3.3301
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | 30 | 0 | 100 | 150.42 | 1270.6 | 1.25 | 1.2426 |  | 30 | 0 | 100 | 89.9 | 524 | 3.22 | 3.1978
0 | 20 | 100 | 8.04 | 24252.4 | 2.14 | 2.1337
0 | 19 | 100 | 9.2 | 27731.8 | 2.44 | 2.4415
0 | 18 | 100 | 10.6 | 31967 | 2.81 | 2.8131 | Measurement time: | 27/05/2023 10:58:20 | [6 empty] | Measurement time: | 27/05/2023 13:06:18
0 | 17 | 100 | 12.29 | 37065.7 | 3.26 | 3.2616 | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator
0 | 16 | 100 | 14.97 | 45157.1 | 3.97 | 3.9728 | Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF)
0 | 15 | 100 | 18.35 | 55327.8 | 4.87 | 4.8698 | Controller: | CCU | [6 empty] | Controller: | CCU
0 | 14 | 100 | 24.16 | 72867.7 | 6.42 | 6.4117 | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
0 | 13 | 100 | 33.86 | 102116.3 | 8.99 | 8.986 | Field detector: | PTW60019 Diamant | [6 empty] | Field detector: | Razor Nano
0 | 12 | 100 | 52.6 | 158625.6 | 13.96 | 13.9593 | Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission
0 | 11.1 | 100 | 99.58 | 300326.6 | 26.44 | 26.4271 | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm
0 | 10 | 100 | 212.24 | 640106.2 | 56.35 | 56.3255 | SSD: | 900 mm |  | Côté G |  | Côté D |  |  | SSD: | 900 mm |  | Côté G |  | Côté D
0 | 9 | 100 | 301.03 | 907859 | 79.95 | 79.8891 | Field size: | 20 x 20 mm |  | 36.906 | 11.1 | 53.4237 | -10 |  | Field size: | 20 x 20 mm |  | 34.6917 | -11 | 58.9324 | 9.9
0 | 8 | 100 | 338.09 | 1019655.2 | 89.76 | 89.7243 | Gantry angle: | 0 |  | 63.246 | 10.1 | 28.8714 | -11 |  | Gantry angle: | 0 |  | 57.7088 | -10 | 36.0498 | 10.9
0 | 7.1 | 100 | 355.45 | 1072010.1 | 94.35 | 94.3314 | Measurement medium: | Water |  | 50 | 10.6029 | 50 | -10.1394 |  | Measurement medium: | Water |  | 50 | -10.3349 | 50 | 10.2904
0 | 6 | 100 | 364.06 | 1097953.2 | 96.68 | 96.6163 | Scan type: | Crossline | [6 empty] | Scan type: | Crossline
0 | 5 | 100 | 369.31 | 1113809.5 | 98 | 98.0096 | Scan mode: | StepByStep |  |  | 20.7423 |  |  |  | Scan mode: | StepByStep |  |  | 20.6253
0 | 4 | 100 | 372.97 | 1124844.8 | 99.01 | 98.9809 | Wedge: | none | [6 empty] | Wedge: | none
0 | 3.1 | 100 | 374.28 | 1128776.3 | 99.42 | 99.3286
0 | 2.1 | 100 | 375.94 | 1133784.6 | 99.81 | 99.7691 | Points [mm]: | [7 empty] | Points [mm]:
0 | 1.1 | 100 | 376.1 | 1134281.7 | 99.92 | 99.8116 | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) |  | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0.1 | 100 | 376.61 | 1135797.8 | 99.99 | 99.9469 | 0 | 30 | 100 | 151.88 | 1283 | 1.26 | 1.2568 |  | 0 | -30 | 100 | 94.09 | 548.4 | 3.39 | 3.375
0 | -0.9 | 100 | 376.81 | 1136408.2 | 100.11 | 100 | 0 | 29 | 100 | 159.62 | 1348.3 | 1.32 | 1.3208 |  | 0 | -29 | 100 | 96.76 | 563.9 | 3.49 | 3.4708
0 | -1.9 | 100 | 375.69 | 1133035.5 | 99.79 | 99.7028 | 0 | 28 | 100 | 165.77 | 1400.3 | 1.37 | 1.3717 |  | 0 | -28 | 100 | 97.91 | 570.6 | 3.53 | 3.512
0 | -2.9 | 100 | 374.78 | 1130283.3 | 99.58 | 99.4613 | 0 | 27 | 100 | 179.85 | 1519.2 | 1.49 | 1.4882 |  | 0 | -27 | 100 | 101.13 | 589.4 | 3.65 | 3.6275
0 | -3.9 | 100 | 373.48 | 1126380.5 | 99.16 | 99.1163 | 0 | 26 | 100 | 190.31 | 1607.5 | 1.57 | 1.5748 |  | 0 | -26.1 | 100 | 99.4 | 579.3 | 3.59 | 3.5655
0 | -4.9 | 100 | 370.06 | 1116059.5 | 98.28 | 98.2086 | 0 | 25 | 100 | 213.39 | 1802.6 | 1.77 | 1.7657 |  | 0 | -25.1 | 100 | 100.53 | 585.9 | 3.63 | 3.606
0 | -5.9 | 100 | 366.94 | 1106654.3 | 97.43 | 97.3806 | 0 | 24.1 | 100 | 210.96 | 1782 | 1.75 | 1.7456 |  | 0 | -24.1 | 100 | 116.19 | 677.2 | 4.19 | 4.1677
0 | -6.9 | 100 | 359.21 | 1083335.9 | 95.37 | 95.3292 | 0 | 22.9 | 100 | 230.68 | 1948.6 | 1.91 | 1.9088 |  | 0 | -23 | 100 | 91.92 | 535.7 | 3.32 | 3.2972
0 | -7.9 | 100 | 349.27 | 1053343.3 | 92.76 | 92.6913 | 0 | 21.9 | 100 | 247.39 | 2089.7 | 2.05 | 2.0471 |  | 0 | -22 | 100 | 111.24 | 648.3 | 4.01 | 3.9902
0 | -9 | 100 | 323.21 | 974760.8 | 85.87 | 85.7753 | 0 | 21 | 100 | 274.78 | 2321.1 | 2.27 | 2.2737 |  | 0 | -21.1 | 100 | 115.64 | 674 | 4.17 | 4.148
0 | -9.9 | 100 | 241.9 | 729549.8 | 64.23 | 64.1968 | 0 | 19.9 | 100 | 299 | 2525.7 | 2.47 | 2.4741 |  | 0 | -20.1 | 100 | 123.24 | 718.2 | 4.44 | 4.4206
0 | -10.9 | 100 | 129.73 | 391255.8 | 34.47 | 34.4285 | 0 | 19 | 100 | 328.77 | 2777.1 | 2.72 | 2.7205 |  | 0 | -19 | 100 | 130.68 | 761.6 | 4.72 | 4.6875
0 | -11.9 | 100 | 65.9 | 198741.7 | 17.5 | 17.4889 | 0 | 18 | 100 | 373.56 | 3155.5 | 3.09 | 3.0911 |  | 0 | -18 | 100 | 144.16 | 840.2 | 5.2 | 5.171
0 | -13 | 100 | 39.43 | 118903 | 10.47 | 10.4642 | 0 | 17 | 100 | 432.28 | 3651.5 | 3.58 | 3.577 |  | 0 | -17 | 100 | 148.3 | 864.3 | 5.35 | 5.3195
0 | -14 | 100 | 27.21 | 82055.8 | 7.23 | 7.2211 | 0 | 16 | 100 | 504.58 | 4262.2 | 4.18 | 4.1752 |  | 0 | -16 | 100 | 166.97 | 973.1 | 6.02 | 5.9892
0 | -15 | 100 | 20.36 | 61393.2 | 5.41 | 5.4033 | 0 | 15 | 100 | 637.36 | 5383.8 | 5.28 | 5.274 |  | 0 | -15 | 100 | 193.31 | 1126.6 | 6.98 | 6.934
0 | -15.9 | 100 | 16.02 | 48329.4 | 4.25 | 4.2515 | 0 | 14.1 | 100 | 834.91 | 7052.6 | 6.91 | 6.9086 |  | 0 | -14 | 100 | 255.63 | 1489.8 | 9.22 | 9.1695
0 | -17 | 100 | 13.01 | 39229 | 3.45 | 3.4527 | 0 | 13.1 | 100 | 1219.63 | 10302.3 | 10.1 | 10.0921 |  | 0 | -13 | 100 | 368.37 | 2146.9 | 13.29 | 13.2135
0 | -17.9 | 100 | 10.99 | 33133 | 2.92 | 2.9166 | 0 | 12 | 100 | 2220.32 | 18755.3 | 18.38 | 18.3725 |  | 0 | -12 | 100 | 568.13 | 3311.1 | 20.48 | 20.3789
0 | -18.9 | 100 | 9.59 | 28919.8 | 2.55 | 2.545 | 0 | 11.1 | 100 | 4460.1 | 37674.8 | 36.9 | 36.906 |  | 0 | -11 | 100 | 967.15 | 5636.6 | 34.87 | 34.6917
0 | -20 | 100 | 8.37 | 25242.5 | 2.22 | 2.2213 | 0 | 10.1 | 100 | 7643.31 | 64563.7 | 63.28 | 63.246 |  | 0 | -10 | 100 | 1608.83 | 9376.3 | 58.07 | 57.7088
[7 empty] | 0 | 9 | 100 | 9969.48 | 84213.1 | 82.53 | 82.4944 |  | 0 | -9 | 100 | 2177.96 | 12693.2 | 78.56 | 78.1236
[7 empty] | 0 | 8.1 | 100 | 11016.91 | 93060.8 | 91.19 | 91.1616 |  | 0 | -8 | 100 | 2471.72 | 14405.2 | 89.24 | 88.6608
[7 empty] | 0 | 7 | 100 | 11519.36 | 97305 | 95.4 | 95.3192 |  | 0 | -7 | 100 | 2609.3 | 15207.1 | 94.16 | 93.5958
[7 empty] | 0 | 6.1 | 100 | 11754.95 | 99295.1 | 97.35 | 97.2686 |  | 0 | -6.1 | 100 | 2683.84 | 15641.5 | 96.76 | 96.2695
[7 empty] | 0 | 5 | 100 | 11895.97 | 100486.4 | 98.48 | 98.4355 |  | 0 | -5 | 100 | 2722.98 | 15869.6 | 98.17 | 97.6735
[7 empty] | 0 | 4 | 100 | 11993.51 | 101310.3 | 99.27 | 99.2426 |  | 0 | -4.1 | 100 | 2739.34 | 15964.9 | 98.77 | 98.2603
[7 empty] | 0 | 3 | 100 | 12047.72 | 101768.2 | 99.75 | 99.6912 |  | 0 | -3 | 100 | 2760.89 | 16090.5 | 99.57 | 99.0333
[7 empty] | 0 | 2 | 100 | 12072.71 | 101979.3 | 99.94 | 99.898 |  | 0 | -2 | 100 | 2764.97 | 16114.3 | 99.78 | 99.1797
[7 empty] | 0 | 1 | 100 | 12081.42 | 102052.9 | 99.99 | 99.97 |  | 0 | -1 | 100 | 2787.84 | 16247.6 | 100.58 | 100
[7 empty] | 0 | 0 | 100 | 12075.73 | 102004.8 | 100 | 99.923 |  | 0 | -0.1 | 100 | 2773.7 | 16165.2 | 99.99 | 99.4928
[7 empty] | 0 | -0.9 | 100 | 12081.22 | 102051.1 | 99.95 | 99.9684 |  | 0 | 0.9 | 100 | 2776.95 | 16184.1 | 100.15 | 99.6094
[7 empty] | 0 | -1.9 | 100 | 12085.04 | 102083.4 | 99.96 | 100 |  | 0 | 1.9 | 100 | 2777.49 | 16187.3 | 100.25 | 99.6287
[7 empty] | 0 | -3 | 100 | 12025.75 | 101582.6 | 99.5 | 99.5094 |  | 0 | 3 | 100 | 2764.27 | 16110.2 | 99.75 | 99.1545
[7 empty] | 0 | -4 | 100 | 11949.78 | 100940.9 | 98.89 | 98.8808 |  | 0 | 4.1 | 100 | 2743.57 | 15989.6 | 98.96 | 98.412
[7 empty] | 0 | -4.9 | 100 | 11882.61 | 100373.4 | 98.41 | 98.325 |  | 0 | 4.9 | 100 | 2712.01 | 15805.7 | 97.8 | 97.28
[7 empty] | 0 | -6 | 100 | 11711.42 | 98927.4 | 96.97 | 96.9084 |  | 0 | 5.9 | 100 | 2681.45 | 15627.5 | 96.71 | 96.1838
[7 empty] | 0 | -6.9 | 100 | 11472.81 | 96911.8 | 94.94 | 94.934 |  | 0 | 6.9 | 100 | 2600.13 | 15153.6 | 93.76 | 93.2668
[7 empty] | 0 | -8 | 100 | 10924.65 | 92281.5 | 90.4 | 90.3981 |  | 0 | 7.9 | 100 | 2452.92 | 14295.7 | 88.51 | 87.9864
[7 empty] | 0 | -9 | 100 | 9393.8 | 79350.3 | 77.76 | 77.7308 |  | 0 | 9 | 100 | 2179.13 | 12700 | 78.59 | 78.1655
[7 empty] | 0 | -10 | 100 | 6456.28 | 54536.8 | 53.41 | 53.4237 |  | 0 | 9.9 | 100 | 1642.94 | 9575.1 | 59.28 | 58.9324
[7 empty] | 0 | -11 | 100 | 3489.12 | 29472.9 | 28.89 | 28.8714 |  | 0 | 10.9 | 100 | 1005.01 | 5857.2 | 36.26 | 36.0498
[7 empty] | 0 | -11.9 | 100 | 1858.89 | 15702.2 | 15.39 | 15.3817 |  | 0 | 12 | 100 | 560.17 | 3264.7 | 20.21 | 20.0933
[7 empty] | 0 | -13 | 100 | 1113.09 | 9402.4 | 9.21 | 9.2105 |  | 0 | 12.9 | 100 | 372.14 | 2168.9 | 13.43 | 13.3487
[7 empty] | 0 | -13.9 | 100 | 770.08 | 6504.9 | 6.37 | 6.3722 |  | 0 | 13.9 | 100 | 268.7 | 1566 | 9.69 | 9.6383
[7 empty] | 0 | -15 | 100 | 584.81 | 4939.9 | 4.84 | 4.8391 |  | 0 | 15 | 100 | 207.86 | 1211.4 | 7.5 | 7.456
[7 empty] | 0 | -15.9 | 100 | 477.78 | 4035.9 | 3.95 | 3.9535 |  | 0 | 15.9 | 100 | 167.3 | 975.1 | 6.04 | 6.0011
[7 empty] | 0 | -17 | 100 | 396.25 | 3347.2 | 3.28 | 3.2788 |  | 0 | 17 | 100 | 152.11 | 886.5 | 5.49 | 5.4562
[7 empty] | 0 | -17.9 | 100 | 350.53 | 2961 | 2.9 | 2.9005 |  | 0 | 18 | 100 | 133.04 | 775.4 | 4.8 | 4.7722
[7 empty] | 0 | -18.9 | 100 | 316.87 | 2676.6 | 2.62 | 2.622 |  | 0 | 18.9 | 100 | 121.09 | 705.7 | 4.37 | 4.3435
[7 empty] | 0 | -20 | 100 | 282.44 | 2385.8 | 2.34 | 2.3371 |  | 0 | 20 | 100 | 116.69 | 680.1 | 4.21 | 4.1857
[7 empty] | 0 | -21 | 100 | 251.97 | 2128.4 | 2.09 | 2.085 |  | 0 | 21 | 100 | 113.77 | 663.1 | 4.11 | 4.0809
[7 empty] | 0 | -22 | 100 | 236.2 | 1995.2 | 1.96 | 1.9545 |  | 0 | 21.9 | 100 | 96.66 | 563.3 | 3.49 | 3.4672
[7 empty] | 0 | -22.9 | 100 | 216.76 | 1831 | 1.8 | 1.7936 |  | 0 | 23 | 100 | 95.89 | 558.8 | 3.46 | 3.4396
[7 empty] | 0 | -23.9 | 100 | 205 | 1731.6 | 1.7 | 1.6963 |  | 0 | 23.9 | 100 | 91.89 | 535.6 | 3.32 | 3.2961
[7 empty] | 0 | -25 | 100 | 196.8 | 1662.4 | 1.63 | 1.6285 |  | 0 | 24.9 | 100 | 90.72 | 528.7 | 3.27 | 3.2541
[7 empty] | 0 | -25.9 | 100 | 181.5 | 1533.1 | 1.5 | 1.5019 |  | 0 | 25.9 | 100 | 83.99 | 489.5 | 3.03 | 3.0127
[7 empty] | 0 | -27 | 100 | 173.62 | 1466.6 | 1.44 | 1.4367 |  | 0 | 27 | 100 | 97.86 | 570.3 | 3.53 | 3.5102
[7 empty] | 0 | -28 | 100 | 162.78 | 1375 | 1.35 | 1.347 |  | 0 | 28 | 100 | 91.7 | 534.4 | 3.31 | 3.2893
[7 empty] | 0 | -29 | 100 | 147.75 | 1248.1 | 1.22 | 1.2226 |  | 0 | 29 | 100 | 74.19 | 432.4 | 2.68 | 2.6612
[7 empty] | 0 | -30 | 100 | 141.02 | 1191.2 | 1.17 | 1.1669 |  | 0 | 30 | 100 | 66.15 | 385.5 | 2.39 | 2.3728

## 3x3
Measurement time: | 12/04/2023 19:15:44 | [5 empty] | Measurement time: | 27/05/2023 11:21:09 | [5 empty] | Measurement time: | 27/05/2023 13:34:18
Radiation device: | Novalis TB Accelerator |  |  | Taille de champ |  |  | Radiation device: | Novalis TB Accelerator |  |  |  | Taille de champ |  | Radiation device: | Novalis TB Accelerator |  |  |  | Taille de champ
Energy: | 6 MV (FFF) |  |  | 30.1152 |  |  | Energy: | 6 MV (FFF) |  |  |  | 30.0419 |  | Energy: | 6 MV (FFF) |  |  |  | 30.1148
Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | PTW60019 Diamant | [5 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission |  | 27.3136 | 15 | 55.3668 | -15 | Reference detector: | Stealth CI transmission | [5 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  | 55.9264 | 14.1 | 26.1743 | -16 | SAD: | 1000 mm |  | Côté G |  | Côté D |  | SAD: | 1000 mm
SSD: | 900 mm |  | 50 | 14.2864 | 50 | -15.1838 | SSD: | 900 mm |  | 49.3145 | -15.1 | 61.2589 | 14 | SSD: | 900 mm |  | Côté G |  | Côté D
Field size: | 30 x 30 mm | [5 empty] | Field size: | 30 x 30 mm |  | 75.639 | -14 | 31.8517 | 15 | Field size: | 30 x 30 mm |  | 44.1454 | -15.1 | 66.0378 | 14
Gantry angle: | 0 |  |  | 29.4703 |  |  | Gantry angle: | 0 |  | 50 | -15.0714 | 50 | 14.3829 | Gantry angle: | 0 |  | 69.228 | -14 | 42.0437 | 14.9
Measurement medium: | Water | [5 empty] | Measurement medium: | Water | [5 empty] | Measurement medium: | Water |  | 50 | -14.8432 | 50 | 14.6016
Scan type: | Inline | [5 empty] | Scan type: | Inline |  |  | 29.4542 |  |  | Scan type: | Inline
Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep |  |  | 29.4448
Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
30 | 0 | 100 | 2.53 | 29960.2 | 2.54 | 2.5285 | -30 | 0 | 100 | 324.34 | 2739.7 | 2.56 | 2.5657 | -30 | 0 | 100 | 134.13 | 781.7 | 4.05 | 4.6085
29.1 | 0 | 100 | 2.69 | 31861.5 | 2.7 | 2.6884 | -29.1 | 0 | 100 | 351.13 | 2966 | 2.78 | 2.7776 | -29 | 0 | 100 | 136.11 | 793.2 | 4.11 | 4.6766
28 | 0 | 100 | 2.86 | 33912.9 | 2.87 | 2.8583 | -28.1 | 0 | 100 | 382.72 | 3232.9 | 3.03 | 3.0275 | -28.1 | 0 | 100 | 154.86 | 902.5 | 4.67 | 5.3208
27 | 0 | 100 | 3.06 | 36297 | 3.07 | 3.0582 | -27.1 | 0 | 100 | 404.92 | 3420.4 | 3.2 | 3.2031 | -27 | 0 | 100 | 171.72 | 1000.8 | 5.18 | 5.9001
26.1 | 0 | 100 | 3.3 | 39177.6 | 3.32 | 3.298 | -26.1 | 0 | 100 | 431.65 | 3646.2 | 3.41 | 3.4145 | -26 | 0 | 100 | 163.64 | 953.7 | 4.94 | 5.6225
25 | 0 | 100 | 3.55 | 42139.1 | 3.57 | 3.5479 | -25.1 | 0 | 100 | 463.89 | 3918.5 | 3.67 | 3.6696 | -25.1 | 0 | 100 | 153.83 | 896.5 | 4.65 | 5.2854
24.1 | 0 | 100 | 3.84 | 45524.4 | 3.85 | 3.8377 | -24 | 0 | 100 | 509.22 | 4301.4 | 4.03 | 4.0281 | -24 | 0 | 100 | 150.35 | 876.2 | 4.54 | 5.1658
23 | 0 | 100 | 4.2 | 49819.6 | 4.22 | 4.1975 | -23.1 | 0 | 100 | 555.09 | 4688.9 | 4.39 | 4.391 | -23 | 0 | 100 | 163.92 | 955.3 | 4.95 | 5.6321
22.1 | 0 | 100 | 4.66 | 55270.6 | 4.68 | 4.6572 | -22 | 0 | 100 | 620.71 | 5243.2 | 4.91 | 4.9101 | -22.1 | 0 | 100 | 175.17 | 1020.9 | 5.29 | 6.0186
21.1 | 0 | 100 | 5.13 | 60868.4 | 5.15 | 5.1269 | -21 | 0 | 100 | 698.78 | 5902.7 | 5.53 | 5.5276 | -21 | 0 | 100 | 189.76 | 1105.9 | 5.73 | 6.5199
20 | 0 | 100 | 5.88 | 69760.5 | 5.91 | 5.8765 | -20.1 | 0 | 100 | 808.12 | 6826.3 | 6.39 | 6.3926 | -20 | 0 | 100 | 228.14 | 1329.6 | 6.89 | 7.8386
19 | 0 | 100 | 6.9 | 81851.7 | 6.93 | 6.8959 | -19.1 | 0 | 100 | 978.23 | 8263.2 | 7.74 | 7.7382 | -19.1 | 0 | 100 | 276.5 | 1611.4 | 8.34 | 9.5002
18 | 0 | 100 | 8.61 | 102123 | 8.65 | 8.6048 | -18.1 | 0 | 100 | 1250.76 | 10565.3 | 9.9 | 9.894 | -18 | 0 | 100 | 345.26 | 2012.2 | 10.43 | 11.8627
17 | 0 | 100 | 10.76 | 127616.7 | 10.8 | 10.7535 | -17 | 0 | 100 | 1785.1 | 15078.9 | 14.13 | 14.1209 | -17 | 0 | 100 | 475.41 | 2770.7 | 14.35 | 16.3345
16 | 0 | 100 | 15.36 | 182192.3 | 15.42 | 15.3508 | -16 | 0 | 100 | 3143.7 | 26555.1 | 24.89 | 24.868 | -16.1 | 0 | 100 | 731.18 | 4261.3 | 22.07 | 25.1224
15 | 0 | 100 | 27.33 | 324069.3 | 27.43 | 27.3136 | -15.1 | 0 | 100 | 6234.12 | 52660.2 | 49.32 | 49.3145 | -15.1 | 0 | 100 | 1284.84 | 7488 | 38.79 | 44.1454
14.1 | 0 | 100 | 55.96 | 663559.5 | 56.17 | 55.9264 | -14 | 0 | 100 | 9561.94 | 80770.6 | 75.65 | 75.639 | -14 | 0 | 100 | 2014.86 | 11742.6 | 60.84 | 69.228
13 | 0 | 100 | 80.19 | 950948.9 | 80.5 | 80.1419 | -13.1 | 0 | 100 | 11156.14 | 94236.9 | 88.23 | 88.2498 | -13.1 | 0 | 100 | 2462.53 | 14351.7 | 74.35 | 84.6094
12 | 0 | 100 | 89.12 | 1056862.4 | 89.47 | 89.0666 | -12.1 | 0 | 100 | 11728.36 | 99070.5 | 92.85 | 92.7763 | -12 | 0 | 100 | 2654.21 | 15468.8 | 80.13 | 91.1952
11.1 | 0 | 100 | 93.03 | 1103210.7 | 93.39 | 92.9742 | -11 | 0 | 100 | 12060.13 | 101873 | 95.46 | 95.4007 | -11.1 | 0 | 100 | 2751.37 | 16035 | 83.1 | 94.5335
10 | 0 | 100 | 95.33 | 1130418 | 95.7 | 95.2728 | -10.1 | 0 | 100 | 12243.9 | 103425.3 | 96.92 | 96.8544 | -10 | 0 | 100 | 2797.99 | 16306.8 | 84.51 | 96.1353
9 | 0 | 100 | 96.41 | 1143310.3 | 96.79 | 96.3522 | -9.1 | 0 | 100 | 12340.32 | 104239.8 | 97.67 | 97.6171 | -9 | 0 | 100 | 2831.29 | 16500.8 | 85.52 | 97.2795
8 | 0 | 100 | 97.56 | 1156875.2 | 97.94 | 97.5015 | -8 | 0 | 100 | 12445.44 | 105127.8 | 98.5 | 98.4487 | -8 | 0 | 100 | 2849.25 | 16605.5 | 86.05 | 97.8966
7.1 | 0 | 100 | 98.27 | 1165369.5 | 98.66 | 98.2111 | -7 | 0 | 100 | 12508.24 | 105658.2 | 99.03 | 98.9455 | -7.1 | 0 | 100 | 2865.28 | 16698.9 | 86.46 | 98.4473
6.1 | 0 | 100 | 98.78 | 1171322.5 | 99.16 | 98.7208 | -6 | 0 | 100 | 12565.22 | 106139.6 | 99.41 | 99.3962 | -6.1 | 0 | 100 | 2884.5 | 16810.9 | 87.07 | 99.1077
5 | 0 | 100 | 99.02 | 1174265.7 | 99.41 | 98.9606 | -5.1 | 0 | 100 | 12580.58 | 106269.3 | 99.53 | 99.5177 | -5 | 0 | 100 | 2905.48 | 16933.2 | 87.73 | 99.8286
4 | 0 | 100 | 99.32 | 1177726.5 | 99.7 | 99.2604 | -4 | 0 | 100 | 12608.71 | 106506.9 | 99.77 | 99.7402 | -4.1 | 0 | 100 | 2910.47 | 16962.3 | 87.85 | 100
3 | 0 | 100 | 99.91 | 1184731.6 | 100.3 | 99.8501 | -3.1 | 0 | 100 | 12623.63 | 106632.9 | 99.92 | 99.8582 | -3 | 0 | 100 | 2909.59 | 16957.1 | 87.8 | 99.9698
2.1 | 0 | 100 | 99.71 | 1182338.9 | 100.09 | 99.6502 | -2 | 0 | 100 | 12641.55 | 106784.3 | 100.05 | 100 | -2 | 0 | 100 | 2905.79 | 16935 | 87.73 | 99.8392
1 | 0 | 100 | 99.57 | 1180734.2 | 99.96 | 99.5103 | -1.1 | 0 | 100 | 12636.34 | 106740.3 | 100 | 99.9588 | -1 | 0 | 100 | 2895.79 | 16876.7 | 87.45 | 99.4956
0.1 | 0 | 100 | 99.59 | 1180936.3 | 99.97 | 99.5303 | 0 | 0 | 100 | 12635.64 | 106734.4 | 100 | 99.9532 | 0 | 0 | 100 | 2904.68 | 16928.5 | 87.74 | 99.8011
-1 | 0 | 100 | 100.06 | 1186498.6 | 100.44 | 100 | 1.1 | 0 | 100 | 12618.01 | 106585.5 | 99.93 | 99.8138 | 1 | 0 | 100 | 2882.65 | 16800.2 | 87.05 | 99.0441
-2 | 0 | 100 | 99.71 | 1182394.6 | 100.1 | 99.6502 | 2 | 0 | 100 | 12606.32 | 106486.7 | 99.77 | 99.7213 | 2 | 0 | 100 | 2880.43 | 16787.2 | 87.02 | 98.9679
-3 | 0 | 100 | 99.89 | 1184528.5 | 100.28 | 99.8301 | 3 | 0 | 100 | 12586.76 | 106321.5 | 99.67 | 99.5666 | 2.9 | 0 | 100 | 2888.82 | 16836.1 | 87.23 | 99.2561
-4 | 0 | 100 | 99.64 | 1181611.1 | 100.03 | 99.5803 | 4.1 | 0 | 100 | 12558.92 | 106086.3 | 99.44 | 99.3464 | 4 | 0 | 100 | 2891.23 | 16850.2 | 87.3 | 99.3389
-5 | 0 | 100 | 99.32 | 1177819.4 | 99.71 | 99.2604 | 5 | 0 | 100 | 12528.97 | 105833.3 | 99.21 | 99.1094 | 5 | 0 | 100 | 2899.47 | 16898.2 | 87.57 | 99.6221
-6 | 0 | 100 | 99.16 | 1175814.3 | 99.54 | 99.1005 | 6 | 0 | 100 | 12502.21 | 105607.3 | 98.97 | 98.8978 | 6 | 0 | 100 | 2886.78 | 16824.2 | 87.17 | 99.186
-7.1 | 0 | 100 | 98.67 | 1170072.5 | 99.05 | 98.6108 | 7 | 0 | 100 | 12435.05 | 105040 | 98.52 | 98.3665 | 7 | 0 | 100 | 2892.7 | 16858.7 | 87.39 | 99.3894
-8 | 0 | 100 | 98.37 | 1166449 | 98.75 | 98.311 | 8 | 0 | 100 | 12375.73 | 104538.9 | 98.01 | 97.8973 | 8 | 0 | 100 | 2873.49 | 16746.7 | 86.78 | 98.7294
-9 | 0 | 100 | 97.62 | 1157574.6 | 98 | 97.5615 | 9 | 0 | 100 | 12252.23 | 103495.6 | 97.03 | 96.9203 | 9 | 0 | 100 | 2842.44 | 16565.8 | 85.84 | 97.6626
-10 | 0 | 100 | 96.74 | 1147138.9 | 97.11 | 96.682 | 10 | 0 | 100 | 12094.59 | 102164.1 | 95.72 | 95.6733 | 9.9 | 0 | 100 | 2804.86 | 16346.8 | 84.72 | 96.3714
-11 | 0 | 100 | 95.22 | 1129166.6 | 95.59 | 95.1629 | 11 | 0 | 100 | 11840.54 | 100018.1 | 93.77 | 93.6637 | 11 | 0 | 100 | 2726.56 | 15890.4 | 82.37 | 93.6811
-12 | 0 | 100 | 93.23 | 1105556 | 93.59 | 93.1741 | 12 | 0 | 100 | 11376.23 | 96096.1 | 90.1 | 89.9908 | 11.9 | 0 | 100 | 2646.35 | 15423 | 79.88 | 90.9252
-12.9 | 0 | 100 | 89.73 | 1064106.5 | 90.08 | 89.6762 | 13 | 0 | 100 | 10275.13 | 86794.9 | 81.37 | 81.2806 | 13 | 0 | 100 | 2409.73 | 14044 | 72.76 | 82.7952
-14 | 0 | 100 | 81.17 | 962486.5 | 81.48 | 81.1213 | 14 | 0 | 100 | 7744.08 | 65414.9 | 61.3 | 61.2589 | 14 | 0 | 100 | 1922.01 | 11201.5 | 58.04 | 66.0378
-15 | 0 | 100 | 55.4 | 656919.7 | 55.61 | 55.3668 | 15 | 0 | 100 | 4026.55 | 34012.6 | 31.89 | 31.8517 | 14.9 | 0 | 100 | 1223.67 | 7131.6 | 36.95 | 42.0437
-16 | 0 | 100 | 26.19 | 310594.1 | 26.29 | 26.1743 | 15.9 | 0 | 100 | 2213.7 | 18699.3 | 17.54 | 17.5113 | 16 | 0 | 100 | 678.84 | 3956.3 | 20.5 | 23.3241
-17.1 | 0 | 100 | 14.83 | 175820.8 | 14.88 | 14.8211 | 17 | 0 | 100 | 1420.86 | 12002.1 | 11.25 | 11.2396 | 17 | 0 | 100 | 437.9 | 2552.1 | 13.22 | 15.0457
-18 | 0 | 100 | 10.73 | 127216.3 | 10.77 | 10.7236 | 17.9 | 0 | 100 | 1062.5 | 8975 | 8.41 | 8.4048 | 18 | 0 | 100 | 333.44 | 1943.3 | 10.07 | 11.4566
-18.9 | 0 | 100 | 8.54 | 101243.2 | 8.57 | 8.5349 | 18.9 | 0 | 100 | 851.35 | 7191.5 | 6.74 | 6.7345 | 18.9 | 0 | 100 | 276.59 | 1612 | 8.35 | 9.5033
-20 | 0 | 100 | 6.99 | 82912.8 | 7.02 | 6.9858 | 19.9 | 0 | 100 | 727.83 | 6148 | 5.76 | 5.7574 | 20.1 | 0 | 100 | 214.44 | 1249.8 | 6.48 | 7.3679
-21 | 0 | 100 | 5.92 | 70167.5 | 5.94 | 5.9165 | 21 | 0 | 100 | 625.96 | 5287.6 | 4.96 | 4.9516 | 20.9 | 0 | 100 | 206.46 | 1203.2 | 6.24 | 7.0937
-22 | 0 | 100 | 5.23 | 62076.2 | 5.26 | 5.2269 | 21.9 | 0 | 100 | 560.11 | 4731.3 | 4.44 | 4.4307 | 22 | 0 | 100 | 189.82 | 1106.3 | 5.73 | 6.522
-23 | 0 | 100 | 4.68 | 55464.5 | 4.7 | 4.6772 | 23 | 0 | 100 | 514.2 | 4343.5 | 4.07 | 4.0675 | 23 | 0 | 100 | 166.13 | 968.2 | 5.02 | 5.708
-23.9 | 0 | 100 | 4.26 | 50477 | 4.27 | 4.2574 | 24 | 0 | 100 | 469.64 | 3967.1 | 3.72 | 3.7151 | 24 | 0 | 100 | 154.42 | 899.9 | 4.66 | 5.3057
-25 | 0 | 100 | 3.85 | 45614.5 | 3.86 | 3.8477 | 25 | 0 | 100 | 425.67 | 3595.7 | 3.37 | 3.3672 | 25 | 0 | 100 | 161.14 | 939.1 | 4.87 | 5.5366
-26 | 0 | 100 | 3.58 | 42415.8 | 3.59 | 3.5779 | 26 | 0 | 100 | 399.44 | 3374.1 | 3.16 | 3.1597 | 26 | 0 | 100 | 140.86 | 820.9 | 4.25 | 4.8398
-27 | 0 | 100 | 3.28 | 38913.9 | 3.29 | 3.278 | 27 | 0 | 100 | 382.76 | 3233.2 | 3.03 | 3.0278 | 27 | 0 | 100 | 142.46 | 830.3 | 4.31 | 4.8947
-28 | 0 | 100 | 3.06 | 36258.1 | 3.07 | 3.0582 | 27.9 | 0 | 100 | 363.91 | 3074 | 2.88 | 2.8787 | 28 | 0 | 100 | 132.35 | 771.4 | 4 | 4.5474
-29 | 0 | 100 | 2.84 | 33680.2 | 2.85 | 2.8383 | 29 | 0 | 100 | 332.24 | 2806.5 | 2.63 | 2.6282 | 29 | 0 | 100 | 125.99 | 734.3 | 3.81 | 4.3289
-30 | 0 | 100 | 2.67 | 31606.2 | 2.68 | 2.6684 | 30 | 0 | 100 | 318.4 | 2689.6 | 2.52 | 2.5187 | 30 | 0 | 100 | 120.82 | 704.2 | 3.65 | 4.1512
Measurement time: | 12/04/2023 19:18:35 | [5 empty] | Measurement time: | 27/05/2023 11:23:54 | [5 empty] | Measurement time: | 27/05/2023 13:37:05
Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator | [5 empty] | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF) | [5 empty] | Energy: | 6 MV (FFF)
Controller: | CCU | [5 empty] | Controller: | CCU | [5 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [5 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele |  | Côté G |  | Côté D |  | Field detector: | PTW60019 Diamant | [5 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission |  | 28.0245 | 16 | 64.8091 | -15 | Reference detector: | Stealth CI transmission | [5 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm |  | 57.1815 | 15 | 36.7646 | -16 | SAD: | 1000 mm | [5 empty] | SAD: | 1000 mm
SSD: | 900 mm |  | 50 | 15.2463 | 50 | -15.5281 | SSD: | 900 mm |  | Côté G |  | Côté D |  | SSD: | 900 mm |  | Côté G |  | Côté D
Field size: | 30 x 30 mm | [5 empty] | Field size: | 30 x 30 mm |  | 39.276 | 16 | 53.8383 | -14.9 | Field size: | 30 x 30 mm |  | 35.8067 | 16.1 | 59.4311 | -15
Gantry angle: | 0 |  |  | 30.7744 |  |  | Gantry angle: | 0 |  | 65.0454 | 15 | 29.4501 | -15.9 | Gantry angle: | 0 |  | 55.9349 | 15 | 39.6086 | -16
Measurement medium: | Water | [5 empty] | Measurement medium: | Water |  | 50 | 15.5838 | 50 | -15.0574 | Measurement medium: | Water |  | 50 | 15.3243 | 50 | -15.4758
Scan type: | Crossline | [5 empty] | Scan type: | Crossline | [5 empty] | Scan type: | Crossline
Scan mode: | StepByStep | [5 empty] | Scan mode: | StepByStep |  |  | 30.6412 |  |  | Scan mode: | StepByStep |  |  | 30.8001
Wedge: | none | [5 empty] | Wedge: | none | [5 empty] | Wedge: | none
Points [mm]: | [6 empty] | Points [mm]: | [6 empty] | Points [mm]:
Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Normalized field | Current field | Ratio | Dose relative (%)
0 | 30 | 100 | 2.66 | 31505.5 | 2.67 | 2.6661 | 0 | 30 | 100 | 338.42 | 2858.7 | 2.68 | 2.676 | 0 | 30 | 100 | 128.32 | 747.8 | 3.88 | 4.3931
0 | 29 | 100 | 2.83 | 33549.7 | 2.84 | 2.8365 | 0 | 29.1 | 100 | 358.4 | 3027.4 | 2.84 | 2.8339 | 0 | 29 | 100 | 138.89 | 809.5 | 4.2 | 4.755
0 | 28.1 | 100 | 3.03 | 35917.4 | 3.04 | 3.037 | 0 | 28 | 100 | 376.09 | 3176.9 | 2.98 | 2.9738 | 0 | 28.1 | 100 | 145.25 | 846.5 | 4.39 | 4.9727
0 | 27.1 | 100 | 3.25 | 38549.8 | 3.26 | 3.2575 | 0 | 27.1 | 100 | 406.96 | 3437.6 | 3.22 | 3.2179 | 0 | 27.1 | 100 | 145.64 | 848.8 | 4.4 | 4.9861
0 | 26.1 | 100 | 3.48 | 41259.3 | 3.49 | 3.488 | 0 | 26 | 100 | 435.67 | 3680.1 | 3.45 | 3.4449 | 0 | 26 | 100 | 160.55 | 935.7 | 4.85 | 5.4965
0 | 25 | 100 | 3.81 | 45149.9 | 3.82 | 3.8188 | 0 | 25 | 100 | 474.96 | 4012.1 | 3.76 | 3.7556 | 0 | 25 | 100 | 169.46 | 987.6 | 5.12 | 5.8016
0 | 24 | 100 | 4.16 | 49320.9 | 4.18 | 4.1696 | 0 | 24 | 100 | 521.18 | 4402.4 | 4.12 | 4.1211 | 0 | 24 | 100 | 182.55 | 1063.9 | 5.52 | 6.2497
0 | 23.1 | 100 | 4.58 | 54284.5 | 4.6 | 4.5906 | 0 | 22.9 | 100 | 586.47 | 4954 | 4.64 | 4.6373 | 0 | 23.1 | 100 | 190.61 | 1110.9 | 5.76 | 6.5257
0 | 21.9 | 100 | 5.18 | 61470.4 | 5.21 | 5.1919 | 0 | 22 | 100 | 634.25 | 5357.5 | 5.02 | 5.0151 | 0 | 22 | 100 | 207.04 | 1206.7 | 6.26 | 7.0882
0 | 21 | 100 | 5.84 | 69230 | 5.86 | 5.8535 | 0 | 21 | 100 | 736.26 | 6219.3 | 5.83 | 5.8218 | 0 | 21 | 100 | 227.32 | 1324.8 | 6.87 | 7.7825
0 | 20.1 | 100 | 6.91 | 81939.7 | 6.94 | 6.9259 | 0 | 19.9 | 100 | 893.73 | 7549.4 | 7.07 | 7.0669 | 0 | 19.9 | 100 | 271.82 | 1584.2 | 8.21 | 9.3059
0 | 19 | 100 | 8.46 | 100350.2 | 8.5 | 8.4795 | 0 | 19 | 100 | 1091.19 | 9217.4 | 8.64 | 8.6283 | 0 | 19 | 100 | 342.23 | 1994.5 | 10.34 | 11.7165
0 | 18 | 100 | 10.93 | 129650 | 10.98 | 10.9552 | 0 | 18 | 100 | 1492.37 | 12606.2 | 11.81 | 11.8005 | 0 | 18 | 100 | 447.96 | 2610.7 | 13.54 | 15.3362
0 | 17 | 100 | 15.79 | 187273.3 | 15.86 | 15.8264 | 0 | 17 | 100 | 2596.19 | 21930.3 | 20.55 | 20.5286 | 0 | 17 | 100 | 667.82 | 3892 | 20.19 | 22.8633
0 | 16 | 100 | 27.96 | 331596.9 | 28.08 | 28.0245 | 0 | 16 | 100 | 4967.12 | 41957.7 | 39.33 | 39.276 | 0 | 16.1 | 100 | 1045.89 | 6095.5 | 31.62 | 35.8067
0 | 15 | 100 | 57.05 | 676469.2 | 57.28 | 57.1815 | 0 | 15 | 100 | 8226.1 | 69486.6 | 65.12 | 65.0454 | 0 | 15 | 100 | 1633.82 | 9521.9 | 49.39 | 55.9349
0 | 14 | 100 | 78.25 | 927943.5 | 78.58 | 78.4304 | 0 | 14 | 100 | 10353.5 | 87456.9 | 81.98 | 81.8672 | 0 | 14 | 100 | 2186.84 | 12744.9 | 66.08 | 74.8679
0 | 13.1 | 100 | 87.64 | 1039301.4 | 88.01 | 87.842 | 0 | 13.1 | 100 | 11342.08 | 95807.5 | 89.8 | 89.6841 | 0 | 13 | 100 | 2499.78 | 14568.8 | 75.58 | 85.5816
0 | 12 | 100 | 92.58 | 1097801.9 | 92.96 | 92.7934 | 0 | 12 | 100 | 11861.22 | 100192.8 | 93.94 | 93.7891 | 0 | 12 | 100 | 2654.83 | 15472.4 | 80.24 | 90.8899
0 | 11 | 100 | 94.9 | 1125399.4 | 95.3 | 95.1188 | 0 | 11 | 100 | 12116.81 | 102351.8 | 95.93 | 95.8101 | 0 | 11 | 100 | 2744.72 | 15996.3 | 82.97 | 93.9673
0 | 10 | 100 | 96.49 | 1144199.1 | 96.89 | 96.7124 | 0 | 10 | 100 | 12294.49 | 103852.6 | 97.34 | 97.215 | 0 | 10 | 100 | 2809.04 | 16371.2 | 84.9 | 96.1694
0 | 9 | 100 | 97.25 | 1153223.2 | 97.65 | 97.4742 | 0 | 9 | 100 | 12390.16 | 104660.8 | 98.11 | 97.9715 | 0 | 9 | 100 | 2836.18 | 16529.3 | 85.73 | 97.0985
0 | 8 | 100 | 98 | 1162096.7 | 98.4 | 98.2259 | 0 | 8 | 100 | 12459.73 | 105248.5 | 98.66 | 98.5216 | 0 | 8.1 | 100 | 2861.67 | 16677.9 | 86.53 | 97.9712
0 | 7.1 | 100 | 98.54 | 1168534.2 | 98.95 | 98.7672 | 0 | 7 | 100 | 12516.93 | 105731.7 | 99.08 | 98.9739 | 0 | 7.1 | 100 | 2869.35 | 16722.7 | 86.74 | 98.2341
0 | 6 | 100 | 99.08 | 1174895.4 | 99.49 | 99.3084 | 0 | 6.1 | 100 | 12554.7 | 106050.7 | 99.41 | 99.2725 | 0 | 6 | 100 | 2895.02 | 16872.3 | 87.44 | 99.113
0 | 5.1 | 100 | 99.57 | 1180703.8 | 99.98 | 99.7995 | 0 | 5 | 100 | 12571.41 | 106191.8 | 99.56 | 99.4047 | 0 | 5.1 | 100 | 2892.8 | 16859.3 | 87.42 | 99.037
0 | 4 | 100 | 99.5 | 1179902.2 | 99.91 | 99.7294 | 0 | 4 | 100 | 12606.17 | 106485.5 | 99.79 | 99.6795 | 0 | 4 | 100 | 2897.87 | 16888.8 | 87.62 | 99.2105
0 | 3 | 100 | 99.49 | 1179831.8 | 99.91 | 99.7194 | 0 | 3.1 | 100 | 12612.21 | 106536.5 | 99.89 | 99.7273 | 0 | 3.1 | 100 | 2904.14 | 16925.4 | 87.81 | 99.4252
0 | 2 | 100 | 99.49 | 1179775.4 | 99.9 | 99.7194 | 0 | 2 | 100 | 12646.7 | 106827.8 | 100.14 | 100 | 0 | 2 | 100 | 2920.71 | 17021.9 | 88.3 | 99.9925
0 | 1 | 100 | 99.77 | 1183129.3 | 100.19 | 100 | 0 | 1 | 100 | 12633.67 | 106717.7 | 100.01 | 99.897 | 0 | 1 | 100 | 2914.67 | 16986.7 | 88.13 | 99.7857
0 | 0 | 100 | 99.58 | 1180883.2 | 100 | 99.8096 | 0 | 0 | 100 | 12629.5 | 106682.5 | 100 | 99.864 | 0 | 0.1 | 100 | 2914.75 | 16987.2 | 88.1 | 99.7884
0 | -0.9 | 100 | 99.67 | 1181917.5 | 100.08 | 99.8998 | 0 | -1 | 100 | 12624.95 | 106644.1 | 99.99 | 99.828 | 0 | -0.9 | 100 | 2915.39 | 16990.9 | 88.17 | 99.8103
0 | -2 | 100 | 99.69 | 1182166.6 | 100.1 | 99.9198 | 0 | -2 | 100 | 12628.4 | 106673.2 | 99.97 | 99.8553 | 0 | -2 | 100 | 2915.4 | 16991 | 88.09 | 99.8107
0 | -3 | 100 | 99.28 | 1177299.2 | 99.69 | 99.5089 | 0 | -2.9 | 100 | 12622.67 | 106624.9 | 99.95 | 99.81 | 0 | -2.9 | 100 | 2920.93 | 17023.2 | 88.26 | 100
0 | -4 | 100 | 99.39 | 1178615.9 | 99.8 | 99.6191 | 0 | -3.9 | 100 | 12567.55 | 106159.3 | 99.56 | 99.3741 | 0 | -4 | 100 | 2909.16 | 16954.7 | 87.93 | 99.597
0 | -4.9 | 100 | 99.22 | 1176569.1 | 99.63 | 99.4487 | 0 | -4.9 | 100 | 12556.29 | 106064.1 | 99.43 | 99.2851 | 0 | -5 | 100 | 2896.03 | 16878.1 | 87.51 | 99.1475
0 | -6 | 100 | 98.84 | 1172042 | 99.25 | 99.0679 | 0 | -6 | 100 | 12527.29 | 105819.1 | 99.19 | 99.0558 | 0 | -5.9 | 100 | 2900.74 | 16905.6 | 87.69 | 99.3088
0 | -7 | 100 | 98.46 | 1167582.8 | 98.87 | 98.687 | 0 | -6.9 | 100 | 12462.81 | 105274.5 | 98.74 | 98.5459 | 0 | -6.9 | 100 | 2879.79 | 16783.4 | 87.07 | 98.5915
0 | -7.9 | 100 | 98.09 | 1163181.7 | 98.5 | 98.3161 | 0 | -8 | 100 | 12408.76 | 104817.9 | 98.3 | 98.1186 | 0 | -8 | 100 | 2862.08 | 16680.2 | 86.54 | 97.9852
0 | -8.9 | 100 | 97.31 | 1153980 | 97.72 | 97.5343 | 0 | -9 | 100 | 12325.93 | 104118.3 | 97.7 | 97.4636 | 0 | -9 | 100 | 2835.35 | 16524.4 | 85.66 | 97.0701
0 | -9.9 | 100 | 96.41 | 1143312.3 | 96.81 | 96.6323 | 0 | -10 | 100 | 12213.94 | 103172.2 | 96.81 | 96.5781 | 0 | -9.9 | 100 | 2800.08 | 16318.9 | 84.68 | 95.8626
0 | -11 | 100 | 95.49 | 1132338.9 | 95.88 | 95.7101 | 0 | -11 | 100 | 12053.01 | 101812.9 | 95.44 | 95.3056 | 0 | -11 | 100 | 2759.79 | 16084.1 | 83.45 | 94.4833
0 | -12 | 100 | 93.65 | 1110544.8 | 94.04 | 93.8659 | 0 | -11.9 | 100 | 11786.93 | 99565.2 | 93.41 | 93.2016 | 0 | -12 | 100 | 2680.76 | 15623.5 | 81.05 | 91.7776
0 | -12.9 | 100 | 90.81 | 1076835.2 | 91.18 | 91.0193 | 0 | -12.9 | 100 | 11230.69 | 94866.6 | 89 | 88.8033 | 0 | -12.9 | 100 | 2566.47 | 14957.4 | 77.58 | 87.8648
0 | -13.9 | 100 | 84.51 | 1002184.3 | 84.86 | 84.7048 | 0 | -14 | 100 | 9614.77 | 81216.8 | 76.15 | 76.0259 | 0 | -13.9 | 100 | 2275.92 | 13264.1 | 68.79 | 77.9176
0 | -15 | 100 | 64.66 | 766706.2 | 64.92 | 64.8091 | 0 | -14.9 | 100 | 6808.77 | 57514.3 | 53.93 | 53.8383 | 0 | -15 | 100 | 1735.94 | 10117.1 | 52.52 | 59.4311
0 | -16 | 100 | 36.68 | 434920.1 | 36.83 | 36.7646 | 0 | -15.9 | 100 | 3724.47 | 31460.9 | 29.52 | 29.4501 | 0 | -16 | 100 | 1156.94 | 6742.7 | 34.99 | 39.6086
0 | -17 | 100 | 19.59 | 232279.2 | 19.67 | 19.6352 | 0 | -17 | 100 | 2086.08 | 17621.3 | 16.54 | 16.4951 | 0 | -17 | 100 | 721.15 | 4202.9 | 21.8 | 24.6891
0 | -17.9 | 100 | 12.65 | 149992.7 | 12.7 | 12.6792 | 0 | -17.9 | 100 | 1379.3 | 11651.1 | 10.93 | 10.9064 | 0 | -17.9 | 100 | 483.74 | 2819.2 | 14.63 | 16.5612
0 | -19 | 100 | 9.26 | 109835.7 | 9.3 | 9.2813 | 0 | -18.9 | 100 | 1019.57 | 8612.4 | 8.08 | 8.0619 | 0 | -19 | 100 | 356.73 | 2079 | 10.79 | 12.2129
0 | -20 | 100 | 7.39 | 87598.7 | 7.42 | 7.407 | 0 | -20 | 100 | 815.68 | 6890.1 | 6.46 | 6.4497 | 0 | -19.9 | 100 | 266.89 | 1555.4 | 8.08 | 9.1372
0 | -20.9 | 100 | 6.17 | 73188.5 | 6.2 | 6.1842 | 0 | -21 | 100 | 685.5 | 5790.5 | 5.43 | 5.4204 | 0 | -21 | 100 | 237.07 | 1381.7 | 7.17 | 8.1163
0 | -21.9 | 100 | 5.37 | 63632 | 5.39 | 5.3824 | 0 | -21.9 | 100 | 606.1 | 5119.7 | 4.8 | 4.7926 | 0 | -22 | 100 | 223.34 | 1301.6 | 6.76 | 7.6462
0 | -23 | 100 | 4.77 | 56510.7 | 4.79 | 4.781 | 0 | -23 | 100 | 537.04 | 4536.4 | 4.25 | 4.2465 | 0 | -23 | 100 | 203.57 | 1186.4 | 6.16 | 6.9694
0 | -23.9 | 100 | 4.25 | 50442.4 | 4.27 | 4.2598 | 0 | -24 | 100 | 489.28 | 4133 | 3.88 | 3.8688 | 0 | -24 | 100 | 188.37 | 1097.8 | 5.7 | 6.449
0 | -24.9 | 100 | 3.89 | 46096.7 | 3.9 | 3.899 | 0 | -24.9 | 100 | 453.59 | 3831.5 | 3.59 | 3.5866 | 0 | -24.9 | 100 | 164.45 | 958.4 | 4.97 | 5.6301
0 | -26 | 100 | 3.58 | 42415.2 | 3.59 | 3.5883 | 0 | -26 | 100 | 413.88 | 3496.1 | 3.28 | 3.2726 | 0 | -25.9 | 100 | 169.9 | 990.2 | 5.14 | 5.8166
0 | -27 | 100 | 3.29 | 39060.3 | 3.31 | 3.2976 | 0 | -26.9 | 100 | 395.23 | 3338.6 | 3.13 | 3.1252 | 0 | -26.9 | 100 | 165.4 | 964 | 5 | 5.6626
0 | -28 | 100 | 3.1 | 36714.8 | 3.11 | 3.1071 | 0 | -27.9 | 100 | 369.53 | 3121.5 | 2.93 | 2.9219 | 0 | -28 | 100 | 159.8 | 931.3 | 4.83 | 5.4709
0 | -29 | 100 | 2.85 | 33811 | 2.86 | 2.8566 | 0 | -29 | 100 | 342.38 | 2892.1 | 2.71 | 2.7073 | 0 | -28.9 | 100 | 155.33 | 905.3 | 4.7 | 5.3178
0 | -30 | 100 | 2.65 | 31448 | 2.66 | 2.6561 | 0 | -30 | 100 | 321.71 | 2717.5 | 2.55 | 2.5438 | 0 | -30 | 100 | 143.31 | 835.2 | 4.33 | 4.9063