## 0,5x0,5
Measurement time: | 12/04/2023 18:07:44 | [6 empty] | Measurement time: | 27/05/2023 10:14:20 | [6 empty] | Measurement time: | 27/05/2023 12:28:57
Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF)
Controller: | CCU | [6 empty] | Controller: | CCU | [6 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele | [6 empty] | Field detector: | PTW60019 Diamant | [6 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm
SSD: | 900 mm | [6 empty] | SSD: | 900 mm | [6 empty] | SSD: | 900 mm
Field size: | 5 x 5 mm | [6 empty] | Field size: | 5 x 5 mm | [6 empty] | Field size: | 5 x 5 mm
Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0
Measurement medium: | Water | [6 empty] | Measurement medium: | Water | [6 empty] | Measurement medium: | Water
Scan type: | Beam | [6 empty] | Scan type: | Beam | [6 empty] | Scan type: | Beam
Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep
Wedge: | none | [6 empty] | Wedge: | none | [6 empty] | Wedge: | none
Points [mm]: | [7 empty] | Points [mm]: | [7 empty] | Points [mm]:
Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 250 | 25 | 819.69 | 286442 | 16.13 | 16.1449 | 0 | 0 | 250 | 25 | 2041.77 | 26500 | 17.7 | 17.8026 | 0 | 0 | 250 | 25 | 730.96 | 4260 | 18.52 | 18.6347
0 | 0 | 245 | 24.5 | 845.62 | 295502.1 | 16.67 | 16.6556 | 0 | 0 | 245 | 24.5 | 2102.98 | 27294.4 | 18.28 | 18.3363 | 0 | 0 | 245.1 | 24.51 | 749.53 | 4368.3 | 19 | 19.1081
0 | 0 | 240 | 24 | 878.37 | 306946 | 17.3 | 17.3007 | 0 | 0 | 240 | 24 | 2171.95 | 28189.5 | 18.97 | 18.9377 | 0 | 0 | 240 | 24 | 788.2 | 4593.6 | 20.03 | 20.0939
0 | 0 | 235 | 23.5 | 907.5 | 317125.8 | 17.85 | 17.8744 | 0 | 0 | 235.1 | 23.51 | 2245.66 | 29146.2 | 19.61 | 19.5804 | 0 | 0 | 235.1 | 23.51 | 805.69 | 4695.6 | 20.48 | 20.5398
0 | 0 | 230 | 23 | 940.27 | 328578.4 | 18.52 | 18.5199 | 0 | 0 | 230 | 23 | 2318.35 | 30089.7 | 20.24 | 20.2142 | 0 | 0 | 230 | 23 | 824.87 | 4807.4 | 21.08 | 21.0288
0 | 0 | 225 | 22.5 | 977 | 341413.7 | 19.21 | 19.2433 | 0 | 0 | 225 | 22.5 | 2401.59 | 31170 | 21 | 20.94 | 0 | 0 | 225 | 22.5 | 857.01 | 4994.7 | 21.79 | 21.8481
0 | 0 | 220 | 22 | 1011.19 | 353359.9 | 19.92 | 19.9167 | 0 | 0 | 220.1 | 22.01 | 2481.16 | 32202.8 | 21.66 | 21.6338 | 0 | 0 | 220.1 | 22.01 | 879.09 | 5123.4 | 22.32 | 22.411
0 | 0 | 215.1 | 21.51 | 1045.77 | 365445.3 | 20.62 | 20.5978 | 0 | 0 | 215.1 | 21.51 | 2566.02 | 33304.2 | 22.39 | 22.3737 | 0 | 0 | 215.1 | 21.51 | 900.15 | 5246.1 | 22.97 | 22.9479
0 | 0 | 210 | 21 | 1081.97 | 378093.9 | 21.32 | 21.3108 | 0 | 0 | 210.1 | 21.01 | 2650.7 | 34403.2 | 23.12 | 23.112 | 0 | 0 | 210.1 | 21.01 | 923.5 | 5382.2 | 23.52 | 23.5432
0 | 0 | 205.1 | 20.51 | 1124.98 | 393125.8 | 22.16 | 22.158 | 0 | 0 | 205.1 | 20.51 | 2747.65 | 35661.5 | 23.95 | 23.9573 | 0 | 0 | 205.1 | 20.51 | 960.18 | 5596 | 24.36 | 24.4783
0 | 0 | 200.1 | 20.01 | 1166.3 | 407562.5 | 22.98 | 22.9718 | 0 | 0 | 200 | 20 | 2833.62 | 36777.3 | 24.62 | 24.7069 | 0 | 0 | 200.1 | 20.01 | 976.85 | 5693.1 | 24.84 | 24.9033
0 | 0 | 195 | 19.5 | 1202.94 | 420368.3 | 23.7 | 23.6935 | 0 | 0 | 195 | 19.5 | 2938.18 | 38134.4 | 25.59 | 25.6186 | 0 | 0 | 195.1 | 19.51 | 1035.82 | 6036.8 | 26.39 | 26.4066
0 | 0 | 190.1 | 19.01 | 1247.41 | 435909.5 | 24.56 | 24.5694 | 0 | 0 | 190.1 | 19.01 | 3039.07 | 39443.9 | 26.53 | 26.4983 | 0 | 0 | 190.1 | 19.01 | 1064.65 | 6204.8 | 27.07 | 27.1416
0 | 0 | 185.1 | 18.51 | 1294.22 | 452265.8 | 25.53 | 25.4914 | 0 | 0 | 185 | 18.5 | 3130.61 | 40631.9 | 27.24 | 27.2964 | 0 | 0 | 185.1 | 18.51 | 1100.51 | 6413.8 | 28.03 | 28.0558
0 | 0 | 180.1 | 18.01 | 1341.53 | 468796.8 | 26.46 | 26.4232 | 0 | 0 | 180.1 | 18.01 | 3243.25 | 42093.9 | 28.28 | 28.2786 | 0 | 0 | 180.1 | 18.01 | 1129.19 | 6580.9 | 28.79 | 28.7869
0 | 0 | 175.1 | 17.51 | 1398.11 | 488569 | 27.54 | 27.5376 | 0 | 0 | 175 | 17.5 | 3367.49 | 43706.4 | 29.38 | 29.3618 | 0 | 0 | 175 | 17.5 | 1180.72 | 6881.2 | 30.09 | 30.1006
0 | 0 | 170.1 | 17.01 | 1448.57 | 506204.9 | 28.52 | 28.5315 | 0 | 0 | 170.1 | 17.01 | 3481.02 | 45179.9 | 30.39 | 30.3517 | 0 | 0 | 170 | 17 | 1225.88 | 7144.4 | 31.23 | 31.2519
0 | 0 | 165.1 | 16.51 | 1505.46 | 526081.9 | 29.64 | 29.652 | 0 | 0 | 165.1 | 16.51 | 3610.49 | 46860.3 | 31.45 | 31.4806 | 0 | 0 | 165.1 | 16.51 | 1265.48 | 7375.3 | 32.12 | 32.2614
0 | 0 | 160.1 | 16.01 | 1557.3 | 544198.7 | 30.69 | 30.6731 | 0 | 0 | 160.1 | 16.01 | 3735.08 | 48477.3 | 32.42 | 32.5669 | 0 | 0 | 160 | 16 | 1290.3 | 7519.9 | 32.82 | 32.8942
0 | 0 | 155.1 | 15.51 | 1617.01 | 565064.5 | 31.84 | 31.8491 | 0 | 0 | 155.1 | 15.51 | 3874.53 | 50287.2 | 33.59 | 33.7828 | 0 | 0 | 155.1 | 15.51 | 1352.79 | 7884.1 | 34.44 | 34.4873
0 | 0 | 150 | 15 | 1680.16 | 587133.5 | 33.06 | 33.093 | 0 | 0 | 150 | 15 | 4011.27 | 52062 | 34.87 | 34.9751 | 0 | 0 | 150.1 | 15.01 | 1390.75 | 8105.3 | 35.54 | 35.455
0 | 0 | 145.1 | 14.51 | 1742.55 | 608935.8 | 34.3 | 34.3218 | 0 | 0 | 145 | 14.5 | 4164.81 | 54054.7 | 36.06 | 36.3138 | 0 | 0 | 145.1 | 14.51 | 1443.13 | 8410.6 | 36.71 | 36.7903
0 | 0 | 140 | 14 | 1809.47 | 632321 | 35.72 | 35.6399 | 0 | 0 | 140.1 | 14.01 | 4307.2 | 55902.8 | 37.34 | 37.5554 | 0 | 0 | 140.1 | 14.01 | 1497.74 | 8728.9 | 38.22 | 38.1825
0 | 0 | 135.1 | 13.51 | 1880.68 | 657203.3 | 37.07 | 37.0425 | 0 | 0 | 135.1 | 13.51 | 4472.71 | 58050.9 | 38.94 | 38.9985 | 0 | 0 | 135.1 | 13.51 | 1547.21 | 9017.2 | 39.42 | 39.4437
0 | 0 | 130.1 | 13.01 | 1954.86 | 683127 | 38.52 | 38.5036 | 0 | 0 | 130 | 13 | 4628.57 | 60073.8 | 40.23 | 40.3575 | 0 | 0 | 130.1 | 13.01 | 1604.01 | 9348.2 | 40.89 | 40.8917
0 | 0 | 125 | 12.5 | 2028.16 | 708741.4 | 39.96 | 39.9473 | 0 | 0 | 125.1 | 12.51 | 4808.12 | 62404.3 | 41.82 | 41.923 | 0 | 0 | 125.1 | 12.51 | 1667.96 | 9720.9 | 42.61 | 42.522
0 | 0 | 120.1 | 12.01 | 2103.13 | 734939.7 | 41.43 | 41.4239 | 0 | 0 | 120.1 | 12.01 | 4987.49 | 64732.2 | 43.33 | 43.487 | 0 | 0 | 120.1 | 12.01 | 1731.99 | 10094.1 | 44.16 | 44.1544
0 | 0 | 115.1 | 11.51 | 2187.41 | 764390.5 | 43.04 | 43.0839 | 0 | 0 | 115.1 | 11.51 | 5179.67 | 67226.5 | 45.08 | 45.1626 | 0 | 0 | 115 | 11.5 | 1796.56 | 10470.4 | 45.68 | 45.8005
0 | 0 | 110.1 | 11.01 | 2275.84 | 795293.3 | 44.77 | 44.8257 | 0 | 0 | 110.1 | 11.01 | 5361.71 | 69589.2 | 46.6 | 46.7499 | 0 | 0 | 110.1 | 11.01 | 1859.84 | 10839.2 | 47.37 | 47.4137
0 | 0 | 105.1 | 10.51 | 2362.57 | 825602.3 | 46.56 | 46.5339 | 0 | 0 | 105.1 | 10.51 | 5566.02 | 72240.9 | 48.28 | 48.5313 | 0 | 0 | 105.1 | 10.51 | 1925.95 | 11224.5 | 48.94 | 49.0991
0 | 0 | 100.1 | 10.01 | 2457.19 | 858664.6 | 48.38 | 48.3976 | 0 | 0 | 100 | 10 | 5767.54 | 74856.5 | 50.1 | 50.2884 | 0 | 0 | 100.1 | 10.01 | 1981.59 | 11548.7 | 50.36 | 50.5175
0 | 0 | 95.1 | 9.51 | 2556.97 | 893533.5 | 50.32 | 50.3629 | 0 | 0 | 95.1 | 9.51 | 5996.28 | 77825.2 | 52.22 | 52.2828 | 0 | 0 | 95.1 | 9.51 | 2062.1 | 12018 | 52.45 | 52.57
0 | 0 | 90.1 | 9.01 | 2653.69 | 927333.8 | 52.32 | 52.2679 | 0 | 0 | 90 | 9 | 6220.37 | 80733.7 | 54.24 | 54.2367 | 0 | 0 | 90.1 | 9.01 | 2142.19 | 12484.7 | 54.5 | 54.6118
0 | 0 | 85 | 8.5 | 2770.43 | 968129.1 | 54.6 | 54.5673 | 0 | 0 | 85.1 | 8.51 | 6469.11 | 83962.1 | 56.26 | 56.4055 | 0 | 0 | 85 | 8.5 | 2231.64 | 13006.1 | 57.01 | 56.8921
0 | 0 | 80 | 8 | 2882.02 | 1007124.3 | 56.77 | 56.7652 | 0 | 0 | 80 | 8 | 6710.1 | 87089.9 | 58.11 | 58.5068 | 0 | 0 | 80.1 | 8.01 | 2319.74 | 13519.5 | 58.97 | 59.1381
0 | 0 | 75 | 7.5 | 2999.47 | 1048165.9 | 59.18 | 59.0785 | 0 | 0 | 75 | 7.5 | 6975.03 | 90528.3 | 60.61 | 60.8167 | 0 | 0 | 75 | 7.5 | 2396.41 | 13966.3 | 61.17 | 61.0927
0 | 0 | 70 | 7 | 3121.72 | 1090885.1 | 61.48 | 61.4864 | 0 | 0 | 70 | 7 | 7251.27 | 94113.6 | 62.79 | 63.2253 | 0 | 0 | 70 | 7 | 2484.59 | 14480.2 | 63.37 | 63.3407
0 | 0 | 65 | 6.5 | 3247.79 | 1134940.1 | 63.93 | 63.9695 | 0 | 0 | 65 | 6.5 | 7531.6 | 97752 | 65.29 | 65.6696 | 0 | 0 | 65 | 6.5 | 2577.38 | 15021 | 65.68 | 65.7062
0 | 0 | 60 | 6 | 3376.25 | 1179831.3 | 66.49 | 66.4997 | 0 | 0 | 60 | 6 | 7843.2 | 101796.3 | 68.39 | 68.3865 | 0 | 0 | 60 | 6 | 2672.51 | 15575.5 | 67.78 | 68.1314
0 | 0 | 55 | 5.5 | 3524.92 | 1231784.6 | 69.47 | 69.428 | 0 | 0 | 55 | 5.5 | 8127.34 | 105484 | 70.47 | 70.864 | 0 | 0 | 55 | 5.5 | 2787.61 | 16246.3 | 70.85 | 71.0657
0 | 0 | 50 | 5 | 3667.95 | 1281768.2 | 72.26 | 72.2451 | 0 | 0 | 49.9 | 4.99 | 8460.37 | 109806.4 | 73.31 | 73.7677 | 0 | 0 | 49.9 | 4.99 | 2892.03 | 16854.8 | 73.35 | 73.7278
0 | 0 | 44.9 | 4.49 | 3827.2 | 1337415.1 | 75.31 | 75.3818 | 0 | 0 | 44.9 | 4.49 | 8789.6 | 114079.4 | 76.28 | 76.6384 | 0 | 0 | 45 | 4.5 | 3016.92 | 17582.7 | 77.02 | 76.9116
0 | 0 | 39.9 | 3.99 | 3989.64 | 1394179.8 | 78.58 | 78.5812 | 0 | 0 | 39.9 | 3.99 | 9141.47 | 118646.4 | 79.44 | 79.7064 | 0 | 0 | 39.9 | 3.99 | 3142.82 | 18316.4 | 79.94 | 80.1212
0 | 0 | 34.9 | 3.49 | 4162.14 | 1454463.3 | 82.01 | 81.9789 | 0 | 0 | 34.9 | 3.49 | 9524.35 | 123615.8 | 83.26 | 83.0448 | 0 | 0 | 34.9 | 3.49 | 3258.98 | 18993.4 | 83.24 | 83.0826
0 | 0 | 29.9 | 2.99 | 4334.49 | 1514688 | 85.36 | 85.3735 | 0 | 0 | 29.9 | 2.99 | 9897.09 | 128453.5 | 86.13 | 86.2948 | 0 | 0 | 29.9 | 2.99 | 3385.07 | 19728.3 | 86.31 | 86.297
0 | 0 | 24.9 | 2.49 | 4513.08 | 1577098.8 | 88.93 | 88.8911 | 0 | 0 | 25.1 | 2.51 | 10301.59 | 133703.5 | 90.06 | 89.8217 | 0 | 0 | 25 | 2.5 | 3517.49 | 20500 | 89.56 | 89.6729
0 | 0 | 20.1 | 2.01 | 4697.01 | 1641371.9 | 92.47 | 92.5138 | 0 | 0 | 24.1 | 2.41 | 10358.98 | 134448.3 | 90.51 | 90.3221 | 0 | 0 | 24.1 | 2.41 | 3545.17 | 20661.3 | 89.76 | 90.3785
0 | 0 | 19 | 1.9 | 4737.67 | 1655582.1 | 93.39 | 93.3147 | 0 | 0 | 23.1 | 2.31 | 10465.61 | 135832.2 | 91.26 | 91.2518 | 0 | 0 | 23.1 | 2.31 | 3567.34 | 20790.5 | 90.85 | 90.9437
0 | 0 | 18.1 | 1.81 | 4781.98 | 1671065.1 | 94.21 | 94.1874 | 0 | 0 | 22.1 | 2.21 | 10522.85 | 136575.2 | 91.43 | 91.7509 | 0 | 0 | 22 | 2.2 | 3605.9 | 21015.3 | 91.81 | 91.9267
0 | 0 | 17.1 | 1.71 | 4809.76 | 1680771 | 94.79 | 94.7346 | 0 | 0 | 21 | 2.1 | 10613.55 | 137752.3 | 92.5 | 92.5418 | 0 | 0 | 21 | 2.1 | 3625.3 | 21128.3 | 92.4 | 92.4213
0 | 0 | 16 | 1.6 | 4864.04 | 1699742.5 | 95.76 | 95.8037 | 0 | 0 | 20.1 | 2.01 | 10697.67 | 138844.1 | 93.19 | 93.2752 | 0 | 0 | 20 | 2 | 3661.28 | 21338 | 93.47 | 93.3386
0 | 0 | 15 | 1.5 | 4897.24 | 1711342.6 | 96.32 | 96.4576 | 0 | 0 | 19 | 1.9 | 10789.65 | 140037.9 | 94.01 | 94.0772 | 0 | 0 | 19 | 1.9 | 3696.32 | 21542.2 | 93.71 | 94.2319
0 | 0 | 14.1 | 1.41 | 4930.09 | 1722822.5 | 97.18 | 97.1046 | 0 | 0 | 18 | 1.8 | 10875.78 | 141155.8 | 95.12 | 94.8282 | 0 | 0 | 18 | 1.8 | 3727.72 | 21725.2 | 95.08 | 95.0324
0 | 0 | 13 | 1.3 | 4961.71 | 1733872.8 | 97.8 | 97.7274 | 0 | 0 | 17 | 1.7 | 10955.38 | 142189 | 95.19 | 95.5223 | 0 | 0 | 17 | 1.7 | 3766.16 | 21949.3 | 95.76 | 96.0123
0 | 0 | 12.1 | 1.21 | 4990.73 | 1744013.3 | 98.15 | 98.299 | 0 | 0 | 16.1 | 1.61 | 11047.24 | 143381.2 | 96.55 | 96.3232 | 0 | 0 | 16.1 | 1.61 | 3775.15 | 22001.6 | 95.87 | 96.2415
0 | 0 | 11.1 | 1.11 | 5037.59 | 1760389.1 | 99.19 | 99.222 | 0 | 0 | 15 | 1.5 | 11116.86 | 144284.9 | 96.83 | 96.9302 | 0 | 0 | 15 | 1.5 | 3801.24 | 22153.7 | 96.18 | 96.9066
0 | 0 | 10.1 | 1.01 | 5047.11 | 1763714.1 | 99.33 | 99.4095 | 0 | 0 | 14 | 1.4 | 11185.79 | 145179.4 | 97.4 | 97.5312 | 0 | 0 | 14 | 1.4 | 3839.33 | 22375.7 | 97.37 | 97.8777
0 | 0 | 9.1 | 0.91 | 5073.04 | 1772777.6 | 100 | 99.9202 | 0 | 0 | 13.1 | 1.31 | 11286.23 | 146483.1 | 98.51 | 98.407 | 0 | 0 | 13 | 1.3 | 3862.17 | 22508.8 | 98.16 | 98.4599
0 | 0 | 8 | 0.8 | 5077.09 | 1774190.6 | 99.91 | 100 | 0 | 0 | 12 | 1.2 | 11346.92 | 147270.8 | 98.64 | 98.9362 | 0 | 0 | 12 | 1.2 | 3895.21 | 22701.4 | 99.07 | 99.3022
0 | 0 | 7 | 0.7 | 5064.11 | 1769655 | 99.79 | 99.7443 | 0 | 0 | 11 | 1.1 | 11404.63 | 148019.8 | 99.37 | 99.4394 | 0 | 0 | 11 | 1.1 | 3894.48 | 22697.1 | 99.17 | 99.2836
0 | 0 | 6.1 | 0.61 | 5012.86 | 1751744.4 | 98.86 | 98.7349 | 0 | 0 | 10 | 1 | 11456.19 | 148688.9 | 100 | 99.8889 | 0 | 0 | 10 | 1 | 3922.58 | 22860.8 | 100 | 100
0 | 0 | 5 | 0.5 | 4938.8 | 1725866.6 | 97.3 | 97.2762 | 0 | 0 | 9 | 0.9 | 11468.93 | 148854.3 | 99.37 | 100 | 0 | 0 | 9 | 0.9 | 3921.19 | 22852.8 | 99.65 | 99.9646
0 | 0 | 4 | 0.4 | 4777.05 | 1669341 | 94.12 | 94.0903 | 0 | 0 | 8 | 0.8 | 11447.09 | 148570.9 | 99.73 | 99.8096 | 0 | 0 | 8 | 0.8 | 3919.32 | 22841.9 | 99.75 | 99.9169
0 | 0 | 3 | 0.3 | 4507.73 | 1575228.3 | 88.84 | 88.7857 | 0 | 0 | 7 | 0.7 | 11399.99 | 147959.5 | 98.9 | 99.3989 | 0 | 0 | 7 | 0.7 | 3909.97 | 22787.4 | 99.49 | 99.6785
0 | 0 | 2 | 0.2 | 4040.66 | 1412009 | 79.62 | 79.5861 | 0 | 0 | 6 | 0.6 | 11288.43 | 146511.7 | 98.05 | 98.4262 | 0 | 0 | 6 | 0.6 | 3853 | 22455.3 | 97.89 | 98.2262
0 | 0 | 1 | 0.1 | 3147.73 | 1099976.8 | 62.09 | 61.9987 | 0 | 0 | 5 | 0.5 | 11065.76 | 143621.5 | 96.26 | 96.4847 | 0 | 0 | 5 | 0.5 | 3775.56 | 22004 | 96.19 | 96.252
0 | 0 | 0 | 0 | 2775.49 | 969896.4 | 54.67 | 54.6669 | 0 | 0 | 4 | 0.4 | 10651.02 | 138238.7 | 92.54 | 92.8685 | 0 | 0 | 4 | 0.4 | 3603.5 | 21001.3 | 92.08 | 91.8656
0 | 0 | -0.5 | -0.05 | 2773.72 | 969277.8 | 54.57 | 54.6321 | 0 | 0 | 3 | 0.3 | 9971.7 | 129421.9 | 86.86 | 86.9453 | 0 | 0 | 3 | 0.3 | 3339.42 | 19462.2 | 84.85 | 85.1333
[8 empty] | 0 | 0 | 2 | 0.2 | 8740.51 | 113442.3 | 76.18 | 76.2103 | 0 | 0 | 2 | 0.2 | 2867.21 | 16710.2 | 72.94 | 73.095
[8 empty] | 0 | 0 | 1 | 0.1 | 7309.28 | 94866.6 | 63.39 | 63.7311 | 0 | 0 | 1 | 0.1 | 2443.37 | 14240 | 61.96 | 62.2899
[8 empty] | 0 | 0 | 0 | 0 | 7280.43 | 94492.1 | 63.33 | 63.4796 | 0 | 0 | 0 | 0 | 2379.89 | 13870.1 | 60.28 | 60.6715
[8 empty] | 0 | 0 | -0.5 | -0.05 | 7277.92 | 94459.5 | 63.34 | 63.4577 | 0 | 0 | -0.5 | -0.05 | 2373.54 | 13833 | 60.43 | 60.5097

## 1x1
Measurement time: | 12/04/2023 18:31:48 | [6 empty] | Measurement time: | 12/04/2023 18:43:02 | [6 empty] | Measurement time: | 12/04/2023 20:01:27 | [6 empty] | Measurement time: | 27/05/2023 10:40:15 | [6 empty] | Measurement time: | 27/05/2023 12:48:22
Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV | [6 empty] | Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF)
Controller: | CCU | [6 empty] | Controller: | CCU | [6 empty] | Controller: | CCU | [6 empty] | Controller: | CCU | [6 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele | [6 empty] | Field detector: | SFD parallele | [6 empty] | Field detector: | SFD parallele | [6 empty] | Field detector: | PTW60019 Diamant | [6 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | <none> | [6 empty] | Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm
SSD: | 900 mm | [6 empty] | SSD: | 900 mm | [6 empty] | SSD: | 900 mm | [6 empty] | SSD: | 900 mm | [6 empty] | SSD: | 900 mm
Field size: | 10 x 10 mm | [6 empty] | Field size: | 10 x 10 mm | [6 empty] | Field size: | 10 x 10 mm | Cône | [5 empty] | Field size: | 10 x 10 mm | [6 empty] | Field size: | 10 x 10 mm
Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0
Measurement medium: | Water | [6 empty] | Measurement medium: | Water | [6 empty] | Measurement medium: | Water | [6 empty] | Measurement medium: | Water | [6 empty] | Measurement medium: | Water
Scan type: | Beam | [6 empty] | Scan type: | Beam | [6 empty] | Scan type: | Beam | [6 empty] | Scan type: | Beam | [6 empty] | Scan type: | Beam
Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep
Wedge: | none | [6 empty] | Wedge: | none | [6 empty] | Wedge: | none | [6 empty] | Wedge: | none | [6 empty] | Wedge: | none
Points [mm]: | [7 empty] | Points [mm]: | [7 empty] | Points [mm]: | [7 empty] | Points [mm]: | [7 empty] | Points [mm]:
Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 250 | 25 | 34.3 | 355518.6 | 34.37 | 17.6024 | 0 | 0 | 250 | 25 | 10.11 | 104818.5 | 40.84 | 20.0317 | 0 | 0 | 250 | 25 | 17.64 | 351239.6 | 0 | 17.64 | 0 | 0 | 250 | 25 | 3940.16 | 33282.9 | 18.84 | 18.8511 | 0 | 0 | 250 | 25 | 925.63 | 5394.6 | 19.51 | 19.4967
0 | 0 | 245 | 24.5 | 35.42 | 367060.7 | 35.52 | 18.1772 | 0 | 0 | 245.1 | 24.51 | 10.53 | 109140.1 | 42.53 | 20.8639 | 0 | 0 | 245.1 | 24.51 | 18.15 | 361275.4 | 0 | 18.15 | 0 | 0 | 245 | 24.5 | 4058.79 | 34285 | 19.4 | 19.4186 | 0 | 0 | 245 | 24.5 | 951.66 | 5546.3 | 20.06 | 20.0449
0 | 0 | 240 | 24 | 36.74 | 380763.2 | 36.85 | 18.8546 | 0 | 0 | 240 | 24 | 10.8 | 111906.1 | 43.57 | 21.3989 | 0 | 0 | 240 | 24 | 18.84 | 375142.6 | 0 | 18.84 | 0 | 0 | 240 | 24 | 4186.39 | 35362.8 | 20.01 | 20.0291 | 0 | 0 | 240.1 | 24.01 | 989.23 | 5765.2 | 20.82 | 20.8363
0 | 0 | 235 | 23.5 | 37.9 | 392811.9 | 38 | 19.4499 | 0 | 0 | 235 | 23.5 | 11.19 | 115931.5 | 45.15 | 22.1716 | 0 | 0 | 235 | 23.5 | 19.47 | 387650.3 | 0 | 19.47 | 0 | 0 | 235 | 23.5 | 4334.03 | 36610 | 20.72 | 20.7355 | 0 | 0 | 235.1 | 23.51 | 1009.86 | 5885.5 | 21.3 | 21.2708
0 | 0 | 230 | 23 | 39.31 | 407400.6 | 39.41 | 20.1735 | 0 | 0 | 230.1 | 23.01 | 11.58 | 119975.7 | 46.74 | 22.9443 | 0 | 0 | 230 | 23 | 20.16 | 401417.3 | 0 | 20.16 | 0 | 0 | 230 | 23 | 4456.26 | 37642.4 | 21.33 | 21.3203 | 0 | 0 | 230 | 23 | 1052.58 | 6134.4 | 22.19 | 22.1706
0 | 0 | 225 | 22.5 | 40.66 | 421446.3 | 40.77 | 20.8663 | 0 | 0 | 225 | 22.5 | 11.93 | 123600.8 | 48.11 | 23.6378 | 0 | 0 | 225 | 22.5 | 20.86 | 415292.4 | 0 | 20.86 | 0 | 0 | 225 | 22.5 | 4627.96 | 39092.8 | 22.13 | 22.1417 | 0 | 0 | 225.1 | 22.51 | 1091.46 | 6361 | 22.99 | 22.9896
0 | 0 | 220 | 22 | 42.13 | 436631.7 | 42.24 | 21.6207 | 0 | 0 | 220.1 | 22.01 | 12.33 | 127788.4 | 49.74 | 24.4304 | 0 | 0 | 220.1 | 22.01 | 21.59 | 429884.7 | 0 | 21.59 | 0 | 0 | 220 | 22 | 4759.24 | 40201.8 | 22.77 | 22.7698 | 0 | 0 | 220.1 | 22.01 | 1125.15 | 6557.4 | 23.69 | 23.6992
0 | 0 | 215.1 | 21.51 | 43.64 | 452288 | 43.77 | 22.3956 | 0 | 0 | 215.1 | 21.51 | 12.72 | 131808.3 | 51.35 | 25.2031 | 0 | 0 | 215.1 | 21.51 | 22.32 | 444419.9 | 0 | 22.32 | 0 | 0 | 215 | 21.5 | 4928.91 | 41635 | 23.56 | 23.5816 | 0 | 0 | 215.1 | 21.51 | 1149.85 | 6701.3 | 24.19 | 24.2195
0 | 0 | 210 | 21 | 44.95 | 465895.6 | 45.08 | 23.0678 | 0 | 0 | 210.1 | 21.01 | 13.22 | 137016.7 | 53.46 | 26.1938 | 0 | 0 | 210.1 | 21.01 | 23.07 | 459230 | 0 | 23.07 | 0 | 0 | 210.1 | 21.01 | 5104.65 | 43119.4 | 24.4 | 24.4224 | 0 | 0 | 210.1 | 21.01 | 1175.91 | 6853.3 | 24.78 | 24.7684
0 | 0 | 205.1 | 20.51 | 46.63 | 483313.8 | 46.75 | 23.93 | 0 | 0 | 205.1 | 20.51 | 13.6 | 140986 | 54.97 | 26.9467 | 0 | 0 | 205.1 | 20.51 | 23.82 | 474244.8 | 0 | 23.82 | 0 | 0 | 205 | 20.5 | 5272.45 | 44536.8 | 25.22 | 25.2252 | 0 | 0 | 205 | 20.5 | 1222.82 | 7126.6 | 25.73 | 25.7564
0 | 0 | 200 | 20 | 48.38 | 501413 | 48.49 | 24.8281 | 0 | 0 | 200.1 | 20.01 | 14.09 | 146079.6 | 56.92 | 27.9176 | 0 | 0 | 200.1 | 20.01 | 24.68 | 491378.6 | 0 | 24.68 | 0 | 0 | 200 | 20 | 5452.95 | 46061.6 | 26.04 | 26.0888 | 0 | 0 | 200.1 | 20.01 | 1267.49 | 7386.9 | 26.7 | 26.6973
0 | 0 | 195 | 19.5 | 49.92 | 517418.5 | 50.06 | 25.6184 | 0 | 0 | 195 | 19.5 | 14.54 | 150721.1 | 58.7 | 28.8092 | 0 | 0 | 195.1 | 19.51 | 25.56 | 508790.5 | 0 | 25.56 | 0 | 0 | 195 | 19.5 | 5635.35 | 47602.3 | 26.94 | 26.9614 | 0 | 0 | 195.1 | 19.51 | 1318.03 | 7681.5 | 27.74 | 27.7619
0 | 0 | 190 | 19 | 51.87 | 537600.6 | 52.04 | 26.6191 | 0 | 0 | 190.1 | 19.01 | 14.99 | 155395.1 | 60.51 | 29.7008 | 0 | 0 | 190.1 | 19.01 | 26.44 | 526435.7 | 0 | 26.44 | 0 | 0 | 190.1 | 19.01 | 5821.42 | 49174.1 | 27.84 | 27.8516 | 0 | 0 | 190.1 | 19.01 | 1351.29 | 7875.4 | 28.43 | 28.4624
0 | 0 | 185.1 | 18.51 | 53.43 | 553819.9 | 53.59 | 27.4197 | 0 | 0 | 185.1 | 18.51 | 15.59 | 161604.4 | 62.84 | 30.8896 | 0 | 0 | 185.1 | 18.51 | 27.42 | 545894 | 0 | 27.42 | 0 | 0 | 185 | 18.5 | 6018.02 | 50834.8 | 28.82 | 28.7922 | 0 | 0 | 185.1 | 18.51 | 1399.38 | 8155.6 | 29.45 | 29.4753
0 | 0 | 180 | 18 | 55.51 | 575366.4 | 55.68 | 28.4871 | 0 | 0 | 180 | 18 | 16.08 | 166692.6 | 65.04 | 31.8605 | 0 | 0 | 180 | 18 | 28.36 | 564511.4 | 0 | 28.36 | 0 | 0 | 180 | 18 | 6231.62 | 52639 | 29.84 | 29.8142 | 0 | 0 | 180.1 | 18.01 | 1440.77 | 8396.8 | 30.34 | 30.3471
0 | 0 | 175.1 | 17.51 | 57.63 | 597271.8 | 57.79 | 29.5751 | 0 | 0 | 175.1 | 17.51 | 16.66 | 172664.2 | 67.22 | 33.0097 | 0 | 0 | 175.1 | 17.51 | 29.4 | 585358.9 | 0 | 29.4 | 0 | 0 | 175.1 | 17.51 | 6449.2 | 54476.9 | 30.86 | 30.8552 | 0 | 0 | 175.1 | 17.51 | 1501.09 | 8748.4 | 31.6 | 31.6177
0 | 0 | 170.1 | 17.01 | 59.62 | 617937.3 | 59.77 | 30.5963 | 0 | 0 | 170 | 17 | 17.19 | 178174 | 69.43 | 34.0598 | 0 | 0 | 170.1 | 17.01 | 30.41 | 605441.2 | 0 | 30.41 | 0 | 0 | 170.1 | 17.01 | 6680.98 | 56434.8 | 31.95 | 31.9641 | 0 | 0 | 170.1 | 17.01 | 1536.75 | 8956.2 | 32.39 | 32.3688
0 | 0 | 165 | 16.5 | 61.72 | 639679.3 | 61.88 | 31.674 | 0 | 0 | 165.1 | 16.51 | 17.81 | 184591.3 | 71.93 | 35.2883 | 0 | 0 | 165.1 | 16.51 | 31.65 | 629998.4 | 0 | 31.65 | 0 | 0 | 165.1 | 16.51 | 6899.43 | 58280.1 | 33.01 | 33.0092 | 0 | 0 | 165 | 16.5 | 1612.24 | 9396.2 | 33.95 | 33.9588
0 | 0 | 160.1 | 16.01 | 64.14 | 664760.7 | 64.29 | 32.9159 | 0 | 0 | 160.1 | 16.01 | 18.37 | 190369.5 | 74.13 | 36.3979 | 0 | 0 | 160.1 | 16.01 | 32.67 | 650324.5 | 0 | 32.67 | 0 | 0 | 160.1 | 16.01 | 7133.02 | 60253.3 | 34.05 | 34.1268 | 0 | 0 | 160.1 | 16.01 | 1665.87 | 9708.7 | 35.11 | 35.0885
0 | 0 | 155.1 | 15.51 | 66.28 | 686954.3 | 66.45 | 34.0142 | 0 | 0 | 155.1 | 15.51 | 19 | 196917 | 76.68 | 37.6461 | 0 | 0 | 155.1 | 15.51 | 33.86 | 674099.8 | 0 | 33.86 | 0 | 0 | 155.1 | 15.51 | 7392.67 | 62446.5 | 35.37 | 35.369 | 0 | 0 | 155.1 | 15.51 | 1701.37 | 9915.6 | 35.81 | 35.8362
0 | 0 | 150 | 15 | 68.63 | 711291.2 | 68.85 | 35.2202 | 0 | 0 | 150.1 | 15.01 | 19.61 | 203299.7 | 79.21 | 38.8548 | 0 | 0 | 150.1 | 15.01 | 35.09 | 698567 | 0 | 35.09 | 0 | 0 | 150.1 | 15.01 | 7652.34 | 64640 | 36.62 | 36.6114 | 0 | 0 | 150.1 | 15.01 | 1763.12 | 10275.5 | 37.16 | 37.1368
0 | 0 | 145.1 | 14.51 | 71.31 | 739144.9 | 71.53 | 36.5955 | 0 | 0 | 145.1 | 14.51 | 20.29 | 210264.4 | 81.88 | 40.2021 | 0 | 0 | 145.1 | 14.51 | 36.34 | 723535.6 | 0 | 36.34 | 0 | 0 | 145.1 | 14.51 | 7917.75 | 66881.9 | 37.86 | 37.8812 | 0 | 0 | 145.1 | 14.51 | 1817.66 | 10593.4 | 38.25 | 38.2856
0 | 0 | 140 | 14 | 73.9 | 765970.9 | 74.12 | 37.9247 | 0 | 0 | 140 | 14 | 20.96 | 217194.1 | 84.7 | 41.5296 | 0 | 0 | 140 | 14 | 37.85 | 753512.5 | 0 | 37.85 | 0 | 0 | 140 | 14 | 8236.07 | 69570.8 | 39.35 | 39.4041 | 0 | 0 | 140.1 | 14.01 | 1904.79 | 11101.2 | 40.09 | 40.1209
0 | 0 | 135 | 13.5 | 76.67 | 794652.7 | 76.89 | 39.3462 | 0 | 0 | 135.1 | 13.51 | 21.8 | 225917.7 | 87.87 | 43.194 | 0 | 0 | 135.1 | 13.51 | 39.12 | 778777.1 | 0 | 39.12 | 0 | 0 | 135.1 | 13.51 | 8500.29 | 71802.7 | 40.62 | 40.6683 | 0 | 0 | 135 | 13.5 | 1955.74 | 11398.1 | 41.2 | 41.194
0 | 0 | 130.1 | 13.01 | 79.49 | 823867 | 79.76 | 40.7934 | 0 | 0 | 130 | 13 | 22.44 | 232618.6 | 90.75 | 44.4621 | 0 | 0 | 130.1 | 13.01 | 40.48 | 805778.5 | 0 | 40.48 | 0 | 0 | 130 | 13 | 8817.3 | 74480.5 | 42.19 | 42.185 | 0 | 0 | 130 | 13 | 2017.38 | 11757.3 | 42.53 | 42.4924
0 | 0 | 125.1 | 12.51 | 82.55 | 855544.6 | 82.8 | 42.3637 | 0 | 0 | 125.1 | 12.51 | 23.19 | 240374.4 | 93.63 | 45.9481 | 0 | 0 | 125 | 12.5 | 42.13 | 838817.2 | 0 | 42.13 | 0 | 0 | 125.1 | 12.51 | 9156.32 | 77344.2 | 43.76 | 43.8069 | 0 | 0 | 125.1 | 12.51 | 2101.15 | 12245.5 | 44.29 | 44.2568
0 | 0 | 120.1 | 12.01 | 85.7 | 888188.7 | 85.95 | 43.9803 | 0 | 0 | 120.1 | 12.01 | 24.16 | 250385.2 | 97.61 | 47.87 | 0 | 0 | 120.1 | 12.01 | 43.74 | 870753.8 | 0 | 43.74 | 0 | 0 | 120.1 | 12.01 | 9461.91 | 79925.6 | 45.22 | 45.269 | 0 | 0 | 120.1 | 12.01 | 2171.34 | 12654.6 | 45.69 | 45.7352
0 | 0 | 115.1 | 11.51 | 89.01 | 922513.7 | 89.3 | 45.6789 | 0 | 0 | 115.1 | 11.51 | 24.88 | 257856.6 | 100.42 | 49.2966 | 0 | 0 | 115.1 | 11.51 | 45.35 | 902822.2 | 0 | 45.35 | 0 | 0 | 115.1 | 11.51 | 9822.65 | 82972.8 | 46.92 | 46.9949 | 0 | 0 | 115 | 11.5 | 2254.11 | 13137 | 47.51 | 47.4786
0 | 0 | 110.1 | 11.01 | 92.48 | 958543.4 | 92.8 | 47.4597 | 0 | 0 | 110.1 | 11.01 | 25.81 | 267513.7 | 104.27 | 51.1393 | 0 | 0 | 110.1 | 11.01 | 47.01 | 935793.5 | 0 | 47.01 | 0 | 0 | 110 | 11 | 10197.76 | 86141.4 | 48.73 | 48.7895 | 0 | 0 | 110.1 | 11.01 | 2322.28 | 13534.3 | 48.97 | 48.9145
0 | 0 | 105 | 10.5 | 95.92 | 994135.3 | 96.24 | 49.2251 | 0 | 0 | 105.1 | 10.51 | 26.73 | 277058.7 | 107.88 | 52.9622 | 0 | 0 | 105.1 | 10.51 | 48.84 | 972280.3 | 0 | 48.84 | 0 | 0 | 105.1 | 10.51 | 10567.56 | 89265.1 | 50.47 | 50.5588 | 0 | 0 | 105.1 | 10.51 | 2409.98 | 14045.4 | 50.81 | 50.7617
0 | 0 | 100 | 10 | 99.51 | 1031374.2 | 99.84 | 51.0674 | 0 | 0 | 100.1 | 10.01 | 27.72 | 287253.8 | 111.98 | 54.9237 | 0 | 0 | 100.1 | 10.01 | 50.72 | 1009654.3 | 0 | 50.72 | 0 | 0 | 100.1 | 10.01 | 10936.53 | 92381.8 | 52.26 | 52.3241 | 0 | 0 | 100.1 | 10.01 | 2502.42 | 14584.1 | 52.72 | 52.7088
0 | 0 | 95 | 9.5 | 103.36 | 1071300.1 | 103.71 | 53.0432 | 0 | 0 | 95.1 | 9.51 | 28.58 | 296213.5 | 115.35 | 56.6277 | 0 | 0 | 95.1 | 9.51 | 52.64 | 1047872 | 0 | 52.64 | 0 | 0 | 95.1 | 9.51 | 11337.41 | 95768.1 | 54.2 | 54.242 | 0 | 0 | 95.1 | 9.51 | 2601.17 | 15159.7 | 54.84 | 54.7888
0 | 0 | 90.1 | 9.01 | 107.29 | 1111985.8 | 107.62 | 55.06 | 0 | 0 | 90.1 | 9.01 | 29.54 | 306211.2 | 119.08 | 58.5298 | 0 | 0 | 90.1 | 9.01 | 54.68 | 1088499.8 | 0 | 54.68 | 0 | 0 | 90 | 9 | 11776.11 | 99473.9 | 56.31 | 56.3409 | 0 | 0 | 90.1 | 9.01 | 2694.03 | 15700.8 | 56.76 | 56.7447
0 | 0 | 85 | 8.5 | 111.47 | 1155313.9 | 111.86 | 57.2052 | 0 | 0 | 85 | 8.5 | 30.75 | 318702.1 | 124.05 | 60.9273 | 0 | 0 | 85 | 8.5 | 56.83 | 1131429.6 | 0 | 56.83 | 0 | 0 | 85 | 8.5 | 12221.93 | 103239.7 | 58.38 | 58.4739 | 0 | 0 | 85.1 | 8.51 | 2789.24 | 16255.8 | 58.77 | 58.7502
0 | 0 | 80 | 8 | 115.61 | 1198238.5 | 115.98 | 59.3298 | 0 | 0 | 80 | 8 | 31.77 | 329310.1 | 128.2 | 62.9483 | 0 | 0 | 80 | 8 | 59.08 | 1176089.6 | 0 | 59.08 | 0 | 0 | 80 | 8 | 12678.03 | 107092.4 | 60.59 | 60.656 | 0 | 0 | 80 | 8 | 2895.52 | 16875.2 | 61.02 | 60.9887
0 | 0 | 75 | 7.5 | 120.6 | 1249946.2 | 121 | 61.8906 | 0 | 0 | 75 | 7.5 | 33.02 | 342277.1 | 133.23 | 65.425 | 0 | 0 | 75 | 7.5 | 61.32 | 1220818 | 0 | 61.32 | 0 | 0 | 75 | 7.5 | 13163.49 | 111193.1 | 62.84 | 62.9786 | 0 | 0 | 75 | 7.5 | 3005.46 | 17515.9 | 63.24 | 63.3044
0 | 0 | 70 | 7 | 125.17 | 1297294.8 | 125.54 | 64.2359 | 0 | 0 | 70 | 7 | 34.18 | 354295.5 | 137.97 | 67.7234 | 0 | 0 | 70 | 7 | 63.76 | 1269418 | 0 | 63.76 | 0 | 0 | 70 | 7 | 13671.4 | 115483.6 | 65.26 | 65.4086 | 0 | 0 | 70 | 7 | 3115.34 | 18156.3 | 65.66 | 65.6188
0 | 0 | 65 | 6.5 | 130.29 | 1350440.8 | 130.73 | 66.8634 | 0 | 0 | 65 | 6.5 | 35.3 | 365862.5 | 142.4 | 69.9425 | 0 | 0 | 65 | 6.5 | 66.34 | 1320731.1 | 0 | 66.34 | 0 | 0 | 65 | 6.5 | 14160.25 | 119612.9 | 67.63 | 67.7474 | 0 | 0 | 65 | 6.5 | 3236.51 | 18862.4 | 68.11 | 68.1711
0 | 0 | 60 | 6 | 135.18 | 1401101.9 | 135.62 | 69.3729 | 0 | 0 | 60 | 6 | 36.54 | 378756.3 | 147.35 | 72.3994 | 0 | 0 | 60 | 6 | 69.1 | 1375596.9 | 0 | 69.1 | 0 | 0 | 60 | 6 | 14736.89 | 124483.8 | 70.41 | 70.5063 | 0 | 0 | 60 | 6 | 3353.87 | 19546.4 | 70.59 | 70.643
0 | 0 | 55 | 5.5 | 140.71 | 1458385.5 | 141.15 | 72.2108 | 0 | 0 | 54.9 | 5.49 | 37.99 | 393733.2 | 153.48 | 75.2724 | 0 | 0 | 54.9 | 5.49 | 71.66 | 1426562.8 | 0 | 71.66 | 0 | 0 | 55 | 5.5 | 15297.12 | 129216.1 | 73.03 | 73.1866 | 0 | 0 | 55 | 5.5 | 3481.52 | 20290.4 | 73.37 | 73.3317
0 | 0 | 49.9 | 4.99 | 146.68 | 1520316 | 147.15 | 75.2746 | 0 | 0 | 50 | 5 | 39.45 | 408902.4 | 159.19 | 78.1652 | 0 | 0 | 50 | 5 | 74.69 | 1486986.2 | 0 | 74.69 | 0 | 0 | 49.9 | 4.99 | 15881.24 | 134150.2 | 75.88 | 75.9812 | 0 | 0 | 49.9 | 4.99 | 3608.72 | 21031.7 | 76.01 | 76.011
0 | 0 | 44.9 | 4.49 | 152.47 | 1580291 | 152.96 | 78.2459 | 0 | 0 | 44.9 | 4.49 | 40.79 | 422743.1 | 164.64 | 80.8203 | 0 | 0 | 44.9 | 4.49 | 77.86 | 1550065.9 | 0 | 77.86 | 0 | 0 | 44.9 | 4.49 | 16504.11 | 139411.7 | 78.88 | 78.9613 | 0 | 0 | 44.9 | 4.49 | 3750 | 21855.1 | 79.09 | 78.9868
0 | 0 | 39.9 | 3.99 | 158.38 | 1641555.2 | 158.95 | 81.2789 | 0 | 0 | 39.9 | 3.99 | 42.26 | 438003.8 | 170.49 | 83.7329 | 0 | 0 | 39.9 | 3.99 | 80.88 | 1610136.8 | 0 | 80.88 | 0 | 0 | 39.9 | 3.99 | 17162.27 | 144971.3 | 82.05 | 82.1101 | 0 | 0 | 39.9 | 3.99 | 3909.63 | 22785.4 | 82.36 | 82.3491
0 | 0 | 34.9 | 3.49 | 165 | 1710095.6 | 165.53 | 84.6762 | 0 | 0 | 34.9 | 3.49 | 43.91 | 455147.9 | 177.37 | 87.0022 | 0 | 0 | 34.9 | 3.49 | 84.08 | 1673811.3 | 0 | 84.08 | 0 | 0 | 34.9 | 3.49 | 17831.74 | 150626.3 | 85.25 | 85.3131 | 0 | 0 | 34.9 | 3.49 | 4043.07 | 23563.1 | 85.15 | 85.1598
0 | 0 | 29.9 | 2.99 | 171.57 | 1778208.3 | 172.19 | 88.0478 | 0 | 0 | 29.9 | 2.99 | 45.38 | 470325.1 | 183.17 | 89.9148 | 0 | 0 | 29.9 | 2.99 | 87.63 | 1744620.9 | 0 | 87.63 | 0 | 0 | 29.8 | 2.98 | 18510.95 | 156363.6 | 88.53 | 88.5627 | 0 | 0 | 29.9 | 2.99 | 4199.47 | 24474.6 | 88.58 | 88.454
0 | 0 | 24.9 | 2.49 | 178.68 | 1851947 | 179.25 | 91.6966 | 0 | 0 | 24.9 | 2.49 | 47.02 | 487355.1 | 189.88 | 93.1643 | 0 | 0 | 24.9 | 2.49 | 91.27 | 1817027.2 | 0 | 91.27 | 0 | 0 | 25.1 | 2.51 | 19250.56 | 162611.2 | 92.09 | 92.1012 | 0 | 0 | 25.1 | 2.51 | 4372.27 | 25481.7 | 92.14 | 92.0937
0 | 0 | 20.1 | 2.01 | 185.35 | 1921047.4 | 185.97 | 95.1196 | 0 | 0 | 20 | 2 | 48.58 | 503460.1 | 196.12 | 96.2552 | 0 | 0 | 20.1 | 2.01 | 94.74 | 1885999.7 | 0 | 94.74 | 0 | 0 | 24.1 | 2.41 | 19380.25 | 163706.7 | 92.69 | 92.7217 | 0 | 0 | 24.1 | 2.41 | 4410.32 | 25703.4 | 92.99 | 92.8952
0 | 0 | 19.1 | 1.91 | 186.57 | 1933744.1 | 187.25 | 95.7457 | 0 | 0 | 19.1 | 1.91 | 48.82 | 505984.3 | 196.86 | 96.7307 | 0 | 0 | 19.1 | 1.91 | 95.31 | 1897468.8 | 0 | 95.31 | 0 | 0 | 23 | 2.3 | 19527.54 | 164950.9 | 93.34 | 93.4264 | 0 | 0 | 23.1 | 2.31 | 4446.63 | 25915.1 | 93.56 | 93.66
0 | 0 | 18.1 | 1.81 | 188.24 | 1950973.1 | 188.94 | 96.6027 | 0 | 0 | 18.1 | 1.81 | 49.27 | 510619.7 | 198.77 | 97.6223 | 0 | 0 | 18.1 | 1.81 | 96.08 | 1912857.3 | 0 | 96.08 | 0 | 0 | 22.1 | 2.21 | 19666.99 | 166128.8 | 94.06 | 94.0935 | 0 | 0 | 22 | 2.2 | 4480.58 | 26112.9 | 94.39 | 94.3751
0 | 0 | 17.1 | 1.71 | 189.28 | 1961833.8 | 189.95 | 97.1364 | 0 | 0 | 17.1 | 1.71 | 49.56 | 513693.9 | 199.9 | 98.1969 | 0 | 0 | 17.1 | 1.71 | 96.8 | 1927058.1 | 0 | 96.8 | 0 | 0 | 21 | 2.1 | 19813.48 | 167366.3 | 94.67 | 94.7944 | 0 | 0 | 21.1 | 2.11 | 4517.08 | 26325.7 | 95.18 | 95.1439
0 | 0 | 16.1 | 1.61 | 190.6 | 1975478 | 191.24 | 97.8138 | 0 | 0 | 16.1 | 1.61 | 49.68 | 514889.3 | 200.65 | 98.4347 | 0 | 0 | 16.1 | 1.61 | 97.45 | 1939968.1 | 0 | 97.45 | 0 | 0 | 20.1 | 2.01 | 19993.08 | 168883.4 | 95.54 | 95.6537 | 0 | 0 | 20.1 | 2.01 | 4537.78 | 26446.3 | 95.54 | 95.5799
0 | 0 | 15.1 | 1.51 | 191.83 | 1988182.3 | 192.51 | 98.445 | 0 | 0 | 15.1 | 1.51 | 50.02 | 518472.6 | 201.98 | 99.1084 | 0 | 0 | 15.1 | 1.51 | 98.26 | 1956085.7 | 0 | 98.26 | 0 | 0 | 19 | 1.9 | 20099.32 | 169780.8 | 96.12 | 96.162 | 0 | 0 | 19 | 1.9 | 4579.23 | 26687.8 | 96.41 | 96.453
0 | 0 | 14.1 | 1.41 | 192.83 | 1998557.5 | 193.59 | 98.9582 | 0 | 0 | 14.1 | 1.41 | 50.28 | 521137.4 | 202.62 | 99.6235 | 0 | 0 | 14.1 | 1.41 | 98.81 | 1967050.1 | 0 | 98.81 | 0 | 0 | 18.1 | 1.81 | 20250.07 | 171054.1 | 96.78 | 96.8832 | 0 | 0 | 18 | 1.8 | 4621.18 | 26932.3 | 97.2 | 97.3366
0 | 0 | 13.1 | 1.31 | 193.81 | 2008732.4 | 194.58 | 99.4612 | 0 | 0 | 13 | 1.3 | 50.35 | 521874.6 | 203.24 | 99.7622 | 0 | 0 | 13.1 | 1.31 | 99.28 | 1976384.9 | 0 | 99.28 | 0 | 0 | 17 | 1.7 | 20361.72 | 171997.3 | 97.48 | 97.4174 | 0 | 0 | 17 | 1.7 | 4641.07 | 27048.3 | 97.76 | 97.7555
0 | 0 | 12.1 | 1.21 | 194.49 | 2015753 | 195.19 | 99.8101 | 0 | 0 | 12 | 1.2 | 50.47 | 523086.9 | 203.39 | 100 | 0 | 0 | 12.1 | 1.21 | 99.62 | 1983253.8 | 0 | 99.62 | 0 | 0 | 16 | 1.6 | 20488.14 | 173065.1 | 98.09 | 98.0222 | 0 | 0 | 16.1 | 1.61 | 4660.48 | 27161.4 | 98.17 | 98.1643
0 | 0 | 11.1 | 1.11 | 194.69 | 2017865 | 195.39 | 99.9128 | 0 | 0 | 11 | 1.1 | 50.3 | 521299.6 | 203.06 | 99.6632 | 0 | 0 | 11.1 | 1.11 | 99.85 | 1987881.1 | 0 | 99.85 | 0 | 0 | 15 | 1.5 | 20631.36 | 174275 | 98.71 | 98.7074 | 0 | 0 | 15 | 1.5 | 4686.98 | 27315.8 | 98.66 | 98.7225
0 | 0 | 10.1 | 1.01 | 194.86 | 2019680.6 | 195.61 | 100 | 0 | 0 | 10 | 1 | 50.11 | 519394.1 | 202.08 | 99.2867 | 0 | 0 | 10.1 | 1.01 | 100 | 1990801.1 | 0 | 100 | 0 | 0 | 14 | 1.4 | 20733.26 | 175135.7 | 99.27 | 99.1949 | 0 | 0 | 14 | 1.4 | 4721.13 | 27514.9 | 99.32 | 99.4418
0 | 0 | 9 | 0.9 | 194.27 | 2013542.3 | 194.97 | 99.6972 | 0 | 0 | 9.1 | 0.91 | 49.74 | 515496.3 | 200.86 | 98.5536 | 0 | 0 | 9 | 0.9 | 99.64 | 1983605.7 | 0 | 99.64 | 0 | 0 | 13 | 1.3 | 20833.26 | 175980.4 | 99.76 | 99.6734 | 0 | 0 | 13 | 1.3 | 4724.49 | 27534.4 | 99.48 | 99.5126
0 | 0 | 8 | 0.8 | 192.67 | 1996900.2 | 193.37 | 98.8761 | 0 | 0 | 8 | 0.8 | 49.14 | 509313 | 198.41 | 97.3648 | 0 | 0 | 8 | 0.8 | 99.19 | 1974644.6 | 0 | 99.19 | 0 | 0 | 12 | 1.2 | 20884.28 | 176411.4 | 99.96 | 99.9175 | 0 | 0 | 12 | 1.2 | 4732.19 | 27579.3 | 99.59 | 99.6748
0 | 0 | 7 | 0.7 | 190.66 | 1976127.2 | 191.41 | 97.8446 | 0 | 0 | 7 | 0.7 | 48.03 | 497779.2 | 193.79 | 95.1654 | 0 | 0 | 7 | 0.7 | 97.86 | 1948193.3 | 0 | 97.86 | 0 | 0 | 11 | 1.1 | 20901.53 | 176557.1 | 100 | 100 | 0 | 0 | 11 | 1.1 | 4747.63 | 27669.3 | 100 | 100
0 | 0 | 6 | 0.6 | 186.78 | 1935894.6 | 187.61 | 95.8534 | 0 | 0 | 6.1 | 0.61 | 46.66 | 483644.4 | 188.37 | 92.451 | 0 | 0 | 6 | 0.6 | 96.09 | 1913028.7 | 0 | 96.09 | 0 | 0 | 10 | 1 | 20883.12 | 176401.6 | 100 | 99.9119 | 0 | 0 | 10 | 1 | 4736.18 | 27602.5 | 99.75 | 99.7588
0 | 0 | 5.1 | 0.51 | 181.55 | 1881681.3 | 182.27 | 93.1695 | 0 | 0 | 5 | 0.5 | 44.43 | 460491.7 | 179.01 | 88.0325 | 0 | 0 | 5 | 0.5 | 93.24 | 1856301.7 | 0 | 93.24 | 0 | 0 | 9 | 0.9 | 20775.85 | 175495.4 | 99.53 | 99.3987 | 0 | 0 | 9 | 0.9 | 4725.03 | 27537.6 | 99.54 | 99.524
0 | 0 | 4 | 0.4 | 172.41 | 1786952.2 | 173.04 | 88.4789 | 0 | 0 | 4 | 0.4 | 41.12 | 426235.5 | 166.01 | 81.4741 | 0 | 0 | 4 | 0.4 | 88.87 | 1769324 | 0 | 88.87 | 0 | 0 | 8 | 0.8 | 20576.48 | 173811.4 | 98.46 | 98.4449 | 0 | 0 | 8 | 0.8 | 4669.53 | 27214.1 | 98.33 | 98.355
0 | 0 | 3 | 0.3 | 159.76 | 1655787.1 | 160.3 | 81.9871 | 0 | 0 | 3 | 0.3 | 36.83 | 381747.8 | 148.5 | 72.974 | 0 | 0 | 3 | 0.3 | 82.07 | 1633882.6 | 0 | 82.07 | 0 | 0 | 7 | 0.7 | 20307.11 | 171536 | 97.26 | 97.1561 | 0 | 0 | 7 | 0.7 | 4590.63 | 26754.3 | 96.72 | 96.6931
0 | 0 | 2 | 0.2 | 141.01 | 1461473.2 | 141.53 | 72.3648 | 0 | 0 | 2 | 0.2 | 30.79 | 319163.8 | 124.31 | 61.0065 | 0 | 0 | 2 | 0.2 | 71.71 | 1427586.1 | 0 | 71.71 | 0 | 0 | 6 | 0.6 | 19825.52 | 167467.9 | 94.93 | 94.852 | 0 | 0 | 6 | 0.6 | 4478.5 | 26100.8 | 94.42 | 94.3313
0 | 0 | 1 | 0.1 | 108.86 | 1128290.5 | 109.27 | 55.8657 | 0 | 0 | 1 | 0.1 | 21.42 | 221969.8 | 86.43 | 42.4411 | 0 | 0 | 1 | 0.1 | 54.14 | 1077751.1 | 0 | 54.14 | 0 | 0 | 5 | 0.5 | 19065.08 | 161044.5 | 91.27 | 91.2138 | 0 | 0 | 5 | 0.5 | 4304.1 | 25084.4 | 90.71 | 90.6579
0 | 0 | 0 | 0 | 97.47 | 1010231.6 | 97.86 | 50.0205 | 0 | 0 | 0 | 0 | 18.5 | 191745.4 | 74.65 | 36.6554 | 0 | 0 | 0 | 0 | 51.49 | 1024999.5 | 0 | 51.49 | 0 | 0 | 4 | 0.4 | 17962.42 | 151730.1 | 86.01 | 85.9383 | 0 | 0 | 4 | 0.4 | 4039.03 | 23539.5 | 85.21 | 85.0747
0 | 0 | -0.5 | -0.05 | 97.67 | 1012313 | 98.05 | 50.1232 | 0 | 0 | -0.5 | -0.05 | 18.51 | 191797 | 74.65 | 36.6753 | 0 | 0 | -0.5 | -0.05 | 51.55 | 1026298.4 | 0 | 51.55 | 0 | 0 | 3 | 0.3 | 16440.86 | 138877.4 | 78.71 | 78.6586 | 0 | 0 | 3 | 0.3 | 3647.13 | 21255.5 | 76.88 | 76.82
[24 empty] | 0 | 0 | 2 | 0.2 | 14113.38 | 119217 | 67.44 | 67.5232 | 0 | 0 | 2 | 0.2 | 3093.52 | 18029.1 | 65.22 | 65.1592
[24 empty] | 0 | 0 | 1 | 0.1 | 11993.23 | 101307.9 | 57.34 | 57.3797 | 0 | 0 | 1 | 0.1 | 2566.23 | 14956 | 54.12 | 54.0529
[24 empty] | 0 | 0 | 0 | 0 | 11947.06 | 100917.9 | 57.2 | 57.1588 | 0 | 0 | 0 | 0 | 2465.49 | 14368.9 | 52.04 | 51.931
[24 empty] | 0 | 0 | -0.5 | -0.05 | 11919.65 | 100686.4 | 57.02 | 57.0276 | 0 | 0 | -0.5 | -0.05 | 2467.99 | 14383.5 | 52.06 | 51.9836

## 2x2
Measurement time: | 12/04/2023 19:01:28 | [6 empty] | Measurement time: | 27/05/2023 11:00:52 | [6 empty] | Measurement time: | 27/05/2023 13:09:05
Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF)
Controller: | CCU | [6 empty] | Controller: | CCU | [6 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele | [6 empty] | Field detector: | PTW60019 Diamant | [6 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm
SSD: | 900 mm | [6 empty] | SSD: | 900 mm | [6 empty] | SSD: | 900 mm
Field size: | 20 x 20 mm | [6 empty] | Field size: | 20 x 20 mm | [6 empty] | Field size: | 20 x 20 mm
Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0
Measurement medium: | Water | [6 empty] | Measurement medium: | Water | [6 empty] | Measurement medium: | Water
Scan type: | Beam | [6 empty] | Scan type: | Beam | [6 empty] | Scan type: | Beam
Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep
Wedge: | none | [6 empty] | Wedge: | none | [6 empty] | Wedge: | none
Points [mm]: | [7 empty] | Points [mm]: | [7 empty] | Points [mm]:
Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 250 | 25 | 130.81 | 394497.1 | 18.48 | 18.4998 | 0 | 0 | 250 | 25 | 4352.31 | 36764.4 | 19.51 | 19.5396 | 0 | 0 | 250 | 25 | 1035.4 | 6034.3 | 20.31 | 20.3299
0 | 0 | 245 | 24.5 | 135.93 | 409950.4 | 19.19 | 19.2239 | 0 | 0 | 245 | 24.5 | 4497.78 | 37993.2 | 20.17 | 20.1927 | 0 | 0 | 245.1 | 24.51 | 1060.83 | 6182.5 | 20.81 | 20.8292
0 | 0 | 240 | 24 | 139.35 | 420271.5 | 19.68 | 19.7075 | 0 | 0 | 240 | 24 | 4645.93 | 39244.6 | 20.83 | 20.8578 | 0 | 0 | 240 | 24 | 1098.57 | 6402.5 | 21.55 | 21.5702
0 | 0 | 235 | 23.5 | 145.04 | 437426.3 | 20.49 | 20.5122 | 0 | 0 | 235 | 23.5 | 4799.86 | 40544.9 | 21.53 | 21.5489 | 0 | 0 | 235.1 | 23.51 | 1125.96 | 6562.1 | 22.09 | 22.108
0 | 0 | 230 | 23 | 149.65 | 451336.2 | 21.13 | 21.1642 | 0 | 0 | 230.1 | 23.01 | 4949.08 | 41805.4 | 22.19 | 22.2188 | 0 | 0 | 230.1 | 23.01 | 1159.4 | 6757 | 22.73 | 22.7646
0 | 0 | 225 | 22.5 | 155.04 | 467598 | 21.9 | 21.9265 | 0 | 0 | 225 | 22.5 | 5104.27 | 43116.2 | 22.89 | 22.9155 | 0 | 0 | 225.1 | 22.51 | 1206.75 | 7033 | 23.67 | 23.6943
0 | 0 | 220.1 | 22.01 | 159.99 | 482500.8 | 22.58 | 22.6265 | 0 | 0 | 220.1 | 22.01 | 5297.96 | 44752.3 | 23.75 | 23.7851 | 0 | 0 | 220.1 | 22.01 | 1245.89 | 7261.1 | 24.45 | 24.4628
0 | 0 | 215.1 | 21.51 | 166.05 | 500779.3 | 23.45 | 23.4836 | 0 | 0 | 215.1 | 21.51 | 5465.56 | 46168.1 | 24.52 | 24.5375 | 0 | 0 | 215 | 21.5 | 1288.13 | 7507.2 | 25.27 | 25.2922
0 | 0 | 210.1 | 21.01 | 171.79 | 518108.3 | 24.26 | 24.2954 | 0 | 0 | 210.1 | 21.01 | 5654.48 | 47763.9 | 25.35 | 25.3857 | 0 | 0 | 210.1 | 21.01 | 1328.91 | 7744.9 | 26.08 | 26.0929
0 | 0 | 205.1 | 20.51 | 177.19 | 534371.5 | 25.03 | 25.059 | 0 | 0 | 205.1 | 20.51 | 5838.19 | 49315.7 | 26.18 | 26.2104 | 0 | 0 | 205 | 20.5 | 1363.35 | 7945.6 | 26.72 | 26.7691
0 | 0 | 200.1 | 20.01 | 183.73 | 554094.7 | 25.94 | 25.984 | 0 | 0 | 200.1 | 20.01 | 6025.44 | 50897.4 | 27.03 | 27.0511 | 0 | 0 | 200.1 | 20.01 | 1405.82 | 8193.1 | 27.57 | 27.603
0 | 0 | 195 | 19.5 | 189.78 | 572349.4 | 26.8 | 26.8396 | 0 | 0 | 195 | 19.5 | 6255.02 | 52836.7 | 28.06 | 28.0818 | 0 | 0 | 195.1 | 19.51 | 1451.88 | 8461.6 | 28.46 | 28.5074
0 | 0 | 190 | 19 | 197.07 | 594351.2 | 27.84 | 27.8706 | 0 | 0 | 190.1 | 19.01 | 6461 | 54576.6 | 28.97 | 29.0065 | 0 | 0 | 190 | 19 | 1504.36 | 8767.4 | 29.52 | 29.5379
0 | 0 | 185.1 | 18.51 | 203.73 | 614439.3 | 28.77 | 28.8125 | 0 | 0 | 185 | 18.5 | 6683.61 | 56457.1 | 29.98 | 30.0059 | 0 | 0 | 185.1 | 18.51 | 1564.54 | 9118.2 | 30.7 | 30.7195
0 | 0 | 180.1 | 18.01 | 211.23 | 637054.7 | 29.84 | 29.8731 | 0 | 0 | 180 | 18 | 6912.66 | 58391.9 | 31.02 | 31.0343 | 0 | 0 | 180 | 18 | 1614.03 | 9406.6 | 31.64 | 31.6912
0 | 0 | 175.1 | 17.51 | 218.74 | 659683.4 | 30.9 | 30.9352 | 0 | 0 | 175 | 17.5 | 7140.61 | 60317.4 | 32.04 | 32.0576 | 0 | 0 | 175.1 | 17.51 | 1668.05 | 9721.4 | 32.73 | 32.7519
0 | 0 | 170.1 | 17.01 | 226.6 | 683398.6 | 32.02 | 32.0468 | 0 | 0 | 170.1 | 17.01 | 7393.65 | 62454.8 | 33.16 | 33.1937 | 0 | 0 | 170 | 17 | 1723.56 | 10044.9 | 33.81 | 33.8418
0 | 0 | 165.1 | 16.51 | 234.26 | 706502.6 | 33.09 | 33.1302 | 0 | 0 | 165.1 | 16.51 | 7656.36 | 64674 | 34.35 | 34.3731 | 0 | 0 | 165.1 | 16.51 | 1785.2 | 10404.2 | 35.03 | 35.0521
0 | 0 | 160.1 | 16.01 | 242.66 | 731842.8 | 34.29 | 34.3181 | 0 | 0 | 160.1 | 16.01 | 7928.62 | 66973.8 | 35.57 | 35.5954 | 0 | 0 | 160.1 | 16.01 | 1829.57 | 10662.7 | 35.91 | 35.9233
0 | 0 | 155 | 15.5 | 251.3 | 757901.2 | 35.5 | 35.54 | 0 | 0 | 155.1 | 15.51 | 8206.1 | 69317.7 | 36.82 | 36.8411 | 0 | 0 | 155.1 | 15.51 | 1890.73 | 11019.2 | 37.08 | 37.1242
0 | 0 | 150.1 | 15.01 | 260.24 | 784865.1 | 36.77 | 36.8044 | 0 | 0 | 150.1 | 15.01 | 8491.5 | 71728.4 | 38.09 | 38.1224 | 0 | 0 | 150.1 | 15.01 | 1952.77 | 11380.8 | 38.3 | 38.3423
0 | 0 | 145.1 | 14.51 | 269.89 | 813952.8 | 38.14 | 38.1691 | 0 | 0 | 145.1 | 14.51 | 8773.12 | 74107.4 | 39.36 | 39.3868 | 0 | 0 | 145.1 | 14.51 | 2027.22 | 11814.7 | 39.78 | 39.8041
0 | 0 | 140 | 14 | 279.78 | 843774 | 39.54 | 39.5678 | 0 | 0 | 140.1 | 14.01 | 9107.87 | 76935 | 40.86 | 40.8896 | 0 | 0 | 140.1 | 14.01 | 2105.85 | 12272.9 | 41.34 | 41.348
0 | 0 | 135.1 | 13.51 | 291.05 | 877785.4 | 41.11 | 41.1617 | 0 | 0 | 135 | 13.5 | 9428.28 | 79641.5 | 42.3 | 42.3281 | 0 | 0 | 135.1 | 13.51 | 2182.39 | 12719 | 42.81 | 42.8509
0 | 0 | 130.1 | 13.01 | 301.14 | 908190.9 | 42.56 | 42.5886 | 0 | 0 | 130.1 | 13.01 | 9763.45 | 82472.7 | 43.82 | 43.8328 | 0 | 0 | 130 | 13 | 2246.43 | 13092.3 | 44.06 | 44.1083
0 | 0 | 125.1 | 12.51 | 312.46 | 942343.4 | 44.15 | 44.1896 | 0 | 0 | 125.1 | 12.51 | 10102.28 | 85334.9 | 45.34 | 45.354 | 0 | 0 | 125.1 | 12.51 | 2330.61 | 13582.8 | 45.71 | 45.7611
0 | 0 | 120.1 | 12.01 | 324.76 | 979452.1 | 45.9 | 45.9291 | 0 | 0 | 120.1 | 12.01 | 10485.5 | 88571.9 | 47.04 | 47.0745 | 0 | 0 | 120 | 12 | 2408.64 | 14037.6 | 47.24 | 47.2932
0 | 0 | 115 | 11.5 | 336.26 | 1014113.4 | 47.52 | 47.5555 | 0 | 0 | 115 | 11.5 | 10854.07 | 91685.3 | 48.68 | 48.7291 | 0 | 0 | 115 | 11.5 | 2490.8 | 14516.4 | 48.9 | 48.9064
0 | 0 | 110.1 | 11.01 | 349.19 | 1053109.3 | 49.34 | 49.3841 | 0 | 0 | 110 | 11 | 11233.44 | 94889.8 | 50.43 | 50.4323 | 0 | 0 | 110 | 11 | 2596.48 | 15132.3 | 50.96 | 50.9814
0 | 0 | 105.1 | 10.51 | 361.44 | 1090072.1 | 51.08 | 51.1165 | 0 | 0 | 105.1 | 10.51 | 11691.03 | 98755.2 | 52.45 | 52.4867 | 0 | 0 | 105.1 | 10.51 | 2689.44 | 15674.1 | 52.79 | 52.8067
0 | 0 | 100.1 | 10.01 | 376.21 | 1134618.6 | 53.17 | 53.2054 | 0 | 0 | 100.1 | 10.01 | 12068.58 | 101944.4 | 54.17 | 54.1817 | 0 | 0 | 100.1 | 10.01 | 2778.28 | 16191.9 | 54.5 | 54.5511
0 | 0 | 95.1 | 9.51 | 390.44 | 1177525.9 | 55.18 | 55.2179 | 0 | 0 | 95.1 | 9.51 | 12521.74 | 105772.3 | 56.16 | 56.2161 | 0 | 0 | 95.1 | 9.51 | 2889.48 | 16839.9 | 56.69 | 56.7345
0 | 0 | 90.1 | 9.01 | 405.06 | 1221607.9 | 57.25 | 57.2855 | 0 | 0 | 90.1 | 9.01 | 12969.42 | 109553.9 | 58.19 | 58.226 | 0 | 0 | 90.1 | 9.01 | 2980.24 | 17368.9 | 58.47 | 58.5165
0 | 0 | 85 | 8.5 | 421.24 | 1270418.4 | 59.53 | 59.5737 | 0 | 0 | 85 | 8.5 | 13482.85 | 113890.8 | 60.5 | 60.531 | 0 | 0 | 85 | 8.5 | 3092.84 | 18025.1 | 60.68 | 60.7274
0 | 0 | 80 | 8 | 436.63 | 1316841.1 | 61.73 | 61.7503 | 0 | 0 | 80 | 8 | 13985.7 | 118138.5 | 62.8 | 62.7885 | 0 | 0 | 80 | 8 | 3213.61 | 18729 | 63.05 | 63.0987
0 | 0 | 75.1 | 7.51 | 453.46 | 1367581.3 | 64.09 | 64.1305 | 0 | 0 | 75 | 7.5 | 14515.67 | 122615.1 | 65.12 | 65.1678 | 0 | 0 | 75 | 7.5 | 3336.45 | 19444.9 | 65.43 | 65.5106
0 | 0 | 70 | 7 | 470.72 | 1419628.5 | 66.55 | 66.5714 | 0 | 0 | 70 | 7 | 15046.98 | 127103.2 | 67.56 | 67.5531 | 0 | 0 | 70 | 7 | 3450.27 | 20108.2 | 67.71 | 67.7455
0 | 0 | 65 | 6.5 | 489.48 | 1476222 | 69.2 | 69.2246 | 0 | 0 | 65 | 6.5 | 15608.38 | 131845.4 | 70.08 | 70.0735 | 0 | 0 | 65 | 6.5 | 3581.51 | 20873.1 | 70.27 | 70.3223
0 | 0 | 60 | 6 | 508.68 | 1534123.9 | 71.94 | 71.9399 | 0 | 0 | 60 | 6 | 16192.39 | 136778.6 | 72.67 | 72.6954 | 0 | 0 | 60 | 6 | 3715.4 | 21653.4 | 72.87 | 72.9513
0 | 0 | 55 | 5.5 | 527.3 | 1590268.3 | 74.56 | 74.5733 | 0 | 0 | 55 | 5.5 | 16781.97 | 141758.8 | 75.33 | 75.3423 | 0 | 0 | 55 | 5.5 | 3859.69 | 22494.3 | 75.73 | 75.7844
0 | 0 | 50 | 5 | 548.69 | 1654774.4 | 77.56 | 77.5983 | 0 | 0 | 50 | 5 | 17456.78 | 147459 | 78.35 | 78.3719 | 0 | 0 | 49.9 | 4.99 | 3997.03 | 23294.8 | 78.45 | 78.481
0 | 0 | 44.9 | 4.49 | 569.99 | 1719032 | 80.58 | 80.6107 | 0 | 0 | 44.9 | 4.49 | 18105.95 | 152942.6 | 81.26 | 81.2863 | 0 | 0 | 44.9 | 4.49 | 4158.19 | 24234 | 81.58 | 81.6454
0 | 0 | 39.9 | 3.99 | 591.42 | 1783651.4 | 83.62 | 83.6414 | 0 | 0 | 39.9 | 3.99 | 18803.16 | 158832 | 84.38 | 84.4164 | 0 | 0 | 39.9 | 3.99 | 4303.6 | 25081.5 | 84.44 | 84.5005
0 | 0 | 34.9 | 3.49 | 614.39 | 1852928 | 86.86 | 86.8899 | 0 | 0 | 34.9 | 3.49 | 19492.48 | 164654.7 | 87.46 | 87.5111 | 0 | 0 | 34.9 | 3.49 | 4459.93 | 25992.6 | 87.54 | 87.57
0 | 0 | 29.9 | 2.99 | 637.85 | 1923695.3 | 90.17 | 90.2078 | 0 | 0 | 29.9 | 2.99 | 20198.31 | 170617 | 90.66 | 90.6799 | 0 | 0 | 29.9 | 2.99 | 4629.01 | 26977.9 | 90.87 | 90.8898
0 | 0 | 24.9 | 2.49 | 661.73 | 1995715.4 | 93.58 | 93.585 | 0 | 0 | 25.1 | 2.51 | 20907.6 | 176608.3 | 93.87 | 93.8643 | 0 | 0 | 25 | 2.5 | 4782.62 | 27873.2 | 93.83 | 93.9059
0 | 0 | 20 | 2 | 684.93 | 2065685.6 | 96.87 | 96.866 | 0 | 0 | 24.1 | 2.41 | 21064.68 | 177935.2 | 94.47 | 94.5695 | 0 | 0 | 24 | 2.4 | 4810.47 | 28035.5 | 94.4 | 94.4528
0 | 0 | 19.1 | 1.91 | 689.6 | 2079761.8 | 97.47 | 97.5265 | 0 | 0 | 23.1 | 2.31 | 21210.77 | 179169.3 | 95.16 | 95.2253 | 0 | 0 | 23.1 | 2.31 | 4842.47 | 28222 | 95.1 | 95.0811
0 | 0 | 18.1 | 1.81 | 693.19 | 2090569.5 | 98.01 | 98.0342 | 0 | 0 | 22 | 2.2 | 21324.14 | 180126.9 | 95.79 | 95.7343 | 0 | 0 | 22 | 2.2 | 4884.73 | 28468.3 | 95.85 | 95.9109
0 | 0 | 17.1 | 1.71 | 696.62 | 2100918.5 | 98.56 | 98.5193 | 0 | 0 | 21.1 | 2.11 | 21485.02 | 181485.9 | 96.43 | 96.4566 | 0 | 0 | 21.1 | 2.11 | 4926.69 | 28712.8 | 96.69 | 96.7347
0 | 0 | 16.1 | 1.61 | 700.68 | 2113176.1 | 99.05 | 99.0935 | 0 | 0 | 20.1 | 2.01 | 21639.65 | 182792.1 | 97.18 | 97.1508 | 0 | 0 | 20.1 | 2.01 | 4962.25 | 28920.1 | 97.36 | 97.4329
0 | 0 | 15.1 | 1.51 | 703.6 | 2121986.4 | 99.47 | 99.5064 | 0 | 0 | 19 | 1.9 | 21770.89 | 183900.7 | 97.73 | 97.74 | 0 | 0 | 19.1 | 1.91 | 4992.31 | 29095.3 | 98.01 | 98.0232
0 | 0 | 14.1 | 1.41 | 705.71 | 2128336.5 | 99.78 | 99.8048 | 0 | 0 | 18 | 1.8 | 21886.54 | 184877.5 | 98.24 | 98.2592 | 0 | 0 | 18 | 1.8 | 5013.13 | 29216.6 | 98.42 | 98.432
0 | 0 | 13.1 | 1.31 | 707.03 | 2132334.9 | 100 | 99.9915 | 0 | 0 | 17.1 | 1.71 | 21963.72 | 185529.5 | 98.6 | 98.6057 | 0 | 0 | 17.1 | 1.71 | 5031.56 | 29324 | 98.79 | 98.7938
0 | 0 | 12.1 | 1.21 | 707.09 | 2132496.7 | 99.96 | 100 | 0 | 0 | 16 | 1.6 | 22065.62 | 186390.3 | 99.13 | 99.0632 | 0 | 0 | 16.1 | 1.61 | 5073.24 | 29567 | 99.57 | 99.6122
0 | 0 | 11.1 | 1.11 | 706.23 | 2129905.9 | 99.85 | 99.8784 | 0 | 0 | 15 | 1.5 | 22152.32 | 187122.6 | 99.47 | 99.4524 | 0 | 0 | 15 | 1.5 | 5091.62 | 29674.1 | 99.92 | 99.9731
0 | 0 | 10.1 | 1.01 | 704.77 | 2125506.1 | 99.61 | 99.6719 | 0 | 0 | 14.1 | 1.41 | 22215.74 | 187658.3 | 99.77 | 99.7371 | 0 | 0 | 14 | 1.4 | 5092.9 | 29681.5 | 99.93 | 99.9982
0 | 0 | 9.1 | 0.91 | 698.29 | 2105965.2 | 98.73 | 98.7555 | 0 | 0 | 13 | 1.3 | 22263.61 | 188062.7 | 100 | 99.9521 | 0 | 0 | 13 | 1.3 | 5092.99 | 29682.1 | 100 | 100
0 | 0 | 8 | 0.8 | 689.43 | 2079233 | 97.51 | 97.5024 | 0 | 0 | 12 | 1.2 | 22274.29 | 188152.9 | 99.96 | 100 | 0 | 0 | 12 | 1.2 | 5085.68 | 29639.5 | 99.86 | 99.8565
0 | 0 | 7 | 0.7 | 676.34 | 2039749.6 | 95.68 | 95.6512 | 0 | 0 | 11 | 1.1 | 22210.39 | 187613.2 | 99.72 | 99.7131 | 0 | 0 | 11 | 1.1 | 5068.65 | 29540.2 | 99.47 | 99.5221
0 | 0 | 6 | 0.6 | 661.52 | 1995062.8 | 93.52 | 93.5553 | 0 | 0 | 10 | 1 | 22040.62 | 186179.1 | 98.97 | 98.9509 | 0 | 0 | 10 | 1 | 5024 | 29280 | 98.63 | 98.6454
0 | 0 | 5 | 0.5 | 637.08 | 1921357.8 | 90.12 | 90.0989 | 0 | 0 | 9 | 0.9 | 21882.97 | 184847.4 | 98.2 | 98.2432 | 0 | 0 | 9 | 0.9 | 4962.8 | 28923.3 | 97.47 | 97.4437
0 | 0 | 4 | 0.4 | 605.88 | 1827272.6 | 85.7 | 85.6864 | 0 | 0 | 8 | 0.8 | 21572.79 | 182227.2 | 96.78 | 96.8506 | 0 | 0 | 8 | 0.8 | 4887.58 | 28484.9 | 95.93 | 95.9668
0 | 0 | 3 | 0.3 | 556.97 | 1679759.1 | 78.8 | 78.7693 | 0 | 0 | 7 | 0.7 | 21125.13 | 178445.9 | 94.79 | 94.8409 | 0 | 0 | 7 | 0.7 | 4768.25 | 27789.4 | 93.58 | 93.6238
0 | 0 | 2 | 0.2 | 486.47 | 1467145.4 | 68.82 | 68.7989 | 0 | 0 | 6 | 0.6 | 20485.48 | 173042.7 | 92 | 91.9692 | 0 | 0 | 6 | 0.6 | 4585.88 | 26726.6 | 90.09 | 90.043
0 | 0 | 1 | 0.1 | 370.24 | 1116591.2 | 52.37 | 52.3611 | 0 | 0 | 5 | 0.5 | 19598.06 | 165546.6 | 87.97 | 87.9851 | 0 | 0 | 5 | 0.5 | 4374.46 | 25494.4 | 85.87 | 85.8918
0 | 0 | 0 | 0 | 343.63 | 1036362.3 | 48.61 | 48.5978 | 0 | 0 | 4 | 0.4 | 18447.34 | 155826.3 | 82.88 | 82.819 | 0 | 0 | 4 | 0.4 | 4082.86 | 23795 | 80.22 | 80.1663
0 | 0 | -0.5 | -0.05 | 343.52 | 1036029.9 | 48.6 | 48.5822 | 0 | 0 | 3 | 0.3 | 16738.45 | 141391.2 | 75.17 | 75.147 | 0 | 0 | 3 | 0.3 | 3652.52 | 21286.9 | 71.71 | 71.7166
[8 empty] | 0 | 0 | 2 | 0.2 | 14495.92 | 122448.3 | 65.09 | 65.0792 | 0 | 0 | 2 | 0.2 | 3037.09 | 17700.2 | 59.63 | 59.6328
[8 empty] | 0 | 0 | 1 | 0.1 | 12391.12 | 104668.9 | 55.6 | 55.6297 | 0 | 0 | 1 | 0.1 | 2558.47 | 14910.8 | 50.27 | 50.2351
[8 empty] | 0 | 0 | 0 | 0 | 12385.15 | 104618.5 | 55.61 | 55.6029 | 0 | 0 | 0 | 0 | 2472.71 | 14411 | 48.52 | 48.5512
[8 empty] | 0 | 0 | -0.5 | -0.05 | 12354.2 | 104357 | 55.46 | 55.4639 | 0 | 0 | -0.5 | -0.05 | 2444.68 | 14247.6 | 48 | 48.0009

## 3x3
Measurement time: | 12/04/2023 19:22:22 | [6 empty] | Measurement time: | 27/05/2023 11:26:24 | [6 empty] | Measurement time: | 27/05/2023 13:39:45
Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator | [6 empty] | Radiation device: | Novalis TB Accelerator
Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF) | [6 empty] | Energy: | 6 MV (FFF)
Controller: | CCU | [6 empty] | Controller: | CCU | [6 empty] | Controller: | CCU
Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41) | [6 empty] | Measurement device: | Blue Phantom 2 (48*48*41)
Field detector: | SFD parallele | [6 empty] | Field detector: | PTW60019 Diamant | [6 empty] | Field detector: | Razor Nano
Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission | [6 empty] | Reference detector: | Stealth CI transmission
SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm | [6 empty] | SAD: | 1000 mm
SSD: | 900 mm | [6 empty] | SSD: | 900 mm | [6 empty] | SSD: | 900 mm
Field size: | 30 x 30 mm | [6 empty] | Field size: | 30 x 30 mm | [6 empty] | Field size: | 30 x 30 mm
Gantry angle: | 0 | [6 empty] | Gantry angle: | 0 | [6 empty] | Gantry angle: | 0
Measurement medium: | Water | [6 empty] | Measurement medium: | Water | [6 empty] | Measurement medium: | Water
Scan type: | Beam | [6 empty] | Scan type: | Beam | [6 empty] | Scan type: | Beam
Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep | [6 empty] | Scan mode: | StepByStep
Wedge: | none | [6 empty] | Wedge: | none | [6 empty] | Wedge: | none
Points [mm]: | [7 empty] | Points [mm]: | [7 empty] | Points [mm]:
Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%) | Inline | Crossline | Depth | Depth (cm) | Normalized field | Current field | Ratio | Dose relative (%)
0 | 0 | 250 | 25 | 34.88 | 413631.7 | 34.88 | 19.187 | 0 | 0 | 250 | 25 | 4567.21 | 38579.7 | 20.09 | 20.1077 | 0 | 0 | 250 | 25 | 1103.89 | 6433.5 | 33.39 | 21.521
0 | 0 | 245 | 24.5 | 36.07 | 427733 | 36.07 | 19.8416 | 0 | 0 | 245.1 | 24.51 | 4715.48 | 39832.1 | 20.73 | 20.7604 | 0 | 0 | 245.1 | 24.51 | 1127.72 | 6572.4 | 34.12 | 21.9856
0 | 0 | 240 | 24 | 37.22 | 441414.7 | 37.22 | 20.4742 | 0 | 0 | 240 | 24 | 4871.62 | 41151 | 21.42 | 21.4479 | 0 | 0 | 240 | 24 | 1170.78 | 6823.3 | 35.44 | 22.8251
0 | 0 | 235 | 23.5 | 38.53 | 456928.4 | 38.53 | 21.1948 | 0 | 0 | 235.1 | 23.51 | 5022.58 | 42426.2 | 22.09 | 22.1125 | 0 | 0 | 235.1 | 23.51 | 1208.94 | 7045.7 | 36.55 | 23.569
0 | 0 | 230 | 23 | 39.79 | 471830.7 | 39.79 | 21.8879 | 0 | 0 | 230 | 23 | 5205.61 | 43972.3 | 22.9 | 22.9183 | 0 | 0 | 230 | 23 | 1240.15 | 7227.6 | 37.51 | 24.1775
0 | 0 | 225 | 22.5 | 41.29 | 489607.4 | 41.29 | 22.713 | 0 | 0 | 225.1 | 22.51 | 5369.54 | 45357 | 23.63 | 23.64 | 0 | 0 | 225.1 | 22.51 | 1275.46 | 7433.4 | 38.57 | 24.8659
0 | 0 | 220.1 | 22.01 | 42.46 | 503545.1 | 42.46 | 23.3566 | 0 | 0 | 220.1 | 22.01 | 5552.18 | 46899.8 | 24.41 | 24.4441 | 0 | 0 | 220.1 | 22.01 | 1314.54 | 7661.2 | 39.76 | 25.6278
0 | 0 | 215.1 | 21.51 | 44.08 | 522768.6 | 44.08 | 24.2478 | 0 | 0 | 215 | 21.5 | 5742.7 | 48509.1 | 25.26 | 25.2829 | 0 | 0 | 215 | 21.5 | 1357.51 | 7911.6 | 41.07 | 26.4655
0 | 0 | 210 | 21 | 45.72 | 542159.8 | 45.72 | 25.1499 | 0 | 0 | 210 | 21 | 5917.15 | 49982.7 | 26.03 | 26.0509 | 0 | 0 | 210.1 | 21.01 | 1394.1 | 8124.8 | 42.19 | 27.1788
0 | 0 | 205.1 | 20.51 | 47.32 | 561079.4 | 47.32 | 26.03 | 0 | 0 | 205.1 | 20.51 | 6128.3 | 51766.3 | 26.97 | 26.9805 | 0 | 0 | 205.1 | 20.51 | 1447.49 | 8436 | 43.79 | 28.2197
0 | 0 | 200.1 | 20.01 | 48.95 | 580410.9 | 48.95 | 26.9267 | 0 | 0 | 200 | 20 | 6333.85 | 53502.6 | 27.86 | 27.8855 | 0 | 0 | 200.1 | 20.01 | 1500.46 | 8744.7 | 45.38 | 29.2524
0 | 0 | 195 | 19.5 | 50.7 | 601241.2 | 50.7 | 27.8893 | 0 | 0 | 195 | 19.5 | 6558.37 | 55399.1 | 28.85 | 28.874 | 0 | 0 | 195.1 | 19.51 | 1547.95 | 9021.5 | 46.85 | 30.1782
0 | 0 | 190.1 | 19.01 | 52.37 | 621028.8 | 52.37 | 28.808 | 0 | 0 | 190.1 | 19.01 | 6772.9 | 57211.3 | 29.79 | 29.8185 | 0 | 0 | 190.1 | 19.01 | 1594.29 | 9291.5 | 48.25 | 31.0817
0 | 0 | 185.1 | 18.51 | 54.08 | 641354.9 | 54.08 | 29.7486 | 0 | 0 | 185.1 | 18.51 | 6988.05 | 59028.7 | 30.74 | 30.7657 | 0 | 0 | 185 | 18.5 | 1651.08 | 9622.5 | 49.96 | 32.1888
0 | 0 | 180.1 | 18.01 | 56.14 | 665748 | 56.14 | 30.8818 | 0 | 0 | 180.1 | 18.01 | 7244.52 | 61195.1 | 31.88 | 31.8948 | 0 | 0 | 180 | 18 | 1711.4 | 9974.1 | 51.78 | 33.3648
0 | 0 | 175.1 | 17.51 | 57.93 | 686905.9 | 57.93 | 31.8664 | 0 | 0 | 175.1 | 17.51 | 7486.14 | 63236.1 | 32.95 | 32.9586 | 0 | 0 | 175.1 | 17.51 | 1762.4 | 10271.3 | 53.36 | 34.3591
0 | 0 | 170.1 | 17.01 | 60.26 | 714583.4 | 60.26 | 33.1481 | 0 | 0 | 170.1 | 17.01 | 7755.56 | 65511.9 | 34.12 | 34.1448 | 0 | 0 | 170.1 | 17.01 | 1824.34 | 10632.3 | 55.19 | 35.5666
0 | 0 | 165.1 | 16.51 | 62.47 | 740830.1 | 62.47 | 34.3638 | 0 | 0 | 165.1 | 16.51 | 8025.91 | 67795.6 | 35.32 | 35.335 | 0 | 0 | 165.1 | 16.51 | 1882.32 | 10970.2 | 56.96 | 36.697
0 | 0 | 160.1 | 16.01 | 64.64 | 766518.6 | 64.64 | 35.5575 | 0 | 0 | 160.1 | 16.01 | 8286.23 | 69994.5 | 36.45 | 36.4811 | 0 | 0 | 160.1 | 16.01 | 1936.29 | 11284.7 | 58.61 | 37.7492
0 | 0 | 155.1 | 15.51 | 66.78 | 791860.9 | 66.78 | 36.7347 | 0 | 0 | 155.1 | 15.51 | 8593.54 | 72590.4 | 37.81 | 37.8341 | 0 | 0 | 155.1 | 15.51 | 2015.07 | 11743.9 | 61.02 | 39.285
0 | 0 | 150.1 | 15.01 | 69.3 | 821750.4 | 69.3 | 38.1209 | 0 | 0 | 150.1 | 15.01 | 8882.5 | 75031.3 | 39.09 | 39.1062 | 0 | 0 | 150.1 | 15.01 | 2071.26 | 12071.4 | 62.7 | 40.3805
0 | 0 | 145.1 | 14.51 | 71.81 | 851554.8 | 71.81 | 39.5016 | 0 | 0 | 145.1 | 14.51 | 9204.44 | 77750.7 | 40.48 | 40.5236 | 0 | 0 | 145.1 | 14.51 | 2141.54 | 12480.9 | 64.83 | 41.7506
0 | 0 | 140.1 | 14.01 | 74.3 | 881068.5 | 74.3 | 40.8713 | 0 | 0 | 140 | 14 | 9535.39 | 80546.3 | 41.94 | 41.9807 | 0 | 0 | 140.1 | 14.01 | 2213.24 | 12898.8 | 66.99 | 43.1485
0 | 0 | 135.1 | 13.51 | 77.21 | 915550.3 | 77.21 | 42.4721 | 0 | 0 | 135.1 | 13.51 | 9853.46 | 83233.1 | 43.36 | 43.381 | 0 | 0 | 135.1 | 13.51 | 2288.79 | 13339.1 | 69.3 | 44.6214
0 | 0 | 130.1 | 13.01 | 79.95 | 948018.6 | 79.95 | 43.9793 | 0 | 0 | 130.1 | 13.01 | 10211.16 | 86254.6 | 44.94 | 44.9558 | 0 | 0 | 130.1 | 13.01 | 2363.23 | 13773 | 71.52 | 46.0726
0 | 0 | 125 | 12.5 | 82.91 | 983151.5 | 82.91 | 45.6076 | 0 | 0 | 125.1 | 12.51 | 10586.27 | 89423.1 | 46.6 | 46.6073 | 0 | 0 | 125.1 | 12.51 | 2448.79 | 14271.6 | 74.12 | 47.7407
0 | 0 | 120.1 | 12.01 | 85.97 | 1019469.4 | 85.97 | 47.2908 | 0 | 0 | 120.1 | 12.01 | 10949.35 | 92490.1 | 48.17 | 48.2058 | 0 | 0 | 120.1 | 12.01 | 2530.84 | 14749.8 | 76.63 | 49.3403
0 | 0 | 115.1 | 11.51 | 89.26 | 1058515.6 | 89.26 | 49.1006 | 0 | 0 | 115 | 11.5 | 11356.22 | 95927 | 50 | 49.9971 | 0 | 0 | 115.1 | 11.51 | 2623.35 | 15288.9 | 79.44 | 51.1438
0 | 0 | 110.1 | 11.01 | 92.39 | 1095555.7 | 92.39 | 50.8224 | 0 | 0 | 110.1 | 11.01 | 11741.9 | 99184.9 | 51.69 | 51.6951 | 0 | 0 | 110.1 | 11.01 | 2719.52 | 15849.4 | 82.38 | 53.0187
0 | 0 | 105.1 | 10.51 | 95.97 | 1137995.6 | 95.97 | 52.7917 | 0 | 0 | 105.1 | 10.51 | 12192.92 | 102994.7 | 53.64 | 53.6807 | 0 | 0 | 105.1 | 10.51 | 2817.27 | 16419.1 | 85.28 | 54.9244
0 | 0 | 100 | 10 | 99.43 | 1179027 | 99.43 | 54.695 | 0 | 0 | 100.1 | 10.01 | 12612.01 | 106534.8 | 55.49 | 55.5258 | 0 | 0 | 100.1 | 10.01 | 2906.58 | 16939.6 | 87.98 | 56.6655
0 | 0 | 95.1 | 9.51 | 103.24 | 1224206.8 | 103.24 | 56.7908 | 0 | 0 | 95.1 | 9.51 | 13091.22 | 110582.7 | 57.6 | 57.6356 | 0 | 0 | 95.1 | 9.51 | 3007.59 | 17528.3 | 91.05 | 58.6348
0 | 0 | 90.1 | 9.01 | 106.98 | 1268568.6 | 106.98 | 58.8481 | 0 | 0 | 90.1 | 9.01 | 13565.62 | 114590 | 59.67 | 59.7242 | 0 | 0 | 90.1 | 9.01 | 3122.75 | 18199.5 | 94.51 | 60.8799
0 | 0 | 85 | 8.5 | 111.13 | 1317761.6 | 111.13 | 61.131 | 0 | 0 | 85 | 8.5 | 14047.46 | 118660.1 | 61.85 | 61.8456 | 0 | 0 | 85 | 8.5 | 3235.83 | 18858.5 | 97.99 | 63.0845
0 | 0 | 80 | 8 | 115.17 | 1365774 | 115.17 | 63.3533 | 0 | 0 | 80.1 | 8.01 | 14550.88 | 122912.6 | 64.08 | 64.0619 | 0 | 0 | 80 | 8 | 3340.91 | 19470.9 | 101.15 | 65.1331
0 | 0 | 75 | 7.5 | 119.59 | 1418152.3 | 119.59 | 65.7847 | 0 | 0 | 75 | 7.5 | 15091.33 | 127477.8 | 66.48 | 66.4413 | 0 | 0 | 75 | 7.5 | 3459.8 | 20163.8 | 104.76 | 67.4509
0 | 0 | 70 | 7 | 124.2 | 1472774.7 | 124.2 | 68.3206 | 0 | 0 | 70 | 7 | 15661.39 | 132293.2 | 68.96 | 68.9511 | 0 | 0 | 70 | 7 | 3584.7 | 20891.7 | 108.57 | 69.8859
0 | 0 | 65 | 6.5 | 128.82 | 1527614.8 | 128.82 | 70.862 | 0 | 0 | 65 | 6.5 | 16237.96 | 137163.5 | 71.44 | 71.4895 | 0 | 0 | 65 | 6.5 | 3723.92 | 21703.1 | 112.76 | 72.6001
0 | 0 | 60 | 6 | 133.57 | 1583929.5 | 133.57 | 73.4749 | 0 | 0 | 60 | 6 | 16803.42 | 141940 | 74 | 73.979 | 0 | 0 | 60 | 6 | 3855.43 | 22469.5 | 116.77 | 75.164
0 | 0 | 55 | 5.5 | 138.73 | 1645072.8 | 138.73 | 76.3133 | 0 | 0 | 55 | 5.5 | 17442.05 | 147334.6 | 76.79 | 76.7906 | 0 | 0 | 55 | 5.5 | 3977.08 | 23178.5 | 120.45 | 77.5356
0 | 0 | 49.9 | 4.99 | 143.87 | 1706103.5 | 143.87 | 79.1408 | 0 | 0 | 49.9 | 4.99 | 18075.36 | 152684.2 | 79.63 | 79.5789 | 0 | 0 | 49.9 | 4.99 | 4125.48 | 24043.4 | 124.92 | 80.4287
0 | 0 | 44.9 | 4.49 | 149.29 | 1770329.5 | 149.29 | 82.1222 | 0 | 0 | 45 | 4.5 | 18757.13 | 158443.2 | 82.6 | 82.5804 | 0 | 0 | 44.9 | 4.49 | 4285.42 | 24975.5 | 129.79 | 83.5469
0 | 0 | 39.9 | 3.99 | 154.76 | 1835207.4 | 154.76 | 85.1312 | 0 | 0 | 39.9 | 3.99 | 19416.32 | 164011.4 | 85.48 | 85.4826 | 0 | 0 | 39.9 | 3.99 | 4436.94 | 25858.5 | 134.35 | 86.5009
0 | 0 | 34.9 | 3.49 | 160.5 | 1903217.7 | 160.5 | 88.2887 | 0 | 0 | 34.9 | 3.49 | 20103.26 | 169814 | 88.54 | 88.5069 | 0 | 0 | 34.9 | 3.49 | 4581.94 | 26703.6 | 138.78 | 89.3277
0 | 0 | 29.9 | 2.99 | 165.87 | 1966919.3 | 165.87 | 91.2426 | 0 | 0 | 29.9 | 2.99 | 20816.95 | 175842.7 | 91.64 | 91.649 | 0 | 0 | 29.9 | 2.99 | 4740.54 | 27628 | 143.59 | 92.4197
0 | 0 | 24.9 | 2.49 | 171.79 | 2037127.7 | 171.79 | 94.4991 | 0 | 0 | 25.1 | 2.51 | 21448.91 | 181180.9 | 94.48 | 94.4313 | 0 | 0 | 25.1 | 2.51 | 4882.94 | 28457.9 | 147.92 | 95.1959
0 | 0 | 20.1 | 2.01 | 177.35 | 2103076.4 | 177.35 | 97.5576 | 0 | 0 | 24.1 | 2.41 | 21643.68 | 182826.1 | 95.34 | 95.2888 | 0 | 0 | 24.1 | 2.41 | 4913.75 | 28637.4 | 148.88 | 95.7966
0 | 0 | 19.1 | 1.91 | 178.09 | 2111876.1 | 178.09 | 97.9647 | 0 | 0 | 23.1 | 2.31 | 21732.07 | 183572.7 | 95.72 | 95.678 | 0 | 0 | 23 | 2.3 | 4962.23 | 28919.9 | 150.31 | 96.7417
0 | 0 | 18.1 | 1.81 | 179.15 | 2124390.7 | 179.15 | 98.5478 | 0 | 0 | 22.1 | 2.21 | 21865 | 184695.6 | 96.36 | 96.2632 | 0 | 0 | 22 | 2.2 | 4986.01 | 29058.6 | 151.03 | 97.2053
0 | 0 | 17.1 | 1.71 | 179.8 | 2132146.7 | 179.8 | 98.9053 | 0 | 0 | 21.1 | 2.11 | 22047.36 | 186236 | 97.09 | 97.0661 | 0 | 0 | 21.1 | 2.11 | 5014.36 | 29223.8 | 151.88 | 97.758
0 | 0 | 16.1 | 1.61 | 180.41 | 2139363.8 | 180.41 | 99.2409 | 0 | 0 | 20.1 | 2.01 | 22156.32 | 187156.4 | 97.55 | 97.5458 | 0 | 0 | 20.1 | 2.01 | 5028.61 | 29306.9 | 152.31 | 98.0358
0 | 0 | 15.1 | 1.51 | 181.34 | 2150432.3 | 181.34 | 99.7525 | 0 | 0 | 19 | 1.9 | 22303.18 | 188396.9 | 98.23 | 98.1923 | 0 | 0 | 19 | 1.9 | 5047.84 | 29418.9 | 152.91 | 98.4107
0 | 0 | 14.1 | 1.41 | 181.41 | 2151191.4 | 181.41 | 99.791 | 0 | 0 | 18 | 1.8 | 22390.39 | 189133.7 | 98.66 | 98.5763 | 0 | 0 | 18 | 1.8 | 5074.19 | 29572.5 | 153.64 | 98.9244
0 | 0 | 13.1 | 1.31 | 181.79 | 2155701.9 | 181.79 | 100 | 0 | 0 | 17.1 | 1.71 | 22490.58 | 189979.9 | 99 | 99.0174 | 0 | 0 | 17.1 | 1.71 | 5091.75 | 29674.8 | 154.23 | 99.2668
0 | 0 | 12 | 1.2 | 181.58 | 2153216.2 | 181.58 | 99.8845 | 0 | 0 | 16 | 1.6 | 22538.07 | 190381.1 | 99.28 | 99.2265 | 0 | 0 | 16.1 | 1.61 | 5108.64 | 29773.2 | 154.79 | 99.5961
0 | 0 | 11.1 | 1.11 | 181.27 | 2149580.5 | 181.27 | 99.714 | 0 | 0 | 15 | 1.5 | 22649.64 | 191323.5 | 99.81 | 99.7177 | 0 | 0 | 15 | 1.5 | 5129.36 | 29894 | 155.43 | 100
0 | 0 | 10.1 | 1.01 | 180.66 | 2142340.2 | 180.66 | 99.3784 | 0 | 0 | 14.1 | 1.41 | 22713.77 | 191865.3 | 100 | 100 | 0 | 0 | 14 | 1.4 | 5127.91 | 29885.6 | 155.36 | 99.9717
0 | 0 | 9.1 | 0.91 | 178.91 | 2121608.9 | 178.91 | 98.4158 | 0 | 0 | 13.1 | 1.31 | 22691.8 | 191679.7 | 99.98 | 99.9033 | 0 | 0 | 13 | 1.3 | 5126.17 | 29875.4 | 155.27 | 99.9378
0 | 0 | 8 | 0.8 | 177.1 | 2100061.1 | 177.1 | 97.4201 | 0 | 0 | 12 | 1.2 | 22634.23 | 191193.3 | 99.74 | 99.6498 | 0 | 0 | 12 | 1.2 | 5110.91 | 29786.5 | 154.8 | 99.6403
0 | 0 | 7 | 0.7 | 173.69 | 2059723.3 | 173.69 | 95.5443 | 0 | 0 | 11 | 1.1 | 22557.9 | 190548.6 | 99.4 | 99.3138 | 0 | 0 | 11 | 1.1 | 5068.06 | 29536.7 | 153.54 | 98.8049
0 | 0 | 6 | 0.6 | 168.78 | 2001443.2 | 168.78 | 92.8434 | 0 | 0 | 10 | 1 | 22409.04 | 189291.2 | 98.67 | 98.6584 | 0 | 0 | 10 | 1 | 5038 | 29361.6 | 152.67 | 98.2189
0 | 0 | 5 | 0.5 | 162.83 | 1930915.1 | 162.83 | 89.5704 | 0 | 0 | 9 | 0.9 | 22171.32 | 187283.1 | 97.68 | 97.6118 | 0 | 0 | 9 | 0.9 | 4962.48 | 28921.4 | 150.29 | 96.7466
0 | 0 | 4 | 0.4 | 154.59 | 1833233.6 | 154.59 | 85.0377 | 0 | 0 | 8 | 0.8 | 21896.91 | 184965.2 | 96.44 | 96.4037 | 0 | 0 | 8 | 0.8 | 4879.19 | 28436 | 147.82 | 95.1228
0 | 0 | 3 | 0.3 | 142.7 | 1692205 | 142.7 | 78.4972 | 0 | 0 | 7 | 0.7 | 21402.12 | 180785.7 | 94.3 | 94.2253 | 0 | 0 | 7 | 0.7 | 4742.64 | 27640.2 | 143.69 | 92.4607
0 | 0 | 2 | 0.2 | 125.15 | 1484111.3 | 125.15 | 68.8432 | 0 | 0 | 6 | 0.6 | 20783.24 | 175557.9 | 91.58 | 91.5006 | 0 | 0 | 6 | 0.6 | 4577.36 | 26676.9 | 138.63 | 89.2384
0 | 0 | 1 | 0.1 | 94.97 | 1126221.8 | 94.97 | 52.2416 | 0 | 0 | 5 | 0.5 | 19881.38 | 167939.8 | 87.56 | 87.5301 | 0 | 0 | 5 | 0.5 | 4322.54 | 25191.9 | 130.93 | 84.2706
0 | 0 | 0 | 0 | 90.04 | 1067711.6 | 90.04 | 49.5297 | 0 | 0 | 4 | 0.4 | 18719.77 | 158127.5 | 82.46 | 82.416 | 0 | 0 | 4 | 0.4 | 3990.29 | 23255.5 | 120.94 | 77.7931
0 | 0 | -0.5 | -0.05 | 89.98 | 1066990.7 | 89.98 | 49.4967 | 0 | 0 | 3 | 0.3 | 17014.9 | 143726.4 | 74.95 | 74.9101 | 0 | 0 | 3 | 0.3 | 3558.93 | 20741.5 | 107.82 | 69.3835
[8 empty] | 0 | 0 | 2 | 0.2 | 14549.96 | 122904.8 | 64.13 | 64.0579 | 0 | 0 | 2 | 0.2 | 2958.95 | 17244.8 | 89.66 | 57.6865
[8 empty] | 0 | 0 | 1 | 0.1 | 12711.14 | 107372.1 | 56.01 | 55.9623 | 0 | 0 | 1 | 0.1 | 2495.77 | 14545.4 | 75.63 | 48.6566
[8 empty] | 0 | 0 | 0 | 0 | 12766.32 | 107838.3 | 56.28 | 56.2052 | 0 | 0 | 0 | 0 | 2404.99 | 14016.3 | 72.9 | 46.8867
[8 empty] | 0 | 0 | -0.5 | -0.05 | 12697.36 | 107255.7 | 55.93 | 55.9016 | 0 | 0 | -0.5 | -0.05 | 2388.81 | 13922 | 72.4 | 46.5713